14861 | 1220648-2024-22967 | 20666571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20666571&pc=OZD | 11/17/2021 | 03/10/2025 |  | 11/12/2024 |  |  | 11/17/2021 | 11/12/2024 |  |  | 05/31/2023 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 AIC PUMP SET US |  |  | ABIOMED, INC. |  |  | US |  | 2022096739 |  | 005042 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 91 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient was being placed on impella 2.5 for hemodynamic support. it was reported that the site was closed with preclose sutures with proglide x2. preclosed unable to close site at the end of case. 8f angioseal was also deployed. site continued to bleed. the cardiac physician at that time decided to go up and over from left femoral artery and do an angiogram off right femoral artery impella access. angiogram showed a dissection flap and some bleed. access site bleeding with the use of sutures to control access site bleeding. | [5 empty] | The impella device was not returned for evaluation. however, upon completion of the investigation, a supplemental report will be submitted. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported access site bleeding ¿ major issue has been completed. the device passed all post-sterile inspection checks. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
14862 | 1220648-2024-22960 | 20666952 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20666952&pc=OZD | 04/13/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 04/13/2021 | 11/12/2024 |  |  | 12/31/2022 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022044184 |  | 005040 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella 2.5 device for mechanical circulatory support. it was reported while on impella 2.5 support, the patient experienced hemolysis. the impella was repositioned, however hemolysis did not resolve and as a result the patient was weaned. the impella 2.5 was successfully explanted the following day. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿if the impella catheter is positioned too far into the left ventricle, the patient will not receive the benefit of impella catheter support. blood will enter the inlet area and exit the outlet area within the ventricle. obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood-colored urine.¿ impella catheter in papillary muscle: ¿if the inlet area of the catheter is lodged adjacent to the papillary muscle, the inflow may be obstructed, resulting in suction alarms. this positioning is also likely to place the outlet area too close to the aortic valve, which can cause outflow at the level of the aortic valve with blood streaming back into the ventricle, resulting in turbulent flow and hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: hemolysis. ¿patient conditions¿including catheter position, pre-existing medical conditions, and small left ventricular volumes¿may also play a role in patient susceptibility to hemolysis." "hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ this report is being filed as part of a retrospective review of historical records.
14863 | 1220648-2024-22973 | 20667307 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20667307&pc=OZD | 10/25/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 10/25/2021 | 11/12/2024 |  |  | 02/28/2023 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022067699 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella 2.5 device for mechanical circulatory support. during implantation, the impella pump was unable to pass through the patient's vasculature, therefore the implantation was aborted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14864 | 1220648-2024-22974 | 20667618 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20667618&pc=OZD | 11/15/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 11/15/2021 | 11/12/2024 |  |  | 10/31/2022 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022028303 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was being placed on impella 2.5 for hemodynamic support. it was reported the procedure was aborted due to the anatomy of the patient. there was no reported harm or injury to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14865 | 1220648-2024-22963 | 20668006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20668006&pc=OZD | 05/07/2024 | 12/16/2024 |  | 11/12/2024 |  |  | 10/17/2024 | 11/12/2024 |  |  |  |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025468676 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured insertion site infection with a definite relationship to the impella device used: ¿patient was admitted to ed (b)(6) 2024 for melena. he had a recent right femoral artery surgical repair of the impella placement site on (b)(6) 2024, due to continued drainage from the right groin vascular surgery placed a wound vac on (b)(6) 2024. the incisional wound vac was removed and cleaned on (b)(6) 2024 and converted to a traditional in-wound vac. wound vac removed on (b)(6) 2024 but surgical site still had a small area of dehiscence. the patient was seen again in office on (b)(6) 2024 where the wound was essentially healed with a small area of hypergranulation at the area of prior dehiscence. silver nitrate was applied to the femoral site where the patient had also acquired intertrigo.¿. the patient recovered with sequelae. additionally, the complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured bleeding with a possible relationship to the impella device used: ¿patient presented to ed on (b)(6) 2024 from primary care referral for profound anemia symptomatic with shortness of breath. at an outlying facility he was found to have profound anemia with a hemoglobin of 6.1 and referred to the hospital. in the ed his hemoglobin was found to be 6.1 with a hematocrit of 20.8. patient was transfused with 2 units of red blood cells. a repeat hemoglobin was found to be 8.3 where patient was found to be stable and discharged on (b)(6) 2024.¿ the patient recovered with sequelae. the patient survived the explant. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist, section: table 3.2 impella catheter components, ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation into infection/sepsis and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g.1 revised reporting contact email since initial manufacturer device report number 1220648-2024-22963 was submitted.
14866 | 1220648-2024-22990 | 20668073 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20668073&pc=OZD | 11/18/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 11/18/2021 | 11/12/2024 |  |  | 10/31/2022 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022032435 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella 2.5 device for mechanical circulatory support. it was reported that due to severe calcification, bleeding and oozing in the groin a balloon tamponade was performed. the case was ultimately aborted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14867 | 1220648-2024-22991 | 20668077 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20668077&pc=OZD | 09/10/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 09/10/2021 | 11/12/2024 |  |  | 10/31/2021 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1467903 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 2.5 device for mechanical circulatory support. during impella placement, a dissection was observed at the impella insertion site. a covered stent and balloon tamponade were utilized to manage the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14868 | 1220648-2024-22958 | 20668108 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20668108&pc=OZD | 10/18/2024 | 12/17/2024 |  | 11/12/2024 |  |  | 10/18/2024 | 11/12/2024 |  |  |  |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025553724 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support and during support, the patient had a hematoma. the hematoma was evacuated and support continued. additionally, there were on going positioning issues. there were issues with the pigtail positioning. it is unknown if the positioning resolved while on support. also, the patient had hemolyzing labs. it is unknown if the hemolysis resolved. the patient remains on impella support. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the hemolysis, positioning issues, and access site bleed has been completed. data logs confirmed pump was in improper position corresponded with clinical description that triggered alarms impella position in aorta. no other issues were found on the logs. product was not returned for review. the root cause of access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. the root cause of positioning issue was most likely use related since the pump was purposefully left shallow. the root cause of hemolysis was most likely patient condition related since patient is listed for heart (and possible kidney) transplant.
14869 | 1220648-2024-22962 | 20668243 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20668243&pc=OZD | 10/18/2024 | 12/05/2024 |  | 11/12/2024 |  |  | 10/18/2024 | 11/12/2024 |  |  | 04/30/2025 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024336980 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  |  |  | Death | Yes | Yes | Pumping Stopped | Infusion or Flow Problem | [5 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support and there was a pump stop. the pump was restarted and upon restarting, there were purge pressures high and flows dropped. big spikes in the left ventricle waveform and appeared to be with pump saturation. the impella was removed and the patient was placed on extracorporeal membrane oxygenation. it was further reported that about twelve hours later, the patient expired. the relationship to the cause of death and the impella is unknown. | [5 empty] | The pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock entering cardiac output (impella cp with smart assist) section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: troubleshooting the purge system unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation into temporary pump stop and saturated flow has been completed. the pump was inspected and biomaterial was found on the impeller. the catheter was found to be completely broken by the kink protector; opening the kink protector shows the motor cables and purge line to be intact. this is likely unrelated to the root cause of the reported failure modes as temporary pump stop and saturated flow did not reproduce in the lab. the returned data logs from the field show motor current spikes with speed deviations after which there was saturated flow and temporary pump stop. this log pattern is likely as a result of biomaterial ingestion. flows were at 5.25l/min while at p7 after motor current spike when the typical range is between 3.9-4.5l/min at that p level. the logs confirm the pump was able to be restarted after the pump stop. the root cause of the temporary pump stop was determined to be due to biomaterial. the root cause of the saturated flow was determined to be due to biomaterial. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-22962.
14870 | 1220648-2024-22977 | 20668288 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20668288&pc=OZD | 06/11/2021 | 11/24/2024 |  | 11/12/2024 |  |  | 06/11/2021 | 11/12/2024 |  |  | 10/31/2022 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022032519 |  | 005042-RO |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella support, the device was observed to be in the incorrect position. it was stated that and observed thrombus was noted in the placement signal line. the medical staff made an effort to aspirate the line was made with no result. weaning was successful, the device was explanted, and the procedural outcome was the patient survived. | [5 empty] | Corrected information for the initial mfr 1220648-2024-22977 was provided in sections b1 and h1 to reflect serious injury. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14871 | 1220648-2024-22979 | 20668434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20668434&pc=OZD | 06/15/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 06/15/2021 | 11/12/2024 |  |  | 10/31/2022 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022021735 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Bradycardia | [11 empty] | Other | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient experienced continuous ectopy during the case. it was reported that tombstone 'st' elevations in both vessels, bradycardia and pressure drops in the patient were observed. it was reported that the medical staff performed a percutaneous coronary intervention, and the impella device was explanted. the procedural outcome was the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14872 | 1220648-2024-22975 | 20668513 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20668513&pc=OZD | 05/11/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 05/11/2021 | 11/12/2024 |  |  | 05/31/2021 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1436214 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 92 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella 2.5 device for mechanical circulatory support. during implantation it was reported that insertion and removal of second perclose during preclosure was difficult. decision made to place only 1 perclose. while advancing the wire to the left ventricle, a significant hematoma and bleeding was noted. the procedure was aborted. no further information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14873 | 1220648-2024-22987 | 20668602 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20668602&pc=OZD | 05/24/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 05/24/2021 | 11/12/2024 |  |  | 01/31/2023 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022051990 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella 2.5 device for mechanical circulatory support. it was stated that the perclose and two sutures were advanced to perform the arteriotomy. slight oozing was noted at the site and manual pressure was held. the impella pump was weaned and removed without issue. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14874 | 1220648-2024-22999 | 20668679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20668679&pc=OZD | 09/16/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 09/16/2021 | 11/12/2024 |  |  | 10/31/2021 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1464963 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 2.5 device for mechanical circulatory support. while on support, during impella placement, heparin was removed from the purge system due to the patient elevated partial thromboplastin time (ptt). device remained implanted until end of support. weaning was successful, the device was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14875 | 1220648-2024-22981 | 20668706 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20668706&pc=OZD | 04/30/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 04/30/2021 | 11/12/2024 |  |  | 12/31/2021 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1477340 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 2.5 device for mechanical circulatory support. while on support, the patient had a gastrointestinal bleed with two units of packed red blood cells being administered. in addition, the patient was septic. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14876 | 1220648-2024-22980 | 20668721 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20668721&pc=OZD | 11/17/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 11/17/2021 | 11/12/2024 |  |  | 06/30/2023 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022111420 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella 2.5 device for mechanical circulatory support. it was reported that sutures were used to control access site bleeding. the device was explanted in the procedural area. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14877 | 1220648-2024-22978 | 20668728 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20668728&pc=OZD | 11/15/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 11/15/2021 | 11/12/2024 |  |  | 06/30/2023 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022120614 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella 2.5 device for mechanical circulatory support. it was reported that three perclose sutures were used to control access site bleeding. in addition, while on support the patient had lost distal pulses. the device was explanted in the procedural area. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14878 | 1220648-2024-22995 | 20668784 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20668784&pc=OZD | 03/15/2022 | 11/12/2024 |  | 11/12/2024 |  |  | 03/15/2022 | 11/12/2024 |  |  | 10/31/2023 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023164352 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that after the impella device was removed, a small dissection of the right femoral artery was observed. to manage the complication, two balloon tamponades with covered stents were performed. the patient was also prescribed two milligrams of protamine and a femstop was placed at the groin site for a duration of four hours. the device was explanted, and the procedural outcome was the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14879 | 1220648-2024-22993 | 20668816 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20668816&pc=OZD | 11/23/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 11/23/2021 | 11/12/2024 |  |  | 03/31/2023 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022090423 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pressure Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella 2.5 device for mechanical circulatory support. it was reported that a purge pressure low alarm occurred. a crack was found in a line that was leaking. the cassette was exchanged. in addition, it was reported that the patient received 1litre of plasma lite and 2 liters of blood. plans were to add heparin to the purge solution and stop peripheral heparin due to chance of bleeding until partial thromboplastin time comes back and ct scan could be done. the device was explanted in the procedural area. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14880 | 1220648-2024-22988 | 20668859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20668859&pc=OZD | 06/27/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 06/27/2021 | 11/12/2024 |  |  | 11/30/2021 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1471862 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 2.5 device for mechanical circulatory support. during support, oozing was noted at the insertion site. a new suture as well as a femostop was placed to resolve the issue. additionally, heparin induced thrombocytopenia was suspected, and the medical staff changed the purge system from heparin to argatroban. the device was removed, and the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14881 | 1220648-2024-23000 | 20668898 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20668898&pc=OZD | 09/20/2021 | 03/13/2025 |  | 11/12/2024 |  |  | 09/20/2021 | 11/12/2024 |  |  | 02/28/2023 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022101338 |  | 005042-RO |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella placement, one perclose suture broke. it was reported that slight oozing was observed. to manage the complication and achieve hemostasis, it was reported that an eight french angioseal was added. the device was explanted, and the procedural outcome was the patient survived. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product for the root cause of this investigation is not determined.
14882 | 1220648-2024-22972 | 20668913 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20668913&pc=OZD | 04/23/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 04/23/2021 | 11/12/2024 |  |  | 10/31/2022 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022032491 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 2.5 device for mechanical circulatory support. while on support, a kink in sheath was evident related to angle of entry, due to body habitus of patient. impella 2.5 loaded onto 0.018 wire, when motor housing entered kink in sheath a hematoma became apparent. unable to pass impella past kink in sheath. decision was made to use impella cp 14 fr sheath. impella 2.5 removed from sheath, 14fr impella cp sheath placed, impella 2.5 reinserted with successful passage through sheath and support started. a crack in 13fr impella 2.5 sheath was evident upon removal. hemostasis was achieved at access site with 14 fr sheath in place. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14883 | 1220648-2024-22971 | 20668972 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20668972&pc=OZD | 10/21/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 10/21/2021 | 11/12/2024 |  |  | 05/31/2023 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022090470 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was implanted with an impella 2.5 device for mechanical circulatory support. during impella placement, the device kinked when the dilator was removed. impella 25 cm sheath was used however it kinked when dilator was removed. a cook 14fr sheath was then used with success. the device remained implanted, and support was successful. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14884 | 1220648-2024-22966 | 20669038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20669038&pc=OZD | 06/05/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 06/05/2021 | 11/12/2024 |  |  | 03/31/2023 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 EU |  |  | OZD |  |  | IMPELLA 2.5 EU |  |  | ABIOMED EUROPE GMBH |  |  | GM |  | 2022059770 |  | 005040 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during impella placement, the medical staff suspected heparin induced thrombocytopenia was present when the patient's platelet count decreased. it is unknown if hemostasis was achieved. weaning was successful, the device was removed, and the procedural outcome was the patient survived explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14885 | 1220648-2024-22982 | 20669116 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20669116&pc=OZD | 06/18/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 06/18/2021 | 11/12/2024 |  |  | 02/28/2023 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022064711 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during impella placement, bleeding was observed at the femoral access site. it was reported that the sheath was removed, and manual pressure was held. it was reported that administration of systemic heparin was stopped. weaning was successful, the device was explanted, and the procedural was the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14886 | 1220648-2024-22998 | 20669171 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20669171&pc=OZD | 08/06/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 08/06/2021 | 11/12/2024 |  |  | 10/31/2022 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022021726 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that after impella weaning and explant, balloon tamponade was utilized after the perclose failed. it was stated that the perclose failed due to calcification of the patient's vessel. the procedural outcome was the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14887 | 1220648-2024-22994 | 20669231 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20669231&pc=OZD | 11/25/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 11/25/2021 | 11/12/2024 |  |  | 10/31/2022 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2021006488 |  | 005040 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock needing an impella 2.5 device for mechanical circulatory support. it was reported that an echocardiogram revealed the patient had a thrombus in the left ventricle. no further information was provided as to the treatment or outcome for the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14888 | 1220648-2024-22976 | 20669252 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20669252&pc=OZD | 04/26/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 04/26/2021 | 11/12/2024 |  |  | 12/31/2022 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022043187 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 2.5 device for mechanical circulatory support. while on support, a bleed confirmed on computed tomography (ct) from the impella to find a retroperitoneal bleeding. perclose x 3, no hematoma. the patient was given volume and two units of blood. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14889 | 1220648-2024-22997 | 20669253 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20669253&pc=OZD | 06/24/2022 | 11/12/2024 |  | 11/12/2024 |  |  | 06/24/2022 | 11/12/2024 |  |  | 06/30/2023 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022142820 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant that access to the femoral artery was achieved and the sheath was inserted without incident. it was reported that the imeplla pigtail was able to cross the left ventricular outflow tract obstruction (lvot). however, the inlet was unable to cross the aortic valve calcification. it was stated that the catheter began to coil in the aortic arch. an additional effort to cross the aortic valve after exchanging the guidewire was unsuccessful. the procedural outcome was the process of delivering the impella was aborted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14890 | 1220648-2024-22985 | 20669266 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20669266&pc=OZD | 05/02/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 05/02/2021 | 11/12/2024 |  |  | 05/31/2021 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1424928 |  | 005048 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 2.5 device for mechanical circulatory support. while on support there was continued bleeding. in addition, there was pulmonary edema remaining in the left lung. to reduce the risk of bleeding, the plan was to remove the device as soon as possible. the patient was successfully weaned from the impella and the device explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14891 | 1220648-2024-22983 | 20669366 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20669366&pc=OZD | 11/17/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 11/17/2021 | 11/12/2024 |  |  |  |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  |  |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk percutaneous coronary insertion was being implanted with an impella 2.5 device for mechanical circulatory support. while on support, the impella stopped alarm, retrograde flow alerted. the physician tried to troubleshoot; however, the alarm would not resolve, and the implantation was aborted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14892 | 1220648-2024-22970 | 20669447 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20669447&pc=OZD | 08/04/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 08/04/2021 | 11/12/2024 |  |  | 09/30/2022 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2021017612 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that after impella weaning and explant, two perclose failed and oozing from the access site was observed. vascular surgery was consulted for a cutdown and femoral artery closure. it was reported that the procedural outcome was the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14893 | 1220648-2024-22969 | 20669499 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20669499&pc=OZD | 08/04/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 08/04/2021 | 11/12/2024 |  |  | 03/30/2023 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022066434 |  | 005040 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the device was noted to be trapped in the papillary muscle and suction was observed. it was reported that the patient was sent to unit for a pump reposition, or removal and placement of a new impella device. the procedural outcome was unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14894 | 1220648-2024-22989 | 20669578 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20669578&pc=OZD | 08/19/2021 | 03/18/2025 |  | 11/12/2024 |  |  | 08/19/2021 | 11/12/2024 |  |  | 11/30/2022 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022076983 |  | 005042 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that bleeding was observed at the impella cut down site. the medical staff held manual pressure during the percutaneous coronary intervention and surgical closure to manage the complication. it was reported that two units of packed red blood cells were transfused to the patient during the case. the pump was explanted, and the procedural outcome was the patient survived. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
14895 | 1220648-2024-22986 | 20669724 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20669724&pc=OZD | 05/21/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 05/21/2021 | 11/12/2024 |  |  | 11/30/2022 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022037597 |  | 005042 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella 2.5 device for mechanical circulatory support. it was reported that during closure, a third perclose was deployed, but oozing was still observed. an angiogram was performed with a balloon tamponade. it was identified that the patient was experiencing a venous bleed and a femstop was placed. the device was explanted in the procedural area. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14896 | 1220648-2024-23001 | 20669733 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20669733&pc=OZD | 10/06/2021 | 03/18/2025 |  | 11/12/2024 |  |  | 10/06/2021 | 11/12/2024 |  |  | 12/31/2022 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022051105 |  | 005040 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella placement, elevated partial thromboplastin time (ptt). it is unknown how the complications were managed. weaning was successful, the device was removed, and the procedural outcome was the patient survived explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
14897 | 1220648-2024-22984 | 20669770 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20669770&pc=OZD | 12/13/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 12/13/2021 | 11/12/2024 |  |  | 07/31/2023 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2022132875 |  | 005040 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella 2.5 device for mechanical circulatory support. it was reported that the procedure was long and complex due to coronary guidewire pericardial effusion with urgent need for draining combined with difficulties to stop bleeding into pericardium. heparin was removed from the purge. ultimately, protamine was used. the impella was explanted as active clotting time (act) was running below 160 sec. patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14898 | 1220648-2024-23002 | 20669789 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20669789&pc=OZD | 10/11/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 10/11/2021 | 11/12/2024 |  |  | 10/31/2021 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 1468453 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 2.5 device for mechanical circulatory support. while on support, the patient experienced intermittent episodes of ventricular tachycardia which would resolve on their own. it is unknown how the complication was managed or if hemostasis was achieved. weaning was successful, and the impella device was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14899 | 1220648-2024-22992 | 20669796 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20669796&pc=OZD | 09/13/2021 | 03/18/2025 |  | 11/12/2024 |  |  | 09/13/2021 | 11/12/2024 |  |  | 06/30/2023 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022112311 |  | 005040 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  |  |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella placement, a kink in the device was noted. it was reported that the impella inlet was placed in the left ventricular outflow tract obstruction for stability. the device was removed, and the procedural outcome was the patient survived explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
14900 | 1220648-2024-22996 | 20669884 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20669884&pc=OZD | 03/28/2022 | 11/12/2024 |  | 11/12/2024 |  |  | 03/28/2022 | 11/12/2024 |  |  | 06/30/2023 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022092176 |  | 005040 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that, while attempting placement of the impella, the medical staff encountered difficulties delivering the device to the left ventricle due to the patient's vasculature. it was reported that the pump was removed, and the impella implantation and percutaneous coronary intervention was aborted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14901 | 1220648-2024-22968 | 20669901 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20669901&pc=OZD | 06/28/2021 | 11/12/2024 |  | 11/12/2024 |  |  | 06/28/2021 | 11/12/2024 |  |  | 02/28/2023 |  | 11/12/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 2.5 |  |  | OZD |  |  | IMPELLA 2.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022063286 |  | 005042-RO |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during impella placement, the patient experienced significant bleeding around the impella repositioning sheath. it was reported that the medical staff performed an emergent femoral artery cutdown and repair at the site where the sheath was located, and hemostasis was achieved. weaning was successful, the device was removed, and the procedural outcome was the patient survived explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
14902 | 1220648-2024-22861 | 20650830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20650830&pc=OZD | 12/14/2022 | 11/11/2024 |  | 11/11/2024 |  |  | 12/14/2022 | 11/11/2024 |  |  |  |  | 11/11/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1542729 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Loss of Data | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported inaccurate placement signal on the automated impella controller (aic). the aic was exchanged to the back-up console for continued support. no further details were provided. however, there was no report or indication of interruption in support or any harm to the unknown age patient. | [5 empty] | The automated impella controller data logs and console were returned for evaluation and analysis. data analysis demonstrated logs were consistent with the reported complaint. an error message of invalid signal occurred and cleared repetitively through the support. however, neither of the error messages lasted for more than two minutes. thus, no user interface alarm was triggered or visible to the user. according to the rtlogs, raw optical placement signal drops from normal value to invalid value (and self-recovered soon) for multiple times, which corresponded to the invalid signal errors in imc logs and was most likely observed from graphical user interface by user that the placement signal (waveform) was not within the expected range. the same pump was transferred to another console for continuous support. however, the backup console logs revealed the same symptom recorded by rtlogs. the console was tested with a demonstration pump and purge cassette, no placement issue was reproduced. inspection of blue receptacle and 5s parameters did not reveal major issues or contaminations that would lead to the alarms. there was no problem found on console. in conclusion, the failure mode was not reproduced during testing. the console functioned/operated to specification. note: given the impella pump was used with both consoles which manifested same optical placement signal issue, a complaint on the pump was created and investigated. this report is being filed as part of a retrospective review of historical records
14903 | 1220648-2024-22862 | 20651249 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20651249&pc=OZD | 10/16/2024 | 12/10/2024 |  | 11/11/2024 |  |  | 10/16/2024 | 11/11/2024 |  |  |  |  | 11/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025528147 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support. while on support, hematoma was noted at the impella site. reported cause was patient being noncompliant and moving arm freely. the patient received one unit of red blood cells and washout of axillary hematoma. the patient was transitioned to palliative/comfort care and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely patient movement since the hematoma at the impella site was formed due to patient being noncompliant and moving arm freely. corrections to the initial mfr report: g1 mdr reporting contact email | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
14904 | 1220648-2024-22863 | 20651333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20651333&pc=OZD | 10/16/2024 | 12/16/2024 |  | 11/11/2024 |  |  | 10/16/2024 | 11/11/2024 |  |  |  |  | 11/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025545030 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. during removal of impella, a generous amount of clot came with it from the peel away. the patient survived the event. | [5 empty] | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
14905 | 1220648-2024-22864 | 20651474 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20651474&pc=OZD | 10/12/2024 | 12/13/2024 |  | 11/11/2024 |  |  | 10/12/2024 | 11/11/2024 |  |  |  |  | 11/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025547629 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella 5.5 support for a 47-year-old male patient, there was excessive jp drain output. axillary washout and exploration was performed. the patient was given one unit of packed red blood cells. the patient expired while on support after arrest on the unit. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation for the reported access site bleed has been completed. the root cause of the access site bleeding issue was not determined since the product was not returned and insufficient clinical details provided. b.2 required intervention was added as it was inadvertently omitted from initial manufacturer device report number 1220648-2024-22864. g.1 revised reporting contact email since initial manufacturer device report number 1220648-2024-22864 was submitted. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
14906 | 1220648-2024-22865 | 20651797 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20651797&pc=OZD | 03/25/2022 | 11/11/2024 |  | 11/11/2024 |  |  | 03/25/2022 | 11/11/2024 |  |  |  |  | 11/11/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1313572 |  | 0042-0000-UK |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  |  | Asian | Death | Yes | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. while on support, the aic showed ac power disconnected alarm frequently even when the cable is plugged in. changed to another backup aic. due to the fact that the patient couldn't wean the support of ecmo and weak native heart, relatives decided to stop support. the survival outcome at explant noted the patient expired. there is not enough information to exclude the aic device as an associated factor in the patient's demise. | [5 empty] | D9: date device returned to manufacturer: unknown. the investigation for the reported power issue has been completed. the product was returned, and the issue was not confirmed. logs from the complaint date and earlier time are consistent with the reported symptom and confirm the failure mode. ac power disconnected, event set: #109 was triggered and cleared instantaneously. field service technician performed a thorough test for the console, however, no ac power/ charging issue was noticed or reproduced. during the power cycles, the console could switch between charging and running on batteries. no abnormality or bad connection was found for the console. per the results of the clinical review and investigation, the root cause was not determined since the failure mode was not reproduced in fs. this report is being filed as part of a retrospective review of historical records.
14907 | 1220648-2024-22866 | 20651831 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20651831&pc=OZD | 12/26/2021 | 11/11/2024 |  | 11/11/2024 |  |  | 12/26/2021 | 11/11/2024 |  |  |  |  | 11/11/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1401769 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. while on support, during transport, a ¿battery temperature high- switch to back-up aic¿ alarm noted which resolved upon plugging in the device. once able, the aic was switched to another aic without incident. the survival outcome at explant noted the patient expired. there is not enough information to exclude the aic device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported battery charging issue has been completed. the product was returned, and the issue was confirmed. data analysis: imc logs are consistent with complaint date. ltpowerdata was reviewed at the time of occurrence according to imclogs however no data points measured meet the triggering criteria for the aic alarm. the high temperature battery alarm could not be reproduced even being simulated in burn-in room for long term testing (24 hours). complaint cause is most likely due to pbm incorrectly reporting battery temperatures. this reading falsely reports the battery temps as something oor triggering the battery alarm. when the next sample is taken 10s later, the pbm reads a correct battery data sample. this can be seen in the logs when the battery temp high alarm is triggered, and exactly 10 second later the alarm clears. per the results of the clinical review and investigation, the root cause was determined to be due to the pbm misreading battery data during a single sample. this report is being filed as part of a retrospective review of historical records.
14908 | 1220648-2024-22867 | 20651853 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20651853&pc=OZD | 10/18/2024 | 11/15/2024 |  | 11/11/2024 |  |  | 10/18/2024 | 11/11/2024 |  |  |  |  | 11/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025528281 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient was to be implanted with an impella 5.5 but the pump was unable to cross at the level of the hemashield graft/axillary artery after multiple attempts. insertion was ultimately aborted. th eecision was made to place an impella cp through the axillary graft instead. the patient¿s outcome is unknown. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | Device not returned to manufacturer for evaluation: should have been left blank on the initial manufacturer device report 1220648-2024-22867 as the device was received, pending investigation.
14909 | 1220648-2024-22868 | 20651891 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20651891&pc=OZD | 07/06/2021 | 11/11/2024 |  | 11/11/2024 |  |  | 07/06/2021 | 11/11/2024 |  |  |  |  | 11/11/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1080533 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Unknown |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Application Program Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy needing an impella device for mechanical circulatory support. it was reported that while on support there was an aic shutdown with aic - controller error (yellow alarm). the alarm resolved by itself after about 10 minutes. the team was able to proceed. additional information was provided that noted the patient expired. | [5 empty] | B2 date of patient death was not provided by the complainant. the investigation for the reported aic - controller error (yellow alarm) and temporary console shut down issue has been completed. the product was returned and the issue was confirmed. logs show that during the reported case, the io-graphics driver crashed. consistent with the complaint intake, an "unexpected controller shutdown" alarm was then issued. after console arrived in danvers, several boot cycles were performed with all normal results. console successfully booted correctly each time. this failure mode, although rare, is known and is manifested in disappearance of graphics interface with a text screen remaining, that is mistakenly interpreted as console rebooting or restarting (software defect in jama: gid-dft-863138) following the software defect mitigation description in jama, as of v8.4, it reads: "system idle process will wait on notification from dumper process when a process dies, and if the process pid matches that of the io graphics process, system idle will put the system in full reset." the io-graphics driver crash is known to occur, this being a rare occurrence. the console behavior reported in the field and seen/verified with console logs show the mitigation against the io graphics driver process actively working by restarting console when system idle detects the process crash. before returning to customer, fs will be instructed to perform fc. this report is being filed as part of a retrospective review of historical records.
14910 | 1220648-2024-22869 | 20651915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20651915&pc=OZD | 05/12/2021 | 11/11/2024 |  | 11/11/2024 |  |  | 05/12/2021 | 11/11/2024 |  |  |  |  | 11/11/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1283305 |  | 0042-0000 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Read Input Signal | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella device for mechanical circulatory support during a coronary artery bypass grafting¿procedure. the 5.0 was not inserted into patient because the physician could not get wire across valve. during the case the aic would not recognize the 5.0. the procedure was aborted due to the patient's anatomy. additional information was provided that noted the patient expired. | [5 empty] | The investigation for the reported pump not detected has been completed. the root cause of the pump not recognized issue was unable to be determined because the console and data were not returned for review. this report is being filed as part of a retrospective review of historical records.
14911 | 1220648-2024-22870 | 20652098 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20652098&pc=OZD | 08/22/2022 | 11/11/2024 |  | 11/11/2024 |  |  | 08/22/2022 | 11/11/2024 |  |  |  |  | 11/11/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1450938 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Read Input Signal | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was noted that the automated impella controller had a malfunctioning purge disc. the aic was replaced and a new console placed for patient use. during the aic exchange. it was noted that the patient's condition continued to decline without improvement on support. additional information noted care was withdrawn and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation into the reported issue has been completed. the device was received for evaluation. logs were downloaded and analyzed, but they were found to be irrelevant for the reported failure. the purge disc retainer was confirmed to be broken. before returning this console back to the customer, field service was instructed to replace the purge disc retainer, validate sensor accuracy via section 12 of the pm procedure (0042-7309), and perform a full functional check on the console and optical system (0042-7316). the cause of the issue was a defective component. the cause of the broken purge retainer was likely due to the wear and tear of the plastic component as well as additional unnecessary force when inserting the purge disc into the purge disc retainer. during prior cases. this report is being filed as part of a retrospective review of historical records.
14912 | 1220648-2024-22872 | 20652122 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20652122&pc=OZD | 10/17/2024 | 12/13/2024 |  | 11/11/2024 |  |  | 10/17/2024 | 11/11/2024 |  |  |  |  | 11/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025548421 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump placed to support a patient admitted in cardiogenic shock and cardiomyopathy. the pump had a purge sidearm leak on the day after the implant. the purge pressure and purge flow were intact and pump remains on for support. there was no patient harm. | [5 empty] | The investigation for the purge leak has been completed. the pump and purge cassette were returned for investigation. no purge-related abnormalities were observed in the data log. a clamp mark was observed near the motor housing of the returned pump. no other physical abnormalities were observed. the pump was primed successfully, and purge was able to come out of the pump. purge ran for 20 minutes. no leak was observed from the side arm, as mentioned in the clinical details. the cause of the purge leak issue was most likely use related. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
14913 | 1220648-2024-22873 | 20652123 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20652123&pc=OZD | 07/14/2021 | 11/11/2024 |  | 11/11/2024 |  |  | 07/14/2021 | 11/11/2024 |  |  |  |  | 11/11/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1155468 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  |  | Death | Yes | Yes | Priming Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock needing an impella device for mechanical circulatory support. it was reported that there was a failure to prime. the impella device was not placed in the body as the catheter lab felt that the impella was not priming correctly and did not insert the pump. the physician decided at that time to use a new aic and a new impella device. additional information was provided that noted the patient expired. | [5 empty] | The investigation for the reported priming issue has been completed. the cause of the failure mode was use related. investigation results provided indicated staff began case start wizard, following steps accordingly. when arriving to step 3, the pump was connected and from the viewpoint of staff, would have appeared as though the console was skipping steps in the case start wizard. the reported complaint of the pumps alleged failure to prime is likely due to the fact that when the purge line was prepped at the site, the yellow leurs were already open, leading to the pump being primed almost immediately. this caused the case start wizard to "skip" or advance ahead through the case start wizard. logs confirmed this to be intentional in console ui software, with rt logs showing valid purge build and pressure. the cause of the issue was not determined since product was not returned.
14914 | 1220648-2024-22871 | 20652127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20652127&pc=OZD | 06/06/2024 | 12/16/2024 |  | 11/11/2024 |  |  | 10/16/2024 | 11/11/2024 |  |  |  |  | 11/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025461249 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 63 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured hemolysis with a "possible" relationship to the impella 5.5 device used: ¿patient with anemia and ldh reading of 2.100, pfhb of 811 and elevated direct bilirubin >1.0. patient transfused multiple units of rbc during admission for anemia, likely component of hemolysis¿. it was reported that the patient/event has not recovered/not resolved. loqi also reported that the patient endured aki with a "possible" relationship to the impella device used: ¿significant increase in creatinine. patient started on cvvhd on (b)(6) 2024¿. it was reported that the patient/event has not recovered/not resolved. | [5 empty] | The investigation into hemolysis and renal failure has been completed. the impella device was not returned for evaluation. the cause of the hemolysis issue was most likely patient condition related, based on data log review and provided clinical details. the cause of the renal failure issue was not determined and the unrelated to impella, based on provided clinical details. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-22871. e2 missing. e3 missing. g1 reporting contact email. g2 report source; study missing. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
14915 | 1220648-2024-22874 | 20652131 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20652131&pc=OZD | 10/17/2024 | 12/12/2024 |  | 11/11/2024 |  |  | 10/17/2024 | 11/11/2024 |  |  |  |  | 11/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025567111 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that impella cp was placed emergently in the setting of hemodynamic compromise. post percutaneous coronary intervention, the medical doctor (md) took 14fr sheath out of the patient and advanced the repositioning sheath. during advancement of the sheath, the cp deep dived into the left ventricle (lv) completely and the inlet curled towards the mitral valve. md pulled cp out of lv, and a kink was noted at the cannula right above the aortic valve, near the motor housing. the md attempted to reposition to release kink, without success. the md made the decision to emergently explant cp and reimplant a new impella cp. the patient remained stable and the procedural outcome was the patient survived surgery. | [5 empty] | The pump was returned for investigation. data logs were not provided for this case run. a cannula kink was found on the returned product. no further investigation was required for this case run. as per the clinician, the pump was inserted too deep towards the mitral valve while advancing the sheath. the doctor explanted the pump due to the kink found during positioning. product damage¿cannula kink failure mode was changed to a positioning issue after investigation. the cannula was found to be kinked due to the position of the device. the cause of the positioning issue was most likely use related. the following have been reported incorrectly on manufacturer device report 1220648-2024-22874. e4 should have been left blank. g1 reporting contact email. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
14916 | 1220648-2024-22875 | 20652134 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20652134&pc=OZD | 12/23/2023 | 12/17/2024 |  | 11/11/2024 |  |  | 10/17/2024 | 11/11/2024 |  |  |  |  | 11/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Myocardial Infarction | Renal Failure | [9 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Abiomed japan forwarded the publication entitled "association of preprocedural synergy between percutaneous coronary intervention with taxus and cardiac surgery (syntax) score with outcomes in impella-assisted high-risk percutaneous coronary intervention" with primary author from nyu langone discussed protect iii patients with syntax 1 scores. despite aggressive revascularization, a high syntax i score was associated with higher rates of 90-day macce and 1-year mortality rates in patients who underwent impella-assisted high-risk percutaneous coronary intervention. paper looked to 90 day major adverse cardiac and cerebrovascular events (macce) rates and 1 year mortality. the paper found that adverse events at discharge demonstrated a greater occurrence of stage 2 or 3 renal injury in patients with a high syntax i score. also, a trend toward higher rates of significant bleeding in patients with a high syntax i score was also noted. furthermore, it was also noted that there were patient's with a stroke upon discharge up to 90 days after discharge. patient's also had mycardial infraction upon discharge up to 90 days after discharge. the paper also found that there were patient deaths upon discharge up to 1 year. reportable due to the potential relationship to the impella and the patient's strokes, myocardial infraction, bleeding, renal injuries, and cause of deaths. | [5 empty] | The investigation for the renal failure, stroke, access site bleed, vascular damage, and ultimately death has been completed. the device was not returned for evaluation. clinical details were not sufficient to determine the root cause of the reported issues. therefore, the root cause of the renal failure, stroke, access site bleed, vascular damage, and ultimately death could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. a.1, a.2, a.3, a.4 and a.5 are unknown. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. h.4 device manufacture date is unknown. medranda, g. a., faraz, h. a., thompson, j. b., zhang, y., bharadwaj, a. s., osborn, e. a., abu-much, a., lansky, a. j., basir, m. b., moses, j. w., o¿neill, w. w., grines, c. l., & baron, s. j. (2024). association of preprocedural syntax score with outcomes in impella-assisted high-risk percutaneous coronary intervention. journal of the society for cardiovascular angiography & interventions, 3(8), 101981. https://doi.org/10.1016/j.jscai.2024.101981.
14917 | 1220648-2024-22877 | 20652148 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20652148&pc=OZD | 05/08/2024 | 12/12/2024 |  | 11/11/2024 |  |  | 10/17/2024 | 11/11/2024 |  |  |  |  | 11/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Abiomed japan forwarded this publication entitled "implications of anemia in patients undergoing pci with impella support: insights from the protect iii study" in which anemia and impella patients were reviewed. results of a logistic regression model for blood transfusion showed that moderate to severe anemia prolonged impella support and peripheral vascular disease were independent factors associated with blood transfusion. this suggests that patients with baseline anemia have more hemolysis during impella support and impaired hemostasis worsening their existing baseline anemia and necessitating blood transfusion. these findings raise the question of whether pre-procedure blood transfusion should be considered in certain cases along with a more meticulous access and closure technique planning to minimize procedural complications. | [5 empty] | The investigation for the hemolysis has been completed. pump was not retuned of investigation. data log review was not returned for investigation. concern over hemolysis during pump use. the cause of the hemolysis issue was not determined since product was not returned and sufficient clinical information was not provided. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. falah, b., redfors, b., zhao, d., bharadwaj, a. s., basir, m. b., thompson, j. b., patel, r. a., schonning, m. j., abu-much, a., zhang, y., batchelor, w. b., grines, c. l., & o¿neill, w. w. (2024). implications of anemia in patients undergoing pci with impella-support: insights from the protect iii study. frontiers in cardiovascular medicine, 11. https://doi.org/10.3389/fcvm.2024.1429900.
14918 | 1220648-2024-22876 | 20652149 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20652149&pc=OZD | 10/20/2021 | 12/17/2024 |  | 11/11/2024 |  |  | 10/17/2024 | 11/11/2024 |  |  |  |  | 11/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [9 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Abiomed japan forwarded 14 publications. one from the team at yale is entitled "mechanical circulatory support in myocardial infarction complicated by cardiogenic shock: impact of sex and timing" was found inclusive of the impella 2.5, impella cp, and impella 5.0 supports in both male and female patients. the team did a retrospective review of patients and found in the patients from 2017-2019 a total of 358 patients to analyze for the effect of gender/sex and timing to therapy and outcomes for patient. the complications noted were:in-hospital mi, stroke, tia, repeat vascularization, major bleeding, and vascular complications requiring surgery, vascular complication without surgery, acute renal dysfunction/failure, new renal replacement therapy required. reportable for the harms. there is no identification of which pumps were attributed to each of the harms. | [5 empty] | The investigation into access site bleeding, stroke/cva, vascular damage - peripheral vessel, renal failure, and hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly on manufacturer device report. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. a.1, a.2, a.3, a.4 and a.5 are unknown. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. h.4 device manufacture date is unknown. shah, t., lansky, a. j., grines, c. l., o¿neill, w. w., moses, j. w., chieffo, a., kapur, n. k., & chou, j. (2022). mechanical circulatory support in myocardial infarction complicated by cardiogenic shock: impact of sex and timing. journal of the society for cardiovascular angiography & interventions, 1(1), 100002. https://doi.org/10.1016/j.jscai.2021.100002
14919 | 1220648-2024-22878 | 20652165 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20652165&pc=OZD | 09/01/2023 | 12/17/2024 |  | 11/11/2024 |  |  | 10/17/2024 | 11/11/2024 |  |  |  |  | 11/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Stroke/CVA | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Abiomed japan forwarded 14 publication, one from memorial hermann entitled "prehabilitation maximizing functional mobility in patients with cardiogenic shock supported on axillary impella" in which 5.0/5.5 patients had physical therapy and noted the following adverse event/outcome information: the primary outcome of this study was in-hospital and 6-month mortality. the secondary outcomes were ambulation-related complications, including bleeding, hemolysis, stroke, pump malfunction, ventricular arrhythmia, and limb ischemia. all patients survived to hospital discharge and remained alive at 6 months. complications related to impella placement included five (13.5%) bleeding events, one (2.7%) case of hemolysis, one (2.7%) stroke, and two (5.4%) pump malfunctions. no ventricular arrhythmias or limb ischemia were reported. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. bhardwaj, a., salas de armas, i. a., bergeron, a., sauer, r. m., gilley, c., reeves, k., patarroyo-aponte, m., akay, m. h., patel, m., kumar, s., patel, j., marcano, j., nathan, s., gregoric, i. d., & kar, b. (2024). prehabilitation maximizing functional mobility in patients with cardiogenic shock supported on axillary impella. asaio journal, 70(8), 661¿666. https://doi.org/10.1097/mat.0000000000002170. | The investigation for the access site bleed, hemolysis, stroke, and pump stop has been completed. the pump was not returned for evaluation. clinical details were not sufficient to determine root cause. therefore the root cause of the access site bleed, hemolysis, stroke, and pump stop could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email.
14920 | 1220648-2024-22879 | 20652171 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20652171&pc=OZD | 05/24/2022 | 02/11/2025 |  | 11/11/2024 |  |  | 10/17/2024 | 11/11/2024 |  |  |  |  | 11/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Abiomed japan forwarded a publication entitled "incidence and severity of thrombocytopenia associated with use of intravascular microaxial ventricular assist devices for treatment of cardiogenic shock" in which 2 years of impella 2.5, cp, rp, and 5.0 were reviewed for thrombocytopenia. there was noted 98% developed thrombocytopenia. | [5 empty] | The investigation for the reported thrombocytopenia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the thrombocytopenia could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. a.1, a.2, a.3, a.4 and a.5 are unknown. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. h.4 device manufacture date is unknown. falah, b., redfors, b., zhao, d., bharadwaj, a. s., basir, m. b., thompson, j. b., patel, r. a., schonning, m. j., abu-much, a., zhang, y., batchelor, w. b., grines, c. l., & o¿neill, w. w. (2024). implications of anemia in patients undergoing pci with impella-support: insights from the protect iii study. frontiers in cardiovascular medicine, 11. https://doi.org/10.3389/fcvm.2024.1429900.
14921 | 1220648-2024-22881 | 20652181 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20652181&pc=OZD | 10/17/2024 | 09/22/2025 |  | 11/11/2024 |  |  | 10/18/2024 | 11/11/2024 |  |  |  |  | 11/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025511478 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported the patient was on impella 5.5 support. while on support, the patient developed hemolysis. heparin was stopped. the patient remains on impella support. | Additional information indicated that, during patient support, the pump experienced placement signal not reliable alarm. audio alarms had been silenced. |  |  |  |  | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. there was no issue found during the case. the device functioned/operated to specification. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 updated with product problem b5 updated with additional information. d6b updated with explant date. h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2024-2288.
14922 | 1220648-2024-22880 | 20652183 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20652183&pc=OZD | 02/21/2024 | 12/13/2024 |  | 11/11/2024 |  |  | 10/16/2024 | 11/11/2024 |  |  | 09/30/2025 |  | 11/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024410527 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Myocardial Infarction | [11 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a cp support in (b)(6) 2024 for just over 5 hours when the pump was explanted and patient survived the percutaneous coronary intervention with the cp support. the pump/patient were in the abiomed's long-term outcome and quality indicator (loqi) impella registry database. the loqi in (b)(6) 2024 reported upon non-st myocardial infarction with an unknown relationship. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. | The investigation for the myocardial infarction has been completed. the impella device was not returned for evaluation. therefore, without sufficient clinical details, the root cause of the reported issues myocardial infarction could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrections to the initial mfr report: g1 mdr reporting contact email h6 clinical code 1721 changed to 1969.
14923 | 1220648-2024-22882 | 20652282 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20652282&pc=OZD | 10/17/2024 | 12/16/2024 |  | 11/11/2024 |  |  | 10/17/2024 | 11/11/2024 |  |  |  |  | 11/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025480669 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp developed oozing at the impella access site. the registered nurse tried to angle match, hold pressure, and redress, but issue continued. the medical doctor came to bedside and added stitch and bleeding subsided. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding issue was most likely use related, md came to bedside and added stitch and the bleeding subsided as per the clinical description provided. corrections to the initial mfr report: g1 mdr reporting contact email.
14924 | 1220648-2024-22883 | 20652295 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20652295&pc=OZD | 10/04/2024 | 12/16/2024 |  | 11/11/2024 |  |  | 10/17/2024 | 11/11/2024 |  |  |  |  | 11/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025546801 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support the patient had possible hemolysis. there were elevated lactate dehydrogenase and free hemoglobin post procedure. blood products were given and heparin was withheld. the patient survived the event. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation into hemolysis has been completed. the impella device was not returned for evaluation. the cause of the hemolysis issue was most likely patient condition related based on given clinical details. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-22883.
14925 | 1220648-2024-22885 | 20652459 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20652459&pc=OZD | 02/18/2022 | 11/11/2024 |  | 11/11/2024 |  |  | 02/21/2022 | 11/11/2024 |  |  |  |  | 11/11/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1298592 |  | 0042-0010-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 35 YR |  | Male |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. while on support, a damaged power cable was noted during transport of the patient. a new console was used. there was no report of patient harm or injury while on support. | [5 empty] | D9 - date device returned to manufacturer: unknown. the investigation for the reported cabling issue has been completed. the product was returned, and the issue was confirmed. data analysis: aic logs were not available for review. device analysis: once received, the console¿s power cord was found to have damaged insulation/ bare wires. per the results of the clinical review and investigation, the root cause was determined to be a defective cable. this report is being filed as part of a retrospective review of historical records.
14926 | 1220648-2024-22886 | 20652488 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20652488&pc=OZD | 10/17/2024 | 02/10/2025 |  | 11/11/2024 |  |  | 10/17/2024 | 11/11/2024 |  |  |  |  | 11/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1126219 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 44 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp with smartassist for six days. an alarm ¿controller error¿ triggered twice and spontaneously resolved. a backup automated impella controller was used but the same alarm triggered on the second aic. the patient survived the explant. | [5 empty] | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. the logs confirmed the reported failure mode given there were multiple instances of a controller error alarm triggered during the clinical procedure. the console was returned for investigation. section 23 of the preventative maintenance was performed and specifications were met. the console operated with a known good pump and purge for a significant time with no anomaly. the controller error is due to an optical bench firmware communication protocol deficiency, resulting in misalignment of data communication packets received by epc. the console software operated as designed. d.2 common device name was revised since the initial manufacturer device report number 1220648-2024-22886 was submitted in accordance with updated procedures. g.1 reporting contact email was revised since the initial manufacturer device report number 1220648-2024-22886 was submitted in accordance with updated procedures. h.4 revised device manufacture date as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-22886. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
14927 | 1220648-2024-22887 | 20652660 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20652660&pc=OZD | 04/25/2024 | 12/30/2024 |  | 11/11/2024 |  |  | 10/17/2024 | 11/11/2024 |  |  |  |  | 11/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | UNKNOWN |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [10 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Abiomed japan forwarded publication entitled "preoperative impella therapy in patients with ventricular septal rupture and cardiogenic shock: hemodynamic and organ function outcomes" in which there were adverse events noted of device related bleeding, nosebleed, gi bleed, hematuria, cva, transfusion, dialysis, stroke, tracheostomy, sternal osteomyelitis, hemorrhagic rectal ulcer, peripheral vascular dissection, reoperation for the impella patients. | [5 empty] | The investigation for the stroke, access site bleed, and vascular damage has been completed. the pump was not returned for investigation. a data log was not provided for the investigation. the patient was reported to have a nosebleed, gi bleed, peripheral vascular dissection, and hemorrhagic rectal ulcer. additionally, cva and stroke were reported during treatment. the cause of the access site bleeding issue was most likely patient condition related since the bleeding was noted from multiple sites. the cause of the vascular damage issue was not determined since product was not returned and sufficient clinical information was not provided. corrections to the initial mfr report:g1 mdr reporting contact email. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. a.1, a.2, a.3, a.4 and a.5 are unknown. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. h.4 device manufacture date is unknown. echieh, c. p., ryan, a., cherian, a., rohilla, y., wang, k., & kazui, t. (2024). preemptive impella 5.5 insertion to reduce operative risk in high-risk cardiac surgery: a case report. international journal of surgery case reports, 121, 109947. https://doi.org/10.1016/j.ijscr.2024.109947.
14928 | 1220648-2024-22888 | 20652692 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20652692&pc=OZD | 08/05/2021 | 11/11/2024 |  | 11/11/2024 |  |  | 08/05/2021 | 11/11/2024 |  |  |  |  | 11/11/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  |  |  | 1504376 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during preparation, the automated impella controller (aic) failed self-check and would not start. the console was replaced with a backup unit, and there was no reported patient harm. | [5 empty] | The device was returned and evaluated, and the investigation for the reported self-check error has been completed. data analysis: the imc logs revealed that the console was failing system self-check due to checkversions failure. this can be a result of an incompatible firmware (fm) version. device analysis: as received, the failure mode could be persistently reproduced. upon further inspection of the systemconfig file on the console, the console was found to have undergone the installation for aic software v8.4 in emfg, but for an undetermined reason, the wrong fw was kept on the sau_powermod_1 and 2 modules. the root cause for the failure of system self-check was due to incompatible fw on the pbm. the reason for the wrong pbm fw was undetermined. this report is being filed as part of a retrospective review of historical records.
14929 | 1220648-2024-22889 | 20652792 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20652792&pc=OZD | 10/17/2024 | 01/02/2025 |  | 11/11/2024 |  |  | 10/17/2024 | 11/11/2024 |  |  |  |  | 11/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025523518 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | Obstruction/Occlusion | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp insertion was not possible due to calcification of the vessel. the case was aborted. the patient¿s outcome is unknown. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue is determined to be patient condition because of the calcification in the vessels and the pump insertion was aborted. corrections to the initial mfr report: g1 mdr reporting contact email
14930 | 1220648-2024-22890 | 20652865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20652865&pc=OZD | 09/29/2021 | 11/11/2024 |  | 11/11/2024 |  |  | 09/29/2021 | 11/11/2024 |  |  |  |  | 11/11/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2022048731 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Battery Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the automated impella controller (aic) exhibited battery failure and battery communication failure. the device was replaced with another aic with no report of patient harm or injury attributed to this malfunction. the procedural outcome noted that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: date of patient death is unknown. the device was returned for evaluation. the cause of the battery failure alarm was due to a defective battery (rapid decline in cell voltage). the exact cause of the defective battery was unable to be determined. this report is being filed as part of a retrospective review of historical records.
14931 | 1220648-2024-22891 | 20652961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20652961&pc=OZD | 10/17/2024 | 12/16/2024 |  | 11/11/2024 |  |  | 10/17/2024 | 11/11/2024 |  |  |  |  | 11/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025533286 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 30 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a 5.5 pump being placed for the support of a patient with systemic and cardiac comorbidities. the pump failed to deliver due to anatomic limitations. the patient¿s outcome is unknown. | [5 empty] | The investigation into delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to patient condition because the pump was unable to pass through the vasculature due to body habitus. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-22891 | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
14932 | 1220648-2024-22892 | 20653553 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20653553&pc=OZD | 10/16/2024 | 11/11/2024 |  | 11/11/2024 |  |  | 10/16/2024 | 11/11/2024 |  |  |  |  | 11/11/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1547850 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the impella cp stopped. investigation of the pump confirmed that the impella experienced a pump stop. the automated impella controller (aic) was returned for functional testing because the customer wanted to be sure that the aic was working as intended. no patient harm was reported. | [5 empty] | The investigation has been completed. the console logs from the reported day of event were consistent with the complaint and show that after approximately five hours of support, the pump suddenly stopped, and did not restart. during testing of console, the issue was not reproduced and it has been concluded that reported pump stop was unrelated to the console. this failure mode was investigated under complaint number (b)(4) (complaint against the pump) and was reported in manufacturer device report number 1220648-2024-18274.
14933 | 1220648-2024-22893 | 20653706 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20653706&pc=OZD | 10/16/2024 | 12/27/2024 |  | 11/11/2024 |  |  | 10/16/2024 | 11/11/2024 |  |  |  |  | 11/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025536008 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support. it was reported that the doctor stated that the patient had hemolysis. plasma free hemoglobin was not measured. the hemolysis was due to an incorrect impella position. after evaluation and repositioning, hemolysis was minimized. it was further reported that care was withdrawn and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation for the hemolysis has been completed. product was not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
14934 | 1220648-2024-22894 | 20653741 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20653741&pc=OZD | 10/17/2024 | 03/10/2025 |  | 11/11/2024 |  |  | 10/17/2024 | 11/11/2024 |  |  |  |  | 11/11/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1140346 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 44 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a controller error alarm triggered twice and the alarm self-resolved. the automated impella controller (aic) was exchanged to the backup aic. there was no patient harm. | [5 empty] | During review of the complaint file, it was identified that the product-specific information for the complaint device investigated and reported on (initial and follow-up 1) was erroneously represented. corrections have been made accordingly to respective fields-section d4: serial number, lot number and unique identifier and section h: device manufacture date. h.6 codes were added as they were inadvertently omitted from manufacturer device report follow up number 1220648-2024-22894. | The investigation for the console (aic) controller alarm-yellow alarm has been completed. data logs were reviewed confirmed the reported failure mode given there was an instance of a controller error alarm (opm comm crc error ¿ event set # 1045) triggered during the clinical procedure. console was returned for evaluation. upon preventative maintenance, specifications were met. the controller error is due to an optical bench fw communication protocol deficiency, resulting in misalignment of data communication packets received by epc. console sw operated as designed. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
14935 | 1220648-2024-22895 | 20654110 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20654110&pc=OZD | 04/20/2023 | 11/11/2024 |  | 11/11/2024 |  |  | 04/20/2023 | 11/11/2024 |  |  |  |  | 11/11/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1540253 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support there was an aic - system self-check error. the aic indicated ao-position alarms. the impella position was confirmed with echocardiogram. the physician later repositioned the pump 3cm deeper in the left ventricle. after the repositioning the readings were better and stable. the perfusionist deactivated all alarms/retrograde flow/signals. it re-activated successfully. there was no harm or injury reported to the patient. | [5 empty] | The investigation for the reported aic - system self-check error issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of rlm connectivity issue was incorrectly configured or inadequate site network infrastructure (includes low wifi signal strength). this report is being filed as part of a retrospective review of historical records.
14936 | 1220648-2024-22896 | 20654157 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20654157&pc=OZD | 05/27/2022 | 11/11/2024 |  | 11/11/2024 |  |  | 05/27/2022 | 11/11/2024 |  |  |  |  | 11/11/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1200770 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) that there was a controller error alarm during preparation of the device. a new console was used with no adverse impact to the patient. | [5 empty] | The investigation for the reported issue has been completed. the product was returned, and the issue was not confirmed. console logs from the reported day of event are consistent with complaint and show controller error alarm (#24) due to kbd comm loss, upon console boot-up. this is preceded by communication error (imc-kbd error: 0xa4 0x01 0x00), and the fact that console¿s boot-up count was not consecutive, indicating prior failed boot-up sequence: after console was manually power-cycled, the issue did not re-occur. console was tested in danvers, and power-cycled several times, but the issue was not reproduced. however, this is a known rare pattern failure which results from momentary communication breakdown during initialization of kbd to 4-in-1 communication. this issue could be caused by either of the communicating devices/microprocessors: msc1210 on the 4-in-1 or at42qt1060 on the kbd assembly. a compromised ribbon cable between these two devices is also a plausible cause. the cause of the issue was not determined since issue was unable to be reproduced. this report is being filed as part of a retrospective review of historical records.
14937 | 1220648-2024-22897 | 20659260 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20659260&pc=OZD | 08/03/2022 | 11/11/2024 |  | 11/11/2024 |  |  | 08/03/2022 | 11/11/2024 |  |  |  |  | 11/11/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1087352 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient on an automated impella controller (aic) for mechanical circulatory support. during start up the blue system self-check failed screen despite power on and off and resetting battery. no further information is available. based on the limited information at hand, as there is no information that aic's inability to set up did not compromise delivery of needed therapy and the aic cannot be excluded as a contributing factor in the noted outcome of the patient. | [5 empty] | The investigation for the reported system self-check error issue has been completed. the product was returned, and the issue was confirmed. console logs from the reported day of event are consistent with complaint and show persistent sau_powermod_1 communication issues throughout the logs, and confirm console¿s inability to properly boot-up. console¿s inspection in danvers revealed corrosion on pbm caused by electrolyte leak from supercaps c69, c70. this is a known pattern failure, addressed by fcn 71, however the pbm sn was outside the range identified by fcn. failures of this type will be further monitored for trends. per the results of the clinical review and investigation, the root cause was determined to be contamination. this report is being filed as part of a retrospective review of historical records.
14938 | 1220648-2024-22898 | 20660182 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20660182&pc=OZD | 08/07/2022 | 11/11/2024 |  | 11/11/2024 |  |  | 08/07/2022 | 11/11/2024 |  |  |  |  | 11/11/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1538036 |  | 0042-0010-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Read Input Signal | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the automated impella controller (aic) exhibited a system error optical sensor message. the staff still connected the pump, and they proceeded with the procedure. it was reported that the procedure was successful. additional information states that the patient died during support. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | D9: the date of the devices return for evaluation is unknown. the device was returned for evaluation. the cause of the issue was use related. the pump connection and further handling was not done properly by the operator at the hospital. this report is being filed as part of a retrospective review of historical records.
14939 | 1220648-2024-22818 | 20648840 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20648840&pc=OZD | 07/15/2021 | 11/10/2024 |  | 11/10/2024 |  |  | 07/15/2021 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1162253 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Fell | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | Field service engineer reported that the automated impella controller (aic) housing was damaged. the console arrived for an annual preventive maintenance service and upon opening the shipping box, the console was found to have significant visible damage to the housing. the shipping box did not show any damage. additional information provided by the customer noted the console was dropped. there was no patient involvement. | [5 empty] | The console was returned for preventive maintenance, and upon examining the automated impella controller (aic), it was discovered that a corner of rear housing was broken. upon inquiry, information was provided by the customer that console was dropped. the console was further inspected and no other damage was observed. after housing replacement, the console passed all testing and operated within specification. the console data logs were received. however, the logs did not contain any information relevant to the reported event. the root cause for the console damage was mishandling during removal from aic cart. this report is being filed as part of a retrospective review of historical records.
14940 | 1220648-2024-22819 | 20648866 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20648866&pc=OZD | 10/25/2022 | 11/10/2024 |  | 11/10/2024 |  |  | 10/28/2022 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1330532 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  | Asian | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was placed onto impella cp support due to acute myocardial infarction/cardiogenic shock. while on support, the automated impella controller (aic) experienced a priming error, which was resolved by replacing the console. the patient ultimately expired but there was no allegation against the aic or impella related to the patient¿s death. | [5 empty] | The investigation for the reported automated impella controller - purge disc/flag issues has been completed. the impella device has been returned for evaluation. the logs from the complaint date revealed that the console issued a purge disc not detected alarm several times through the case start wizard. re-insertion of purge cassette/purge disc was observed, yet the console still had problems in detecting the purge disc. the reported priming error was consistent with the disk detection issue during case start. the console was tested. the issue with properly detecting the purge pressure disc was reproduced and the purge flag was confirmed to be broken (missing at blue disc retainer). the cause of the issue was a broken/ missing purge flag. this report is being filed as part of a retrospective review of historical records.
14941 | 1220648-2024-22820 | 20649104 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649104&pc=OZD | 12/08/2021 | 11/10/2024 |  | 11/10/2024 |  |  | 12/08/2021 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1244247 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automatic impella controller (aic) had a controller failure alarm without the controller turning off. another console was used. there was no reported harm or injury to the patient. | [5 empty] | The investigation for the reported aic - controller failure (red alarm) issues has been completed. the device has been returned for evaluation. the cause of the controller failure (red alarm) was defective pud. this report is being filed as part of a retrospective review of historical records.
14942 | 1220648-2024-22821 | 20649137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649137&pc=OZD | 09/09/2022 | 11/10/2024 |  | 11/10/2024 |  |  | 09/09/2022 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1213826 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Application Program Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the automated impella controller (aic) exhibited controller error message followed by a system check failure message and a change console immediately message. the device was replaced with another aic. the procedural outcome is patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: the date of patient death is unknown the device was returned for evaluation. the console was power-cycled, and within 10 boots the system self-check error was reproduced. it was observed through the device logs that the same pmd/cpld communication errors were occurring. the cause of the impella stopped/ controller failure and subsequent system self-check failure are due to a defective impellatronics board. this report is being filed as part of a retrospective review of historical records.
14943 | 1220648-2024-22822 | 20649144 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649144&pc=OZD | 01/17/2023 | 11/10/2024 |  | 11/10/2024 |  |  | 01/17/2023 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1266272 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Application Program Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. it was reported that the pump stopped while utilizing the aic. the device was replaced with another aic and the procedure was successful. the procedural outcome is patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: the date of patient death is unknown the device was returned for evaluation. console logs from the reported day of event are consistent with complaint and show that after that after the pump was connected to the console, it stopped immediately after start, and the console issued alarm ¿impella stopped¿. few restart attempts were made (that also involved disconnecting the pump and then reconnecting) and were all unsuccessful. logs from the backup console confirmed that the same pump was successfully started and the support continued. device analysis the issue was reproduced, and neither pump simulator or an engineering pump was able to start. after temporarily replacing the impellatronic board with a known-good one, the issue was resolved, and console operated within specification. this report is being filed as part of a retrospective review of historical records.
14944 | 1220648-2024-22823 | 20649146 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649146&pc=OZD | 09/10/2021 | 11/10/2024 |  | 11/10/2024 |  |  | 09/10/2021 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1100964 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic | White | Death | No | Yes | Application Program Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [37 empty] | A |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. during support, the patient expired, after the pump was removed the aic would not turn off and was stuck on the generic screen. after approximately 20 minutes, the staff performed a hard stop using the off button under the aic. then the aic appeared to turn on and off as normal. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the device logs were reviewed and showed that initial attempts at restarting the pump were successful, resulting in periods of stable hemodynamic support. however, the alarms cleared previously repeatedly triggered. these alarms would continue to trigger and clear from 2am to 10am, which were then followed by: pmd has detected a pump speed deviation. logs show that console properly detected pump removal, also evident by loss of optical signal algorithms typically seen when a pump is removed. device analysis: aic shutdown issues: the cause of the failure mode was use related. pump pause: console was run with a pump and purge setup in the burn-in chamber for 5 days without any issues. failure mode exhibited by the console was most likely due to a failing/failed pump. as the pump was continuously restarted after the initial failure, pump stops were occurring more frequently as the case continued to progress. the pump stop was not investigated, due to the product not being returned. this report is being filed as part of a retrospective review of historical records.
14945 | 1220648-2024-22824 | 20649170 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649170&pc=OZD | 08/04/2021 | 11/10/2024 |  | 11/10/2024 |  |  | 08/04/2021 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1120449 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) placement signal went blank, and the console shows an alarm indicating to switch to a backup controller. this aic was replaced with another aic resulting in no issues with the replacement aic. there was no patient harm reported. | [5 empty] | The investigation into the pump not detected/ loss of opm data has been completed. the impella automated controller was not received from the customer and therefore, an evaluation of the device by the manufacturer was not possible. the console logs obtained from the cloud, although incomplete, allowed for identification of the failure mode that was consistent with observations and problem description: "transient loss of optical data". this pattern failure is characterized by all signals from optical bench having values of (b)(4). for this particular event signals were missing for approx. 17 seconds, after which the placement signal recovered. the cause of the controller alarm was transient loss of optical data. the cause of optical data loss was unable to be determined because the issue was unable to be reproduced. this report is being filed as part of a retrospective review of historical records.
14946 | 1220648-2024-22825 | 20649171 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649171&pc=OZD | 05/03/2023 | 11/10/2024 |  | 11/10/2024 |  |  | 05/03/2023 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1416992 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Read Input Signal | [6 empty] | Hemodynamic instability | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Us complainant reported that the patient was placed onto impella cp support due to acute myocardial infarction / cardiogenic shock. while on support, the automated impella controller (aic) experienced a purge disc/flag issue that was resolved by changing the aic. the patient ultimately expired as care was withdrawn, but there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and visual inspection of the purge assembly area confirmed a ripped/ missing blue retainer. the failure mode was due to broken/ missing purge retainer. this report is being filed as part of a retrospective review of historical records.
14947 | 1220648-2024-22826 | 20649172 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649172&pc=OZD | 04/05/2023 | 11/10/2024 |  | 11/10/2024 |  |  | 04/05/2023 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1160667 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Failure to Read Input Signal | [6 empty] | Hemodynamic instability | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Us complainant reported that the patient was placed onto impella cp support due heart failure / cardiogenic shock. while on support, the automated impella controller (aic) experienced purge disc/flag issues, which were resolved by swapping to another aic. the patient ultimately expired, but there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and visual inspection of the purge assembly area confirmed a broken blue retainer. the failure mode was due to broken purge retainer. this report is being filed as part of a retrospective review of historical records.
14948 | 1220648-2024-22827 | 20649184 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649184&pc=OZD | 11/18/2021 | 11/10/2024 |  | 11/10/2024 |  |  | 11/18/2021 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1129885 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Key or Button Unresponsive/not Working | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was placed on impella 5.5 for hemodynamic support. while on support, the aic experienced kbd/rotary knob issues. additional information was provided that noted the patient expired. | [5 empty] | The device was returned for evaluation. the purge flag was confirmed to be broken, the cause of the broken purge flag is unable to be determined. this report is being filed as part of a retrospective review of historical records.
14949 | 1220648-2024-22828 | 20649185 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649185&pc=OZD | 11/10/2021 | 11/10/2024 |  | 11/10/2024 |  |  | 11/10/2021 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2022125917 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was needing an impella device for mechanical circulatory support. the complainant reported that the console was turned on and did not display waveforms. the staff switched to a backup console and the same pump was used with no issues. additional information was provided that noted the patient expired. | [5 empty] | The investigation for the reported optical signal issue has been completed. the product was returned for evaluation. this known pattern failure, which is characterized by a loss of optical data (placement, snr, contrast, lamp, gain) and triggering of a controller error alarm, has a low probability of reoccurrence. if needed, patient hemodynamics and impella position can be monitored with imaging. the root cause of the issue could not be determined since the issue could not be reproduced. this report is being filed as part of a retrospective review of historical records.
14950 | 1220648-2024-22829 | 20649186 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649186&pc=OZD | 07/28/2021 | 11/10/2024 |  | 11/10/2024 |  |  | 07/28/2021 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER JP |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER JP |  |  | ABIOMED, INC. |  |  | US |  | 1379919 |  | 0042-0000-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The quality professional reported that the automated impella controller (aic) displayed an error message after the pump cable was connected to the controller during support preparation. it was reported the message "please connect the connection cable" was displayed on the controller screen. reportedly, after replacing the controls, the same pump continued to be used and was able to be used without problems thereafter. there was no patient harm reported. | [5 empty] | E1: name of the initial reporter is unknown. the investigation into the case start issue has been completed. the impella automated controller was not received from the customer. the data analysis of the logs were partially consistent with complaint and show that pump detection was delayed by 1m27s, but the pump was eventually recognized and used on current console ic6610. during the delay, the epm circuitry was power cycled after pump was not detected for 20 seconds after the case start wizard step 3 prompt. this behavior is a mitigation mechanism in response to a rare pattern failure "epm crystal issue". when the console was tested by the field service engineer, the issue was not reproduced, and console was able to detect pump connection every time. the cause of the delayed pump detection was an issue with the crystal of the epm on the impellatronic board. this report is being filed as part of a retrospective review of historical records.
14951 | 1220648-2024-22830 | 20649204 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649204&pc=OZD | 07/01/2022 | 11/10/2024 |  | 11/10/2024 |  |  | 07/01/2022 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1614917 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Failure to Read Input Signal | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was on an automated impella controller (aic) for mechanical circulatory support. it was noted that the device exhibited a purge disc not detected alarm, the team changed the cassettes and replaced the console. the procural outcome is patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | This device was not returned for evaluation. review of the device logs found the purge disc not detected alarm clears and is followed by several log lines of purge pressure disc detected and no purge pressure disc detected. the cause of the issue was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records.
14952 | 1220648-2024-22831 | 20649206 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649206&pc=OZD | 06/27/2022 | 11/10/2024 |  | 11/10/2024 |  |  | 06/27/2022 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 116828 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Failure to Read Input Signal | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. it was reported that the console alarmed purge pressure high and purse blocked alarms. the staff tried to troubleshoot the issue by tracing the tubing for kinks, de airing, and also changing out the purge cassette. after starting to go through the prompts to change the purge cassette, the console did not recognize the new purge cassette even though the tubing was being bolused. the device was replaced with another aic and the procedure was successful. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The device was returned for evaluation. the device logs were reviewed and showed that during patient support, purge pressure high and purge system block were triggered sequentially as purge pressure continuously building up. the log confirmed alarms due to abnormally high purge pressure value before the purge cassette was exchanged with no expected gradual decrease in pressure. it was observed that the staff attempted to change the purge cassette and bag and perform purge cassette de-airing without success. console issued piston blocked with 87% range remaining (unlikely to affect de-airing), but purge fluid could be not detected, but there were no issues detecting the purge cassette. device analysis: visual inspection of the purge assembly area confirmed a broken stand on purge blue retainer, causing the purge flag to fall out of retainer. the root cause high purge pressure and purge system blocked alarms was not determined. the high purge pressure prompted users to replace for new purge cassette, during which the broken purge retainer and purge flag issue appeared and prevented user from completing procedure. this report is being filed as part of a retrospective review of historical records.
14953 | 1220648-2024-22832 | 20649208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649208&pc=OZD | 06/23/2022 | 11/10/2024 |  | 11/10/2024 |  |  | 06/23/2022 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2023210391 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  |  | Malfunction | Yes | Yes | Application Program Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the automated impella controller (aic) exhibited controller error / switch to back up. the device was replaced with another aic and the procedure was successful. additional information noted the patient became asystolic and expired on support. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The device was returned for evaluation. a review of the device logs found loss of optical data. device analysis: the console was run on pump and purge for a brief and long-term test to reproduce the issue exhibited in the field. the issue was not able to be reproduced. this report is being filed as part of a retrospective review of historical records.
14954 | 1220648-2024-22833 | 20649209 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649209&pc=OZD | 06/12/2022 | 11/10/2024 |  | 11/10/2024 |  |  | 06/12/2022 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1313336 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Failure to Read Input Signal | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the automated impella controller (aic) exhibited pp low alarm. the device was replaced with another aic and the procedure was successful. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The device was returned for evaluation. console purge assembly was inspected, and it was observed that the blue purge disc retainer had snapped off the purge assembly. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
14955 | 1220648-2024-22834 | 20649211 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649211&pc=OZD | 06/07/2022 | 11/10/2024 |  | 11/10/2024 |  |  | 06/07/2022 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1373062 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Decrease in Pressure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the automated impella controller (aic) exhibited purge pressure low alarm. the device was replaced with another aic and the procedure was successful. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: the date of patient death is unknown. the device was returned for evaluation. review of the device logs revealed the reported purge pressure low alarm, and purge system open alarm after purge cassette and disk was inserted. device analysis: visual inspection on purge system did not reveal broken purge flag/ blue retainer. however, it was found that the purge disk housing was sticky (potentially glucose residue). during testing, the pc and purge disc were fully inserted, and the intermittent disk detection symptom was reproduced. the purge disc detection was based on optical sensor built in purge pressure sensor. therefore, with glucose residue blocking/ interfering the light-based purge disc detection. the cause of this issue is contamination. this report is being filed as part of a retrospective review of historical records.
14956 | 1220648-2024-22835 | 20649213 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649213&pc=OZD | 06/04/2022 | 11/10/2024 |  | 11/10/2024 |  |  | 06/04/2022 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1535397 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Fracture | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the automated impella controller (aic) purge disc broke off. it was reported that the nurse tried to advance the purge disc into the aic purge disc slot multiple times but still was getting error messages that the disc was not detect. on the final attempt the purge disc slot on the aic broke off. the device was replaced with another aic. additional information states the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The device returned for evaluation. console purge assembly was inspected. upon opening the purge assembly door, it was confirmed the purge disc retainer had broken/snapped off of the purge assembly. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
14957 | 1220648-2024-22836 | 20649214 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649214&pc=OZD | 05/29/2023 | 11/10/2024 |  | 11/10/2024 |  |  | 05/29/2023 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1458318 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the automated impella controller (aic) experienced purge disc/flag issues that were resolved by changing the controller. the patient ultimately expired as care was withdrawn. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the product was returned and the aic - purge disc/flag issues was confirmed. device analysis: when the console was returned for investigation, the purge motor shaft confirmed to be blocked. upon inspection, the gasket had significant portions of dry glucose that, when removed, allowed the purge system to function as expected. the symptoms described in the complaint description could all be caused by a glucose leak into the purge system. the cause of the aic issue was glucose build up in the purge assembly. this report is being filed as part of a retrospective review of historical records.
14958 | 1220648-2024-22837 | 20649216 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649216&pc=OZD | 01/27/2023 | 11/10/2024 |  | 11/10/2024 |  |  | 01/27/2023 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1221065 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high risk surgery was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the automated impella controller (aic) experienced a rotary knob issue with no resolution specified. the patient ultimately expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported aic - kbd/rotary knob issues has been completed. the product was returned and the aic - kbd/rotary knob issues was confirmed. the rotary knob when tested did not bounce back and needed to be pulled back manually confirming the failure mode. replacing the encoder eliminated the failure. the cause of the issue was a defective encoder. this report is being filed as part of a retrospective review of historical records.
14959 | 1220648-2024-22838 | 20649219 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649219&pc=OZD | 01/26/2023 | 11/10/2024 |  | 11/10/2024 |  |  | 01/26/2023 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1673731 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  |  | Asian | Death | Yes | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the automated impella controller (aic) experienced a purge issue that was resolved by replacing the controller. the patient ultimately expired. there is not enough information to exclude the device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported aic - purge issue - other has been completed. the product was returned and the aic - purge issue - other was confirmed. console was returned for investigation. purge system was disassembled with no obvious flaws found. console was able to deliver the necessary torque and there was no notable glucose build up. the console was able to function as expected. the most likely cause of failure was the purge cassette being incompletely pushed in which would trigger the rfid repeatedly and prevent the piston from having its full range. this report is being filed as part of a retrospective review of historical records.
14960 | 1220648-2024-22839 | 20649220 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649220&pc=OZD | 12/14/2022 | 11/10/2024 |  | 11/10/2024 |  |  | 12/14/2022 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1545370 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Read Input Signal | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, it was noted that there was no audible alarm on the automated impella controller (aic). also, lv waveform was not adjustable. additional information noted that the patient expired on support. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported aic - buzzer/ alarm inaudible has been completed. the product was returned and the aic - buzzer/ alarm inaudible was confirmed. console logs from the reported day of event were analyzed and showed multiple alarms presenting, that would be associated with audio signals, however the data does not allow us to infer whether the alarms were audible. console data (imc log and ltlog) also showed that the lv adjustment reminders were turned on, and the operator successfully performed signal adjustment twice. inspection of the console revealed that the speaker cable was not plugged into the 4-in-1 carrier, explaining lack of audio alarms. there was no issue found regarding lv signal adjustment. the cause of the aic issue (inaudible alarms) was a loose connector. this report is being filed as part of a retrospective review of historical records.
14961 | 1220648-2024-22840 | 20649221 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649221&pc=PYX | 06/08/2021 | 11/10/2024 |  | 11/10/2024 |  |  | 06/08/2021 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1181833 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Female |  |  | White | Death | Yes | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with post-op left ventricular assist device (lvad) was implanted with an impella rp device for mechanical circulatory support. while on support, the automated impella controller (aic) experienced a controller failure with no resolution specified. the patient ultimately expired. there is not enough information to exclude the device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported aic - controller failure (red alarm) has been completed. the product was returned and the aic - controller failure (red alarm) was not confirmed. during this first bootup (after receiving console from fs), imc logs show there is no heartbeat from audioserver. then the watchdog triggers a partial reset. it does that 3 times, and eventually it triggers a full reset. after every reset, we see a core dump for audioserver. after the full reset, there are no more resets triggered by the watchdog and the console boots up in english (instead of german). the bootup after full reset, there are log entries suggesting read error when reading ucfg.ini file ((b)(6) 2021 21:59:37). this file contains the language configuration. if the ucfg.ini file is not read, console will bootup in the default language (english). the case was started after the first partial reset. the pump paused on every reset (partial and full), and then it resumed after console bootup. partial resets result in a pump pause and the screen goes blank for about 5 seconds. the complaint mentioned ac power was disconnected before the console restarted, and the logs confirm that ac power was disconnected at (b)(6) 2021 21:58:10, and the next partial reset (reset #2) happened at 21:58:15. but this is completely unrelated, as the previous partial reset (reset #1) happened while ac was connected. issue was not reproduced in fs. console was power cycled 5 times with ac plugged in, and 5 times without. upon every bootup, console was kept powered on for about 5 minutes. the cause of the resets was unable to be determined. this report is being filed as part of a retrospective review of historical records.
14962 | 1220648-2024-22841 | 20649223 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649223&pc=OZD | 02/01/2022 | 11/10/2024 |  | 11/10/2024 |  |  | 02/01/2022 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1422863 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  |  | Asian | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was noted at the time of priming, an alarm for an optical sensor system error occurred on the automated impella controller (aic). the connection cable was reconnected and watched for a few minutes with no improvement. the device was used in a situation where the ps waveform and suction were not functioning. the patient was in poor condition and died after withdrawal. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported optical signal issue has been completed. the product was returned and the optical signal issue was not confirmed. data analysis: aic logs from the complaint date revealed that the console triggered the optical sensor system error window, and the user was able to bypass these windows to start the case with a placement signal not reliable alarm (opm invalid signal ¿ event set #1036) alarm. the case was continued with an invalid placement signal. device analysis: the console was returned for further investigation. before replacing the optical cable, the contrast was less than 10% and bulb power was 0 ¿w. the optical cable was replaced with kit (just in case), and the parameters were within specifications. the cause of the reported aic issue was a defective blue receptacle. this report is being filed as part of a retrospective review of historical records.
14963 | 1220648-2024-22842 | 20649229 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649229&pc=OZD | 05/27/2023 | 11/10/2024 |  | 11/10/2024 |  |  | 05/27/2023 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1117649 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Hispanic |  | Death | Yes | Yes | No Pressure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Us complainant reported that the patient was placed onto impella 5.5 due to cardiomyopathy. while on support, the automated impella controller (aic) experienced purge disc/flag issues that were resolved by changing the console. the patient ultimately expired as care was withdrawn, but there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported issue has been completed. the console was tested after arriving in danvers. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. the cause of the issue was broken/ missing purge flag. this report is being filed as part of a retrospective review of historical records.
14964 | 1220648-2024-22843 | 20649230 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649230&pc=OZD | 12/03/2022 | 11/10/2024 |  | 11/10/2024 |  |  | 12/03/2022 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1129486 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  |  | White | Malfunction | Yes | No | Defective Device | [6 empty] | Hemodynamic instability | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Us complainant reported that the patient was placed onto impella cp support for an unspecified indication. damage to the automated impella controller (aic) console was discovered prior to therapy, which was resolved by switching to another controller. the patient ultimately expired as care was withdrawn, but there was no allegation against the aic or impella related to the patient¿s death. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported issue has been completed. the device was returned for investigation. device analysis confirmed on visual inspection the broken purge retainer. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
14965 | 1220648-2024-22844 | 20649231 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649231&pc=OZD | 12/05/2021 | 11/10/2024 |  | 11/10/2024 |  |  | 12/05/2021 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1324937 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | No Pressure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock needing an impella cp device for mechanical circulatory support. there was a purge failure which prompted the team to replace the aic. the ifu was followed and a backup console utilized for support. additional information was provided that noted the patient expired. | [5 empty] | The investigation for the reported aic - purge disc/flag issue has been completed. the product was returned, and the issue was confirmed. the cause fluid ingress in the purge assembly most likely played a role, but there was no visible buildup of glucose on the purge pressure optical sensor. this report is being filed as part of a retrospective review of historical records.
14966 | 1220648-2024-22845 | 20649234 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649234&pc=OZD | 11/24/2021 | 11/10/2024 |  | 11/10/2024 |  |  | 11/24/2021 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1162662 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Battery Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the automated impella controller (aic) exhibited a battery failure and battery low message. the device was replaced with another aic. the procedural outcome noted that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | This device was returned for evaluation. the cause of the battery failure was due to a defective battery (rapid decline in cell voltage). the exact cause of the defective battery was unable to be determined. this report is being filed as part of a retrospective review of historical records.
14967 | 1220648-2024-22846 | 20649237 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649237&pc=OZD | 11/15/2022 | 11/10/2024 |  | 11/10/2024 |  |  | 11/15/2022 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1226869 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  |  |  | Death | Yes | Yes | Failure to Read Input Signal | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. during preparation the aic could not detect the purge system. the aic also showed a crack in the blue purge disc retainer. the device was replaced with another aic and the procedure was successful. it was noted that the patient expired at explant. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The device was returned for evaluation. a review of the device logs found that the console began to repeatedly trigger the purge disc not detected alarm. a physical inspection of the purge assembly was performed. it was observed in the purge disc retainer that the purge disc flag was missing/broken. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
14968 | 1220648-2024-22847 | 20649238 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649238&pc=OZD | 09/01/2021 | 11/10/2024 |  | 11/10/2024 |  |  | 09/01/2021 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1321788 |  | 0042-0000-JP |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the automated impella controller (aic) exhibited a controller yellow alarm. when the alarm occurs, there is no problem with the operation of the pump, and the alarm disappears naturally. the same alarm appeared the next day and disappeared spontaneously. at that time, the pump position monitoring stop alarm is also activated. no particular action was taken, and the patient's hemodynamics were not affected. | [5 empty] | D9: the date of the devices return for evaluation is unknown. this device was returned for investigation. the device logs confirmed the reported failure mode given there were two controller error alarms (opm comm crc error ¿ event set # (b)(4)) triggered during the clinical procedure. the console was evaluated: this known pattern failure, which is characterized by a temporary loss of optical data (placement, snr, contrast, lamp, gain) and triggering of a controller error alarm, has a low probability of reoccurrence. if needed, patient hemodynamics and impella position can be monitored with imaging. the cause of the transient loss of optical data was unable to be determined because the issue was unable to be reproduced. this report is being filed as part of a retrospective review of historical records.
14969 | 1220648-2024-22848 | 20649240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649240&pc=OZD | 09/02/2022 | 11/10/2024 |  | 11/10/2024 |  |  | 09/02/2022 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2023222111 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  |  |  | Malfunction | Yes | Yes | No Pressure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support.the complainant reported a loss in purge pressure from the automated impella controller (aic). although no leaks were identified, the purge disk, holder, and accompanying springs fell out of the cassette compartment when the purge cassette door was opened for inspection. the aic was removed and replaced with a new console. it was noted that the family withdrew care, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation into the reported issue has been completed. the device was received for evaluation. logs were reviewed and deemed not relevant for this failure mode. console purge assembly was inspected. upon opening the purge assembly door, it was confirmed the purge disc retainer had broken/snapped off of the purge assembly. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
14970 | 1220648-2024-22849 | 20649242 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649242&pc=OZD | 03/06/2023 | 11/10/2024 |  | 11/10/2024 |  |  | 03/06/2023 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2023265731 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. during support the controller shutdown. it was reported that the customer had another aic to utilize, and the procedure was successful with no reported harm to the patient. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation into the reported issue has been completed. the device was returned for evaluation. the cause of the aic issue was likely a defective kbd/display making the user force a shutdown. this report is being filed as part of a retrospective review of historical records.
14971 | 1220648-2024-22850 | 20649243 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649243&pc=OZD | 02/13/2023 | 04/03/2025 |  | 11/10/2024 |  |  | 02/13/2023 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1663431 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Application Program Problem | [6 empty] | Hemodynamic instability | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported that the patient was placed onto impella 5.5 support due to cardiomyopathy. while on support, the aic experienced a controller error yellow alarm with no resolution specified. the patient ultimately expired, there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | This report is being filed as part of a retrospective review of historical records. h1 type of reportable event was incorrectly reported as malfunction in the initial mfr 1220648-2024-22850 and has been corrected to death. | The investigation into the reported issue has been completed. the device was received for evaluation. logs confirmed that the controller error coincided with multiple subsequent placement signal not reliable errors. following the set of errors, the optical bench was forced to reset, which resolved the issue. real time (rt) logs confirmed that immediately prior to the controller error, the optical bench reported an oscillating contrast value far outside the normal range of contrast. this oscillating contrast issue is suspected to be result of a hardware fault on the optical bench. console ic9750 was sent back to fs for further investigation. the issue was not able to be reproduced in the lab. the optical signal issue was most likely due a defective optical bench. the console is designed to automatically reset the optical bench when this issue emerges, which resolves the issue. the optical signal issue was most likely due a defective optical bench. this report is being filed as part of a retrospective review of historical records.
14972 | 1220648-2024-22851 | 20649264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649264&pc=OZD | 07/04/2021 | 11/10/2024 |  | 11/10/2024 |  |  | 07/04/2021 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1139054 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  |  | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) experienced battery alarm. the console was replaced to address the issue. no patient harm reported. | [5 empty] | The device was returned for evaluation. console arrived in danvers and an interior and exterior inspection was performed, all normal/clear. console was placed on pump simulator for several days, with no reproduction of the events that occurred in the field and led to complaint cause. console behavior was normal as expected throughout testing. this report is being filed as part of a retrospective review of historical records. complaint cause is due to erratic loss of communication from the opm. this is a known failure mode as all instances of this failure mode that have occurred have ranged in length of time from 2 mins to 10 mins of communication loss. communication in most cases is reestablished and in this instance was as well. evidence of the known failure mode is also clear in the value presented from the raw optical signal value when the signal bottoms out and reads (b)(4), on both instances of the controller error. at this moment in time, it is undetermined what the root cause of this failure mode is. this report is being filed as part of a retrospective review of historical records.
14973 | 1220648-2024-22852 | 20649265 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649265&pc=OZD | 06/24/2021 | 09/30/2025 |  | 11/10/2024 |  |  | 06/24/2021 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1563183 |  | 0042-0010-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported the automatic impella controller would not start although the battery was completely charged, and the console was connected to the alternating current power outlet. several attempts were made to start the aic without success. the battery switch was checked, confirmed in the correct position, and the power button was pressed for at least 20 seconds with no reaction. the aic was exchanged, and no harm to a patient was reported or indicated as result of this product malfunction. | [5 empty] | The device was returned for evaluation. the issue was reproduced, and it was observed that the console was receiving power, as the amber led was lit. replacing the power button did not resolve the issue, but when pbm assembly was replaced, the console could be successfully powered on. this report is being filed as part of a retrospective review of historical records. | D2 common device name revised on manufacturer device report 1220648-2024-22852 in accordance with updated procedures. d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2024-22852. e4 should have been left blank on the initial manufacturer device report number 1220648-2024-22852 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact email revised on manufacturer device report 1220648-2024-22852 in accordance with updated procedures.
14974 | 1220648-2024-22853 | 20649266 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649266&pc=OZD | 06/16/2021 | 11/10/2024 |  | 11/10/2024 |  |  | 06/16/2021 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1098590 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The console failed self-test, another device was used, and there was no reported patient involvement, harm, or injury. | [5 empty] | E1, e3, e4: the name and occupation of the initial reporter is unknown. the investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the system self-check failure seen in the field was caused by corrosion from the leaking of the super caps. this report is being filed as part of a retrospective review of historical records.
14975 | 1220648-2024-22854 | 20649268 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649268&pc=OZD | 06/18/2021 | 11/10/2024 |  | 11/10/2024 |  |  | 06/18/2021 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1300701 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant had an automatic impella controller (aic) that failed to power down correctly. no report of patient injury or harm. | [5 empty] | The investigation for the reported aic - a/c power issues has been completed. the device has been returned for evaluation. however, the cause of the failure to shutdown/display alarms was unable to be determined because the issue was unable to be reproduced. this report is being filed as part of a retrospective review of historical records.
14976 | 1220648-2024-22855 | 20649269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649269&pc=OZD | 06/15/2021 | 11/10/2024 |  | 11/10/2024 |  |  | 06/15/2021 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1183257 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced an alarm suction and decreasing the pressure level did not resolve. the complainant replaced this aic and the issue resolved. there was no patient harm reported. | [5 empty] | The investigation into the placement signal issue has been completed. the data analysis of the logs from the reported day of event confirmed the suction alarms and also indicated there were numerous placement signal issues and alarms. the real time logs showed that the suction alarm presented right when the pump was started at pressure level p-9. reducing the pressure level to p-1 removed the suction alarm. then the pressure level was raised to p-3 with no suction alarms. then it was raised to p-5 and the suction alarm came back. the log analysis suggests that the suction alarm was placement/patient related and not a product malfunction. the five signal values were found to be within specification. a visual inspection revealed no significant contamination found on the pump connector and no issues were found with the console. the placement signal issue was use related due to the end user¿s placement of the device and/or the patient condition. this report is being filed as part of a retrospective review of historical records.
14977 | 1220648-2024-22856 | 20649272 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649272&pc=OZD | 06/14/2021 | 11/10/2024 |  | 11/10/2024 |  |  | 06/14/2021 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1129228 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) experienced a "optical sensor system error" alarm after connecting the pump to the aic. the console was replaced to address the issue. no patient harm reported. | [5 empty] | The device was returned for evaluation. console logs on the day of reported event were consistent with complaint and showed two failed attempts to start a case using impella with smartassist, both times receiving notification of "optical sensor system error". the device logs show that optical signal was invalid, with very low values of snr and contrast, indicating optical losses between pump sensor and optical bench detector (ccd). inspection of the optical fiber in aic pump plug receptacle revealed some contamination, but also some mechanical damage to the ferrule surrounding the fiber. when an attempt was made to reproduce the issue, it's been observed that test pump plug could not be inserted securely all the way in. further inspection of the blue receptacle revealed a piece of pump catheter connector that broke off and was lodged inside the housing of the receptacle (see photos attached), preventing the connector from being fully engaged. the resulting connection was such that the metal pins were engaged creating electrical contact sufficient for pump detection, but the optical fibers were not close enough for the sensor to operate properly. the cause of such heavy mechanical damage was not determined, however was most likely use-related. this report is being filed as part of a retrospective review of historical records.
14978 | 1220648-2024-22857 | 20649273 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649273&pc=OZD | 06/08/2021 | 11/10/2024 |  | 11/10/2024 |  |  | 06/08/2021 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1086076 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) console shut down and alarmed placement signal unreliable. the staff reported the controller failure alarms were appearing when plugged in. the clinician also observed the same alarm upon inspection and turned it off and on and plugged it into a different outlet as well and the alarm appeared and would come and go. the staff also reported placement signal unreliable alarms on previous cases from this console. they were unable to recall what asset numbers were associated with the alarms. there was no patient harm reported. | [5 empty] | E1: name of the initial reporter is unknown the investigation into the controller failure shut down/placement signal issues have been completed. the impella automated controller was returned for investigation. the data logs confirmed the controller failure on the reported complaint date. specifically, the voltage from the purge pressure sensor was out-of-specification and triggered a controller failure. the placement signal issues could not be confirmed in the logs. there were no instances of placement signal unreliable alarms, relating to when the optical placement signal would go invalid. the returned console was tested thoroughly on an engineering bench, but either of the failure modes could not be reproduced. during power cycling the console more than 50 times, no controller failure alarms were triggered, and the logs contained no entries associated with a defective purge pressure sensor. the optical bench was tested. simulated placement signal of 110 mmhg was normal. optical power and 5s signal parameters were within specification. as a final check, the console ran an optical pump with purge for 15 mins. no issues related to the failure mode were reproduced. the controller failure (purge pressure sensor defective) was most likely due to a defective purge pressure sensor pca. the placement signal issues could not be confirmed nor reproduced during testing this report is being filed as part of a retrospective review of historical records.
14979 | 1220648-2024-22858 | 20649274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649274&pc=OZD | 04/17/2021 | 11/10/2024 |  | 11/10/2024 |  |  | 04/17/2021 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1152333 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the automated impella controller (aic) produced a "connect black arrows" screen. a new aic was used for the procedure, and there was no reported patient harm. | [5 empty] | The device was returned and evaluated, and the investigation for the reported "pump did not start" has been completed. when attempting to plug pump into console, console fails to recognize pump due to known epm crystal issue. console successfully power cycles epm, however is still unable to detect/ recognize the cp pump. the console then prompts the user with the old catheter platform. complaint cause was due to known epm crystal timing issue. this report is being filed as part of a retrospective review of historical records.
14980 | 1220648-2024-22859 | 20649275 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649275&pc=OZD | 01/20/2022 | 11/10/2024 |  | 11/10/2024 |  |  | 01/20/2022 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1606424 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Alarm Not Visible | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The user facility reported the patient was implanted with an impella 5.5 device for mechanical circulatory support and while on support, there was an inaudible alarm with automated impella controller (aic). a loaner aic was arranged for return of the complaint aic for evaluation and analysis. no further details were provided; there was no report or indication of any harm to the patient. | [5 empty] | The automated impella controller (aic) data logs and console were returned for evaluation and analysis. data analysis data analysis of the aic logs from the reported complaint occurrence date showed there were a few placements signal-related alarms that occasionally triggered. there were no obvious errors in the logs that would indicate an alarm malfunction. however, the analysis does show the user interacted with the interface to turn the suction audio off and then back on, which could indicate the user was attempting to troubleshoot an inaudible alarm. the console was tested thoroughly on an engineering lab bench. when a suction alarm was triggered, the aic alarmed appropriately. there were no issues observed over a long period of time with the ability to produce an audible tone. the internal cable associated with the speaker was properly connected/ aligned. product history review was completed, and no action was required as a result of the review. the alarm triggering and clearing behavior did not match the known software jama defect which causes issues with alarm clearing. regardless, this was revised in sw v10.2, and the console was upgraded with this software before returning to service. the root cause of the inaudible aic alarm was unable to be determined due to the inability to reproduce. this report is being filed as part of a retrospective review of historical records.
14981 | 1220648-2024-22860 | 20649278 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20649278&pc=OZD | 02/12/2022 | 11/10/2024 |  | 11/10/2024 |  |  | 02/12/2022 | 11/10/2024 |  |  |  |  | 11/10/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1111984 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  |  | White | Death | Yes | No | Battery Problem | [6 empty] | Hemodynamic instability | [12 empty] | Death | [24 empty] | R |  |  | No | Initial submission |  |  |  |  | 02/12/2022 |  |  |  |  | The us complainant had a battery failure which prompted the team to replace the automated impella controller (aic) with the backup console. patient care was withdrawn and there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and console interior connections were checked and verified. when visually inspecting the batteries and pbm, it was observed that one of the batteries status leds were completely blank. the battery failure alarm was due to a defective battery 1 (decline of individual cell voltage). the root cause of the battery cell failure was undetermined.
14982 | 1220648-2024-22757 | 20646747 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20646747&pc=OZD | 11/30/2021 | 11/08/2024 |  | 11/09/2024 |  |  | 11/30/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1084840 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) had a frozen screen, which was not resolved by restarting the console. there was no patient involvement. | [5 empty] | The device was returned and evaluated, and the investigation for the reported display issue has been completed. console logs do not have any entries on the reported day of event, but later there was a controller error due to kbd communication loss, which might have resulted in kbd being unresponsive even after a power-cycle (although the alarm did not re-occur). visual inspection of the console performed in danvers revealed some irregularities, most notably: missing tamper evident seal from the top of console. degraded/discolored ribbon cable between kbd and 4-in-1, which would be responsible for glass kbd to become occasionally unresponsive. the ribbon cable between impellatronic and 4-in-1 was found to be severely kinked and likely to affect impellatronic operation in the future. this report is being filed as part of a retrospective review of historical records.
14983 | 1220648-2024-22758 | 20646763 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20646763&pc=OZD | 11/10/2021 | 11/08/2024 |  | 11/09/2024 |  |  | 11/10/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1426097 |  | 0042-0000-JP |  | Yes |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 47 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. during use, the automated impella controller (aic) was found to have noise on the screen. the procedure was still completed with this device, and there was no reported patient harm. | [5 empty] | D8/9: product return date is unknown. the device was returned and evaluated, and the investigation for the reported display issue has been completed. device analysis: replacing the module resolved the issue. however, the backstrap cable was also replace as a precaution. the cause of the noisy video was due to a defective impella connect module. the cause of the defective impella connect module was undetermined. this report is being filed as part of a retrospective review of historical records.
14984 | 1220648-2024-22759 | 20646863 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20646863&pc=OZD | 04/17/2022 | 11/08/2024 |  | 11/09/2024 |  |  | 04/17/2022 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1320441 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported automated impella controller console failed to recognize a purge disc when transferring pump to a new console. a back console was used without any delay in purge solution being delivered to the device. the patient ultimately expired on support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the product was returned and the aic - purge disc/flag issues was confirmed. visual inspection of the blue purge disk retainer was not fixed in place. the retainer was not attached on one side to purge unit ( as it was supposed to be) due to broken mounting post. therefore, the purge disc detection would not be well-functioning. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
14985 | 1220648-2024-22762 | 20646907 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20646907&pc=OZD | 01/06/2022 | 11/09/2024 |  | 11/09/2024 |  |  | 01/06/2022 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1233233 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Application Program Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock needing an impella device for mechanical circulatory support. it was reported that during startup, the aic (automated impella controller) displayed a serious "controller error" alarm. the aic was switched and the procedure completed. additional information was provided that noted the patient expired. | [5 empty] | The investigation for the reported aic - controller error (yellow alarm) issue has been completed. the product was returned; however, the root cause of the issue could not be determined since the issue could not be reproduced. this report is being filed as part of a retrospective review of historical records.
14986 | 1220648-2024-22763 | 20646920 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20646920&pc=OZD | 12/19/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 12/19/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1181074 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Application Program Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock needing an impella cp device for mechanical circulatory support. it was reported that the rn was doing a bag change and during the procedure the aic (automated impella controller) screen froze and was unable to go to next step. red alarm flashing complete procedure. the aic was switched to a new controller and support continued. additional information was provided that noted the patient expired. | [5 empty] | The investigation for the reported issue has been completed. the product was returned, and the issue was confirmed. the root cause of the unresponsive keyboard (frozen) was most likely fluid spilled on the softkeys during the purge bag change. this report is being filed as part of a retrospective review of historical records.
14987 | 1220648-2024-22761 | 20646948 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20646948&pc=OZD | 01/05/2022 | 11/09/2024 |  | 11/09/2024 |  |  | 01/05/2022 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1392197 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Display or Visual Feedback Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock needing an impella device for mechanical circulatory support. the staff reported that the console revealed a abnormal placement signal and lv waveform ¿ at the top of the box, causing difficulties in monitoring the placement of the impella. the aic console was switched and the issue was solved. additional information was provided that noted the patient expired. | [5 empty] | The investigation for the reported aic - display issue has been completed. the product was returned, and the issue was confirmed. the root cause of the aic issue was use related. this report is being filed as part of a retrospective review of historical records.
14988 | 1220648-2024-22760 | 20647023 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20647023&pc=OZD | 03/15/2022 | 11/08/2024 |  | 11/09/2024 |  |  | 03/15/2022 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1169581 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  |  | Malfunction | Yes | No | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The us complainant had an aic (automated impella controller) with a purge disc failure. the team replaced the aic. reportable due to purge disc failure and the potential relationship between the impella and cause of death. | [5 empty] | He issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. data review: imc logs from the complaint date revealed that the console issued the purge disc not detected alarm (event set #(b)(4)) several times. device analysis: the console was tested after arriving in fs. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken (missing at blue disc retainer - see attached image). before returning this console back to the customer, fs was instructed to replace the purge flag, retainer (hairline crack observed ¿ unrelated to failure mode), and perform a full functional check on the console and optical system (b)(4).
14989 | 1220648-2024-22766 | 20647380 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20647380&pc=OZD | 10/24/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 10/24/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1542953 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  |  | Death | Yes | Yes | Application Program Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Us complainant reported that the patient was placed onto impella 5.5 support due to cardiomyopathy. while on support, the automated impella controller (aic) experienced a controller error yellow alarm that self-resolved. it was further reported that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the imc logs confirmed the reported failure mode given there were two controller error alarms triggered during the clinical procedure. this report is being filed as part of a retrospective review of historical records.
14990 | 1220648-2024-22767 | 20647383 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20647383&pc=OZD | 10/20/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 10/20/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1313339 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Complainant reported that after a week of support, the automated impella controller (aic) console was unable to reflect placement signal. per rn, console was swapped out. support noted that there were no placement issues with the new console. it was further reported that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported issue has been completed. data logs confirmed the invalid placement signal and controller error occurred on the complaint date. console was tested thoroughly by the post market engineering team, and the failure was not reproduced. this known pattern failure, which is characterized by a temporary loss of optical data and triggering of a controller error alarm, has a low probability of reoccurrence. this report is being filed as part of a retrospective review of historical records.
14991 | 1220648-2024-22765 | 20647388 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20647388&pc=OZD | 09/16/2022 | 11/09/2024 |  | 11/09/2024 |  |  | 09/16/2022 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1423376 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Read Input Signal | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that during preparation for a case the automated impella controller (aic) exhibited an optical sensor alarm. the staff decided to proceed with the case. however, the plug was not inserted all the way and was unable to be pushed in further. the device was replaced with another aic and the procedure was successful. no reported harm to the patient due to this malfunction. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The device was returned for evaluation. device logs confirmed a placement signal not reliable alarm (event set #(b)(4)) due to an invalid placement signal. device analysis: failure mode was reproduced. known-good test pump could not be fully connected. there was a small white plastic piece in the pump connector which caused the difficulty plugging in the pump connector, creating an air gap. this air gap led to an invalid placement signal with very low snr. once it was removed the console performed as expected. this report is being filed as part of a retrospective review of historical records.
14992 | 1220648-2024-22768 | 20647469 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20647469&pc=OZD | 10/18/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 10/18/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1140261 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Death | Yes | Yes | Application Program Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Us complainant reported patient was placed on impella cp for hemodynamic support. while on support, the aic experienced a controller error (yellow alarm). issue was resolved by changing out console. it was further reported that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported issue has been completed. complaint cause is a known pattern failure characterized by temporary loss of placement signal. no repair will be necessary as the issue has a low probability of reoccurrence, lasts only up to 10 minutes. if needed, patient hemodynamics and impella position can be monitored with imaging the console will not need to be returned from the field. the transient loss of placement signal could not be reproduced. this report is being filed as part of a retrospective review of historical records.
14993 | 1220648-2024-22764 | 20647476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20647476&pc=OZD | 11/08/2022 | 11/09/2024 |  | 11/09/2024 |  |  | 11/08/2022 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1107437 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a controller error. the aic was removed and replaced with another aic, resulting in a reportedly successful procedure. there was no patient harm reported. | [5 empty] | The investigation into the controller error/loss of opm data has been completed. the impella automated controller was returned for investigation. the data logs revealed multiple instances of an invalid placement signal. this is a known pattern failure in which the majority of signals received from optical bench (placement, contrast, lamp, gain) are interpreted as -1638.4 values and the calculated placement signal is maxed out (railed at 3000 mm hg). this lasts between 2 to 10 minutes, after which all signals go back to normal. unrelated to the reported placement signal issues, there were was an interruption in data streaming functionality, but it was recovered via the device handshake. the returned device was tested with no optical signal failure reproduction. the complaint cause is a known pattern failure characterized by temporary loss of placement signal. no repair will be necessary as the issue has a low probability of reoccurrence, lasts only up to 10 minutes. if needed, patient hemodynamics and impella position can be monitored with imaging. the cause of the transient loss of optical data could not be determined due to the inability to reproduce. this report is being filed as part of a retrospective review of historical records.
14994 | 1220648-2024-22771 | 20647594 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20647594&pc=OZD | 09/06/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 09/06/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1383039 |  | 0042-0040-JP |  | No |  |  | No |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced optical bench check noted that the contrast percentage decreased significantly. the optical kit was replaced to address the issue. no patient was involved. | [5 empty] | The investigation has been completed. the console was not returned. the root cause of the failed 5s spec was contamination on the blue receptacle fiber. this report is being filed as part of a retrospective review of historical records.
14995 | 1220648-2024-22770 | 20647605 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20647605&pc=OZD | 09/07/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 09/07/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1376689 |  | 0042-0000-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) experienced position waveform and a controller abnormality alarm. the console was replaced to address the issue. no patient harm reported. | [5 empty] | The investigation has been completed. the console was not sent back. data logs revealed that this is a known pattern failure in which all signals received from optical bench. the transient loss of placement signal could not be reproduced. this report is being filed as part of a retrospective review of historical records.
14996 | 1220648-2024-22773 | 20647609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20647609&pc=OZD | 06/29/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 06/29/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1230747 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The physician reported that the automated impella controller (aic) experienced a case start issue. it was reported that during pump insertion, the aic would not recognize the impella cp w/smart assist. the aic screen stated 'defective impella catheter. disconnect and replace impella catheter'. the aic was turned off and the team restarted the procedure, including blowing into the aic connector plug area (even with compressed air from the room) with the same result. reportedly a second pump was attempted with the same result. the impella placement was aborted due to another aic not being available and an intra-aortic balloon pump placed instead. there was no patient harm reported. | [5 empty] | The investigation into the pump not detected issue has been completed. the impella automated controller (aic) was returned for investigation. the data analysis of the logs showed the console failed to read pump eeprom multiple times. the failure was reproduced on the returned device by field service. swapping the impellatronic board with a 'known-good' fixed the failure mode and the console successfully read pump eeprom. the cause of the impella defective alarm was a defective impellatronic board. the cause of the defective impellatronic board is unknown. this report is being filed as part of a retrospective review of historical records.
14997 | 1220648-2024-22774 | 20647615 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20647615&pc=OZD | 05/29/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 05/29/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1410591 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  |  | Black Or African American | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a placement signal loss. it was reported that the placement was confirmed by the signal was 0/0. this aic was replaced with another aic resulting in the placement signal returning to normal. there was no patient harm reported. | [5 empty] | The investigation into the optical signal/placement signal issues has been completed. the impella automated controller (aic) was returned for investigation. the data analysis of the logs showed a position alarm followed by a placement signal alar, which stated persistent for nearly eight hours when the alarm briefly cleared. the console then began to alternate between triggering "placement signal low" and "impella position unknown". when the staff attempted to remove the pump from the console, the console was unable to shut down as the console software could not detect the pump was disconnected due to a loss of communication to the opm. during this period of time, it was confirmed through the real time logs that the placement signal did indeed flatline and all optical data was lost as well. this is a known pattern failure to occur on consoles with aic software v7.3 up to and including v8.2. the returned device was visually inspected all interior connections were found to be acceptable, with particular attention to the flex ribbon cable from the opm to the epc. the root cause was unable to be determined. this report is being filed as part of a retrospective review of historical records.
14998 | 1220648-2024-22776 | 20647616 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20647616&pc=OZD | 04/23/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 04/23/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1314432 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced a controller error yellow alarm. there was no reported patient involvement. | [5 empty] | The device was returned and evaluated, and the investigation for the reported controller error (yellow alarm) has been completed. imc logs confirmed the reported failure mode. there were multiple occurrences of the purge system open, air in the purge system and controller failure (defective purge pressure sensor) alarms. the rt logs highlighted a loss in pressure, which could not be resolved despite troubleshooting. the stepper motor ticked forward moving the cassette piston, but the pressure would only increase slightly or not at all. this behavior triggered the controller failure (defective purge pressure sensor #414) alarm given there is a forward stroke with no pressure change. the console was tested for roughly 12 hours with a pump and purge. the cause of the purge pressure leak was not determined. the console performed as expected. this report is being filed as part of a retrospective review of historical records.
14999 | 1220648-2024-22775 | 20647617 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20647617&pc=OZD | 05/12/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 05/12/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1107436 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced a controller error yellow alarm that was resolved by replacing the console. there was no reported patient involvement. | [5 empty] | The device was returned and evaluated, and the investigation for the reported controller error (yellow alarm) has been completed. on the complaint date, the imc logs revealed that there was a kbd read failure which resulted in a controller error (event set #24 - loss of comm with keypad). the failure mode was not reproduced after power cycling the console 150 times. there were also no underlying kbd errors found in the log analysis afterwards. there was also no discoloration on the ribbon cable between the 4-in-1 and kbd. the cause of the controller error (loss of communication to kbd) was unable to be determined because the issue was unable to be reproduced. this report is being filed as part of a retrospective review of historical records.
15000 | 1220648-2024-22772 | 20647622 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20647622&pc=OZD | 07/01/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 07/01/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED, INC. |  |  | US |  | 1379919 |  | 0048-0034-JP |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella for mechanical circulatory support. during support, the automated impella controller (aic) produced a controller error (yellow alarm). no resolution was specified, and no patient harm was reported. | [5 empty] | Clinical review: clinical staff reported that console ic6605 issued a controller error during transport by ambulance when plugged into portable ac inverter 300w. this alarm did not affect the console¿s ability to run the cp c7 impella. once the console was unplugged from the inverter, the alarm resolved. data analysis: imc logs confirmed the reported controller error alarm (pbm voltage out of specification ¿ event set 501) and showed that the vpurge and vpump1 voltages had triggered this alarm. see the logs in the attachments section. this report is being filed as part of a retrospective review of historical records.
15001 | 1220648-2024-22777 | 20647623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20647623&pc=OZD | 04/30/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 04/30/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1085768 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced a controller failure red alarm. there was no reported patient involvement. | [5 empty] | The device was returned and evaluated, and the investigation for the reported controller failure (red alarm) has been completed. imc logs revealed that a controller failure ¿ purge pressure sensor defective was triggered. before the alarm was triggered, there was also another entry in the logs indicating the pressure sensor¿s voltage signal was out-of-range. the failure mode was reproduced with the original hw. after replacing the purge pressure sensor, the alarm resolved. the suspect purge pressure sensor was installed in the burn-in fixture, and the same alarm was reproduced. the controller failure alarm (purge pressure sensor defective) was due to a defective purge pressure sensor. this report is being filed as part of a retrospective review of historical records.
15002 | 1220648-2024-22778 | 20647625 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20647625&pc=OZD | 04/29/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 04/29/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1132986 |  | 0042-0000-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a battery communication failure yellow alarm with no resolution specified. there was no patient involvement. | [5 empty] | D9: the date of the devices return for evaluation is unknown. e1, e3, e4: the name of the initial reporter and occupation is unknown. the investigation for the reported issue has been completed. the device was returned for evaluation. the cause of battery communication alarm was a defective pbm. the cause of the defective pbm is unknown. this report is being filed as part of a retrospective review of historical records.
15003 | 1220648-2024-22769 | 20647644 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20647644&pc=OZD | 09/14/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 09/14/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1207017 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) experienced a error alarm. the console was replaced to address the issue. no patient harm reported. | [5 empty] | The investigation for the reported issue has been completed. the console and data logs were returned. device analysis confirmed controller error observed during procedure was associated with a loss of placement signal. controller error observed during procedure was associated with a loss of placement signal. this is a known pattern failure caused by breakdown of communication with optical bench. abiomed is aware of the issue and is working on appropriate corrective action. this is a low probability event and no repairs are needed as the condition resolves after power-cycle. if needed, monitor patient hemodynamics and impella position with imaging. however, this malfunction also resulted in inability to shut down the console, and the issue had been resolved on site by disconnecting ac power and disengaging batteries, which, in turn, caused corruption of files and/or cf card. the root cause of the transient loss of opm-related signals could not be determined due to the inability to reproduce. this report is being filed as part of a retrospective review of historical records.
15004 | 1220648-2024-22779 | 20647878 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20647878&pc=OZD | 08/11/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 08/11/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1500243 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 92 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Application Program Problem | [6 empty] | Hemodynamic instability | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a controller failure that was resolved with replacing the console with a backup. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and console was booted with no issues. the cause of the controller failure is impellatronics epm firmware corruption. this report is being filed as part of a retrospective review of historical records.
15005 | 1220648-2024-22780 | 20647879 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20647879&pc=OZD | 07/27/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 07/27/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1394733 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  | Hispanic |  | Death | Yes | Yes | Application Program Problem | [6 empty] | Hemodynamic instability | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The us complainant had an optical sensor and display issue with the automated impella controller (aic). the aic was replaced and support proceeded with the 2nd aic. ultimately patient expired on support, and there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | Console logs from the reported day of event are consistent with complaint and show that optical placement signal from pump was invalid during case start. logs also show that multiple attempts were made to reconnect the pump, however case wizards prompts were not followed in order .the cause of the failure mode was most likely use related: pump connector not plugged all the way in, combined with frequent disconnecting/reconnecting of the pump performed out of synch from case start wizard prompts. this report is being filed as part of a retrospective review of historical records.
15006 | 1220648-2024-22781 | 20647916 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20647916&pc=OZD | 07/28/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 07/28/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1249747 |  | 0042-0000-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  |  | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Hemodynamic instability | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Us complainant reported that a call was received call from rn that automated impella controller lost placement signal and lv waveforms, and also received controller error alarm. it was noted that alarms resolved, and waveforms returned. it was agreed that swapping console would be best, decided to switch to new aic. removed console from use. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | D9: the date of the devices return for evaluation is unknown. the investigation for the reported issue has been completed. the impella device was not received from the customer. only the logs were received. console logs from the reported day of event are consistent with complaint. the cause of the loss of placement signal was loss of communication to optical bench. the reason for communication loss was unable to be determined because the issue was unable to be reproduced. this report is being filed as part of a retrospective review of historical records.
15007 | 1220648-2024-22782 | 20647931 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20647931&pc=OZD | 07/18/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 07/18/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1276246 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Defective Device | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The us complainant had the automated impella controller (aic) fall during transport from medical center to tertiary medical center. procedure outcome was that patient expired, there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and console was inspected in order to assess the extent of mechanical damage. the cause of the aic damage was that it had been dropped. this report is being filed as part of a retrospective review of historical records.
15008 | 1220648-2024-22783 | 20647932 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20647932&pc=OZD | 07/14/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 07/14/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1469546 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic |  | Death | Yes | Yes | Application Program Problem | [6 empty] | Hemodynamic instability | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The us complainant had an automated impella controller (aic) fail with a red alarm at case start, and the team replaced the aic. the procedure outcome was that patient expired, and there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and failure mode was reproduced on the first boot. console alarmed for controller failure. failure mode is due to impellatronics epm sw corruption. this report is being filed as part of a retrospective review of historical records.
15009 | 1220648-2024-22784 | 20647933 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20647933&pc=OZD | 07/21/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 07/21/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1329519 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 90 YR |  | Male |  | Hispanic | White | Death | Yes | Yes | Application Program Problem | [6 empty] | Hemodynamic instability | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The us complainant had a cp supporting an elderly patient with a cp when the automated impella controller (aic) failed. the patient remained stable and on pump support. family withdrew care, but there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and this is a known pattern failure in which all signals received from optical bench. the root cause of the transient loss of opm-related signals could not be determined due to the inability to reproduce. this report is being filed as part of a retrospective review of historical records.
15010 | 1220648-2024-22785 | 20647935 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20647935&pc=OZD | 07/13/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 07/13/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1200772 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  |  | White | Death | Yes | Yes | Unintended Application Program Shut Down | [6 empty] | Hemodynamic instability | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automated impella controller (aic) experienced a shut down and would not show the date or time. the aic was replaced, and the case continued. it was reported that the procedural outcome was the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | Section a5 patient ethnicity is unknown the investigation for the reported issue has been completed. the impella device was received from the customer and the unexpected controller shutdown failure mode was not reproduced. no indication of a full or partial reset was seen in the logs. the root cause of the unexpected controller shutdown was most likely due to the 4-in-1 but could not be verified due to the inability to reproduce. this report is being filed as part of a retrospective review of historical records.
15011 | 1220648-2024-22786 | 20647936 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20647936&pc=OZD | 01/27/2022 | 11/09/2024 |  | 11/09/2024 |  |  | 01/27/2022 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1452267 |  | 0042-0010-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  |  | White | Death | Yes | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. while on support, automated impella controller (aic) failed to recognize the purge cassette. the aic was replaced, and the case continued. additional information noted that the patient expired on support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The date of this devices return for evaluation is unknown. the investigation for the reported aic - purge disc/flag issues has been completed. the product was returned and the aic - purge disc/flag issues was confirmed. console logs from the reported day of event are consistent with complaint and show multiple "piston blocked" events, however no evidence of purge disk not being detected (to the contrary - disk was detected during case start). as a result, de-airing procedure could not be completed. after console was received and inspected, it's been observed that purge motor shaft was immobilized due to dextrose contamination. after the assembly was cleaned, the issue was resolved. the cause of the aic issue was contamination. this report is being filed as part of a retrospective review of historical records.
15012 | 1220648-2024-22787 | 20647941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20647941&pc=OZD | 12/10/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 12/10/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1240266 |  | 0042-0000-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  |  | Death | Yes | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the automated impella controller (aic) was associated with a stoppage of the pump, which was resolved by replacing the console. the patient ultimately expired as care was withdrawn. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported aic - pump stop has been completed. the product was returned and the aic - pump stop was confirmed. device analysis: console was tested with several known good pumps where the same issue occurred across all pumps. field service swapped in a known good impellatronics board and tested with known good pumps. no issues occurred. affected aic was ran with pump and purge after an impellatronics board replacement where no issues occurred. the cause of the repeated impella stopped alarms was a defective impellatronic board. the cause of the defective impellatronic board is unknown. this report is being filed as part of a retrospective review of historical records.
15013 | 1220648-2024-22788 | 20647952 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20647952&pc=OZD | 12/01/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 12/01/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1298592 |  | 0042-0010-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Death | Yes | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the automated impella controller (aic) had a purge disc failure. a backup console was utilized for support. additional information noted that the patient expired on support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The date of this devices return for evaluation is unknown. the investigation for the reported aic - purge disc/flag issues has been completed. the product was returned and the aic - purge disc/flag issues was confirmed. data analysis: imc logs at the complaint date are consistent with the complaint. the alarm, purge disc not detected - alarm cleared ¿ event #402, was issued after the insertion of pc. the pc/ purge disc was removed and reinserted, the purge disc alarm was persistent with the original and demo pc/ purge disc set. attempt of rebooting the console failed in resolving this alarm. device analysis: the alarm was reproduced in fs aachen. upon open-up the purge lid, a plenty amount of glucose was found on the purge unit and inside the console. and the purge disc detector flag was glued to the protective film, so it did not move when the purge disc was inserted. the cause of the purge disc not detection/ acknowledge was the stuck purge flag, causing from the glucose built-up between flag and protective film. this report is being filed as part of a retrospective review of historical records.
15014 | 1220648-2024-22789 | 20648177 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20648177&pc=OZD | 09/30/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 09/30/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1087214 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Application Program Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Us complainant reported that the patient was placed onto impella 5.5 support due to acute myocardial infarction / cardiogenic shock. while on support, the automated impella controller (aic) experienced a controller failure red alarm that was resolved by switching out the aic. the patient ultimately expired, there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported issue has been completed. the device and date logs were returned for investigation. several power cycles were unsuccessful in reproducing the controller failure that occurred in the field. console was placed into the burn-in room with a pump and purge setup and allowed to run over night, no errors occurred. a purge cassette was placed into the console and forced pud communication was established, no errors occurred over the course of several hours. this report is being filed as part of a retrospective review of historical records.
15015 | 1220648-2024-22791 | 20648291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20648291&pc=OZD | 10/18/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 10/18/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1220363 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The clinical consultant reported that the automated impella controller (aic) silver button on the front of the aic was not working properly. there was no patient harm reported. | [5 empty] | E1: name of the initial reporter is unknown. the investigation into the keyboard rotary knob issue has been completed. the impella automated controller (aic) was returned for investigation. the data analysis of the logs showed that there was either 0 network device detected, or very weak signal detected during the complaint date, which could potentially be due to either the weak wi-fi signal in the hospital or the degraded sensation for wi-fi of this impella connect module. the returned console was investigated the silver button malfunction complaint was confirmed once the console was booted up, the rotary knob could not navigate thoroughly to all the entries within the menu. after the impellatronic board was momentarily replaced with a known-good one, the rotary knob malfunction was resolved. the issue resulted from a defective impellatronic. the cause of the rotary knob issue was a defective impellatronic board. the cause of the defective impellatronic board is unknown. this report is being filed as part of a retrospective review of historical records.
15016 | 1220648-2024-22790 | 20648296 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20648296&pc=OZD | 09/03/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 09/03/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1119122 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Application Program Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automated impella controller ( aic) emitting a controller error alarm during support with intermittent placement signal issues. the initial aic was replaced, and the case continued. the patient survived the explant, bur expired at the end of procedure. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported issue has been completed. complaint cause is a known pattern failure characterized by temporary loss of placement signal. no repair will be necessary as the issue has a low probability of reoccurrence, lasts only up to 10 minutes. if needed, patient hemodynamics and impella position can be monitored with imaging the console will not need to be returned from the field. the transient loss of placement signal could not be reproduced. this report is being filed as part of a retrospective review of historical records.
15017 | 1220648-2024-22792 | 20648309 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20648309&pc=OZD | 08/30/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 08/30/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1210719 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  |  | Death | Yes | Yes | Application Program Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The us complainant had an automated impella controller (aic) that failed/alarmed during support. the aic was replaced and support proceeds on the cp. patient outcome at explant was patient survived but the procedure outcome was patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported issue has been completed. complaint cause is a known pattern failure characterized by temporary loss of placement signal. no repair will be necessary as the issue has a low probability of reoccurrence, lasts only up to 10 minutes. if needed, patient hemodynamics and impella position can be monitored with imaging the console will not need to be returned from the field. the transient loss of placement signal could not be reproduced. this report is being filed as part of a retrospective review of historical records.
15018 | 1220648-2024-22793 | 20648316 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20648316&pc=OZD | 03/14/2023 | 11/09/2024 |  | 11/09/2024 |  |  | 03/14/2023 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1174511 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  | No | Initial submission | [9 empty] | The us complainant had placed an impella cp for support of a patient admitted with stemi and vf arrest. additionally, during support, the automated impella controller (aic) alarm and failed which prompted the cp to stop. the team replaced the aic on the pump patient who was pump dependent. the patient was supported appropriately after the aic was replaced per ifu recommendations. additional information noted that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported aic - system self-check error has been completed. the product was returned, and the aic - system self-check error was not confirmed. controller error/ failure was not reproduced during testing, but upon further inspection, corrosion of the rs232 traces next to the battery failure buzzer super capacitors was identified. the root cause of the controller failure/ error was due to a corroded power battery manager (pbm). this report is being filed as part of a retrospective review of historical records.
15019 | 1220648-2024-22794 | 20648317 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20648317&pc=OZD | 09/02/2022 | 11/09/2024 |  | 11/09/2024 |  |  | 09/02/2022 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2023214337 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Failure to Read Input Signal | Application Program Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  | No | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the automated impella controller (aic) exhibited a controller error. it was reported that placement signal ao and lv waveforms spontaneously disappeared and returned. the device was replaced with another aic. no report of patient harm or injury related to this malfunction. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The device was returned for evaluation. console logs from the reported day of event are consistent with complaint and show controller alarm due to optical bench communication crc error as well as multiple errors indicating ambient pressure out of range. device analysis: the reported issues were not reproduced. no communication issues were observed, and optical bench ¿5s¿ parameters were within specification. it is most likely that some unspecified malfunction (affecting serial communication) of the 4-in-1 assembly contributed to this issue. this report is being filed as part of a retrospective review of historical records.
15020 | 1220648-2024-22795 | 20648320 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20648320&pc=OZD | 04/18/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 04/18/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1080527 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. during support, the automated impella controller (aic) unexpectedly stopped. multiple attempts to restart were unsuccessful. the patient ultimately expired, but there was no allegation that this was related to the device. | [5 empty] | The device was returned and evaluated, and the investigation for the reported pump stop has been completed. console logs from the reported day of event are consistent with complaint and show multiple "impella stopped" events and alarms. multiple attempts to restart pump, also after power-cycling the console, were unsuccessful. after being received at the lab, the aic had been tested, but the issue was not reproduced and console operated within specification. all electrical and mechanical aspects of the console operation were verified during pm testing, which console passed. there is no evidence suggesting that the console contributed to pump stops events. a separate investigation into the pump device revealed some bearing wear, which is consistent with pump stop events evidenced in logs. therefore, pump stop events have been determined to be caused by the pump, not the console. this report is being filed as part of a retrospective review of historical records.
15021 | 1220648-2024-22796 | 20648324 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20648324&pc=OZD | 08/02/2022 | 11/09/2024 |  | 11/09/2024 |  |  | 08/02/2022 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2023188900 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Failure to Read Input Signal | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  | No | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the automated impella controller (aic) exhibited purge disc failure. the device was replaced with another aic and the procedure was successful. no report of patient harm or injury related to this malfunction. additional information noted that the patient was turned down for a heart transplant and expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The device was returned for evaluation. device logs from the complaint date of (b)(6) 2022 revealed that the console issued the purge disc not detected alarm. device analysis: the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken/ missing. this report is being filed as part of a retrospective review of historical records.
15022 | 1220648-2024-22797 | 20648326 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20648326&pc=OZD | 07/25/2022 | 11/09/2024 |  | 11/09/2024 |  |  | 07/25/2022 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1102528 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Failure to Read Input Signal | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  | No | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the automated impella controller (aic) exhibited a purge disc not detected alarm. the device was replaced with another aic and the procedure was successful. no report of patient harm or injury related to this malfunction. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The device was returned for evaluation. review of the device logs reported several "purge disc not detected" alarms during case start step 2: insert purge cassette and disc. device analysis: console purge assembly was inspected. upon opening the purge assembly door, it was confirmed the purge disc flag had broken/snapped off of the purge assembly. this report is being filed as part of a retrospective review of historical records.
15023 | 1220648-2024-22798 | 20648331 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20648331&pc=OZD | 09/06/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 09/06/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1411424 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) upon changing out the purge cassette the aic did not recognize the purge sensor. this aic was replaced with another aic resulting in no issues with the replacement aic. there was no patient harm reported. | [5 empty] | The investigation into purge disc/flag issues has been completed. the impella automated controller was returned for investigation. the data analysis of the logs from the complaint date revealed that the console issued the purge disc not detected alarm (event set #(b)(4)) several times. the returned console was tested and the issue with properly detecting the purge pressure disc was reproduced, and the purge flag was confirmed to be broken (missing at blue disc retainer). the root cause of the broken purge flag is due to repeated force and wear and tear of the plastic component. this report is being filed as part of a retrospective review of historical records.
15024 | 1220648-2024-22799 | 20648511 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20648511&pc=OZD | 11/16/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 11/16/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1322583 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | Yes | Battery Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the automated impella controller (aic) exhibited a battery failure alarm. the resolution for this issue is unknown. the procedural outcome noted that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The device was sent back for evaluation. the device logs confirmed the reported failure. there was a battery failure alarm and a battery 1 communication failure alarm due to low pack voltage which could be a result of the battery failing to request charging. device analysis: the failure mode was reproduced on an engineering bench. the battery 1 pack voltage was measured to be 0.6 v rather than the expected 16 v, and the battery lcd indicator was blank. battery 1 was replaced to resolve the failure alarm. this report is being filed as part of a retrospective review of historical records.
15025 | 1220648-2024-22800 | 20648513 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20648513&pc=OZD | 11/02/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 11/02/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2022119297 |  | 0042-0000-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Battery Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the automated impella controller (aic) was plugged into red outlet and alarm notification came on battery failure. the device was replaced with another aic. the procedural outcome noted that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: the date of the patient's death is unknown. d9: the date of the devices return for evaluation is unknown. the device was returned for evaluation. the device logs from the reported day of event are consistent with complaint. the device exhibited a battery failure alarm, followed by a battery failure. also, after the ac power disconnected manually, a battery level low alarm and a battery failure was issued, due to the depletion/lack of connection of battery 1. device analysis: the cause of the battery failure was due to a defective battery (rapid decline in cell voltage). the exact cause of the defective battery was unable to be determined. this report is being filed as part of a retrospective review of historical records.
15026 | 1220648-2024-22801 | 20648514 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20648514&pc=OZD | 10/27/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 10/27/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1313442 |  | 0042-0000-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Application Program Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. it was reported that the console did not alarm with suction issues were experienced, it was found that the option to restore sound was not available. the procedural outcome noted that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The device was returned for evaluation. the cause of the no alarm sound was the speaker cable not plugged in, the cause of the speaker cable not plugged in is unable to be determined. this report is being filed as part of a retrospective review of historical records.
15027 | 1220648-2024-22802 | 20648515 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20648515&pc=OZD | 10/01/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 10/01/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1598659 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Application Program Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the automated impella controller (aic) exhibited a yellow warning. the device was replaced with another aic. no report of patient harm or injury related to this malfunction. there is not enough information to exclude the impella device as an associated factor in the patient's demise. the procedural outcome noted that the patient expired. | [5 empty] | The device was returned for evaluation. device logs were reviewed and are consistent with complaint. there were numerous messages of purge-related issues, piston blocked with 0% range remaining, purge fluid not detected error. at some point, a yellow alarm of #501 controller error present, implicating insufficient purge voltages device analysis: the issue of yellow alarm, controller error, and purge-related issues were successfully reproduced during investigation. the motor shaft was found much difficult to spin by hand, compared to other console (with consoles power off). once the stepper motor was removed and disassembled, the gasket of stepper motor was found sticky on the side facing cassette. it was confirmed that the dextrose dried up on gasket increased the friction, drew extra current to overcome the friction and attempt to spin. the extra current supplied to drive stepper motor would ultimately lower and trigger the #501 controller error alarm. this report is being filed as part of a retrospective review of historical records.
15028 | 1220648-2024-22803 | 20648517 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20648517&pc=OZD | 06/23/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 06/23/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1134039 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Application Program Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Us complainant reported that the patient was placed onto impella cp support due to acute myocardial infarction / cardiogenic shock. while on support, the automated impella controller (aic) experienced a controller error yellow alarm that was resolved by changing the console. the patient ultimately expired as care was withdrawn, there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and the console was thoroughly tested on an engineering bench. the failure mode could not be reproduced by power cycling. there were no imc-kbd errors in the logs. the root cause of the controller error could not be determined because the failure mode could not be reproduced. this report is being filed as part of a retrospective review of historical records.
15029 | 1220648-2024-22804 | 20648518 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20648518&pc=OZD | 06/22/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 06/22/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1146386 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  | Black Or African American | Death | Yes | Yes | No Pressure | [6 empty] | Hemodynamic instability | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Us complainant reported that the patient was placed onto impella cp support due to acute myocardial infarction / cardiogenic shock. while on support, the automated impella controller (aic) experienced purge disc/flag issues that were resolved by exchanging the console. the patient ultimately expired as care was withdrawn, there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and console arrived in house and it was observed that the blue purge retainer had broken off of the console. complaint cause is due to console purge disc retainer being broken off of purge assembly. this can happen due to the repeated horizontal force the retainer is subjected to when staff attempt to insert purge discs. complaint cause is due to a broken purge disc retainer. this is likely due to the wear and tear of the plastic component as well as additional unnecessary force when inserting the purge disc into the purge disc retainer. this report is being filed as part of a retrospective review of historical records.
15030 | 1220648-2024-22805 | 20648519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20648519&pc=OZD | 06/11/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 06/11/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1570900 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Failure to Read Input Signal | Obstruction of Flow | [5 empty] | Hemodynamic instability | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Us complainant reported that the patient was placed onto impella rp support due to acute myocardial infarction / cardiogenic shock. while on support, the automated impella controller (aic) experienced a purge issue that was resolved by exchanging the console. the patient ultimately expired as care was withdrawn, but there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and its purge system was thoroughly tested, but no anomalies were discovered, however the purge motor seemed a bit louder than usual, and its shaft harder to spin (but still within specification). although no direct evidence of any hardware malfunction was found, the purge assembly will be replaced as a precaution, due to prior purge issues reported. the cause of the console being unable to prime was unable to be determined because the issue was unable to be reproduced, and neither purge cassette or pump were returned for analysis. this report is being filed as part of a retrospective review of historical records.
15031 | 1220648-2024-22806 | 20648520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20648520&pc=OZD | 05/26/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 05/26/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1140265 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  |  | White | Death | Yes | Yes | Defective Device | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The us complainant had a impella rp supporting a patient when the automated impella controller fell to the floor and caused sensor damage. there was no reported harm or injury to the patient from the automated impella controller damage. the automated impella controller should be swapped with a backup. it was reported that the procedural outcome was the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | After console was received in danvers, the damage had been confirmed. in addition to damaged pump plug connector, the optical pump plug connector cover was missing, but also the impella connect module had dislodged from its mounting bracket and needed re-seating. visual inspection of the inside of console did not reveal any additional damage. this report is being filed as part of a retrospective review of historical records.
15032 | 1220648-2024-22807 | 20648521 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20648521&pc=OZD | 04/19/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 04/19/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1556637 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Complainant reported that the automated impella controller (aic) was not operating properly. the aic was requesting the staff to plug in gray-to-gray connection at case start when using a cp. console switched and case continued. it was reported that the procedural outcome was the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer. while in the field, console failed to detect a pump during case start. imc logs show that the console was able to detect pump connection but failed to read pump eeprom. the root cause of optical signal issue was due to bugs or defects in software (excludes firmware) this report is being filed as part of a retrospective review of historical records.
15033 | 1220648-2024-22808 | 20648523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20648523&pc=OZD | 05/10/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 05/10/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  |  |  | 1158456 |  | 0042-0000-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Read Input Signal | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the automated impella controller (aic) exhibited optical failure. during set up the team received a message that the aic could not use the optical sensor. the staff tried to unplug and plug in catheter with no resolution. the device was replaced with another aic. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The device was returned for evaluation. device logs show that after connecting the purge and pump, the console displayed placement signal unreliable alarm (event #1036). the staff had not started the pump motor at this point. the staff removed both pump and purge and then performed a power-cycle of the console. after bootup, during case start, they received "optical sensor system error: disconnect cable" error. at this point they switched to a backup console and had no issues. device analysis: the error was not reproduced. the cause of the placement signal unreliable error was unable to be determined. this report is being filed as part of a retrospective review of historical records.
15034 | 1220648-2024-22809 | 20648524 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20648524&pc=OZD | 05/04/2021 | 11/09/2024 |  | 11/09/2024 |  |  | 05/04/2021 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1493604 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  |  | Death | Yes | Yes | Application Program Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the automated impella controller (aic) exhibited a controller failure, this occurred after restarting the device three times. the device was replaced with another aic. the procedural outcome noted that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The device was returned for evaluation. a review of the device logs confirmed the reported problem. device analysis: the cause of the multiple pump stops was malfunctioning pmd circuitry. this report is being filed as part of a retrospective review of historical records.
15035 | 1220648-2024-22810 | 20648525 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20648525&pc=OZD | 03/10/2023 | 11/09/2024 |  | 11/09/2024 |  |  | 03/10/2023 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1371242 |  | 0042-0010 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | Yes | Fracture | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. it was reported that at the end of a case, the controllers had a broken component. the procedure was successfully completed using this device. no report of patient harm or injury related to this malfunction. survival outcome at explant noted the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The device was returned for evaluation. visual inspection of the purge assembly area confirmed a ripped/ missing blue retainer. the failure mode was due to broken purge retainer. this report is being filed as part of a retrospective review of historical records.
15036 | 1220648-2024-22811 | 20648526 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20648526&pc=OZD | 12/17/2022 | 11/09/2024 |  | 11/09/2024 |  |  | 12/17/2022 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1413896 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Fracture | Failure to Read Input Signal | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the automated impella controller's (aic) pressure transmitter holder was broken. the device exhibited a low purge alarm when the staff opened the purge cassette door to troubleshoot this issue. the device was replaced with another aic and the procedure was successful. no report of patient harm or injury related to this malfunction. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: the date of the device evaluation is unknown. the device was returned for evaluation. device analysis: a physical inspection of the purge assembly revealed that the purge disc retainer had broken off. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15037 | 1220648-2024-22812 | 20648527 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20648527&pc=OZD | 06/13/2022 | 11/09/2024 |  | 11/09/2024 |  |  | 06/13/2022 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1109488 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Decrease in Pressure | Failure to Read Input Signal | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the automated impella controller (aic) exhibited an air in the purge alarm. the staff tried to de-air the device and replace the tubing; however, this did not resolve the issue. the device also would not recognize the cassette. the device was replaced with another aic and the procedure was successful. no report of patient harm or injury related to this malfunction. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | Device logs from the complaint date revealed the air in purge system alarm and confirmed that no purge pressure increased after the forward movement of purge piston. clinical staff performed the de-air wizard accordingly, and successfully de-aired/ cleared the alarm during on-site troubleshooting. later, the console issued the purge disc not detected alarm occurred several times. reinsertion of pc/ purge disc was observed, yet the console still had problem in detecting purge disc. device analysis: the issues with properly detecting the purge pressure disc were reproduced, and purge flag was confirmed to be broken (missing at blue disc retainer). also, dextrose residue/ contamination was found in purge insert. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15038 | 1220648-2024-22813 | 20648533 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20648533&pc=OZD | 05/09/2022 | 11/09/2024 |  | 11/09/2024 |  |  | 05/09/2022 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1450415 |  | 0042-0000-UK |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  |  |  | Death | Yes | Yes | Decrease in Pressure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Eu complainant reported that the patient was placed onto impella cp support due to acute myocardial infarction / cardiogenic shock. while on support, the automated impella controller (aic) experienced a purge issue that was resolved by replacing the console. the patient ultimately expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | A5: patient race and ethnicity is unknown. d9: device was returned to regional engineers and evaluated and that the return date is unknown. the investigation for the reported issue has been completed. the impella device was received from the customer and console arrived at fs aachen for troubleshooting and repairing. fs tested with pc and pump without problem. the issue was not reproduced at fs office. during inspection, the purge insert was found contaminated with glucose residue and bound the purge motor. this report is being filed as part of a retrospective review of historical records.
15039 | 1220648-2024-22814 | 20648534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20648534&pc=OZD | 04/09/2022 | 11/09/2024 |  | 11/09/2024 |  |  | 04/11/2022 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1581781 |  | 0042-0040-JP |  | Yes |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Death | Yes | Yes | Application Program Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant in japan reported that the automated impella controller (aic) alarmed during the support to 'switch to backup' which was done. it was later reported that the patient expired on support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | Patient age, gender and ethnicity is unknown. d9: device was returned to regional engineers and evaluated and the return date is unknown the investigation for the reported issue has been completed. the impella device was not received from the customer. testing of the console performed at abiomed jp did not reveal any irregularities. "transient loss of optical data" issue was not reproduced during testing. optical bench parameters ("5s") were within specification even before blue receptacle replacement. this report is being filed as part of a retrospective review of historical records.
15040 | 1220648-2024-22815 | 20648535 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20648535&pc=OZD | 02/06/2022 | 11/09/2024 |  | 11/09/2024 |  |  | 02/07/2022 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1303838 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  | White | Death | Yes | Yes | Failure to Read Input Signal | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automated impella controller (aic) did not recognize the purge cassettes. the initial aic was replaced, and the case continued. ultimately care was withdrawn and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer in aachen fs, it's been observed that purge disk retainer was broken, which explains intermittent disk detection. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15041 | 1220648-2024-22816 | 20648536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20648536&pc=OZD | 01/28/2022 | 11/09/2024 |  | 11/09/2024 |  |  | 01/28/2022 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1117563 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Read Input Signal | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automated impella controller (aic) did not recognize the purge cassette. the aic was replaced, and the case continued. it was reported that the procedural outcome was the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and purge assembly was tested with purge and pump, with no issues occurring with the de-airing process or purge cassette recognition. stepper motor was disassembled and visually inspected, with no signs of dextrose or resistance when attempting to rotate motor shaft. the root cause of the lack of purge cassette recognition could not be determined due to the inability to reproduce. this report is being filed as part of a retrospective review of historical records.
15042 | 1220648-2024-22817 | 20648537 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20648537&pc=OZD | 04/25/2022 | 11/09/2024 |  | 11/09/2024 |  |  | 04/25/2022 | 11/09/2024 |  |  |  |  | 11/09/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1303838 |  | 0042-0010-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was noted during prep of impella pump, the connection from the cable to the aic was not recognized, so failure optical sensor not detected occurred. patient expired on support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The date of this devices return for evaluation is unknown. the investigation for the reported optical signal issue has been completed. the product was returned and the optical signal issue was confirmed. console was tested with pump and purge to simulate case start. however, the pt cable was intentionally not fully snapped into the front panel optical connector to simulate the likely event that occurred in the field. test results show that snr was also below optical pump detection threshold, causing the same result that occurred on the complaint date. the root cause of the case start issue is due to not fully seating the pt cable into the console. this report is being filed as part of a retrospective review of historical records.
15043 | 1220648-2024-22505 | 20636767 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636767&pc=OZD | 08/29/2022 | 11/07/2024 |  | 11/08/2024 |  |  | 08/29/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1481995 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | During an in service, the device exhibited a battery failure alarm. no patient involvement. | [5 empty] | The console (aic) was returned for evaluation. upon evaluation of the console, battery 2 was depleted (blank lcd power indicator screen). it provided only 5.0 v (open terminals). therefore, the root cause of the battery failure was most likely due to a depleted battery. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15044 | 1220648-2024-22504 | 20636785 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636785&pc=OZD | 10/28/2021 | 11/07/2024 |  | 11/08/2024 |  |  | 10/29/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1420740 |  | 0042-0000-JP |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. while on support, the screen went black and briefly stopped functioning. the patient's blood pressure decreased, and no alarm history could be identified. a temporary error occurred with the aic software program, which was restarted to resolve the issue. | [5 empty] | D9 - date device returned to manufacturer: unknown. the investigation for the reported power on issues/blank screen issue has been completed. the product was returned, and the issue was confirmed. console logs from the reported day of event are consistent with complaint and indicate that a software process (calculator - part of aic sw) became unresponsive around 4:41:49 on 2021-10-29, and console software detected it, and performed system restart, as a designed response (mitigation mechanism) to such event - that¿s when aic screen went blank for about 10-15 seconds. during the restart, pump 340776 might have been paused for less than 27 seconds, but then the console resumed full operation (pump restarted on its own). there was no "unexpected controller shutdown" alarm (#603) after software restart. this type of event is very rare, and unlikely to reoccur ¿ as evidenced by inability to reproduce the issue during testing in jp fs. console passed all functional tests and was returned to the customer. per the results of the clinical review and investigation, the root cause was determined to be due to console's software mitigation mechanism reacting to an unresponsive software process. this report is being filed as part of a retrospective review of historical records.
15045 | 1220648-2024-22514 | 20636828 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636828&pc=OZD | 04/27/2022 | 11/07/2024 |  | 11/08/2024 |  |  | 04/27/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1567887 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. during support, an automated impella controller (aic) lost power upon pump transfer from a previous console. the issue was resolved with a new console. there was no harm to the patient and support continued. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, ic8758 had no problems with reading eeprom from pump simulator or a demo/service pump. an attempt to run a demo/test impella 5.5 pump on ic8758 was unsuccessful (motor current high alarms, and pump speed deviation alarm), but same pump ran successfully on a known-good engineering console. there was no evidence of corruption of either epm or pmd firmware, therefore any issues were dissimilar to those occurring during the case. nevertheless, it's recommended for the impellatronic assembly to be replaced as a precaution. therefore, the root cause of the shutdown issues was due to the(calculator process crash) was not determined, as the exact symptoms were not reproduced. the impellatronic assembly was replaced and a functional check was performed. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15046 | 1220648-2024-22439 | 20636835 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636835&pc=OZD | 05/28/2022 | 11/07/2024 |  | 11/08/2024 |  |  | 05/28/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1413896 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  |  | Black Or African American | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on an automated impella controller (aic) for mechanical circulatory support. while on support, the purge disc retainer had snapped off. there was no report of patient harm or injury. | [5 empty] | The console (aic) was returned for evaluation. data analysis: imc logs are not relevant for this failure mode. device analysis: console arrived in danvers. console purge assembly was inspected. upon opening the purge assembly door, it was confirmed the purge disc retainer had broken/snapped off of the purge assembly. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15047 | 1220648-2024-22468 | 20636862 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636862&pc=OZD | 08/06/2021 | 11/07/2024 |  | 11/08/2024 |  |  | 08/06/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1263481 |  | 0042-0000-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the automated impella controller (aic) did not start, system check did not function. the customer had another aic onsite they were able to utilize. there was no reported harm to the patient. | [5 empty] | The investigation for the reported issue has been completed. the impella device was returned from the customer (exact date of product return, unknown) . the cause of inability to boot up was a defective pbm. the cause of the defective pbm is unknown. this report is being filed as part of a retrospective review of historical records.
15048 | 1220648-2024-22420 | 20636881 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636881&pc=OZD | 06/01/2022 | 11/07/2024 |  | 11/08/2024 |  |  | 06/01/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1464426 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility¿s biomed reported for an automated impella controller (aic) that the purge clip is broken. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported issue has been completed. the product was returned, and the issue was confirmed. data analysis: imc logs are not relevant for this failure mode. device analysis: console arrived in danvers. console purge assembly was inspected. upon opening the purge assembly door, it was confirmed the purge disc retainer had broken/snapped off of the purge assembly. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15049 | 1220648-2024-22436 | 20636896 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636896&pc=OZD | 12/30/2021 | 11/07/2024 |  | 11/08/2024 |  |  | 12/30/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1154176 |  | 0042-0000-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported an automatic impella controller (aic) battery failed to hold a charge. no additional information provided. no report of patient injury or harm. | [5 empty] | The investigation for the reported aic - battery failure (red alarm) issues has been completed. the impella device was returned for evaluation. however, clinical details were sufficient to determine the root cause. the root cause of the battery charging issues was due to a defective power supply. this report is being filed as part of a retrospective review of historical records.
15050 | 1220648-2024-22454 | 20636946 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636946&pc=OZD | 01/03/2022 | 11/07/2024 |  | 11/08/2024 |  |  | 01/03/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1467028 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient with high-risk percutaneous coronary intervention was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) experienced failing to prime the pump and/or cassette during preparation of the device. no clinical details regarding resolution. no patient harm reported. | [5 empty] | The investigation has been completed. the console (b)(6) was sent back for further investigation. visual inspection revealed contamination of the purge insert with glucose residue, and the purge motor shaft was unable to spin freely, due to fluid ingress and contamination. after purge motor assembly was cleaned, the issue was resolved, and console was able to purge. the cause of the failure mode was contamination. this report is being filed as part of a retrospective review of historical records.
15051 | 1220648-2024-22446 | 20636981 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636981&pc=OZD | 01/26/2022 | 11/07/2024 |  | 11/08/2024 |  |  | 01/26/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1431282 |  | 0042-0010-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient with high-risk percutaneous coronary intervention was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) experienced optical signal issues, which resolved through console replacement. no patient harm reported. | [5 empty] | The investigation has been completed. the product was not returned. the root cause of the failure to detect an optical pump was due to lack of programming of the pmd bit. this report is being filed as part of a retrospective review of historical records.
15052 | 1220648-2024-22491 | 20637005 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20637005&pc=OZD | 09/01/2022 | 11/07/2024 |  | 11/08/2024 |  |  | 09/01/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1401648 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a controller error alarm immediately after bootup. this aic was replaced with a new aic resolving the issue, and the procedure was successfully completed. there was no patient harm reported. | [5 empty] | E1: name of the initial reporter is unknown. the investigation into the controller error has been completed. the impella automated controller was returned for investigation. the data log reports from the reported day of event are consistent with complaint and show apparent failure of optical bench lamp (manifested by unusually low value of "light" parameter, consistent with light bulb not being powered), that remained unresolved after bench self-reset, and also after power-cycling of the console. the returned console was tested, and the issue was not reproduced, despite multiple boot-up cycles. monitoring the optical bench current during aic start-up did not reveal any anomalies. the cause of the issue was not determined because issue was unable to be reproduced. this report is being filed as part of a retrospective review of historical records.
15053 | 1220648-2024-22437 | 20637019 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20637019&pc=OZD | 05/26/2022 | 11/07/2024 |  | 11/08/2024 |  |  | 05/26/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1223029 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) had a broken transducer component. a loaner was sent to the customer to initiate return of this device. there was no report of patient harm or injury. | [5 empty] | The investigation into the damage has been completed. the impella automated controller was returned for investigation. the returned console was inspected the purge pressure sensor retainer was confirmed to be broken. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15054 | 1220648-2024-22434 | 20637051 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20637051&pc=OZD | 05/13/2022 | 11/07/2024 |  | 11/08/2024 |  |  | 05/31/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1261012 |  | 0042-0000-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) that the console did not recognize the pump after the pump cable was connected, following the preparation procedure. to resolve the issue, the console was exchanged for a backup and the procedure completed. there was no report of patient harm or injury. | [5 empty] | The investigation into the aic - pump not detected issue has been completed. the automated impella controller (aic) was returned for investigation. console logs from the reported day of event are consistent with complaint and show multiple prompts to connect the catheter (case start wizard step 3) but the pump was never detected. there were no alarms or errors recorded in device log. console had not been power-cycled between attempts to re-connect the pump. the issue was not reproduced later with a demo pump at the distributor¿s office, nor did it happen before at the hospital (during prior clinical use on 2022-01-12, a pump was successfully detected and recognized). the issue was not reproduced in aachen fs. the symptoms (pump not detected) and their infrequent occurrance, are consistent with a known pattern failure ¿epm crystal issue¿ where due to timing issue on impellatronic, the i2c communication between epm circuitry and pump¿s eeprom cannot be established. the mitigation mechanism has been introduced in aic_v8.4.1. to minimize chances of reoccurrence, ic4658 software is to be upgraded from current v8.2 to newest release (v8.5). repair: execute fcn 84 (aic sw upgrade to v8.5) and perform functional check. the cause of the aic issue was most likely a design flaw on impellatronic board, which had been addressed by software mitigation mechanism implemented in aic_v8.4.1. this report is being filed as part of a retrospective review of historical records.
15055 | 1220648-2024-22487 | 20637071 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20637071&pc=OZD | 04/08/2021 | 11/07/2024 |  | 11/08/2024 |  |  | 04/08/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1154603 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. when the automated impella controller (aic) was turned on, it showed the message "system self-check failed. change console immediately" and would not proceed any further. resetting the console multiple times did not resolve the issue, so the console was replaced. there was no reported patient harm. | [5 empty] | The device was returned and evaluated, and the investigation for the reported controller error (yellow alarm) has been completed. the device logs showed pmd communication errors which resulted in a failure of system self check. failure mode was reproduced during testing. evaluation of the device found that the connection at j14 on the pbm was not fully seated, which affected the 20v supply to the pmd circuitry. once it was fully seated, the console operated as expected. the cause of the failure of system self-check was a loose cable. the cause of the loose cable was unable to be determined. this report is being filed as part of a retrospective review of historical records.
15056 | 1220648-2024-22480 | 20637124 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20637124&pc=OZD | 06/09/2021 | 11/07/2024 |  | 11/08/2024 |  |  | 06/09/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1303854 |  | 0042-0000-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during an annual preventative maintenance it was observed that the coil on the pbm board for the automated impella controller (aic) was not soldered. no patient involvement. | [5 empty] | D9: the date of the devices return for evaluation is unknown. e1, e3, e4: the name of the initial reporter and occupation is unknown. the device was returned for evaluation. the loose component was identified at l123 (common-mode choke). according to information obtained from vendor, this component is installed manually during "secondary a" fabrication step of pbm pca. results of visual inspection of the component and its designated pcb location, indicate that the part has not been installed correctly (2 out of 4 contacts had insufficient amount of solder). this report is being filed as part of a retrospective review of historical records.
15057 | 1220648-2024-22453 | 20637180 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20637180&pc=OZD | 08/30/2021 | 11/07/2024 |  | 11/08/2024 |  |  | 08/30/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMAUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1166648 |  | 0042-0000-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility¿s biomedical engineer reported for an automated impella controller (aic) that during preventive maintenance the display failed. there was no patient involvement. | [5 empty] | The investigation for the reported display issue has been completed. the product was returned, and the issue was confirmed. data analysis: relevant logs are unavailable due to the failure of the 4-in-1. imclog server was unresponsive in console failed state. device analysis: complaint cause was reproduced and console booted into failed state. in failed state, the console seemed on, however the screen was completely blank. console interior was inspected and connections were verified. voltage measurements were taken from the pbm to the epc and the pbm voltage supply is valid, epc likely malfunctioning. per the results of the clinical review and investigation, the root cause was determined to be due to the 4-in-1. this report is being filed as part of a retrospective review of historical records.
15058 | 1220648-2024-22461 | 20637229 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20637229&pc=OZD | 08/27/2021 | 11/07/2024 |  | 11/08/2024 |  |  | 08/27/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1164194 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) a failure of the aic to display when booting up. the issue was eventually resolved with a reboot. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported power on issues/blank screen issue has been completed. the product was returned, and the issue was confirmed. data analysis: aic logs from the reported complaint date showed that the console booted up normally, but the user was pressing the softkeys sporadically which could indicate there were issues with the display. the unexpected reboot was due to the clinical staff toggling the battery switch. device analysis: the console was tested thoroughly on an engineering lab bench, but the failure mode could not be reproduced. testing included power cycling the console, and upon each power on, the display was photographed. no issues were observed that were related to a blank screen. the backlight supply cable and the lvds video cable were wiggled to exploit a possible loose connection, and no display issues were observed during this time. per the results of the clinical review and investigation, the root cause could not be determined due to inability to reproduce. this report is being filed as part of a retrospective review of historical records.
15059 | 1220648-2024-22451 | 20637231 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20637231&pc=OZD | 12/10/2021 | 11/07/2024 |  | 11/08/2024 |  |  | 12/10/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1186354 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a battery failure occurred with the automatic impella controller (aic) during preventative maintenance (pm). no patient was involved. | [5 empty] | The investigation for the reported aic - battery failure (red alarm) issues has been completed. the device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the cause of the battery failure alarm was due to a defective battery (rapid decline in cell voltage). the exact cause of the defective battery was unable to be determined. this report is being filed as part of a retrospective review of historical records.
15060 | 1220648-2024-22476 | 20637233 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20637233&pc=OZD | 07/19/2021 | 11/07/2024 |  | 11/08/2024 |  |  | 07/19/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1276246 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The biomedical engineer reported that the automated impella controller (aic) optical cover needed replacement due to missing/broken cover. there was no patient harm reported. | [5 empty] | E1: name of the initial reporter is unknown. the investigation into the damage has been completed. the impella automated controller was returned for investigation. the returned console was inspected on an engineering bench, and the blue pump receptacle was confirmed to be broken. the pump receptacle dust cover was missing. after momentarily replacing the pump receptacle, the console was tested with a known-good pump and purge. the field service engineer was instructed to replace the front panel optical connector and perform a full functional check on the console prior to returning to the customer. no issues were observed during testing. the root cause of the damaged blue receptacle and cover was undetermined. this report is being filed as part of a retrospective review of historical records.
15061 | 1220648-2024-22463 | 20637290 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20637290&pc=OZD | 08/27/2021 | 11/07/2024 |  | 11/08/2024 |  |  | 08/27/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1284911 |  | 0042-0010-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) failing self-check during preparation for a procedure. a new aic was used. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported power on issues/blank screen issue has been completed. the product was returned, (exact product return date unknown) and the issue was confirmed. data analysis: aic logs from the complaint date confirmed the issues properly starting up, which were due to communication issues with the pbm, sau_powermod_1. device analysis: the failure mode was reproduced while in fs aachen. the cables, connectors, and pbm pca were inspected thoroughly but nothing apparent was found. given the start up issue could not be resolved with the original hw, fs replaced the pbm pca so that console would power on normally. there were no communication issues afterwards. per the results of the clinical review and investigation, the root cause was determined to be due to the pbm pca. the cause of the defective pbm is unknown. this report is being filed as part of a retrospective review of historical records.
15062 | 1220648-2024-22507 | 20637344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20637344&pc=OZD | 08/23/2022 | 11/07/2024 |  | 11/08/2024 |  |  | 08/25/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1194348 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a controller error alarm immediately after bootup. it was reported that the aic was restarted, and the procedure was successfully completed with this aic. there was no patient harm reported. | [5 empty] | E1: name of the initial reporter is unknown. the investigation into the controller error has been completed. the impella automated controller was returned for investigation. the data log reports analysis noted the yellow controller alarm corresponding to #24 (loss of communication to the keypad for up-to 5 secs (includes communication loss between the embedded passive components (epc) and keyboard (kbd) software for all kbd hardware (hw) revisions or between the kbd and i2c for kbd hw revision 1 as defined in 0042-9650)) is present in the logs corresponding to the complaint. the returned console was tested, and the failure reproduced on the lab bench during testing. the ribbon cable between the 4 in1 and kbd was replaced with a known good cable, but the failure was still reproduced. in the next step, the i2c communication between the kbd and 4in1 was monitored using an oscilloscope. the oscilloscope was configured to monitor i2c communication and trigger was set to i2c start, single event capture. while monitoring the i2c protocol communication, it was found that the kbd initialization sequence captured on the i2c line was as expected (as per the kbd ic specifications). as further testing the 4in1 was replaced with a known good 4 in1, this resolved the issue, and the failure was not reproduced in the next 70 power cycles. the root cause of keyboard freeze was the malfunctioning of msc1210 present on 4-in-1. the cause of the aic issue was a defective 4in1 board this report is being filed as part of a retrospective review of historical records.
15063 | 1220648-2024-22449 | 20637382 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20637382&pc=OZD | 01/11/2022 | 11/07/2024 |  | 11/08/2024 |  |  | 02/27/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1541270 |  | 0042-0010-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) experienced a purge disc/flag issue, which was resolved by changing the console. no patient harm reported. | [5 empty] | The investigation has been completed. the original complaint description states "console has not recognized the purge cassette". although this cannot be confirmed via aic logs, it is likely that a different message was presented to the user: "purge fluid not detected" which can be caused by an immobilized purge spindle. the contamination was the root cause of the purge issues during case start. this report is being filed as part of a retrospective review of historical records.
15064 | 1220648-2024-22517 | 20637428 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20637428&pc=OZD | 04/21/2022 | 11/07/2024 |  | 11/08/2024 |  |  | 04/21/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1277973 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Unintended Application Program Shut Down | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, an automated impella controller (aic) error and unexpected shutdown during the night was noted. aic was able to be restated with no further issues. no patient harm reported. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, the failure mode was not reproduced on engineering lab bench. long term testing revealed no abnormalities with the system. epc supply was monitored but was normal for when console was operating off battery and ac power. cables and connections were not intermittent. therefore, root cause the unexpected console restart was not determined due to the inability to reproduce. this report is being filed as part of a retrospective review of historical records.
15065 | 1220648-2024-22481 | 20637476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20637476&pc=OZD | 10/28/2021 | 11/07/2024 |  | 11/08/2024 |  |  | 10/28/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1557936 |  | 0042-0010-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automatic impella controller (aic) had a console failure. no additional information provided. there was no report of patient harm or injury. | [5 empty] | D9 - date device returned to manufacturer: unknown. the investigation for the reported controller failure issue has been completed. the product was returned, and the issue was confirmed. data analysis: logs are consistent with complaint intake. subsequent boots following the complaint date trigger these alarms. long term log analysis of the battery data reveals that beginning before the complaint date of occurrence, all cell voltages of battery 1 had been declining for an extended period of time. device analysis: console arrived in aachen. it was confirmed that persistent battery failure (transport connection blown/missing) was not due to a disengaged battery switch. due to low/ flatlined voltage, battery 1 data is unavailable through ata or bq software. when console was booted for the first time in aachen, console alarms are consistent with what was seen in the field. console interior connections were checked and verified. when visually inspecting the batteries and pbm, it was observed that one of the batteries status leds were completely blank. per the results of the clinical review and investigation, the root cause was determined to be due to a defective battery (rapid decline in cell voltage). the exact cause of the defective battery was utd. this report is being filed as part of a retrospective review of historical records.
15066 | 1220648-2024-22459 | 20637496 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20637496&pc=OZD | 08/24/2021 | 11/07/2024 |  | 11/08/2024 |  |  | 08/24/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER EU |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER EU |  |  | ABIOMED, INC. |  |  | US |  | 1108441 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) did not recognize the purge disc when inserted. it was reported the purge system was changed out and the error occurred again with the replacement purge system. this aic was replaced, the same purge system was successfully used with the replacement aic resulting in no issues. there was no patient harm reported. | [5 empty] | The investigation into purge disc/flag issues has been completed. the impella automated controller was not received from the customer and therefore, an evaluation of the device was not possible. the data analysis of the logs from the complaint date revealed that the console issued the purge disc not detected alarm (event set #402) several times. the field service representative tested the aic and noted the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. the root cause of the broken purge flag is likely due to repeated force and wear and tear of the plastic component. this report is being filed as part of a retrospective review of historical records.
15067 | 1220648-2024-22506 | 20637543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20637543&pc=OZD | 08/24/2022 | 11/07/2024 |  | 11/08/2024 |  |  | 08/24/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1550078 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | Battery failure was reported. patient information is unknown; however, it was noted that there was no harm to patient. | [5 empty] | The console (aic) was returned for evaluation. upon evaluation of the console, it was confirmed that persistent battery failure (transport connection blown/missing) was not due to a disengaged battery switch. when the console was booted, the console alarms were consistent with what was seen in the field. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15068 | 1220648-2024-22531 | 20637585 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20637585&pc=OZD | 09/15/2021 | 11/07/2024 |  | 11/08/2024 |  |  | 09/15/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1087949 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automatic impella controller (aic) experienced a controller failure red alarm. the resolution for this issue is unknown. no report of patient involvement, harm, or injury. | [5 empty] | The investigation for the reported aic - battery failure (red alarm) issues has been completed. the device has been returned for evaluation. failure mode was due to low purge pressure causing no change in pressure when the purge cassette piston stroked. however, the overall root cause of the low purge pressure is undetermined. this report is being filed as part of a retrospective review of historical records.
15069 | 1220648-2024-22526 | 20637638 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20637638&pc=OZD | 06/15/2022 | 11/07/2024 |  | 11/08/2024 |  |  | 06/15/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1145434 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Electrical Power Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a ground fault error with an aic (automated impella controller) console. there was no reported harm as a result. | [5 empty] | The impella device was received from the customer on 20-jun-2022. data analysis: aic logs were not relevant to this failure mode. device analysis: console was tested on an engineering lab bench. the protective earth resistance test (part of iec62353 failed, surpassing the threshold of .3 ohms with a measurement of .628 ohms. this device under test would intermittently fail this portion of the electrical safety test. after replacing the power cord, the console would not fail. the root cause of the failure of electrical safety was most likely due to the ac power cord. before sending this console back to the customer, fs was instructed to replace the cordset (6000-0138), power entry module (2100-0086) as a precaution, and perform the electrical safety test from the pm procedure (0042-7309). this report is being filed as part of a retrospective review of historical records.
15070 | 1220648-2024-22528 | 20637641 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20637641&pc=OZD | 09/29/2021 | 11/07/2024 |  | 11/08/2024 |  |  | 09/29/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1095119 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced a controller error yellow alarm. no clinical details regarding resolution or patient condition were provided. no patient was involved. | [5 empty] | The investigation has been completed. the console (b)(6) was sent back for further investigation. the failure mode was reproduced during testing. momentarily replacing the 4-in-1 with a known-good resolved the issue during power cycling. controller alarms upon boot-up were caused by communication failure between kbd and 4-in-1, due to 4-in-1 malfunction. this report is being filed as part of a retrospective review of historical records.
15071 | 1220648-2024-22530 | 20637647 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20637647&pc=OZD | 09/13/2021 | 11/07/2024 |  | 11/08/2024 |  |  | 09/13/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1288953 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automatic impella controller (aic) experienced a controller failure red alarm. when the staff was placing the purge cassette in the console during start up, the customer received an error code of "purge cassette error, try new console". field service went to the customer site and tried to replicate the error; this was unsuccessful. the resolution for this issue is unknown. no report of harm or injury to the patient. | [5 empty] | E1, e3, e4: the name of the initial reporter and occupation is unknown. the investigation for the reported aic - battery failure (red alarm) issues has been completed. the device has not been returned for evaluation. however, investigation details were sufficient to determine the root cause. the root cause of the purge system failure (piston blocked due to reduced range) was most likely due to an undetermined issue during initial setup (when connecting the cassette and pressure disc). some of the possible issues are: piston blocked due to a defective cassette. purge cassette not inserted correctly. blockage on the purge line (line blocked or kinked). however, the root cause of this issue is unable to be determined. this report is being filed as part of a retrospective review of historical records.
15072 | 1220648-2024-22529 | 20637650 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20637650&pc=OZD | 09/17/2021 | 11/07/2024 |  | 11/08/2024 |  |  | 09/17/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1084503 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. while on support, the automatic impella controller (aic) experienced a controller failure red alarm with no resolution specified. no report of patient harm or injury. | [5 empty] | The investigation for the reported aic - battery failure (red alarm) issues has been completed. the device has been returned for evaluation. the root cause of the controller failure was due to a defective purge assembly. the exact subcomponent of the purge assembly was not determined. this report is being filed as part of a retrospective review of historical records.
15073 | 1220648-2024-22500 | 20637690 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20637690&pc=OZD | 11/10/2022 | 11/07/2024 |  | 11/08/2024 |  |  | 11/10/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1603745 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) immediately showed a battery failure upon startup. there was no reported patient involvement. | [5 empty] | The investigation for the reported battery failure (red alarm) has been completed. the device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. as the users were able to successfully troubleshoot the device in the field, the cause of the issue was determined to be use related. this report is being filed as part of a retrospective review of historical records.
15074 | 1220648-2024-22516 | 20637712 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20637712&pc=OZD | 04/08/2022 | 11/07/2024 |  | 11/08/2024 |  |  | 04/08/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1089165 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The us complainant had an automated impella controller (aic) purge disc/door malfunction. there was no harm or injury and the instructions for use (ifu) was followed by using a backup controller for patient use. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, no issues found. console purge door latch button was pressed, door opened with no issues. door was opened and closed repeatedly with no issues. console interior latch mechanism for purge door was inspected, no issues were found. this report is being filed as part of a retrospective review of historical records.
15075 | 1220648-2024-22513 | 20637762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20637762&pc=OZD | 10/27/2021 | 11/07/2024 |  | 11/08/2024 |  |  | 10/27/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1560022 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced loose power supply, missing screws within the unit. no clinical details regarding resolution or patient condition were provided. no patient was involved. | [5 empty] | The investigation has been completed. the dismounted power supply module was most likely due to an internal manufacturing error. this report is being filed as part of a retrospective review of historical records.
15076 | 1220648-2024-22515 | 20637770 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20637770&pc=OZD | 10/20/2021 | 11/07/2024 |  | 11/08/2024 |  |  | 10/26/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1126092 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced a purge disc issue. no clinical details regarding resolution or patient condition were provided. no patient was involved. | [5 empty] | The investigation has been completed. the console (b)(6) was sent back for further investigation. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broke. the root issue with properly detecting the purge pressure disc was due to a broken purge flag. this report is being filed as part of a retrospective review of historical records.
15077 | 1220648-2024-22520 | 20637772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20637772&pc=OZD | 03/18/2022 | 11/07/2024 |  | 11/08/2024 |  |  | 03/18/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1324777 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The clinical consultant reported that the automated impella controller (aic) loaner had a cracked purge disc holder. a loaner was sent to the customer to initiate return of this device. there was no patient harm reported. | [5 empty] | The investigation into the damage has been completed. the impella automated controller was returned for investigation. the visual inspection of the returned purge assembly area confirmed the retainer was cracked. the cause of the issue was a defective component.
15078 | 1220648-2024-22510 | 20637807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20637807&pc=OZD | 06/09/2022 | 11/07/2024 |  | 11/08/2024 |  |  | 06/09/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1326823 |  | 0042-0000-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility¿s biomedical engineer reported for an automated impella controller (aic) a broken purge clip. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported purge disc/flag issue has been completed. the product was returned, and the issue was confirmed. data analysis: imc logs are not relevant for this failure mode. the failure was visually observed, not involved in clinical procedure. device analysis: fs technician visually confirmed the broken purge clip on the console. per the results of the clinical review and investigation, the root cause was determined to be a defective component. this report is being filed as part of a retrospective review of historical records.
15079 | 1220648-2024-22488 | 20637835 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20637835&pc=OZD | 04/12/2021 | 11/07/2024 |  | 11/08/2024 |  |  | 04/12/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1456795 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the console was in for preventative maintenance (pm) and on initial start up the console had a system self-check failure. no patient involvement. | [5 empty] | The name of the initial reporter and occupation is unknown. the investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. pbm components were shorted due to psu screws coming loose. a screw likely fell underneath the pbm and caused a short. this report is being filed as part of a retrospective review of historical records.
15080 | 1220648-2024-22512 | 20637841 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20637841&pc=OZD | 05/02/2022 | 11/07/2024 |  | 11/08/2024 |  |  | 05/02/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2023175083 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. the complainant reported an automated impella console (aic) failing to recognize a new cassette, upon routine change. a new console was used without any injury/harm. | [5 empty] | The investigation into the aic -purge disc/ flag issue has been completed. the automated impella controller (aic) was returned for investigation. console logs on the complaint date were retrieved. imc logs from the complaint date were consistent with the ci. after few days of support, purge cassette was about to be changed via purge cassette change wizard. however, console reported for flush fluid not detected despite of attempts of de-airing. staff also attempted to troubleshoot by removing and re-inserting the same pc, yet the issue persisted. therefore, the procedure could not be completed. visual inspection of the purge assembly area confirmed a broken blue retainer at the screw post. the failure mode was due to broken/ missing purge retainer. see attached image. before returning this console back to the customer, field service was instructed to replace the purge flag and retainer, clean the purge assembly, and perform preventive maintenance section 10-12 (0042-7309) and functional check on console and optical system. this report is being filed as part of a retrospective review of historical records.
15081 | 1220648-2024-22519 | 20637879 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20637879&pc=OZD | 03/22/2022 | 11/07/2024 |  | 11/08/2024 |  |  | 03/22/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1133147 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) would not read the purge cassette. this aic was replaced, and the same purge cassette used with no further issues. there was no patient harm reported. | [5 empty] | The investigation into the purge issue has been completed. the impella automated controller was returned for investigation. the data analysis of the logs indicated the complaint date is consistent with complaint intake. on (b)(6) 2022, the logs reported several "purge disc not detected" alarms. the returned console was investigated, and the purge disc retainer was visually inspected, it was observed that the purge disc flag was missing/ broken. it was also observed that the console had unreported physical damage to the housing and purge assembly area. the console interior components were inspected for physical damage, no issues were observed. the damage to the console was most likely due to the aic housing falling, however it is undetermined. the cause of the purge issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15082 | 1220648-2024-22518 | 20637920 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20637920&pc=OZD | 10/11/2021 | 11/07/2024 |  | 11/08/2024 |  |  | 10/11/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1156148 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced a purge disc issue. no clinical details regarding resolution or patient condition were provided. no patient was involved. | [5 empty] | The investigation has been completed. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. the lack of purge disc recognition is due to a broken purge flag. this report is being filed as part of a retrospective review of historical records.
15083 | 1220648-2024-22523 | 20637926 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20637926&pc=OZD | 09/24/2021 | 11/07/2024 |  | 11/08/2024 |  |  | 09/24/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1566931 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced a purge disc issue. no clinical details regarding resolution or patient condition were provided. no patient was involved. | [5 empty] | The investigation has been completed. the console (b)(6) was sent back for further investigation. the issue with properly detecting purge fluid was reproduced, and purge disc housing was confirmed to be broken off of the purge assembly. failure mode was reproduced in house. the root cause of the failure mode is the blue disc retainer breaking off the purge assembly. this report is being filed as part of a retrospective review of historical records.
15084 | 1220648-2024-22525 | 20637933 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20637933&pc=OZD | 09/29/2021 | 11/07/2024 |  | 11/08/2024 |  |  | 09/29/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1080780 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with high-risk percutaneous coronary intervention was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) experienced a purge disc issue, which was resolved by changing the console. no patient harm reported. | [5 empty] | The investigation has been completed. the console (ic1137) was sent back for further investigation. the purge disc detection failure was confirmed by visual inspection of the purge system. purge disc detection flag was missing. the root cause of the purge disc not detected was the missing flag. this report is being filed as part of a retrospective review of historical records.
15085 | 1220648-2024-22522 | 20637982 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20637982&pc=OZD | 10/03/2021 | 11/07/2024 |  | 11/08/2024 |  |  | 10/06/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1313446 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) experienced a purge disc issue. no clinical details regarding resolution. no patient harm reported. | [5 empty] | The investigation has been completed. the console (ic5545) was sent back for further investigation. the defective console ic5545 was inspected in fs, finding the purge disk holder broken. with the loose blue retainer and missing/ broken purge disk, the purge disk was unable to be tied to the purge pressure sensor. hence, the purge fluid pressure cannot be detected. also, dextrose residue was observed around the purge disk holder area. the root cause of the console in ability to prime was damage on purge disk holder of the purge pressure assembly. this report is being filed as part of a retrospective review of historical records.
15086 | 1220648-2024-22532 | 20638000 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20638000&pc=OZD | 09/10/2021 | 11/07/2024 |  | 11/08/2024 |  |  | 09/10/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1320563 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. during support, a large crash was heard, and console was found to be face down on the floor causing the pump to stop. the pump and aic were replaced. no report of patient harm related to the aic. | [5 empty] | The device was returned for evaluation. no damage to the exterior was observed. when lifting the connector cover, it can be seen that the optical receptacle broke and was forced into the console, likely from when the console had fallen. this report is being filed as part of a retrospective review of historical records.
15087 | 1220648-2024-22533 | 20638858 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20638858&pc=OZD | 01/04/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 01/04/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1207018 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | Battery Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility¿s biomedical engineer reported an automated impella controller failed to boot up. there was no report of patient harm or injury. additional information was found during investigation of the device. there was a controller error and a battery failure observed in the logs. | [5 empty] | The investigation has been completed. the logs from the complaint date are filled with communication errors from the purge system. there was a controller error and a battery failure observed in the logs. the system self-check failure was reproduced on an engineering bench. as received, the console would not boot up (failed system self-check). upon further analysis, "fwcheck" revealed that the hardware revision and software version for the pud module was ¿n.a.¿. the cause of the battery failure was due to battery depletion. the cause of the controller failure was due to defective optical bench. the exact root cause of the defective optical bench was unable to be determined. the defective pud was temporarily replaced with a known good one, which solved the system self-check failure. the syscall error is due to pud corruption. the failure cause was likely linked to the software installation process and not the console¿s original hardware. this report is being filed as part of a retrospective review of historical records.
15088 | 1220648-2024-22534 | 20638883 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20638883&pc=OZD | 02/10/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 02/10/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1109700 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an automated impella controller (aic) for mechanical circulatory support. while on support, there was a controller error during use of the device. the aic was switched for a backup controller without any adverse impact on the patient. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported controller failure issue has been completed. the console was returned for evaluation. the console logs from the reported day of event are consistent with the complaint and showed alarms (optical bench communication crc invalid) and (optical bench ambient pressure out of range). log analysis also showed a similar event throughout support. however, each time the issues were resolved within less than two minutes. therefore alarms were not triggered. the cause of the issue was not determined since the issue could not be reproduced. this report is being filed as part of a retrospective review of historical records.
15089 | 1220648-2024-22535 | 20638921 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20638921&pc=OZD | 02/11/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 02/11/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1314952 |  | 0042-0010-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Nurse | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | Key or Button Unresponsive/not Working | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported that at the initial start, the automated impella controller starts with a blue screen. the soft keys are illuminated but did not buzz when touched, and the rotating knob does not work. there was no report of patient harm or injury. | [5 empty] | The exact date the device was returned for evaluation is unknown. the investigation has been completed. the console was returned. the logs showed keyboard communication issues after fcn update. the rotary push button was non-functional. a known-good 4-in-1 was momentarily swapped to resolve the system self check failure. the cause of the automated impella controller issue was a defective 4in1 board. this report is being filed as part of a retrospective review of historical records.
15090 | 1220648-2024-22536 | 20638977 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20638977&pc=OZD | 03/31/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 03/31/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1248780 |  | 0042-0000-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was on an automated impella controller (aic) for mechanical circulatory support. the aic does not charge even though it is plugged in. the aic turns off and smoke came out. there was a smell noted. there was no noted harm or injury. another aic was used for care. the smoke was immediately noted and the product was removed immediately. | [5 empty] | D.9 the exact date the device was returned for evaluation is unknown. the investigation into the alternating current (a/c) power issue has been completed. the console logs from the reported day of event showed the console being power-cycled six times and each time the a/c was reported to be disconnected. the console was received and inspected. this revealed that the internal power supply was damaged and burnt. after the power supply was replaced, the console was able to be powered from the a/c and needed to change batteries. the source of smoke and cause for the console being unable to charge was damaged power supply. the cause for power supply damage was not determined. this report is being filed as part of a retrospective review of historical records.
15091 | 1220648-2024-22537 | 20639059 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639059&pc=OZD | 04/06/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 04/06/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1245285 |  | 0042-0000-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a battery failure alarm upon boot up during preventative maintenance. no patient was involved. | [5 empty] | D.9 the exact date the device was returned for evaluation is unknown. the investigation for the reported issue has been completed. clinical review stated that during preventative maintenance, the console was issued a battery failure alarm upon boot up. the console was set aside for further analysis. logs were reviewed and showed a battery failure alarm was issued upon startup, along with other battery related alarms. the console was evaluated and the liquid crystal diode indicator on battery 2 showed zero bars, while battery 1 had full bars. batttest showed channel 2 was missing fu. this is likely a malfunction of the power battery manager¿s (pbm) charging subcircuit for channel 2, affecting charging of battery 2. the cause of battery failure alarm was a defective pbm. the cause of the defective pbm is unknown. this report is being filed as part of a retrospective review of historical records.
15092 | 1220648-2024-22538 | 20639088 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639088&pc=PYX | 04/06/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 04/06/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1438552 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  | White | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella rp for mechanical circulatory support. while on support, the automated impella controller (aic) was fully charged and was unplugged from the power source and caused the aic to shut down while the pump was still connected. no clinical details regarding resolution were noted. no patient harm was reported. | [5 empty] | The investigation has been completed. the console logs from the reported day of event are consistent with the complaint and showed an alarm occurring multiple times and each time right after the pump was started. the logs also revealed that a second pump was used as a replacement with similar results. the console was returned and was extensively tested but no issues were observed and the console operated within specification. based on available information regarding use of the console as well as both pumps, it' has been determined that there was another console used (and associated complaint (b)(4)) that shared the failure mode. the cause of the failure mode was use related. this report is being filed as part of a retrospective review of historical records.
15093 | 1220648-2024-22539 | 20639146 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639146&pc=OZD | 03/30/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 03/31/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1548488 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a not charging issue. the hospital staff noted the aic was not charging when plugged in. they checked to ensure the battery switch was "on" and tried once flipping to "off" and then back on but no changes. they tried different surge protected outlets but no changes. the controller battery died and was unable to power on. no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | The investigation for the reported issue has been completed. clinical review stated that the hospital staff noted the automated impella controller (aic) was not charging when plugged in. they checked to ensure the battery switch was "on" and tried once flipping to "off" and then back on but no changes. different surge protected outlets were tried but no changes. the controller battery died and was unable to power on. the console was returned for investigation. the logs from the day of the complaint showed a battery level low alarm and alternating current (ac) power disconnected alarm. battery level low alarm was issued upon boot up (29% battery level). the console was plugged into the ac but it did not resolve the ac power disconnected alarm (ac power was not recognized). output of ac/dc power supply was 0v. the ac fuses were intact. power supply only had an 8 volt ac input. it was determined that the ac power cord was defective. the blue neutral wire was discontinuous throughout the power cord. once the power cord was replaced, the aic charged the batteries when connected to ac power. the aic battery charging issues were due to a defective ac power cord. this report is being filed as part of a retrospective review of historical records.
15094 | 1220648-2024-22540 | 20639228 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639228&pc=OZD | 04/09/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 04/09/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1075486 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the console would not recognize the catheter. they originally used an automated impella controller and the system would not recognize the pump. there was troubleshooting including removing the white plug and reinstalling, turning the console off and rebooting. it was noted that the team ultimately used a different console which was reported to have worked well. there was no reported harm or injury to the patient. | [5 empty] | The investigation for the reported issue has been completed. clinical review stated that the clinical staff originally used a console, but it would not recognize the pump. for troubleshooting, they had attempted removing the white plug and reinstalling, turning console off and rebooting. the team ended up grabbing a different console which worked. the console was returned for investigation. the logs revealed that while in the field, the console failed to detect a pump during case start. the logs showed that the console was able to detect a pump connection but failed to read pump eeprom. the console booted-up normally, but the optical bench state machine (ossiopsens) got stuck in post_start_state and did not reach post_ok_state. this prevented the software from reading pump eeprom. the cause of the failure to read the pump eeprom was due to bugs or defects in software (excludes firmware). this report is being filed as part of a retrospective review of historical records.
15095 | 1220648-2024-22541 | 20639272 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639272&pc=OZD | 04/10/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 04/13/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1142490 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported an automated impella controller console glitch. the patient had a second impella cp device inserted and it appeared that the console would not recognize the pump. it stayed on the screen. the console was powered off and the case was restarted with the same error. the console was changed and the pump insertion went smoothly. there was no reported harm or injury to the patient. | [5 empty] | The investigation for the reported issue has been completed. while in the field, the console failed to detect a pump during case start. the logs showed that the console was able to detect the pump connection but failed to read pump eeprom. in the logs, it is seen that during console boot-up, the optical bench state machine (ossiopsens) got stuck in post_start_state and did not reach post_ok_state. this prevents the software from reading pump eeprom. the root cause is bugs or defects in the software (excluding firmware). this report is being filed as part of a retrospective review of historical records.
15096 | 1220648-2024-22542 | 20639541 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639541&pc=OZD | 11/14/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 11/14/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1302709 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the automated impella controller (aic) produced a battery failure alarm. the console was replaced, and the procedure continued normally. no patient harm was reported. | [5 empty] | The device was returned and evaluated, and the investigation for the reported battery failure (red alarm) has been completed. long term log analysis of the battery data revealed that beginning before the complaint date of occurrence, cell 4 voltage of battery 1 had been declining for an extended period of time. it was confirmed that the persistent battery failure (transport connection blown/missing) was not due to a disengaged battery switch. when the console was booted for the first time in the laboratory, the console alarms were consistent with what was seen in the field. the battery failure alarm was due to a defective battery 1 (decline of individual cell voltage). the root cause of the battery cell failure was undetermined. this report is being filed as part of a retrospective review of historical records.
15097 | 1220648-2024-22543 | 20639542 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639542&pc=OZD | 10/02/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 10/06/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1126092 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The biomedical engineer reported that the automated impella controller (aic) experienced a controller error alarm. there was no patient harm reported. | [5 empty] | E1: name of the initial reporter is unknown. the investigation into the controller error has been completed. the impella automated controller was returned for investigation. the data log reports indicated that the controller error occurred on 10/6/22, however the issue occurs as early as 10/2/22. upon boot up, ic2163 alarms for a controller error for loss of communication with the keyboard pca. the returned console interior connections were verified. the 5-inch ribbon cable from the keyboard to the 4-in-1 was visually inspected with no issues observed. the console was powered on and the alarm was not present/ reproduced. the console was then powered cycled ~25 times before the error occurred again. the controller error is very intermittent between boots. normal boots were left on, and the keyboard was tested extensively with no issues occurring. during this error state, the soft keys on the keypad were unresponsive. the cause of the aic issue was a defective keyboard display. this report is being filed as part of a retrospective review of historical records.
15098 | 1220648-2024-22544 | 20639552 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639552&pc=OZD | 09/20/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 09/20/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1300700 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Nurse | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Use of Device Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The nurse reported that the automated impella controller (aic) would not allow entry of the sodium bicarbonate concentrations as 12.5 milliequivalent (meq) during the purge fluid change as it was not an available option in the selections. the nurse felt that the pharmacy was mixing 12.5 meq in a 500 cubic centimeter (cc) bag of dextrose 5% in water and expressed if they were to enter it as 25 meq in 1,000cc, proper notification of low purge volume would not occur. there was no patient harm reported. | [5 empty] | The investigation into the purge issue has been completed. the impella automated controller was not received from the customer and therefore, an evaluation of the device was not possible. the data analysis of the logs from the reported day of event are consistent with the complaint and show that purge fluid bag change was conducted daily, between 2022-09-13 and 2022-10-03. on 6 occasions (including 2022-09-20), bicarbonate concentration was selected as ¿25¿. as per software (sw) design (jama: aic151-srs-983 / gid-srs-516010) for aic_v11.1 and higher, bicarbonate concentration list of the purge fluid entry table only contains 3 values: 0, 25, 50. there was no product malfunction (hardware or software). customer feedback was communicated to the sw team for possible sw enhancement. there was no issue found during the case. the device functioned/operated to specification. this report is being filed as part of a retrospective review of historical records.
15099 | 1220648-2024-22545 | 20639589 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639589&pc=OZD | 02/25/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 02/25/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1136925 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the pump was started, and an optical sensor failure alarm occurred with no placement signal waveform. the automatic impella controller (aic) was switched, and the issue was resolved. no harm or injury noted. | [5 empty] | The investigation for the reported optical signal issue has been completed. the device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the cause of the aic issue was likely due to a work instruction / technician error. this report is being filed as part of a retrospective review of historical records.
15100 | 1220648-2024-22546 | 20639599 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639599&pc=OZD | 02/23/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 02/23/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | ABIOMED, INC. |  |  | US |  | 1216409 |  | 0042-0000-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced a purge error upon startup. the aic was replaced with a backup console, and there was no reported patient involvement. | [5 empty] | D8/9: device returned to field service, date unknown the device was returned and evaluated, and the investigation for the reported purge issue has been completed. aic logs revealed the console issued a controller failure (purge pressure sensor defective). the purge unit driver was the root cause of the controller failure. fs replaced the entire purge assembly because of inability to just replace pud (only assembly was available). the purge unit driver was the root cause of the controller failure. this report is being filed as part of a retrospective review of historical records.
15101 | 1220648-2024-22547 | 20639653 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639653&pc=OZD | 02/11/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 02/11/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1168278 |  |  |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Fell | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The biomedical engineer reported that the automated impella controller (aic) fell and broke. a loaner was sent to the customer to initiate return of this device. there was no patient harm reported. | [5 empty] | E1: name of the initial reporter is unknown. the investigation into the damage has been completed. the impella automated controller was returned for investigation. the returned console visual inspection found cracks and damages on multiple places. the front and rear housing were broken, part of the console enclosure was missing/ broken. the broken housing also led to deformation of the liquid crystal display (lcd) mount and 4 in1 board. on the front side of the console, the purge unit was also found broken. additionally, the console was found with an intermittent flickering lcd screen. the root cause of the console damage was unable to be determined, but given the result of console condition, it was likely due to the drop. this report is being filed as part of a retrospective review of historical records.
15102 | 1220648-2024-22548 | 20639677 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639677&pc=OZD | 02/16/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 02/16/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1218087 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Fell | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) was dropped during transfer of the patient. there was no patient harm reported. | [5 empty] | E1: name of the initial reporter is unknown. the investigation into the damage has been completed. the impella automated controller was returned for investigation. the returned console visual inspection confirmed the damage. the blue pump receptacle was damaged because the console was dropped. the interior of the console was inspected for damage, and there was no damage found. the cause of the issue was use related. this report is being filed as part of a retrospective review of historical records.
15103 | 1220648-2024-22549 | 20639698 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639698&pc=OZD | 12/10/2020 | 11/08/2024 |  | 11/08/2024 |  |  | 12/10/2020 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1126566 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced a purge door that would not close on an aic. a new console was used without any issues. no patient was involved. | [5 empty] | The investigation has been completed. the console (ic2190) was sent back for further investigation. console purge door was repeatedly tested empty and with a purge cassette and purge disc in place where no issue with the door was observed. door latch mechanism was inspected and no dextrose or damage was observed. the reported issue with the purge door unable to remain closed could not be reproduced. this report is being filed as part of a retrospective review of historical records.
15104 | 1220648-2024-22550 | 20639703 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639703&pc=OZD | 02/14/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 02/14/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | ABIOMED, INC. |  |  | US |  | 1413258 |  | 0042-0000-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient of unknown race/ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. while preparing for a case, the automated impella controller (aic) became stuck at the "connect the 2 gray cables" step and could not continue with setup. the resolution is unknown, but there was no reported patient harm. | [5 empty] | D9/9: device returned to field service, date unknown. the device was returned and evaluated, and the investigation for the reported optical signal issue has been completed. data analysis: imc logs confirmed the reported that the failure mode occurred during case start step 3. according to clinical staff, impella cp was utilized for case but imc logs from the complaint date revealed that console did not recognize the pump being connected. device analysis: test impella cp pump was utilized but the failure mode was not reproduced. console was opened, cable connection between blue receptacle and impellatronic board was confirmed to be properly seated. no abnormality was found within the console. the intermittent pump detection issue was likely because of malfunction of impellatronic board, after replacement of impellatronic, console passed annual preventive maintenance tests and was sent back to customer. the case start issue/ frozen screen issue was likely due to pump detection problem caused by defective impellatronic board, the exact root cause of defective impellatronic was unable to be determined. this report is being filed as part of a retrospective review of historical records.
15105 | 1220648-2024-22551 | 20639715 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639715&pc=OZD | 02/14/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 04/14/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, JP |  |  | ABIOMED, INC. |  |  | US |  | 1439042 |  | 0042-0000-JP |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an unknown primary indication was implanted with an impella for mechanical circulatory support. during support, the automated impella controller (aic) had a power failure and auditory beeping. there was no resolution specified, but there was no reported patient harm. | [5 empty] | D9/9: device returned to field service, date unknown. the device was returned and evaluated, and the investigation for the reported a/c power issue has been completed. data analysis: imc logs from the complaint occurred date was consistent with the complaint content. console was booted up successfully, and ac power supply was confirmed to be disconnected before power down the console. device analysis: console was shipped back to field service (fs) tokyo for investigation. fs team reproduced the failure mode and recorded as videos. according to fs team and the videos, the beep sound lasted for 22 minutes. at the moment ac power was unplugged, led indicator on console was showing green, indicating the console was not fully shut down for few seconds (yet with dark screen and dimmed softkeys). from the second video, console indicator led dimmed with ac unplugged and batteries engaged. it was speculated that both batteries depleted without ac power (ltlogs unavailable due to short operation time around the complaint date). fs team temporarily replaced the pbm and performed pm without issue. this report is being filed as part of a retrospective review of historical records.
15106 | 1220648-2024-22552 | 20639717 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639717&pc=OZD | 05/23/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 05/23/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER EU |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER EU |  |  | ABIOMED, INC. |  |  | US |  | 1622563 |  | 0042-0010-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Misassembled | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) had an object moving inside the controller. it was reported that the customer received this brand new console and realized during a quick inspection that something was rolling inside the console when he tilted it from side to side. there was no patient harm reported. | [5 empty] | E1: name of the initial reporter is unknown . d9: the date of the device return for evaluation is unknown. the investigation into the damage has been completed. the impella automated controller was not received from the customer. the field service engineer (fse)opened the console and found two screws lying inside the housing. the engineer then checked the controller for missing screws and saw that one screw was missing in the speaker plate. the engineer tried to reinstall the screw, but it was not possible, because the brass thread in the housing was damaged assumedly due to overtightening. the engineer then checked whether the other screws are tight enough and realized that a screw was also over cranked but had not fallen off yet. during the visual inspection the fse realized that there was a cut in the insulation of the vga-cable. the fse replaced the screws, rear housing, and cable then performed preventive maintenance. the cause of the aic issue was a loose fastener as a result of the manufacturing process. this report is being filed as part of a retrospective review of historical records.
15107 | 1220648-2024-22553 | 20639718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639718&pc=OZD | 02/08/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 02/08/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, EU |  |  | ABIOMED, INC. |  |  | US |  | 1308460 |  | 0042-0010-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 37 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient of unknown race/ethnicity with an unknown primary indication was implanted with an impella device for mechanical circulatory support. during use, the automated impella controller (aic) had a sudden console failure. the console was replaced, and there was no reported patient harm. | [5 empty] | D8/9: device returned field service, date unknown. the device was returned and evaluated, and the investigation for the reported controller error (red alarm) has been completed. data analysis: imc logs revealed there was a loss of communication to the optical bench given a controller error alarm (opm comm. loss - event set #1043) was triggered. this presented issues getting through the case start procedure. the user attempted to disconnect the pump, but it was not recognized. an emergency shutdown had to be performed to turn the console off. data analysis: the console was tested thoroughly at field service lab, but the failure mode could not be reproduced. there were no optical bench-related alarms triggered during more than 24 hours of testing with an optical pump simulator and purge. the root cause of the controller error (opm comm. loss) could not be determined because the failure mode was not reproduced during testing. this report is being filed as part of a retrospective review of historical records.
15108 | 1220648-2024-22554 | 20639730 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639730&pc=OZD | 05/03/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 05/03/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER EU |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER EU |  |  | ABIOMED, INC. |  |  | US |  | 1571192 |  | 0042-0040-EU |  | Yes |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Misassembled | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) had an object moving inside the controller. it was reported that the customer found two screws in the transport case. there was no patient harm reported. | [5 empty] | E1, e3, e4: name of the initial reporter and occupation is unknown. d9: the date of the devices return for evaluation is unknown. the investigation into the damage has been completed. the impella automated controller was not received from the customer. the field service engineer (fse) observed that one of the rear housing screws was taped to the housing after being found in the shipping box when being returned from preventive maintenance. the console was opened for a physical inspection, it was observed that the fan assembly was missing a screw. the missing screws were found within the console and fastened to their respective spots. the root cause of the rattling was due to loose fasteners as a result of the manufacturing process. this report is being filed as part of a retrospective review of historical records.
15109 | 1220648-2024-22555 | 20639757 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639757&pc=OZD | 01/17/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 01/17/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1559446 |  | 0042-4201 |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) exhibited battery failure. a loaner device was sent to the customer. no indication of patient involvement, harm, or injury. | [5 empty] | D9: the date of the return of the device for evaluation is unknown. the console (aic) was returned for evaluation. upon evaluation of the console, the reported failure mode was duplicated on an engineering bench. the battery 1 lcd indicator was blank (fully charged). however, battery 1 was replaced to resolve the failure alarm. therefore, the root cause of battery failure to a defective battery 1 and the root cause of the defective battery 1 was due to single cell failure. during evaluation, the engineer also found that the optical bench was defective. the root cause of this issue is unable to be determined. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15110 | 1220648-2024-22556 | 20639772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639772&pc=OZD | 08/27/2024 | 11/18/2024 |  | 11/08/2024 |  |  | 08/27/2024 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025525327 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported that impella was successfully removed, patient tolerated well. unfortunately, it was noted that the patient had a thrombus in the right femoral artery [rfa] down to the superficial femoral artery [sfa] after impella was removed. staff noted that the patient may have peripheral vascular disease [pvd] as well. patient then consulted with vascular surgery and a plan was made to perform a cut-down and a thrombectomy. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details.
15111 | 1220648-2024-22557 | 20639777 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639777&pc=OZD | 09/04/2024 | 09/26/2025 |  | 11/08/2024 |  |  | 09/04/2024 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025508978 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 21 YR |  | Female |  |  |  | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the patient on impella 5.5 support was taken to the operation room to remove a hematoma beside the impella site. the bleeding was resolved. it was further reported that the placement signal was not showing on the aic. the placement control and alarms were deactivated. the staff will monitor the motor current. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | Access site bleeding - major: the clinical states that there was a hematoma around the access site. surgery stopped the bleeding. due to lack of clinical details provided, the root cause of the access site bleeding - major cannot be determined. optical signal issue: due to lack of product returned and no data logs provided, the root cause of the optical signal issue cannot be determined. | H6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-22557. d10 concomitant medical products and therapy dates updated. e4 should have been left blank on the initial report. g1 reporting contact email updated.
15112 | 1220648-2024-22559 | 20639781 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639781&pc=OZD | 10/15/2024 | 01/24/2025 |  | 11/08/2024 |  |  | 10/15/2024 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1300887 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) had a battery communication failure even while plugged into the ac power. there was no patient involvement. | [5 empty] | The investigation into the automated impella controller (aic) has been completed. the aic was returned for investigation. after reviewing the console logs, the reported battery communication failure issue with ic5318 was confirmed. there was an instance of battery 2 comm. failure alarm was observed in the imc logs. the aic was tested after arriving in field service. the reported battery communication failure issue was able to be reproduced. the root cause of this issue was an internal cell imbalance in battery 2.
15113 | 1220648-2024-22558 | 20639783 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639783&pc=OZD | 10/08/2024 | 12/10/2024 |  | 11/08/2024 |  |  | 10/11/2024 | 11/08/2024 |  |  | 05/31/2025 |  | 11/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024351998 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Disability | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and the patient had limb ischemia. the left lower extremity was red above the knee extending to the ankle. the patient required an above knee amputation. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email.
15114 | 1220648-2024-22560 | 20639787 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639787&pc=OZD | 10/15/2024 | 12/24/2024 |  | 11/08/2024 |  |  | 10/15/2024 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025537358 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella cp support. while on support, impella puncture site bled heavily. a hematoma has formed on the groin. femstop was placed and blood products were given. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. corrections to the initial mfr report: d4 updated udi number g1 mdr reporting contact email. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
15115 | 1220648-2024-22561 | 20639792 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639792&pc=OZD | 10/11/2024 | 12/12/2024 |  | 11/08/2024 |  |  | 10/15/2024 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025480298 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Thecomplainant reported that after the impella cp insertion, there was kink in the cannula. staff attempted to reposition the pump to resolve the issue to no avail. the pump was explanted. | [5 empty] | The investigation into product damage (cannula kink) has been completed. the impella device was not returned for evaluation. the cause of the product damage (cannula kink) most likely was use related due to improper technique. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-22561. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 & cp impella cp with smartassist system & impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: cautions. ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
15116 | 1220648-2024-22562 | 20639808 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639808&pc=OZD | 10/15/2024 | 12/11/2024 |  | 11/08/2024 |  |  | 10/15/2024 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532306 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The us complainant had placed a cp pump to support a patient admitted in an acute myocardial infarction / cardiogenic shock. the pump did stop after purge alarms and ineffective troubleshooting. the troubleshooting was for 5 minutes, and the purge cassette and automated impella controller were replaced. then eventually the pump itself was replaced. support proceeded on the 2nd pump. the patient was transferred to a tertiary center and survived the event. | [5 empty] | The investigation for the pump stop has been completed. the pump was returned and evaluated. biomaterial and blood was found in and around the purge gap and impeller. the high purge pressure reproduced in the lab while the purge was running. there was blood ingress found in the purge line in the catheter near the motor housing. when the catheter was cut at this location, the high purge pressure resolved. log review showed an increase in motor current, pump flow, and purge pressure before "purge pressure high" alarms triggered. "impella stopped - motor current high" alarms followed shortly after. the root cause of the pump stop was blood ingress. the cause of the ingress is unknown. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure : which triggers the ¿purge pressure high¿ alarm message: could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
15117 | 1220648-2024-22565 | 20639816 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639816&pc=OZD | 10/15/2024 | 12/11/2024 |  | 11/08/2024 |  |  | 10/15/2024 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531515 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The us complainant had a cp support ongoing for a patient admitted in acute myocardial infarction / cardiogenic shock and found to have coronary artery disease. the patient on day 9 had the team deliver high dose heparin to resolve thrombus. the pump remained on and discussions were ongoing for the possible escalation to the 5.5 pump. the cp was explanted after the doctor saw the patient¿s ejection fraction had recovered. the patient survived the event. | [5 empty] | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can led to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can led to systemic embolization, serious injury, or death."
15118 | 1220648-2024-22563 | 20639819 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639819&pc=OZD | 10/15/2024 | 11/16/2024 |  | 11/08/2024 |  |  | 10/15/2024 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1156148 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported that the automated impella controller (aic) is not detecting purge disk. no patient harm has been reported. | [5 empty] | The investigation for the console "not detecting purge disk" has been completed. data logs were reviewed finding the reported issue that purge disc couldn't be detected was confirmed. there were multiple purge disk not detected alarms on that date. the console was returned for evaluation. the purge disc flag was found to be broken (missing at blue retainer). the root cause of the issue was a broken purge disc flag. | The impella device was returned and evaluated. the root cause of the issue was a broken purge disc flag.
15119 | 1220648-2024-22564 | 20639826 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639826&pc=PYX | 10/15/2024 | 12/04/2024 |  | 11/08/2024 |  |  | 10/15/2024 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025502620 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that after dialysis was started on the 74-year-old male patient on impella rp support, the impella flows started to decrease. the rp was unable to flow above 2.4l. heart team decided to explant, as the patient right ventricle function had improved. there was no patient harm reported. | [5 empty] | The investigation for the low pump flow has been completed. product and data logs were not returned for review. the root cause of low flow was unable to be determined as limited clinical details were provided and no product was returned for review. added h6 impact code 4627. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
15120 | 1220648-2024-22570 | 20639829 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639829&pc=OZD | 01/20/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 01/20/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1561819 |  | 0042-0040-JP |  | Yes |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) exhibited a battery failure alarm was issued during start-up inspection. the battery did not exceed 50% even though the power cable was already plugged into an outlet. no patient involvement. | [5 empty] | D9: the date of the device return is unknown. e1: the name of the initial reporter and occupation is unknown. the console (aic) was returned for evaluation. upon evaluation of the console logs, the reported day of event are consistent with complaint and show battery failure alarm upon boot-up. the log battery data shows cell imbalance of two cells within battery 1: cell 3 and cell 2 were gradually declining until a rapid failure of cell 2 that resulted in battery 1 internal shutdown. therefore, the root cause was a defective battery. the console's batteries were replaced, and a functional check was conducted. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15121 | 1220648-2024-22566 | 20639835 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639835&pc=OZD | 01/20/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 01/20/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1557944 |  | 0042-0040-JP |  | Yes |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report |  | Other | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) exhibited a battery failure alarm, and a battery status unknown alarm occurred. the battery level display at that time was 50%. the console was replaced. no report of patient harm or injury. | [5 empty] | D9: the date of the devices return for evaluation is unknown. e1, e3, e4: the name of the initial reporter and occupation is unknown. the console (aic) was returned for evaluation. upon evaluation of the console, the reported issue of battery failure was duplicated on an engineering bench. the battery 1 package voltage was measured to be 3.8 v, rather than the expected 16 v and the battery lcd indicator was blank (fully discharged). battery 1 was replaced to resolve the failure alarm. therefore, the root cause of the battery failure alarm was due to a defective battery and the root cause of the defective battery was due to single cell failure. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15122 | 1220648-2024-22568 | 20639837 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639837&pc=OZD | 08/16/2024 | 03/10/2025 |  | 11/08/2024 |  |  | 08/16/2024 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025479722 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | Intraoperative Cardiac Injury | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported that impella was loaded on the wire and advanced into the graft through the anastomosis. as device was being advanced, the graft separated from the artery, pulling a large piece of artery away with it. the tissue was believed to be very friable and of extremely poor quality. the physician immediately placed clamps proximately and distally to control the vessel, and the vessel was repaired with a 7 mm hemashield platinum graft. due to being out of access locations, the decision was made to optimize the patient on cp support. the patient survived the event. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, the tissue was believed to be very friable and extremely poor quality. the root cause of the vascular damage issue was most likely patient condition related. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
15123 | 1220648-2024-22569 | 20639841 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639841&pc=OZD | 10/15/2024 | 12/16/2024 |  | 11/08/2024 |  |  | 10/15/2024 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025573844 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, heparin was held due to gastrointestinal bleed. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation into vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly on manufacturer device report 1220648-2024-22569. e4 should have been left blank. g1 reporting contact email. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations "assess access site for bleeding and hematoma." section: entering cardiac output: use of the repositioning sheath and peel-away introducer: "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site." section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)."
15124 | 1220648-2024-22567 | 20639845 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639845&pc=OZD | 10/18/2024 | 09/30/2025 |  | 11/08/2024 |  |  | 10/15/2024 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025490626 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 7 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Pumping Stopped | Device Sensing Problem | [5 empty] | Insufficient Information | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | Required Intervention |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported wrong red alarms of impella cp in ventricle were occurring. the impella position was verified, and the cp was in good position. additionally, the impella stopped suddenly, and it would not restart. the patient was taken to the operation room to remove the impella. it was further reported that the patient expired after the pump was removed. the relationship between the impella and cause of death is unknown. | Additional information received that the patient did not expire upon removal on the impella. the patient was successfully bridged to a left ventricular assist device and right ventricular assist device. | The investigation uncovered optical signal issues with the impella cp. |  |  |  | The investigation for the pump stop has been completed. the pump was not returned for evaluation. log review showed a motor current spike before the pump stop and "impella stopped - motor current high" alarms. there was an increase in purge flow before the pump stop. the impella cp was used beyond the 8 day design life. the root cause of the pump stop was not determined since product was not returned and insufficient clinical detail was provided. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high-risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The product was returned so an investigation of the product returned was completed. the returned pump had cracks in the sidearm filter. when running the pump, the purge pressure matched the range in the field (350-400 mmhg) and the purge flow was occasionally up near 30 millitlieters per hour, which correlates with the sidearm filter leak. log review showed a motor current spike before the pump stop and "impella stopped - motor current high" alarms. there was an increase in purge flow before the pump stop. no biomaterial was found on the returned pump and no blood ingress was found in the purge line. however, with the increased purge flow leading to the high motor current pump stop. there was most likely biomaterial/blood entering the purge gap, leading to the pump stop. the high purge flow was most likely caused by the sidearm filter leak from the crack that may have been due to ipa interaction. the root cause of the pump stop was most likely the damaged sidearm filter. the impella cp was used beyond the 8 day design life. d.9 date device was returned/information was added. b.1 revised to product due to new information received since initial manufacturer device report number 1220648-2024-22567 was submitted. b.2 revised to reflect blank due to new information received since initial manufacturer device report number 1220648-2024-22567 was submitted. b.3 revised date of event as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-22567. b.5 information was added since the initial manufacturer device report number 1220648-2024-22567 was submitted. e.4 should have been left blank on the initial manufacturer device report number 1220648-2024-22567 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.3 date of awareness on the initial manufacturer device report number 1220648-2024-22567 should of been 18-oct-2024. h.1 revised selection due to new information received since initial manufacturer device report number 1220648-2024-22567 was submitted. h.6 revised health effect clinical code due to new information received since initial manufacturer device report number 1220648-2024-22567 was submitted. a new code was added. health effect impact code 1802 should not have been on the initial manufacturer device report number 1220648-2024-22567 per information received. added new type of investigation code due to investigation being completed on the returned product. investigation conclusions since the first follow up manufacturer device report number 1220648-2024-22567 was submitted due to the investigation being completed on the returned product. h.10 investigation results were changed due to returned product being investigated. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the product was returned and evaluated. no abnormalities were found with the 5s parameters. the pump passed the laser continuity test, no biomaterial was found, and no abnormalities were found while running the pump for 15 minutes. the patient was 7 years old and likely small in size. the 5s parameters were in spec. the root cause of the optical signal issue was most likely patient condition related. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-22567. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated.
15125 | 1220648-2024-22571 | 20639854 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639854&pc=OZD | 04/14/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 04/14/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1103519 |  | 0042-0000-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during preventive maintenance (pm), the automated impella controller (aic) experienced a battery alarm going off. no patient was involved. | [5 empty] | The investigation for the reported issue has been completed. clinical review stated that while in for preventative maintenance, the console triggered a battery failure alarm during the battery test. the console was set aside for further investigation. the logs showed a battery failure alarm. the console was investigated and as part of troubleshooting, the battery connectors were switched at the power battery manager (pbm). the failure mode followed the pbm¿s channel 1. the pbm was inspected and there was visible damage to component q23. the battery failure alarm was due to a defective pbm. the cause of the defective pbm was undetermined. this report is being filed as part of a retrospective review of historical records.
15126 | 1220648-2024-22572 | 20639877 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639877&pc=OZD | 01/14/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 01/14/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1173718 |  | 0042-0000-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) exhibited battery failure. a replacement was sent to the customer. no report of injury or harm. | [5 empty] | D9: the date of the devices return for evaluation is unknown. the console (aic) was returned for evaluation. upon evaluation of the console, the reported failure mode battery failure was reproduced on an engineering bench. the battery 2 lcd indicator was blank (fully discharged). battery 2 was replaced to resolve the failure alarm, and discharging/ charging was validated. therefore, the root cause of battery failure was due to a defective battery 2. the root cause of the defective battery 2 was due to single cell failure. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15127 | 1220648-2024-22574 | 20639905 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639905&pc=OZD | 01/10/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 01/10/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1164194 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced intermittent controller error alarm. a loaner was sent to the customer. no patient harm reported. | [5 empty] | The device was returned for evaluation, and the failure mode could not be reproduced. the root cause is unable to be determined. this report is being filed as part of a retrospective review of historical records.
15128 | 1220648-2024-22573 | 20639919 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639919&pc=OZD | 02/21/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 02/21/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1320577 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that "purge fluid not detected" message was displayed after pump was primed. exchanging the purge cassette did not resolve the issue. a replacement console was successfully used. there was no reported harm or injury to the patient. | [5 empty] | The investigation for the reported aic - purge issue has been completed. console logs from the reported day of event are mostly consistent with complaint and show multiple "purge fluid not detected" messages, however, contrary to the complaint intake description, there were multiple purge issues (e.g. "piston blocked", or "home not found") from the start, and the pump was never fully primed. console imc logs as well as rt logs indicate that purge cassette piston was not advancing. inspection of the console revealed glucose contamination between purge motor gasket and purge insert, which resulted in motor shaft being immobilized. after the affected area was cleaned, the motor shaft was able to spin, and issue was resolved. this report is being filed as part of a retrospective review of historical records.
15129 | 1220648-2024-22575 | 20639935 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639935&pc=OZD | 01/17/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 01/17/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1300416 |  | 0042-0000-US |  | No answer provided |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. the complainant experienced a purge cassette (pc) failure, and the aic was constantly alarming for air in purge despite troubleshooting. the staff changed the purge cassette, and the alarm subsided. she saved faulty purge cassette, and support proceeded without any reported harm or injury. | [5 empty] | The controller was not returned for evaluation, the customer returned the purge cassette and data stick. the investigation into the purge issue has been completed. data logs showed the sudden occurrence of the "air in purge" alarm several hours into support. multiple rapid de-airing attempts were made to resolve the issue. the purge flow ticks were seen to be very rapid and erratic at this time. no damage or abnormalities were found on the returned purge cassette. when connecting the purge cassette to the console, the "purge system open" and "purge pressure low" alarms triggered briefly before eventually resolving. purge metrics were within specification and purge fluid exited the purge gap. therefore, it was determined that the cause of the priming issue was use related. this report is being filed as part of a retrospective review of historical records.
15130 | 1220648-2024-22579 | 20639979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639979&pc=OZD | 05/25/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 05/25/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1402264 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) the console had a broken purge clip. there was no report of patient harm or injury. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, the purge assembly area confirmed a broken blue retainer. before returning to customer, the purge disc retainer was replaced, the purge assembly was cleaned, and a functional check was performed. the failure mode will be monitored and trended. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15131 | 1220648-2024-22576 | 20639985 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639985&pc=OZD | 02/05/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 02/05/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1136980 |  | 0042-0000-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. during startup, the automated impella controller (aic) produced an optical sensor failure with no placement signal or flow on the screen. the console was replaced, and there was no reported patient harm. | [5 empty] | D8/9: device returned to field service, date unknown. the device was returned and evaluated, and the investigation for the reported optical signal issue has been completed. data analysis: device logs confirmed the reported failure mode. despite pump was connected, aic was unable to establish optical connection with pump and alarming for event. device analysis: the failure mode was neither reproduced during testing on engineer bench nor during clinical use for the past few months. based on the data analysis from complaint occurred date, console manifested optical related failures. console was initially unable to make proper connection with pump despite pump was plugged all the way in according to clinical staff. after a self-reset of optical bench, a well-known optical pattern failure occurred, causing all optical placement signal spiked instantaneously. the cause of the patterned optical signal issue was not determined. this report is being filed as part of a retrospective review of historical records.
15132 | 1220648-2024-22578 | 20639992 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20639992&pc=OZD | 01/27/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 02/03/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1178827 |  | 0042-0000-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella for mechanical circulatory support. when the purge cassette was attached, the automated impella controller (aic) read "disc not detected," and this issue persisted with a second purge cassette. the aic was replaced with a new console, and there was no reported patient harm. | [5 empty] | The device was returned for service, and a functional check was performed. the information about the purge disk issue was not reported prior to the device return, therefore the device was not evaluated for the purge disk issue. console logs from the reported day of event are consistent with complaint and show purge disk being intermittently detected, as evidenced by repeated sequence of alarm #402 (purge disk not detected) being set then cleared. purge cassette was correctly detected via rfid mechanism. the symptoms observed in the field and confirmed in logs are consistent with either purge disk detect flag being broken or purge retainer being cracked. however, there are no records of either the console failing any functional check testing (at least two such tests were performed since then) or having any parts replaced. it is most likely that defective part was replaced without being documented. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15133 | 1220648-2024-22577 | 20640002 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640002&pc=OZD | 01/14/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 01/14/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1376689 |  | 0042-0000-JP |  | Yes |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Female |  |  | Asian | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the automated impella controller (aic) exhibited battery alarm failure. the device was replaced. no report of patient injury or harm. | [5 empty] | D9: the date of the devices return for evaluation is unknown. the console (aic) was returned for evaluation. upon evaluation of the console, the reported failure mode was persistent battery failure (transport connection blown/missing) was not due to a disengaged battery switch. due to low/ flatlined voltage, battery 1 data is unavailable through ata or bq software. when visually inspecting the batteries and pbm, it was observed that one of the batteries statuses leds were completely blank. therefore, the root cause of the failure battery was due to a defective battery 1 (decline of individual cell voltage). the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15134 | 1220648-2024-22581 | 20640013 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640013&pc=OZD | 01/28/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 01/28/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1254076 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Alarm Not Visible | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) was producing inaudible alarms. the console was replaced with a loaner aic. there was no reported patient involvement, injury, and harm. | [5 empty] | The device was returned and evaluated, and the investigation for the reported inaudible buzzer/alarm has been completed. device analysis: the console was powered on with the battery switch disengaged to purposely trigger an alarm. the alarm was inaudible. the console was disassembled, and the interior connections were inspected. it was found that the speaker cable was unplugged. the cause of the failure mode was a loose connector. this report is being filed as part of a retrospective review of historical records.
15135 | 1220648-2024-22582 | 20640015 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640015&pc=OZD | 08/21/2024 | 07/18/2025 |  | 11/08/2024 |  |  | 08/29/2024 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025517781 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 55 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Flutter | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | Heart Block | Renal Impairment | [7 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complaint team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured hemolysis with a definite relationship to the impella device used: (b)(6) 2024 hgb 6.4, plt 45, pfh 120, ldh 2233. it was reported that the patient/event has not recovered/not resolved. the patient survived the event. | Additionally, the complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured systemic organ failure with a "definite" relationship to the impella device used: ¿(b)(6) 2024 intermittent complete heart block and a-flutter vs accelerated junctional; lle fasciotomy, removal of cp and placement of impella 5.5 (b)(6) 2024 lfts continue to elevate (ast 783, alt 157, tbili 7.4); (system organ failure) (b)(6) 2024 hgb 6.4, plt 45, pfh 120, ldh 2233 (hemolysis) (b)(6) 2024 de-escalation call for cshock late afternoon, however arterial cannula not pre-closed, plan for tomorrow however hypotensive through the night and now requiring 3 pressors, lactic acid 1.4 >2.2>2.9 (b)(6) 2024 patient likely positive for ttp d/t adamsts13 level and thrombocytopenia, plan to start daily plex for goal plt >150; cr slowly increasing, now 2.4 ((b)(6) 2024 cr 0.8...renal failure aki), gib (b)(6) 2024 ecmo turn down yesterday to ecmo flows ~ 3 lpm. impella increased to p6, gib, albumin, blood products pretty much daily, levo gtt, vaso gtt (b)(6) 2024 ecmo removed and femoral artery repair (not listed) completed (b)(6) 2024: extubated on 9/1. epi weaned off on (b)(6) 2024. vascular surgery met with patient and family regarding possible left aka (b)(6) 2024: left leg aka completed by vascular surgery, impella @ p8 ~ 4.8 l (b)(6) 2024: impella turned down to p5 ~ 3.5l flow, remains under oht eval. (b)(6) 2024: impella wean to p2 with echo and hemodynamics completed. cshock called for de-escalation, decision to remove impella 5.5. (b)(6) 2024: impella 5.5 removal tomorrow, colonoscopy on friday¿. the patient recovered with sequelae. | Additional information received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured atrial flutter and accelerated junctional rhythm with a "probable" relationship to the impella device used: ¿(b)(6) 2024 complete heart block and atrial flutter vs accelerated junctional, the patient recovered with no sequelae. |  |  |  | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. the cause of the hemolysis issue cannot be determined since the complaint device not returned for investigation and insufficient clinical data provided. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | No product was received for investigation, should any new information or investigation is received regarding thrombocytopenia, renal failure, or vascular damage - peripheral vessel a supplemental report will be submitted. b5 updated with additional information received from the clinical site h6 revised to reflect the additional failure mode received from the clinical site. g1 reporting contact email revised on manufacturer device report 1220648-2024-22582 in accordance with updated procedures. g2 report source; study was missing from manufacturer device report 1220648-2024-22582. | The investigation of hemolysis, thrombocytopenia, renal failure, vascular damage - peripheral vessel, and vf/vt/af/at has been completed. the impella device was not returned for evaluation. data log reviewed: no motor current (mc) spike observed outside of p-level changes. intermittent suction alarms and positioning alarms throughout support. no details on pump positioning or other troubleshooting in clinical. hemolysis: the cause of the hemolysis issue cannot be determined since the complaint device not returned for investigation and insufficient clinical data provided. thrombocytopenia: the cause of thrombocytopenia issue cannot be determined due to insufficient clinical details. renal failure: the cause of renal failure issue cannot be determined due to lack of clinical information. vascular damage - peripheral vessel: the cause of vascular damage - peripheral vessel issue cannot be determined due to lack of clinical information. vf/vt/af/at: the root cause of the heart block / atrial flutter / accelerated junctional rhythm was unable to be determined due to insufficient clinical details. b5 updated with additional information received from the clinical site h6 updated to include additional failure modes.
15136 | 1220648-2024-22584 | 20640019 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640019&pc=OZD | 12/23/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 12/23/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1565140 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | It was reported that there was a battery failure, attempts to troubleshoot by switching the on off battery button on the bottom but this was not successful, and there was no mention of patient harm or medical intervention. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the battery failure was due to a defective battery. the root cause of the defective battery was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
15137 | 1220648-2024-22583 | 20640030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640030&pc=OZD | 01/07/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 01/07/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1142488 |  | 0042-0000-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) exhibited a battery failure alarm and battery communication failure. battery #1 cell #2 voltage is too low. the battery set was replaced. no report of patient harm or injury. | [5 empty] | D9: the date of the devices return is unknown. the console (aic) was returned for evaluation. upon evaluation of the console, battery 1 of the console was confirmed for depletion with blank lcd screen. additionally, the serial number of the battery was found within the defective batch of batteries that presented single cell decline failure. the root cause of battery failure was unable to determine, but highly likely to be single cell decline failure mode. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15138 | 1220648-2024-22585 | 20640034 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640034&pc=OZD | 01/07/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 01/07/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1561819 |  | 0042-0040-JP |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Female |  |  | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the automated impella controller (aic) exhibited a console failure alarm. the device was replaced. no report of patient harm or injury. | [5 empty] | The date of the device return for evaluation is unknown. the console (aic) was returned for evaluation. upon inspection of the console, the known pattern failure, which is characterized by a temporary loss of optical data and triggering of a controller error alarm, has a low probability of reoccurrence. if needed, patient hemodynamics and impella position can be monitored with imaging. therefore, the root cause of the controller failure alarm (optical signal) was unable to be determined because the issue was unable to be reproduced. the optical bench was replaced a full functional check on the console and optical system was performed; before the console was returned to the customer. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15139 | 1220648-2024-22580 | 20640050 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640050&pc=OZD | 01/12/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 01/12/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1565544 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) exhibited battery failure. no further information is available. no report of patient involvement, injury, or harm. | [5 empty] | The console (aic) was not returned for evaluation. firmed the reported failure. there was a battery failure alarm and a battery 1 communication failure alarm due to low pack voltage. the failure mode will be monitored and trended. the data from the complaint date showed a cell voltage decline in only one of battery 1¿s cells which occurred as the battery failure alarm triggered. this report is being filed as part of a retrospective review of historical records.
15140 | 1220648-2024-22586 | 20640054 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640054&pc=OZD | 12/24/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 12/24/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, US |  |  | ABIOMED, INC. |  |  | US |  | 1177184 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the aic (automated impella controller) was not charging, multiple outlets used with power cord but the aic was not charging, another aic was used. there was no reported harm or injury to the patient. | [5 empty] | The investigation into the a/c power issue has been completed. the automated impella controller (aic) was returned for investigation. since the power-on, the console did not detect ac cord being plugged in. ltlog power data form the same time frame showed gradually decreasing state of charge and remaining capacity of both batteries, as well as negative current indicating that batteries are being discharged. when console was tested in danvers, it's been observed that the pbm was not receiving power from the power supply, but also there was no ac voltage on the input of power supply. inspection of the power entry module revealed blown 2 inline fuses on the ac input, internally compromised pem, and a compromised grounding wire between pem (loose wire) and the gnd stud on the housing (broken strands inside the wire). when the pem was temporarily by-passed to verify power supply operation, it's been discovered that he power supply was defective as well and did not generate 24v dc output. after all defective components were replaced, console operated within specification and pbm was able to charge batteries. it is most likely that the loose ground wire caused failure of pem, resulting in malfunction of power supply and blown fuses. this report is being filed as part of a retrospective review of historical records.
15141 | 1220648-2024-22587 | 20640082 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640082&pc=OZD | 04/09/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 04/13/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1084302 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on an automated impella controller (aic) for mechanical circulatory support. it was reported that the aic experienced a system error. the console was removed from service. no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | The investigation has been completed. the logs confirmed a controller error related to purge issues. the logs also showed that the case had numerous suctions alarms and then the staff changed the purge cassette after they received a purge pressure low alarm (which cleared in 20 seconds). they changed the purge cassette. a month later they received a purge system blocked alarm, which then cleared and then a controller alarm was set. throughout the case there were short instances of purge pressure high, followed by purge pressure low and in some instances even a purge system open alarm. these continue even after the purge cassette was changed during the case. this suggests bio-material buildup on the purge line at pump end, creating a blockage for the purge and eventually causing high purge pressure. no issue was found with the console. the console ran a purge and a pump simulator for over five days with no errors. the root cause was not determined. this report is being filed as part of a retrospective review of historical records.
15142 | 1220648-2024-22588 | 20640146 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640146&pc=OZD | 11/18/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 11/18/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED. INC. |  |  | US |  | 1228108 |  | 0042-0000-US |  | Yes |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report |  | Other | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | Key or Button Unresponsive/not Working | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during preventative maintenance the automated impella controller (aic) exhibited a controller error. the error showed for a few seconds then disappeared followed by a system forced reboot, when the console powered and booted back up as normal, attempts where then made to go through the console menu without success. no patient involvement. | [5 empty] | D9: the date of the device evaluation is unknown. e1, e3, e4: the initial reporter and occupation is unknown. the device was returned for evaluation. the cause of the rpb malfunction was a defective impellatronic board. the cause of the defective impellatronic board is unknown. the cause of the controller error yellow alarm was unable to be determined because the issue was unable to be reproduced.
15143 | 1220648-2024-22589 | 20640153 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640153&pc=OZD | 11/15/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 11/15/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1567887 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) was not holding battery. it only goes to 50% and has an error "battery failure". confirmed that the unit was plugged in to the ac. a loaner was sent to the customer. no report of patient harm or injury. | [5 empty] | E1: the name and occupation of the initial reporter is unknown. the investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the battery failure alarm was due to a defective battery (rapid decline in cell voltage). the exact cause of the defective battery was unknown. this report is being filed as part of a retrospective review of historical records.
15144 | 1220648-2024-22590 | 20640173 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640173&pc=OZD | 11/30/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 11/30/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1557944 |  | 0042-0040-JP |  | No |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | During software update, power up, there was a battery failure alarm that was triggered, log collection and investigation. one of the two batteries was dead. logs and surveyed data are uploaded to the target ci. there was no mention of patient involvement resulting in harm or medical intervention. | [5 empty] | Console logs from the reported day of event are consistent with complaint and show persistent battery failure alarms (#360) and battery 2 communication failure alarms (#353) upon every boot-up on that day. bat test result indicates battery 2 being internally disconnected due to cell imbalance (failed cell 3), which also disabled communication with battery. after batteries were replaced, console operated within specification and no alarms were present. this report is being filed as part of a retrospective review of historical records.
15145 | 1220648-2024-22591 | 20640208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640208&pc=OZD | 11/19/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 11/19/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1317756 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was on an automated impella controller (aic) for mechanical circulatory support. the complainant had a 5.5 support ongoing when the aic was swapped due to purge issues. there was no report of patient harm or injury. | [5 empty] | The investigation into the purge issue has been completed. the automated impella controller (aic) was returned for investigation. as stated in the disposable investigation, high purge pressure found during support with pump is likely due to the kink found on the catheter, and the blood ingress under the impeller. the controller failure that occurred during that support was likely due to the kinking on the catheter. out of an abundance of safety, the purge assembly will be replaced. before returning to customer, field service will be instructed to replace purge assembly. this report is being filed as part of a retrospective review of historical records.
15146 | 1220648-2024-22592 | 20640216 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640216&pc=OZD | 05/19/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 05/19/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1203421 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) that when opening the purge assembly door for the console, it was observed that the blue disc retainer was broken. there was no report of patient harm or injury. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, the purge assembly area confirmed a broken blue retainer. the purge disc retainer was replaced and a functional check was conducted. the failure mode will be monitored and trended. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15147 | 1220648-2024-22593 | 20640368 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640368&pc=OZD | 10/18/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 10/18/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1567887 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was placed on an impella cp for hemodynamic support. while on support, the automatic impella controller (aic) experienced a controller failure (red alarm). the issue was resolved by changing out the console. there was no reported harm or injury to the patient. | [5 empty] | The investigation for the reported aic - battery failure (red alarm) issues has been completed. the device has been returned for evaluation. clinical details were sufficient to determine the root cause. the reported issue was related to a defective dc power supply. this report is being filed as part of a retrospective review of historical records.
15148 | 1220648-2024-22594 | 20640400 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640400&pc=OZD | 04/10/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 04/13/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1142490 |  | 0042-0000-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | Connection Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) console had a glitch. the console got stuck at one screen once the pump connected and would not recognize that it was connected. the console was powered on and off and still the same error. the console was changed out and the pump was successfully inserted. no patient harm was reported. | [5 empty] | D.9 the exact date the device was returned for investigation is unknown. the investigation has been completed. during a case, the console displayed motor current high alarms. the pump was investigated for clots, which caused motor current high alarms, and eventually pump stop. according to the rep: "a clot was visualized in the impella's inlet cage, when removed." this failure is captured in 21-05634-1. the staff power cycled the console and tried to start the pump again, but the console failed to detect the pump. this failure is captured in 21-05607-1. logs confirmed the above narrative, and console did not fail. the motor current high alarms were due to the clot in the pump. the console was upgraded to remedy the 'pump not detected' failure mode. a functional check was also performed on the console. the automated impella controller alarms for high motor current were due to a clot found in the pump inlet cage. this report is being filed as part of a retrospective review of historical records.
15149 | 1220648-2024-22595 | 20640420 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640420&pc=OZD | 10/15/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 10/15/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2022067927 |  | 0042-0000 |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 41 YR |  | Female |  |  | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella for mechanical circulatory support. when the automated impella controller (aic) was started, it failed the self-check. another aic was used for the procedure, and no patient harm was reported. | [5 empty] | D8/9: device returned to field service, date unknown. the device was returned and evaluated, and the investigation for the reported syscall error has been completed. data analysis: aic logs from the complaint date show pbm communication issues. device analysis: the console was investigated by field service. the pbm was replaced to resolve the communication issues. the root cause of the system self-check failure was due to communication issues with a defective pbm. the exact root cause of the defective pbm was not determined. this report is being filed as part of a retrospective review of historical records.
15150 | 1220648-2024-22596 | 20640425 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640425&pc=OZD | 11/08/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 11/08/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1556637 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported the automated impella controller (aic) experienced a battery failure (red alarm), and battery comm. failure (yellow alarm). no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, the device logs were reviewed and confirmed battery failure. the cause of the failure mode was a defective battery. this report is being filed as part of a retrospective review of historical records.
15151 | 1220648-2024-22597 | 20640444 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640444&pc=OZD | 10/07/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 10/07/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1202890 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) generated a yellow alarm ¿placement signal not reliable'. the alarm continued to sound despite troubleshooting efforts. it was reported that decreasing the flows would resolve the alarm, but only for a few minutes. this aic was replaced with another aic resulting in no issues with the replacement aic. there was no patient harm reported. | [5 empty] | The investigation into the controller error/placement signal issue has been completed. the impella automated controller was not received from the customer and therefore, an evaluation of the device was not possible. the data analysis of the logs revealed the day before the complaint date, the console issued a placement signal not reliable (optical snr ¿ event set #130) alarm in response to an abnormally low signal not reliable (snr). the real time logs revealed a gradual decline in snr (labeled as raw placement signal) during the functional check for the returned console the field service engineer replaced the optical bench to resolve the issue. the cause of the optical placement signal issues was due to an optical bench with marginally low snr. this report is being filed as part of a retrospective review of historical records.
15152 | 1220648-2024-22598 | 20640445 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640445&pc=OZD | 08/18/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 08/18/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1167410 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an automatic impella controller (aic) failed to function during a procedure. the motor did not run and there was no pressure. the customer had another aic onsite to utilize, no report of patient harm or injury. | [5 empty] | The investigation for the reported aic - a/c power issues has been completed. the impella device has been returned for evaluation. the cause of the optical placements signal issue was most likely due to invalid g0 gauge factor. the cause of the invalid (g0=0) gauge factor was unable to be determined because the issue was unable to be reproduced. this report is being filed as part of a retrospective review of historical records.
15153 | 1220648-2024-22599 | 20640468 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640468&pc=OZD | 11/08/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 11/08/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1144656 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | Battery Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automated impella controller (aic) experienced a battery failure (red alarm), and aic - other. no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | The investigation into the battery power issue has been completed. the automated impella controller (aic) was returned for investigation. the cause of the battery failure (red alarm) was due to a defective battery (rapid decline in cell voltage). the exact cause of the defective battery was undetermined. the cause of charger communication error was a defective charger. the cause of the defective charger is unknown. this report is being filed as part of a retrospective review of historical records.
15154 | 1220648-2024-22601 | 20640510 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640510&pc=OZD | 11/08/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 11/08/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1143010 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported the automated impella controller (aic) experienced a battery communication failure (yellow alarm). no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the decline in cell 3 voltage was due to a defective battery (rapid decline in cell voltage). this report is being filed as part of a retrospective review of historical records.
15155 | 1220648-2024-22600 | 20640523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640523&pc=OZD | 12/01/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 12/01/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1557111 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had the automated impella controller (aic) in for functional check and found a battery failure. there was no patient use and no harm or injury. should an aic fail in patient use the ifu states to use a backup. | [5 empty] | The investigation into the battery power issue has been completed. the automated impella controller (aic) was returned for investigation. the cause of the battery failure and battery communication error alarms was due to a defective battery (rapid decline in cell voltage). the exact cause of the defective battery was utd. this report is being filed as part of a retrospective review of historical records.
15156 | 1220648-2024-22602 | 20640579 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640579&pc=OZD | 10/15/2024 | 12/30/2024 |  | 11/08/2024 |  |  | 10/15/2024 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025516944 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 47 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support when there was high purge pressure. the lysis protocol was initiated. however, this was unsuccessful and the pump stopped. the pump was explanted and the patient was stable. | [5 empty] | The investigation of high purge pressure and pump stop has been completed. there were no data logs returned. the pump was returned and there was biomaterial wrapped around the impeller and in the purge gap. there was a kink right at the motor housing, however when the coating on the catheter was cut back, there was blood on both sides of the kink. the pump passed electrical testing. according to the clinical details, they got high purge pressure on 10/15 and tpa was started but it was unsuccessful. the pump then stopped on 10/16. based on the returned product and clinical details, the high purge pressure is a cascading event from a buildup of biomaterial that lead to a pump stop. the root cause of the pump stop is most likely due to biomaterial as there was high purge pressure that didn't resolve with tissue plasminogen activator (tpa). g1 reporting contact email revised on manufacturer device report 1220648-2024-22602 in accordance with updated procedures. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
15157 | 1220648-2024-22603 | 20640619 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640619&pc=OZD | 10/16/2024 | 12/16/2024 |  | 11/08/2024 |  |  | 10/16/2024 | 11/08/2024 |  |  | 08/31/2025 |  | 11/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024392560 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  |  | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient had an attempted impella 5.5 placement. however, it was aborted due to patient anatomy. the patient¿s outcome is unknown. | [5 empty] | The investigation into delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-22603. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist: section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
15158 | 1220648-2024-22604 | 20640622 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640622&pc=OZD | 07/30/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 07/30/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1193561 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  |  | White | Malfunction | No | Yes | Battery Problem | [6 empty] | Insufficient Information | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. while on support, there was a battery failure notice appeared on the aic when it was turned on. the patient was left on the aic from transferring hospital until loaner arrived. there was no report of patient harm or injury. | [5 empty] | The investigation into the battery power issue has been completed. the automated impella controller (aic) was returned for investigation. the battery experienced permanent failure and one of four cells was damaged. by swapping battery 1 battery 2 connections, we were able to determine that charger operated correctly and the undamaged battery was able to be charged and discharged on both channels. the cause of the battery alarm was damaged cell in battery 1. the cause of battery failure was not determined. this report is being filed as part of a retrospective review of historical records.
15159 | 1220648-2024-22605 | 20640647 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640647&pc=OZD | 08/31/2021 | 12/11/2025 |  | 11/08/2024 |  |  | 09/02/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1147484 |  | 0042-0000-EU |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | White | Injury | Yes | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [11 empty] | Other | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient noted as high risk percutaneous coronary insertion was on an automated impella controller (aic) for mechanical circulatory support. while on support, the aic experienced a battery failure. the initial aic was replaced, and the case continued. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported battery charging issue has been completed. the product was returned, and the issue was confirmed. data analysis: aic logs revealed that a battery level low (event set #23) alarm was issued and cleared within three seconds. device analysis: the console¿s batteries were tested while in fs, and no issues were observed. per the results of the clinical review and investigation, the root cause was determined to be due to a momentary communication glitch between battery 2 and power battery management (pbm), where a single sample from the battery data (in this case value of batteryremainingcapacity) was corrupt. this report is being filed as part of a retrospective review of historical records. | Corrected information provided in b1, b1, d2, h1, h1 and h6.
15160 | 1220648-2024-22606 | 20640684 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640684&pc=OZD | 09/02/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 09/02/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2022050044 |  | 0042-0010-EU |  | No |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male |  |  | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. the aic battery could not be charged, and the pump was transferred to another aic. no patient harm was reported. | [5 empty] | The investigation into the a/c power issue has been completed. the automated impella controller (aic) was returned for investigation. the unexpected controller shutdown and loss of battery charge exhibited in the field was due to a failed power supply. the exact root cause of failure for the power supply is undetermined. before returning to customer, aachen fs was instructed to: replace the charger. this report is being filed as part of a retrospective review of historical records.
15161 | 1220648-2024-22609 | 20640705 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640705&pc=OZD | 10/09/2024 | 12/17/2024 |  | 11/08/2024 |  |  | 10/09/2024 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025481556 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | Yes | Restricted Flow rate | Pumping Stopped | [5 empty] | Hemolysis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported during impella cp support for patient, suction issues arose which required administration of heparin to resolve. the patient developed hemolysis which was not successfully managed. history in automated impella controller (aic) showed short pump stop during implant due to high motor current. patient expired while on support. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation into hemolysis, low pump flow, and pump stop has been completed. the "pump stop" failure mode was changed to "temporary pump stop" as the pump was able to be restarted. insufficient data logs were returned for analysis. the returned pump was inspected and biomaterial was found in the purge gap. there were no other physical abnormalities and the pump passed electrical testing. the data logs confirm "impella stopped - motor current high" alarms. there was a motor current spike with speed deviation before temporary pump stop which is characteristic of biomaterial ingestion. the logs show continuous suction alarms across various p-levels and a motor current spike before reported hemolysis. the placement signal was also trending downward. the root cause of the temporary pump stop was determined to be due to biomaterial. the root cause of the low pump flow was determined to be due to biomaterial. the root cause of the hemolysis was determined to be due to biomaterial. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-22609. d6b explant date was reported incorrectly on manufacturer device report 1220648-2024-22609. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
15162 | 1220648-2024-22608 | 20640708 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640708&pc=OZD | 09/23/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 10/11/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1200770 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced data streaming issues. it was reported that aic disconnected from the wifi multiple times during the day and connected as a new pump. an attempt to troubleshoot was unsuccessful. the console was replaced to address the issue. no patient was involved. | [5 empty] | The investigation for the reported "aic - loss of opm data" has been completed. the product was returned and the "aic - loss of opm data" was confirmed. the root cause of the "aic - loss of opm data" was determined to be "backstrap" cable on the back of automated impella controller (aic). this report is being filed as part of a retrospective review of historical records.
15163 | 1220648-2024-22613 | 20640721 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640721&pc=OZD | 11/29/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 11/29/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1126219 |  | 0048-0002 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant re;ported a battery failure. there was no patient harm or injury related to this as there was no patient use at the time. the battery is backup for the a/c power. the instructions for use states to use the backup console for any failures. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. batteries were depleted because the console was not routinely charged as recommended. this report is being filed as part of a retrospective review of historical records.
15164 | 1220648-2024-22607 | 20640743 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640743&pc=OZD | 10/16/2024 | 12/12/2024 |  | 11/08/2024 |  |  | 10/16/2024 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532521 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp had low flow with suction alarms at p2. repositioning improved. the flow. however, suction was still persistent above p5 and p6. the staff did note that the impella cp kinked during sheath exchange. the physician elected to exchange the impella cp due to an inability to increase flows. the patient survived the explant. | [5 empty] | The investigation for the low pump flow has been completed. the clinical team confirmed the cannula kink. data logs were reviewed which confirmed the low pump flow when the pump started and run for two minutes and remained to the rest of the case. the product was not returned for review. based on clinical description and data analysis, the root cause of low flow was most likely a kinked cannula. g.1 revised reporting contact email since the initial manufacturer device report 1220648-2024-22607 was submitted. h.6 added code to health effect - impact code as it was inadvertently omitted from initial manufacturer device report 1220648-2024-22607. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
15165 | 1220648-2024-22610 | 20640772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640772&pc=OZD | 11/08/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 11/08/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1391515 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported an aic purge system failure that could not be resolved with a new purge cassette. the console was taken out of service and replaced with a backup. there was no reported harm or injury to the patient. | [5 empty] | The investigation into the purge issue has been completed. the automated impella controller (aic) was returned for investigation. data analysis: aic logs confirmed the reported failure mode. the purge cassette¿s piston had 0% of range once inserted, which is related to the stepper motors inability to move. eventually, the console triggered a purge system failure alarm. device analysis: upon arrival for investigation, the console was found to have glucose build up on the stepper motor¿s gasket which did not allow for the spindle to move. this was confirmed by trying to move the spindle when the console was not powered on. the cause of the purge system failure was due to build up of glucose on the stepper motor. the cause of the glucose build up was unable to be determined. this report is being filed as part of a retrospective review of historical records.
15166 | 1220648-2024-22616 | 20640775 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640775&pc=OZD | 11/03/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 11/03/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1129486 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) did not detect the purge disc after multiple attempts, the same purge cassette was used on a different aic without issue. also, when turned on the aic did not connect to impella connect. no indication of patient involvement, harm, or injury. | [5 empty] | The device was returned for evaluation. the device logs from the complaint date revealed that the console issued the purge disc not detected alarm several times. device analysis: the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. additionally, when connecting to the module via serial terminal, it was confirmed that there were no wifi configurations on the module. the module was never configured to the hospital network. this report is being filed as part of a retrospective review of historical records.
15167 | 1220648-2024-22611 | 20640782 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640782&pc=OZD | 12/01/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 12/01/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1129883 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) blue purge disc was broken and needed replacement. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported damage issue has been completed. the product was returned, and the issue was confirmed. the damage was confirmed upon initial inspection. per the results of the clinical review and investigation, the root cause could not be determined. this report is being filed as part of a retrospective review of historical records.
15168 | 1220648-2024-22614 | 20640846 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640846&pc=OZD | 10/29/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 10/29/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1570740 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was on an automated impella controller (aic) for mechanical circulatory support. upon start up, the device exhibited battery failure. the resolution for this issue is unknown. no report of patient harm or injury. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the battery failure alarm was due to a defective battery (rapid decline in cell voltage). this report is being filed as part of a retrospective review of historical records.
15169 | 1220648-2024-22612 | 20640895 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640895&pc=OZD | 10/29/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 10/29/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1093083 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) was unable to hold charge. actions taken to resolve the issue are unknown. no patient involvement. | [5 empty] | The investigation into the battery power issue has been completed. the automated impella controller (aic) was returned for investigation. the cause of the decline of the voltage of cell 3 was due to a defective battery (rapid decline in cell voltage). the exact cause of the defective battery was undetermined. this report is being filed as part of a retrospective review of historical records.
15170 | 1220648-2024-22620 | 20640904 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640904&pc=OZD | 11/02/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 11/02/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1478158 |  | 0042-0000-JP |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) exhibited a battery issue. when the ac power was reconnected, the display indicating charging was temporarily displayed, but the charging display stopped within a few seconds, and the display was 50% (battery yellow). there was no reported harm or injury possible. | [5 empty] | E1, e3, e4: the name and occupation of the initial reporter is unknown. d9: the date of the device return for evaluation is unknown. log files are consistent with complaint intake date, the device alarmed for "battery level low". the device was evaluated, plugging the console into ac power to recharge the batteries and the console would not charge past 50%. switching the battery connectors resulted in the same communication failure on the same channel, indicating an issue with the pbm. the cause of the issue with charging the battery set past 50% was defective charging sub circuity on the pbm. this report is being filed as part of a retrospective review of historical records.
15171 | 1220648-2024-22615 | 20640911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640911&pc=OZD | 11/29/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 11/29/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1413896 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported the automated impella controller (aic) battery failing to charge past 50%, despite ample time plugged in. there was no harm or impact to a patient or support. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the battery failure was due to a defective battery (rapid decline in cell voltage). this report is being filed as part of a retrospective review of historical records.
15172 | 1220648-2024-22621 | 20640914 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640914&pc=PYX | 10/15/2024 | 10/17/2025 |  | 11/08/2024 |  |  | 10/15/2024 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025543778 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that during support for 74-year-old male patient, the impella rp flex flows slowly decreased over the course of the night without alarm or change in motor current. only notable change was that the placement signal pressures were significantly elevated. after consultation with team, it was determined this was most likely thrombus. the rp flex was removed and visible thrombus was seen upon removal at inlet area. the patient is noted to be stable. | [5 empty] | The investigation for the thrombus and low pump flow has been completed. clinical details state that ps changing with no changes in motor current. the logs were observed and a drift was noted. there was no product returned. the data logs showed that throughout the whole case there were spikes in motor current and then decreases in the flows without a change in p-level. this is most likely biomaterial interacting with the impeller. post explant, the clinical stated that thrombus was observed. the root cause of the low pump flow is most likely due to biomaterial. in addition to the low pump flow experienced, the clinical stated that there was a notable change in the ps and flows without a change in the motor current. after observing the logs there is dp sensor drift that is causing the notable change. the root cause of the placement signal issues is most likely due to sensor drift based on the log pattern. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella rp system with automated impella controller & impella rp flex with smart assist system with automated impella controller instructions for use & clinical reference manual section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the pump was returned. 5s parameters showed lower than expected signal-to-noise ratio (snr) and contrast with higher-than-expected cavity length. biomaterial was found wrapped around the impellor blades. the data logs showed that throughout the whole case there were spikes in motor current and then decreases in the flows without a change in p-level. this is most likely biomaterial interacting with the impeller. post explant, the clinical stated that thrombus was observed. low pump flow: the root cause of the low pump flow is most likely due to biomaterial. placement signal issues: the fm placement signal issue was added due to the ps changing with no changes in motor current. the logs were observed, and a drift was noted. clinical stated that there was a notable change on 10/15 in the ps and flows without a change in the motor current. after observing the logs there is dp sensor drift that is causing the notable change. the root cause of the placement signal issues is most likely due to sensor drift based on the log pattern. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated to include placement signal coding. d10 concomitant medical products and therapy dates updated.
15173 | 1220648-2024-22618 | 20640925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640925&pc=OZD | 11/02/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 11/02/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1291396 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) was powered up, but not loading. no report of patient involvement, injury, or harm. | [5 empty] | E1, e3, e4: the initial reporters name and occupation is unknown. the investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the battery alarm was a defective battery (rapid decline in cell voltage). the exact cause of the defective battery was unknown. this report is being filed as part of a retrospective review of historical records.
15174 | 1220648-2024-22617 | 20640928 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640928&pc=OZD | 10/16/2024 | 12/23/2024 |  | 11/08/2024 |  |  | 10/16/2024 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025478115 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Crack | Fluid/Blood Leak | Pumping Stopped |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp purge was leaking. the luer at sidearm was damaged. the doctor decided against a purge filter bypass. as a result, the pump stopped after several hours. the impella was exchanged. the patient remained stable. the patient survived the explant. | [5 empty] | The investigation for the purge leak and pump stop has been completed. the product was returned and evaluated. there was a crack down the length of the luer on the sidearm that connects to the cassette tubing. there was also biomaterial found around the purge gap. log review showed a drop in purge pressure shortly after a bag change. after about 30 minutes, purge pressure increased suddenly and then leveled around 600 mmhg. purge pressure slowly increased after that until a motor current spike occurred and the pump stopped. the "purge leak - pump" fm was removed because it was incorporated into the root cause of the pump stop. the cause of the pump stop was a purge leak due to sidearm yellow luer damage most likely occurring during use due to the high tensions in combination with disinfectants and certain drug ingredients (example oil, alcohol, lipids, etc.). the lack of purge pressure likely caused biomaterial buildup, which ultimately led to a pump stop. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention: section: using the automated impella controller with the impella catheter. ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock. section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention: section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention: section: cleaning. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
15175 | 1220648-2024-22619 | 20640929 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640929&pc=OZD | 11/23/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 11/23/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1134268 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported an automated impella controller (aic) failing to charge. there was no patient involved or impacted. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the inability to charge past 50% was due to a defective battery (rapid decline in cell voltage). the exact cause of the defective battery was unknown. this report is being filed as part of a retrospective review of historical records.
15176 | 1220648-2024-22622 | 20640932 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640932&pc=OZD | 07/28/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 07/28/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1172893 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | Key or Button Unresponsive/not Working | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The biomedical engineer reported that during preventive maintenance (pm) the automated impella controller (aic) experienced power on issues, reportedly the light-emitting diodes (led) was amber, then green, and back to amber during bootup. the engineer also reported that once the console was fully booted up, the rotating knob(encoder) was not responding. there was no patient involvement. | [5 empty] | The investigation into the keyboard rotary knob (kbd) and power on issue issues have been completed. the impella automated controller (aic) was returned for investigation. the investigation of the returned console determined the complaint was reproduced on first boot of ic3441, the console booted with the front panel led lit amber and remained amber after boot. the console interior connections were inspected, all connections were normal. the console was then taken apart and known good spare parts were placed in, in an attempt to find a faulty component. the power battery manager was then inspected. fluid ingress was discovered on the led power sub circuitry. a known good pbm was placed into ic3441. on the next boot, the led first lit amber, and transitioned to green, which is the expected behavior. the issue was not reproduced, and the console was found to be within specification and passing all testing. it is most likely that some undetermined electrostatic discharge (esd) event occurred, however the aic did not receive any permanent damage. the cause of the kbd being unresponsive was unable to be determined because the issue could not be reproduced. the failures observed by the field service engineer during preventive maintenance were caused by a malfunctioning pbm due to fluid ingress and a malfunctioning 4-in-1 due to the failure of the msc1210 microcontroller. the failure of the rpb encoder was due to the 4-in-1. the root cause of the microcontroller failure was unable to be determined. the cause of the pbm failure is due to fluid ingress. this report is being filed as part of a retrospective review of historical records.
15177 | 1220648-2024-22626 | 20640933 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640933&pc=OZD | 10/15/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 10/15/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1232684 |  | 0042-0010-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that there was a problem with the battery on the automated impella controller (aic), noticed during habitual check of all the controllers. no patient involvement. | [5 empty] | E1: name of initial reporter is unknown. e2/e3: unknown. the investigation into the battery power issue has been completed. the automated impella controller (aic) was returned for investigation. analysis from the battery test 1 was not providing voltage confirmed by the missing "fu" character, also revealed, cell 4 of battery 1 was completely depleted. the cause of the battery failure alarms was due to a defective battery (rapid decline in cell voltage). the exact cause of the defective battery was undetermined. this report is being filed as part of a retrospective review of historical records.
15178 | 1220648-2024-22624 | 20640985 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20640985&pc=PYX | 10/29/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 10/29/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1086076 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the automated impella controller (aic) exhibited an alarm for purge related issues. the patient had air in the purge system. the customer utilized a replacement they had onsite. no report of harm or injury related to this issue. | [5 empty] | The investigation for the reported purge issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the purge issue and its relation to the impella automated controller device was not determined since the product, data logs, and sufficient clinical information were not provided. this report is being filed as part of a retrospective review of historical records.
15179 | 1220648-2024-22623 | 20641042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641042&pc=OZD | 10/22/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 10/22/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1481065 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The clinical consultant reported that the automated impella controller (aic) was chipped on the top right corner. it was reported the chip was observed after opening the box that the aic was returned in after having preventive maintenance performed. there was no patient involvement. | [5 empty] | The investigation into the damage has been completed. the impella automated controller was returned for investigation. the console damage, chipped on the top right corner, was confirmed by visual inspection of the returned console. the purge cassette and pump were used for simulation, and it was verified the functionality of this console was not affected by the exterior damage. according to the checklist of preventive maintenance and information provided by customer, it was most likely the console was mishandled during transportation from the manufacturer to the field (i.e. dropped/ knocked on the top right corner of console). the cause of the cracked front housing was most likely mishandling during transportation/ moving. this report is being filed as part of a retrospective review of historical records.
15180 | 1220648-2024-22625 | 20641055 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641055&pc=OZD | 11/10/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 11/10/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1313879 |  | 0042-0000-EU |  | Yes |  |  | No |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced a purge disc issue. purge door not working properly. no patient was involved. | [5 empty] | D9: the date when device was returned is unknown. the investigation has been completed. aachen fs noted that the purge door hinge was coated in dextrose. dextrose was cleaned out of purge door hinge. purge door was able to be opened without issue. the root cause of the issue opening the purge door in the field is due to dextrose build up on the door hinges. this report is being filed as part of a retrospective review of historical records.
15181 | 1220648-2024-22628 | 20641083 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641083&pc=OZD | 10/24/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 10/25/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1542953 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the automated impella controller experienced a controller error yellow alarm. no clinical details regarding resolution or patient involvement/condition were provided. no patient was involved. | [5 empty] | The investigation has been completed. the console (ic8468) was sent back for further investigation. the root cause of the transient loss of opm-related signals could not be determined due to the inability to reproduce. this report is being filed as part of a retrospective review of historical records.
15182 | 1220648-2024-22627 | 20641131 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641131&pc=OZD | 10/26/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 10/26/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1439042 |  | 0042-0000-JP |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced battery charging issues during maintenance. it was noted that the device was only used for training, not for clinical use. no patient involvement. | [5 empty] | E1, e3, e4: the name of the initial reporter and occupation is unknown. d9: the date of the return of the device for evaluation. the investigation for the reported issue has been completed. the device was returned for evaluation. the cause of the battery failure alarm was due to a defective battery (rapid decline in cell voltage). the exact cause of the defective battery was unknown. the cause of the battery failure alarm was due to a defective battery (rapid decline in cell voltage). the exact cause of the defective battery was unable to be determined. this report is being filed as part of a retrospective review of historical records.
15183 | 1220648-2024-22629 | 20641155 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641155&pc=OZD | 10/18/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 10/18/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1218788 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a continuous battery alarm was observed on the automated impella controller (aic). it was reported that the issue did not resolve, therefore the initial aic was replaced, and the case continued. | [5 empty] | The investigation for the reported display issue has been completed. the product was returned, and the display issue was confirmed. per the results of the clinical review and investigation: battery failure alarm is due to one malfunctioning battery. the exact root cause of the malfunctioning battery is unknown. this report is being filed as part of a retrospective review of historical records.
15184 | 1220648-2024-22630 | 20641181 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641181&pc=OZD | 06/14/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 06/14/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | ABIOMED, INC. |  |  | US |  | 1183895 |  | 0042-0000 |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced a sudden pump stop during maintenance. there was no reported patient involvement. | [5 empty] | D9: the date of the devices return for evaluation is unknown. e4: the name of the initial reporter and occupation is unknown. the investigation for the reported pump stop has been completed. aic logs revealed that an impella stopped (current mismatch) alarm would trigger at p-level 1. the device was returned for evaluation. the cause of the issues starting the service pump was due to a defective impellatronics board. this report is being filed as part of a retrospective review of historical records.
15185 | 1220648-2024-22631 | 20641197 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641197&pc=OZD | 11/02/2021 | 03/27/2025 |  | 11/08/2024 |  |  | 11/02/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1385451 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | Biomed engineer reports that purge cassette is not being recognized and not latching in the automated impella controller (aic). it doesn't fit tightly and does not go forward to the next step. the aic was returned for investigation. no patient is involved. | [5 empty] | Data analysis: event logs showed no content after (b)(6) 2021, therefore, no data available on the complaint occurred date (b)(6) 2021). however, the event logs from (b)(6) 2021 are relevant to the complaint. during the case start wizard, the purge disc was not inserted (or fully inserted) within 10 seconds of the purge cassette implement, which triggered the purge disc not detected alarm. the alarm was cleared as soon as the disk was inserted and sensed. however, there was another alarm of purge system open which occurred during the case start step 6, indicating purge pressure lower than specification. device analysis: demo pump and two purge cassettes were utilized through the investigation, no latching/ connecting issues found. the original complaint of ¿purge cassette not being recognized nor latching¿ and ¿purge disc not connecting¿ could be caused from not fully insertion of the cassette/ disk during operation. the purge system open alarm was reproduced during investigation, with extreme low purge pressure presented on the user interface. from the dissemble of purge assembly, it was found that the ribbon cable connecting pressure transducer and purge unit driver was loose. the loose cable caused incomplete connections for sensation of pressure, ultimately giving inaccurate measurement/ reading of low purge pressure. after re-insert the ribbon cable and ensure the connection, the purge pressure value was recovered to normal, and no alarm was issued. therefore, it confirmed that the purge system open alarm (low purge pressure) was due to an unaligned ribbon cable, the exact root cause of the loosen connection was not determined. this report is being filed as part of a retrospective review of historical records. | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records.
15186 | 1220648-2024-22633 | 20641213 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641213&pc=OZD | 07/26/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 07/26/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1139980 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) placement signal was not reliable. it was reported that the first catheter lost placement signal with the same alarm, the physician completed the procedure. the physician took the pump out and the patient decompensated. a new pump was placed with the same error resulting in the aic being switched out with a replacement. there was no report of patient harm or injury. | [5 empty] | The investigation into the optical signal/placement signal issues has been completed. the impella automated controller (aic) was returned for investigation. data analysis: imc logs revealed that there was a placement signal unreliable alarm issued on the complaint date. further examination of the rt logs with the first pump showed periodic spiking of the raw placement signal which eventually railed to max value. upon switching the pump, the same placement signal issues persisted. device analysis: the console was tested thoroughly on an engineering bench. upon arrival, the 5s optical parameters and optical power were within specification according to the pm procedure (b)(4). there was contamination found on the pump connector¿s ferrule which could not be removed with wet cleaning. however, even with a known good optical pump connector, the failure mode (placement signal spiking) could be reproduced, and 5s parameters were still within specification. the root cause of the placement signal issues was due to a defective optical bench. the cause of the defective optical bench was not determined. this report is being filed as part of a retrospective review of historical records.
15187 | 1220648-2024-22632 | 20641221 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641221&pc=OZD | 09/16/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 09/16/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1470310 |  | 0042-0000-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported an automatic impella controller (aic) alarm occurred after a pump stopped. there was no reported harm to the patient as a result of the reported issue. no additional information provided. | [5 empty] | The investigation for the reported aic - controller failure (red alarm) issues has been completed. the impella device has not been returned for evaluation. the cause of the purge fluid leak was not determined as the console and data were not returned for review. this report is being filed as part of a retrospective review of historical records.
15188 | 1220648-2024-22634 | 20641267 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641267&pc=OZD | 10/15/2024 | 12/16/2024 |  | 11/08/2024 |  |  | 10/15/2024 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025512621 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a 5.5 pump placed for care escalation from the cp for a patient admitted in acute myocardial infarction / cardiogenic shock. after 8 days there was a pump stop that prompted explant. prior to the explant the patient had been without purge for 30 minutes as the purge cassette had malfunctioned and the patient had no purge flowing for that duration. the patient survived the event. | [5 empty] | Upon visual inspection, blood was present in the purge gap and in purge line just proximal of the motor housing. pump was attempted to be run in the lab and pump stop reproduced. data logs were reviewed and lt log shows pump did not have purge for approximately 30 minutes while purge cassette change was occurring. once purge cassette was completed 30 minutes later, high purge pressure alarms were triggered. additionally, at time of purge cassette change completion, motor current began to rise and pump flow became saturated before high motor current pump stop alarms were triggered. clinical details noted that high purge pressure alarms and eventual pump stop occurred after a 30 minute period with no purge running in pump due to a long cassette change when a "defective purge cassette" alarm was triggered. the root cause of the pump was due a use issue as the pump did not have any purge for 30 minutes due to prolonged cassette change. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-22634. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: section: direct aortic insertion: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
15189 | 1220648-2024-22636 | 20641270 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641270&pc=OZD | 12/10/2023 | 12/16/2024 |  | 11/08/2024 |  |  | 10/15/2024 | 11/08/2024 |  |  | 10/31/2025 |  | 11/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024407788 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 83 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured acute renal failure with a "possible" relationship to the impella cp device used: ¿(b)(6) 2023: aki is worsening - likely due to hemolysis from impella. discharge summary: overnight on (b)(6), patient had increasing pressor requirements, renal failure and worsening lactic acidosis despite escalation of impella/pressors. her son was called to discuss further escalations of care, he did not want to pursue intubation, cpr or dialysis for his mother. with her worsening shock state, creatinine: baseline 1.05 peak 3.28 final 3.28¿. the patient outcome is unknown. another notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured hemolysis with a "possible" relationship to the impella device used: ¿(b)(6) 20233: aki is worsening - likely due to hemolysis from impella. ldh 878 d.bil 1.5 hemoglobin (baseline 12.4, nadir 9.4)¿. the patient outcome is unknown. | [5 empty] | The investigation for the hemolysis and renal failure has been completed. the pump was not returned for investigation. the data log shows no trend in motor current or any positioning issues during impella support. some placement signal trends were reproduced during the case but did not impact parameters related to the pump's performance. as per the clinical details, there were complications potentially linked to the impella device, where hemolysis and renal failure worsened during treatment. no information provided on patient condition after pump removal. the cause of the hemolysis issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the renal failure issue was not determined since product was not returned and sufficient clinical information was not provided. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
15190 | 1220648-2024-22635 | 20641287 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641287&pc=OZD | 09/08/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 09/09/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1480871 |  | 0042-0010-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced that a fuse was damaged. the fuse was replaced, and the case continued. no patient was involved. | [5 empty] | The investigation has been completed for aic - damage during therapy. there were no issues running the pump/ purge, and the logs looked normal afterward via a log review. the root cause of the power supply mounting issues was most likely due to a technician error during manufacturing. this report is being filed as part of a retrospective review of historical records.
15191 | 1220648-2024-22638 | 20641295 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641295&pc=OZD | 07/22/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 07/22/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1162407 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) that during preventive maintenance there was a battery failure. there was no patient involvement. | [5 empty] | The investigation into the battery power issue has been completed. the automated impella controller (aic) was not returned for investigation. however, log file details provided were sufficient to determine a root cause. console logs from the reported day of event are consistent with complaint and show alarms. the exact nature and root cause of charger malfunction were not investigated because charger was not available for analysis. the cause of battery alarms was a defective pbm. the cause of the defective pbm is unknown. this report is being filed as part of a retrospective review of historical records.
15192 | 1220648-2024-22639 | 20641347 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641347&pc=OZD | 10/13/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 10/18/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1122628 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support, the automated impella controller (aic) experienced a controller failure. the console was replaced to address the issue. no patient harm reported. | [5 empty] | The investigation for the reported aic - controller failure (red alarm) has been completed. the product was returned and the aic - controller failure (red alarm) was confirmed. imc logs were consistent with the controller error alarms on the complaint occurred date. upon boot-up, the purge sensor defective error message appeared, followed by the controller failure (red alarm). according to the logs, attempts of reinserting the pc and rebooting the console failed in solving the persistent alarm. console was investigated on engineering bench. visual inspection of purge assembly revealed that ribbon cable for purge pressure sensor was mechanically compromised. the root cause of the controller failure (red alarm) was due to the compromised pps ribbon cable. this report is being filed as part of a retrospective review of historical records.
15193 | 1220648-2024-22637 | 20641385 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641385&pc=OZD | 10/16/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 10/16/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1122628 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automated impella controller (aic) experienced a battery failure (red alarm). no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | E1: name of initial reporter is unknown. e2/e3: unknown. the investigation for the reported battery failure (red alarm) issue has been completed. the impella device was received from the customer. it was confirmed that persistent battery failure (transport connection blown/missing) was not due to a disengaged battery switch. due to low/ flatlined voltage, battery 2 data is unavailable through ata or bq software. when console was booted for the first time, console alarms are consistent with what was seen in the field. console interior connections were checked and verified. when visually inspecting the batteries and charger, it was observed that one of the batteries status leds were completely blank. failed battery was checked, ocv was measured to be ~v2.9. compared to the other battery where the measured voltage was ~16v. the decline in all cell voltages of battery 2 is undetermined. this report is being filed as part of a retrospective review of historical records.
15194 | 1220648-2024-22640 | 20641471 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641471&pc=OZD | 10/12/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 10/12/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1566737 |  | 0042-0040-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 41 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported a patient was placed on impella cp for hemodynamic support. while on support, the automated impella controller (aic) experienced a battery failure (red alarm). another aic was used. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported battery issue has been completed. the impella device was not received from the customer. the clinical details provided were sufficient to determine a root cause. the cause of the battery failure alarm was due to a defective battery (rapid decline in cell voltage). the exact cause of the defective battery was undetermined. this report is being filed as part of a retrospective review of historical records.
15195 | 1220648-2024-22641 | 20641482 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641482&pc=OZD | 09/06/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 09/06/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1168284 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support, the automated impella controller (aic) experienced not recognizing the pump connected. no clinical details regarding resolution. no patient harm reported. | [5 empty] | The investigation has been completed. the console (ic3372) was sent back for further investigation. the root cause of the data streaming issues could not be determined. this report is being filed as part of a retrospective review of historical records.
15196 | 1220648-2024-22646 | 20641487 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641487&pc=OZD | 09/01/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 09/01/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2022102507 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. the console exhibited purge pressure issues. no information available about the resolution. no report of patient harm or injury. | [5 empty] | The investigation for the reported aic - controller failure (red alarm) issues has been completed. the device has been returned for evaluation. root cause of the controller failure is due to a defective purge pressure sensor board. this report is being filed as part of a retrospective review of historical records.
15197 | 1220648-2024-22648 | 20641517 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641517&pc=OZD | 07/06/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 07/06/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1112621 |  | 0042-0000-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) did not recognize the purge cassette. there was no patient harm reported. | [5 empty] | Note: date device returned to manufacturer (d9) is unknown. the investigation into purge disc/flag issues has been completed. the impella automated controller was returned for investigation. the data analysis of the logs from the reported day of event are mostly consistent with complaint and showed purge-related errors (sensor voltage out of range) and an alarm (418 - purge pressure sensor defective). the returned device was investigated, and the issue was not reproduced during 300 power-cycles, and the aic did not present with any purge alarms upon boot-up. the visual inspection of the console revealed scratches on the purge assembly insert around its mounting screws, as well as deformed screw heads, which may be a result of tampering. there was also a crease on the purge sensor foil, and dextrose residue in the cassette retainer area. no evidence of damage was found inside the purge assembly. the cause of sensor malfunction could not be determined, because the issue was not reproduced. the cause of the controller alarm was a defective purge pressure sensor. however, the root cause is unknown. this report is being filed as part of a retrospective review of historical records.
15198 | 1220648-2024-22642 | 20641518 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641518&pc=OZD | 10/15/2024 | 12/12/2024 |  | 11/08/2024 |  |  | 10/15/2024 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025557883 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support. while finishing up positioning, left ventricle perforation noted was. the patient was placed back on bypass, while hole was stitched close. the patient was cannulated centrally for venoarterial extracorporeal membrane oxygenation, chest left open and sent to cardiovascular intensive care unit for management. the family decided to withdraw care and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation for the ventricle perforation has been completed. pump was not returned for investigation. clinical information provided is limited and the reason for the ventricle perforation is unclear. therefore, the root cause of the ventricle perforation issue was not determined since product was not returned and insufficient clinical information was provided. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
15199 | 1220648-2024-22643 | 20641550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641550&pc=OZD | 10/15/2024 | 12/24/2024 |  | 11/08/2024 |  |  | 10/15/2024 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025536039 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant in japan reported that during impella cp implantation for 79-year-old male patient, when it was time to turn on the impella, it did not start. the cp was therefore replaced. there was no harm to the patient. | [5 empty] | The investigation pump did not start has been completed. the cause of the pump did not start issue was most likely open electrical line. the location of the electrical open line was not determine without returned product and sufficient clinical information was not provided. g1 reporting contact email revised on manufacturer device report 1220648-2024-22643 in accordance with updated procedures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
15200 | 1220648-2024-22645 | 20641571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641571&pc=OZD | 07/12/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 07/12/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1276245 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) did not recognize the purge disc. it was reported the staff followed the instructions from the aic however, the screen never progressed despite multiple attempts. this aic was replaced with another aic, the original purge disc was reportedly recognized with the replacement aic. there was no patient harm reported. | [5 empty] | The investigation into purge disc/flag issues has been completed. the impella automated controller was returned for investigation. the returned console was investigated for not properly detecting the purge pressure disc and was reproduced, and the purge flag was confirmed to be broken (missing at blue disc retainer). the root cause of the issue was due to the broken purge flag. this report is being filed as part of a retrospective review of historical records.
15201 | 1220648-2024-22649 | 20641590 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641590&pc=OZD | 07/06/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 07/07/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1566790 |  | 0042-0000-JP |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 25 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) soft key button was not responding. this aic was replaced with another aic. there was no patient harm reported. | [5 empty] | Note: date device returned to manufacturer (d9) is unknown. the investigation into the keyboard rotary knob (kbd) has been completed. the impella automated controller (aic) was returned for investigation. console operated within specification, and successfully passed functional check. a console's keyboard being unresponsive for certain amount of time is consistent with console experiencing an esd event, or being exposed to e-m radiation, or having liquid spilled on display (visual inspection of console did not reveal any evidence of liquid spill or fluid ingress). however, the root cause is unable to be determined. this report is being filed as part of a retrospective review of historical records.
15202 | 1220648-2024-22647 | 20641596 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641596&pc=OZD | 10/08/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 10/08/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1195175 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported the automated impella controller (aic) experienced a battery failure (red alarm). no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | E1: name of initial reporter is unknown. e2/e3: unknown. the investigation for the reported battery issue has been completed. the impella device was not received from the customer. the cause of the battery failure was due to a defective battery (rapid decline in cell voltage). the exact root cause of the defective battery 2 was due likely due to a single cell failure. this report is being filed as part of a retrospective review of historical records.
15203 | 1220648-2024-22653 | 20641607 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641607&pc=OZD | 08/26/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 08/31/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, JP |  |  | ABIOMED, INC. |  |  | US |  | 1325612 |  | 0048-0024-JP |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 35 YR |  | Female |  |  | Asian | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella for mechanical circulatory support. during priming, the automated impella controller (aic) produced a "purge system failure" alarm. a purge system failure alarm occurred during priming, and when priming was performed again after restarting, the procedure screen was skipped and the screen suddenly switched to the pump position screen. the device was replaced, no report of harm or injury to the patient. | [5 empty] | D9: the date of the devices return for evaluation is unknown. the device was returned and evaluated, and the investigation for the reported display issue has been completed. console logs from the reported day of event are mostly consistent with complaint and show that despite "piston blocked" alarm, the priming was eventually successful. operators might have been unaware of that fact and attempted to repeat priming procedure - but since this step was already concluded, it may have appeared as if priming procedure was not being executed. aic was power-cycled, and pump re-connected to console several times, as the means of troubleshooting of a perceived problem. after the first connection, pump was initialized and operational parameters recorded in its eeprom. these events were subject of separate investigation which concluded that there was no malfunction of either pump or purge cassette, and any perceived malfunction of aic could be explained by a combination of operator error (not inserting purge cassette fully into purge assembly) and use issue. this report is being filed as part of a retrospective review of historical records.
15204 | 1220648-2024-22644 | 20641667 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641667&pc=OZD | 07/31/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 07/31/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1202890 |  | 0042-0000-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the automated impella controller (aic) alarmed that the purge cassette was not recognized by the console. the console was replaced with no reported harm or injury to the patient. | [5 empty] | Note: date device returned to manufacturer (d9) is unknown. the investigation into the purge disc issue has been completed. the device was returned for evaluation. the failure mode could not be reproduced. glucose buildup was found on the gasket of the stepper motor and on the purge insert tray which was cleaned immediately. the root cause of the inability to deair was most likely due to buildup of glucose on the purge stepper motor gasket. this report is being filed as part of a retrospective review of historical records.
15205 | 1220648-2024-22651 | 20641692 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641692&pc=OZD | 07/21/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 07/21/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1261275 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) knob was loose and unable to change the screen or make selections. this occurred during completing inservice at the site and there was no patient involvement. | [5 empty] | The device was returned and evaluated, and the investigation for the reported display issue has been completed. data analysis: log entries from complaint date show several menus navigated successfully, despite complaint that rpb would not select menu items: device analysis: console arrived in danvers. rpb was inspected, and it was observed that the inner piece of the rpb that mates to the encoder is considerably worn out. when using the rpb on the console, it is noticeable when compared to a new rpb that there is play when using it. no issues were observed when attempting to select menu items. a new rpb was placed next to the old to compare the wear and tear. there is significant wear to the originally rpb. difficulty in selecting menu items with the rpb is due to the rpb being worn out, allowing play in the knob. this report is being filed as part of a retrospective review of historical records.
15206 | 1220648-2024-22652 | 20641698 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641698&pc=OZD | 10/04/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 10/04/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1223031 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported the automated impella controller (aic) experienced a battery communication failure (yellow alarm). no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the battery failure was due to a defective battery one. this report is being filed as part of a retrospective review of historical records.
15207 | 1220648-2024-22650 | 20641709 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641709&pc=OZD | 08/29/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 08/29/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1145118 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the automated impella controller (aic) experienced a failure alarm during support. the device started having purge flow and pressure issues leading to a purge system stopped alarm. the initial aic was replaced, and the case continued. no harm or injury noted. | [5 empty] | The investigation for the reported aic - controller failure (red alarm) issues has been completed. the impella device has been returned for evaluation. failure mode is due to air in purge line that was not properly de-aired preventing proper purge pressure from building. the root cause of the air in the purge is undetermined. this report is being filed as part of a retrospective review of historical records.
15208 | 1220648-2024-22654 | 20641717 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641717&pc=OZD | 09/28/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 09/28/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1563183 |  | 0042-0010-EU |  | No |  |  | No |  | No |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that, during testing, the automated impella controller (aic) experienced a battery communication failure yellow alarm with no resolution specified. no patient involvement. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, the device logs were reviewed. the cause of the battery failure alarms was due to a defective battery (rapid decline in cell voltage). this report is being filed as part of a retrospective review of historical records.
15209 | 1220648-2024-22655 | 20641734 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641734&pc=OZD | 08/23/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 08/23/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1571192 |  | 0042-0040-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Fell | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) had fallen during transferring the patient, and the plug broke. there was no patient harm reported. | [5 empty] | D9: the date of the devices return for evaluation is unknown. e1, e3, e4: name of the initial reporter and occupation is unknown the investigation into the damage has been completed. the cause of the damage to aic was due to mishandling of the console by the end user prior to shipping to transfer the patient. this report is being filed as part of a retrospective review of historical records.
15210 | 1220648-2024-22656 | 20641736 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641736&pc=OZD | 10/15/2024 | 12/05/2024 |  | 11/08/2024 |  |  | 10/15/2024 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025519225 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support, the 58-year-old male patient arrived with the impella cp out of position. an attempt to reposition the device was unsuccessful. therefore, the impella cp was removed with a plan to replace with an additional impella cp. | [5 empty] | The investigation into positioning issues has been completed. the impella device was not returned for evaluation. the root cause of positioning issue was most likely patient movement as it occurred during transferring to new hospital. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-22656. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
15211 | 1220648-2024-22659 | 20641757 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641757&pc=OZD | 09/26/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 09/26/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1083048 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a battery communication failure yellow alarm and a battery failure red alarm. no details regarding resolution were provided. no indication of patient involvement. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the battery failure was due to a defective battery (rapid decline in cell voltage). this report is being filed as part of a retrospective review of historical records.
15212 | 1220648-2024-22657 | 20641758 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641758&pc=OZD | 09/28/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 09/28/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1370139 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | Connection Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced aic purge disc/flag issues. no clinical details regarding resolution or patient condition were provided. no patient harm reported. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the product was returned and the aic - purge disc/flag issues was confirmed. imc logs from the complaint date revealed that the console issued the purge fluid not detected error through the case start wizard. the console was tested. the issue with properly detecting purge fluid was reproduced, and purge disc housing (blue retainer) was confirmed to be broken off of the purge assembly. failure mode was reproduced in house. the root cause is the broken blue purge disc retainer. this report is being filed as part of a retrospective review of historical records.
15213 | 1220648-2024-22660 | 20641772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641772&pc=OZD | 09/20/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 09/20/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1144655 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was placed onto impella cp support for high-risk percutaneous coronary intervention. while on support, the automated impella controller (aic) experienced a battery failure red alarm with no resolution specified. no patient harm was reported. | [5 empty] | The investigation into the battery power issue has been completed. the automated impella controller (aic) was returned for investigation. the cause of the red battery failure alarm was due to a defective battery (rapid decline in cell voltage). the exact cause of the defective battery was undetermined. this report is being filed as part of a retrospective review of historical records.
15214 | 1220648-2024-22658 | 20641782 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641782&pc=OZD | 08/12/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 08/18/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1261289 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) did not recognize the purge cassette and would not move past the implant screen. there was no patient harm reported. | [5 empty] | E1: the name of the initial reporter and occupation is unknown. the investigation into purge disc/flag issues has been completed. the impella automated controller was returned for investigation. the data analysis of the logs validates complaint intake of lack of purge disc detection. the returned console was tested, the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken (missing at blue disc retainer). the root cause of the issue was a broken/missing purge flag. this report is being filed as part of a retrospective review of historical records.
15215 | 1220648-2024-22662 | 20641822 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641822&pc=OZD | 09/14/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 09/14/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1095480 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the automated impella controller (aic) skipped the prime screen. the purge tubing was connected, and no purge fluid was coming out of the end of the catheter. the resolution for this issue is unknown. no reported harm or injury reported. | [5 empty] | The investigation into the purge issue has been completed. the automated impella controller (aic) was returned for investigation. data analysis: the device logs revealed that the priming was completed before the user had confirmed that the purge fluid issue. this results in the user interface to seemingly skip the priming step, but the screen image (priming progress bar) was not displayed since physical priming of the tubing had already been completed. device analysis: the suspected skipping of the purge priming step during case start, which resulted in perceived abnormal aic behavior (displaying the priming progress bar), was likely due to a combination of events: a kinked purge line after the purge pressure disc which met criteria for completion of priming. user related issue, not following case start procedure adequately. this report is being filed as part of a retrospective review of historical records.
15216 | 1220648-2024-22661 | 20641853 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641853&pc=OZD | 07/19/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 07/19/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER EU |  |  | ABIOMED, INC. |  |  | US |  | 1569167 |  | 0042-0010-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The physician reported that the automated impella controller (aic) would not finish the preparation setup. it was reported that the aic screen showed "please connect power cable to impella" however the aic was already connected to the power cable. this aic was replaced with another aic resulting in no issues with the replacement aic. there was no report of patient harm or injury. | [5 empty] | The investigation into the pump not detected/controller failure issue has been completed. the impella automated controller (aic) was returned for investigation. data analysis: imc logs revealed that while in the field, the console was not able to detect a pump during case start: this is a known failure mode to occur with aic software versions v8.2 and up. while v8.4.1 has implemented a mechanism against this issue, the mechanism is known to not always resolve the issue. it is generally known that when this issue occurs, it is best to boot cycle the console, as in most cases this resolves the issue. the staff at the hospital did not attempt to boot cycle the console after the pump detection issue. device analysis: console arrived with germany fs. germany fs was asked to test pump detection by plugging in and unplugging a sa pump. germany fs reported that the console was able to detect sa pumps without issue after repeated testing. the root cause of the pump not being detected was an issue with the crystal of the epm on the impellatronic board. this report is being filed as part of a retrospective review of historical records.
15217 | 1220648-2024-22663 | 20641866 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641866&pc=OZD | 04/27/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 04/27/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1290418 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) had a self-test failure on startup. there was no reported patient involvement. | [5 empty] | The device was returned and evaluated, and the investigation for the reported system self-check error has been completed. complaint intake was consistent with complaint date. it was observed that when attempting to power on the console, the system self-check fail occurred again. the console was opened and physically inspected. it was observed that there was corrosion on the power battery manager (pbm), near super capacitors c69 and c70. the cause of the aic issue was contamination. this report is being filed as part of a retrospective review of historical records.
15218 | 1220648-2024-22664 | 20641890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641890&pc=OZD | 10/15/2024 | 11/08/2024 |  | 11/08/2024 |  |  | 10/15/2024 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1089162 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an automated impella controller (aic) had a battery failure error. this issue was discovered during routine check of the aic. therefore, there was no patient involvement. | [5 empty] | The investigation has been completed. the automated impella controller was returned for investigation. after reviewing the logs, the reported battery failure issue was confirmed. there were numerous instances of battery failure alarms (transport connection blown/missing) observed on the reported complaint date. the console was tested. the battery failure issue was able to be reproduced. results of running the batt test utility revealed that cell 4 in battery 1 had a voltage that was significantly less than the rest of the cells in the battery. the reported battery failure issue was determined to be caused by internal battery cell imbalance (found in cell 4 of battery 1). the root cause of this issue was an internal cell imbalance found in battery 1.
15219 | 1220648-2024-22669 | 20641893 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641893&pc=OZD | 10/10/2024 | 12/09/2024 |  | 11/08/2024 |  |  | 10/10/2024 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025454497 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient had an impella 5.5 implanted and during the implant, the patient had ventricular tachycardia requiring defibrillation as the catheter crossed the aortic valve. the patient returned to sinus rhythm. the patient survived the explant. | [5 empty] | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
15220 | 1220648-2024-22670 | 20641902 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641902&pc=OZD | 09/09/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 09/09/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1322578 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a battery failure red alarm. no clinical details regarding resolution or patient involvement / condition were provided. | [5 empty] | E1: name of initial reporter is unknown. e2/e3: unknown. the investigation into the a/c power issue has been completed. the automated impella controller (aic) was returned for investigation. the cause of the battery failure alarm was due to a defective battery (cell imbalance). the exact cause of the defective battery was undetermined. this report is being filed as part of a retrospective review of historical records.
15221 | 1220648-2024-22665 | 20641926 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641926&pc=OZD | 09/15/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 09/15/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1148815 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 21 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient with myocarditis was placed onto impella cp support for mechanical circulatory support. while on support, the automated impella controller (aic) experienced a battery failure red alarm with no resolution specified. no patient harm was reported. | [5 empty] | The investigation into the battery power issue has been completed. the automated impella controller (aic) was returned for investigation. the cause of the battery failure and battery communication error was due to a defective battery (rapid decline in cell voltage). the exact cause of the defective battery was undetermined. this report is being filed as part of a retrospective review of historical records.
15222 | 1220648-2024-22667 | 20641929 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641929&pc=OZD | 09/07/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 09/07/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1085080 |  | 0042-0000-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was placed onto impella support due to acute myocardial infarction / cardiogenic shock. while on support, the automated impella controller (aic) experienced battery charging issues with no resolution specified. the controller permanently toggles between a few seconds of battery failure alarm and battery charging. no report of patient harm or injury. | [5 empty] | The investigation for the reported issue has been completed. the root cause of the battery failure alarm was due to a defective battery. the exact root cause of the defective battery was not determined. the cause of the controller failure is impellatronics epm firmware corruption. this report is being filed as part of a retrospective review of historical records.
15223 | 1220648-2024-22668 | 20641990 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20641990&pc=OZD | 09/07/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 09/07/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1249368 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 72 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. while on support, they received a purge pressure alarm, then they moved to a new unit. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported purge issue has been completed. the product was returned, and the issue was confirmed. data analysis: aic logs confirmed the reported complaint and showed multiple different purge pressure alarms were triggered. despite changing the cassette, purge pressure would surpass the threshold (pmax > 1050 mm hg) to trigger the purge pressure high alarm (event set #408) and purge system blocked alarm (event set #409). there were also errors indicating the purge system¿s piston was blocked (92% range remaining). these alarms and errors could occur if the purge line was physically blocked or kinked, and the purge solution could not flow. device analysis: the console was tested thoroughly on an engineering bench with a known good pump and purge. there were no issues observed throughout the test. the issues maintaining adequate purge pressure was unable to be reproduced. per the results of the clinical review and investigation, the root cause was not determined. this report is being filed as part of a retrospective review of historical records.
15224 | 1220648-2024-22666 | 20642014 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20642014&pc=OZD | 09/06/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 09/06/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1185328 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a battery communication failure yellow alarm. no clinical details regarding resolution or patient involvement / condition were provided. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the battery failure alarm was due to a defective battery (rapid decline in cell voltage). the exact cause of the defective battery was unable to be determined. this report is being filed as part of a retrospective review of historical records.
15225 | 1220648-2024-22673 | 20642029 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20642029&pc=OZD | 10/14/2024 | 12/12/2024 |  | 11/08/2024 |  |  | 10/15/2024 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025529151 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemolysis | [11 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, the patient developed hemolysis. additionally, impella was repositioned and the patient cardioverted. coronary angiogram revealed small clot distal left anterior descending (lad). the patient remains on impella support. | [5 empty] | The investigation for the hemolysis and the arrhythmia has been completed. data logs confirmed pump was in improper position corresponded with clinical description that triggered alarms impella position in ventricle/aorta. no other issues were found on the logs. product was not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. the root cause of the arrhythmia could not be determined without additional clinical details. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
15226 | 1220648-2024-22671 | 20642035 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20642035&pc=OZD | 07/15/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 07/15/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1276767 |  | 0042-0000-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. while on support, the aic experienced a total screen failure (screen was black) with no restart possible. the aic was switched with a backup unit, and no report of patient harm or injury. | [5 empty] | D9 - date device returned to manufacturer: unknown. the device was returned and evaluated, and the investigation for the reported display issue has been completed. data analysis: the imc logs showed that an emergency shutdown was performed to turn the console off. this confirmed that there were difficulties viewing the screen. device analysis: the console was tested in aachen fs but the failure mode could not be reproduced. testing included: running a pump/ purge, looking for any issues with the display. opening the console and wiggling the video cables to see if it affects the display. the root cause of the lcd video display issues could not be determined because the failure mode could not be reproduced. this report is being filed as part of a retrospective review of historical records.
15227 | 1220648-2024-22672 | 20642229 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20642229&pc=OZD | 02/11/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 02/11/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1349162 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) outlet for the white power pug was ripped off. there was no patient harm reported. | [5 empty] | The investigation into the damage has been completed. the impella automated controller was not received from the customer. the field service engineer (fse) confirmed the damage to the aic and replaced the related parts. the cause of the issue was use related. this report is being filed as part of a retrospective review of historical records.
15228 | 1220648-2024-22674 | 20642315 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20642315&pc=OZD | 07/08/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 07/08/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1072117 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  |  | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The physician reported that the automated impella controller (aic) froze during catheter preparation. it was reported that the aic was rebooted, and the problem was resolved. there was no patient harm or injury reported. | [5 empty] | The investigation into the pump shutdown/power on issue has been completed. the impella automated controller (aic) was returned for investigation. the data analysis of the logs were inconsistent with the complaint, and do not support the claim. ¿ there are no entries on the reported day of event (2021-07-08), which indicates (b)(6) hadn¿t been used that day ¿ there¿s no evidence indicating that pump 323877 was ever used with (b)(6), as the sn is not referenced in available logs since 2021-01-26 ¿ the only pumps used with (b)(6) around the given time frame were 331244 and 330207, both used on 2021-08-12, and we are unaware of any issues during their use (and no complaints were entered against them) ¿ between (b)(6) 2021, console (b)(6) has never been rebooted (booted twice) on the same day. multiple requests for clarification or additional information remained unanswered. the returned console operated within specification. the complaint was not confirmed as the issue was not able to be reproduced, additional information was not supplied, and clinical data was insufficient to determine root cause of the alleged issue. this report is being filed as part of a retrospective review of historical records.
15229 | 1220648-2024-22675 | 20642327 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20642327&pc=OZD | 10/15/2024 | 12/12/2024 |  | 11/08/2024 |  |  | 10/15/2024 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025512862 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient had an attempted impella 5.5 implant and it was aborted due to patient anatomy. the staff attempted the insertion through the right axillary. the patient's anatomy made the implant impossible and it was aborted. no patient harm has been reported. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since the implant was aborted due to patient's anatomy. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
15230 | 1220648-2024-22676 | 20642406 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20642406&pc=OZD | 09/17/2024 | 11/08/2024 |  | 11/08/2024 |  |  | 09/17/2024 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025525314 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a positioning issue involving an impella cp pump. during patient support, the patient had an episode of sustained ventricular tachycardia (vt) while local team was on site. the care team attempted to break the vt using lidocaine and adenosine before ultimately defibrillating. the patient went asystolic after defibrillation with cardiopulmonary resuscitation. intrinsic rhythm was returned as bradycardia. the care team placed venous and arterial sheaths in the left groin. the patient survived the explant. | [5 empty] | The investigation has been completed. no product was returned. as the necessary clinical information was not provided and the device was not returned, the root cause of the ventricular tachycardia was not determined.
15231 | 1220648-2024-22677 | 20642419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20642419&pc=OZD | 07/14/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 07/14/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1371809 |  | 0042-0000-JP |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 40 YR |  | Female |  |  | Asian | Malfunction | No | Yes | Battery Problem | [6 empty] | Insufficient Information | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy on an automated impella controller (aic) for mechanical circulatory support. that when attempting to get the aic up and running when there was a "battery failure, unknown battery status alarm" and another controller was used. | [5 empty] | D89/d9 - date device returned to manufacturer: unknown. the investigation into the a/c power issue has been completed. the automated impella controller (aic) was returned for investigation. data analysis: imc logs confirmed the reported failure mode. the console persistently issued a battery failure alarm (transport connection blown/ missing ¿ event set 360). intermittently, the console would also issue a battery 2 communication failure alarm (smbus comm. loss or sda short ¿ event set 353). device analysis: the console was tested in fs japan in collaboration with the post market engineering team. the failure mode was reproduced upon boot up. battery (sn 4232) was measured with a dmm to have only a voltage of 3.966vdc across the +/- terminals. the root cause of the battery failure alarm was due to a defective battery. the pbm pca was also tested to validate normal charging and discharging of the replacement battery set (per sections 18-20 of the pm procedure (0042-7309). the root cause of the battery failure alarm was due to a defective battery. this report is being filed as part of a retrospective review of historical records.
15232 | 1220648-2024-22678 | 20642491 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20642491&pc=OZD | 10/15/2024 | 12/13/2024 |  | 11/08/2024 |  |  | 10/15/2024 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025524759 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support had oozing in all access sites. the activated clotting (act) time was over 300. systemic anticoagulation on hold until act trends down. additionally, when the physician was repositioning the cp, the patient experienced a sustained ventricular tachycardia requiring cardioversion. it was further reported that the patient expired while on impella support. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation into access site bleeding and vf/vt/af/at has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was most likely patient condition related, oozy in all access sites as per the clinical description provided. as the necessary information was not provided, the root cause of the vt/vf was not determined. instructions for use: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-22678. g1 reporting contact email. h10 instructions for use missing. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
15233 | 1220648-2024-22679 | 20642501 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20642501&pc=OZD | 09/01/2024 | 01/30/2025 |  | 11/08/2024 |  |  | 09/01/2024 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025515155 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support, the placement signal malfunctioned. the pump was then removed. the procedure outcome was successful. | [5 empty] | The investigation for the optical signal issue has been completed. the product was returned and evaluated. the "placement signal not reliable" alarm was reproduced when plugged into the lab console. contrast was slightly under spec on the returned pump, but the other 5s parameters showed no abnormalities. the pump passed the laser continuity test, and no bond issues, kinks in the catheter, or ingested biomaterial were found. log review showed placement signal going out of range (with a "placement signal not reliable" alarm) while signal-to-noise ratio gradually decreased from approximately 10000 to approximately 4000 over the course of three days. shortly before the placement signal went out of range, placement signal stepped down approximately 50 mmhg in about an hour before decreasing more. contrast slightly decreased during this time, but the other 5s parameters were unaffected. when the placement signal went out of range, the pump had been used for 790 hours (approximately 33 days), which is below the design life of 60 days per design traceability matrix 0550-9900. the root cause of the optical signal issue was most likely sensor silicone wear. d.6b explant date was added as it was inadvertently omitted from the initial manufacturer device report number 1220648-2024-22679. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2024-22679 was submitted in accordance with updated procedures. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
15234 | 1220648-2024-22680 | 20642612 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20642612&pc=OZD | 07/12/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 07/12/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1327996 |  | 0042-0000-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) displayed a yellow console failure alarm at start up. ¿switch to backup console. ¿ error upon startup. there was no report of patient harm or injury. | [5 empty] | The device was returned and evaluated, and the investigation for the reported controller error (yellow alarm) has been completed. console logs from the reported day of event are consistent with complaint and show controller error alarm due to defective optical sensor lamp (#1039), immediately after first boot-up since console arrived at customer site. the alarm was preceded by two bad_lamp errors, due to optical bench failing self-check (as evidenced by light parameter of 0.4v instead of usual 2.4v-2.7v). analysis of console logs from the preventive maintenance work performed 2 weeks prior, revealed same error which was however resolved after optical bench self-reset, and therefore did not result in any alarm that would be visible to technician. once the console arrived in danvers, visual inspection of the optical bench revealed incorrect routing of one of optical fibers, that was tightly wrapped around housing of ambient pressure sensor, compromising the fiber (likely cracking or breaking it) and causing optical power loss. after optical bench was replaced, console operated within specification and no further bad_lamp errors occurred. the root cause of the controller alarm was compromised optical fiber resulting in failed self-test of optical bench upon aic boot-up. the fiber was most likely mishandled during pm procedure. this report is being filed as part of a retrospective review of historical records.
15235 | 1220648-2024-22681 | 20642671 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20642671&pc=OZD | 08/27/2024 | 11/27/2024 |  | 11/08/2024 |  |  | 08/27/2024 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025512738 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported oozing issue involving an impella cp pump. during patient support, there was some groin oozing and nurse to redress and angle match to help reduce oozing. the next day, the patient was given one unit of packed red blood cells for lower hemoglobin, but it seems as though oozing from impella site has stopped. the pump was successfully weaned and the patient survived the explant. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
15236 | 1220648-2024-22683 | 20643652 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20643652&pc=OZD | 10/15/2024 | 12/24/2024 |  | 11/08/2024 |  |  | 10/15/2024 | 11/08/2024 |  |  | 10/31/2025 |  | 11/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024413882 |  | 1000361 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | EXTRA CORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant in australia had placed a 5.5 pump to support a patient admitted in with cardiogenic shock and multiple organ failure. the pump was placed but there was an observation made of an access site bleed on the day of insertion. the team troubleshot by stitching, giving 3 units of blood, and platelets infused. the team states there is a coagulopathy underlying and the team gave sodium bicarbonate in the purge and stopped heparin. the pump remained on for support with the 5.5 and ecmo. the patient survived the event. | [5 empty] | The investigation access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding is determined to be patient condition because of the bleeding from all access sites. g1 reporting contact email revised on manufacturer device report 1220648-2024-22683 in accordance with updated procedures. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
15237 | 1220648-2024-22684 | 20643686 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20643686&pc=OZD | 08/20/2024 | 11/08/2024 |  | 11/08/2024 |  |  | 08/20/2024 | 11/08/2024 |  |  | 06/30/2025 |  | 11/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024375072 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 18 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that after handoff to the intensive care unit, the staff noted a hematoma at the impella site. a fellow applied manual pressure and removed dressing. re-angle was matched and a purse string suture was placed. it was noted that bleeding resolved along with the hematoma. the procedure outcome was successful with this device. | [5 empty] | The investigation has been completed. the impella device has not been received from the customer. the root cause of the access site bleeding issue was most likely use related. re-angle matched and purse string suture was placed. bleeding resolved along with hematoma as per the clinical description provided.
15238 | 1220648-2024-22692 | 20643720 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20643720&pc=OZD | 06/09/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 06/09/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1205390 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Battery Problem | [5 empty] | Insufficient Information | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. while on support, the aic experienced a controller failure red alarm and battery failure red alarm that was resolved by switching to a backup console. there was no report of patient harm or injury. | [5 empty] | The investigation into the a/c power issue has been completed. the automated impella controller (aic) was returned for investigation. device analysis: the console was booted and the same behavior occurred in house. console booted with the triggering alarms noted above. console battery level was at 50%. console was opened and interior connections were inspected, all connections were ok. console was further taken apart and the pbm was inspected, and corrosion was found leaking from the super capacitors. this corrosion damaged both sides of the board. pbm sn: (b)(6). this sn is within range of the known board affected by leaking super capacitors. complaint cause was due to corrosion near the c69 and c70 supercapacitors on the charger. this report is being filed as part of a retrospective review of historical records.
15239 | 1220648-2024-22689 | 20643943 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20643943&pc=OZD | 01/06/2023 | 11/08/2024 |  | 11/08/2024 |  |  | 01/06/2023 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1129485 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | No | Unintended Application Program Shut Down | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) is shutting down/restarting during support two times. there was never any harm to the patient and support continued. prognosis was poor for this patient as he has a history of drug abuse, including methamphetamines with recent positive tic screen. care was ultimately withdrawn and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported issue has been completed. data analysis: as can be seen in the imc logs, the rtlogger error triggered constantly throughout the procedure and previous procedure and the watchdog triggered two unexpected controller shutdowns on 12/23 and 1/6 device analysis: the rtlogger errors were able to be reproduced and the unexpected shutdown was even triggered once by the watchdog. the compact flashes were replaced by known goods and the errors went away with the console functioning as expected. this report is being filed as part of a retrospective review of historical records.
15240 | 1220648-2024-22687 | 20644084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20644084&pc=OZD | 02/04/2023 | 11/08/2024 |  | 11/08/2024 |  |  | 02/04/2023 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1456617 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Hispanic | White | Malfunction | Yes | Yes | Failure to Read Input Signal | [6 empty] | Hemodynamic instability | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Complainant reported that staff attempted to use two different automated impella controllers ( aic's) when transferring patient from kingwood to houst meth. both aics took about 40 seconds to read that there was an error with the catheter and wouldn't go any further. staff noted decision to keep kingwood's aic for now for rp flex without issue. (bipella). care was ultimately withdrawn and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | As per date analysis on the date of the failure, (b)(6) running on v12.1 ((b)(6) sw12.1.pdf) had been using pump (b)(4) for three four days which is an rp flex. (b)(6) was attempted to be swapped for (b)(6) which was incompatible because it was on v11.1.1. (b)(6) was attempted next but also failed due to being v10.2. at this point, console (b)(6) was reconnected to the pump and was completed successfully. the cause of the aic issue was a software issue due to incompatible sw. this report is being filed as part of a retrospective review of historical records.
15241 | 1220648-2024-22688 | 20644298 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20644298&pc=OZD | 01/24/2023 | 11/08/2024 |  | 11/08/2024 |  |  | 01/24/2023 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1236845 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Application Program Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Us complainant reported that the customer was placed onto impella cp support due to acute myocardial infarction / cardiogenic shock. while on support, the automated impella controller (aic) experienced a battery failure red alarm with no resolution specified. the patient ultimately expired, there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported issue has been completed. aic logs confirmed the reported failure. there was a battery failure alarm (transport connection blown/ missing ¿ event set #360) due to low pack voltage. the lt powerdata from the complaint date shows a cell voltage decline in battery2 v3 which occurred as the battery failure alarm triggered. device analysis revealed the failure mode was reproduced on an engineering bench. batttest showed battery2 cell v3 having diminished voltages. the root cause of the defective battery was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
15242 | 1220648-2024-22686 | 20644339 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20644339&pc=OZD | 10/15/2024 | 12/12/2024 |  | 11/08/2024 |  |  | 10/15/2024 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025566099 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during support for 61-year-old female patient, the impella cp was repositioned due to the pump diving too deep. the pump going too deep caused the patient to have ventricular tachycardia (vt). the patient had to be shocked. the impella was moving due to the touhy not being tight enough. the patient is stable post-explant. | [5 empty] | The investigation for the reported ventricular fibrillation/ventricular tachycardia (vf/vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf/vt could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
15243 | 1220648-2024-22695 | 20644520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20644520&pc=OZD | 01/05/2023 | 11/08/2024 |  | 11/08/2024 |  |  | 01/05/2023 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1086076 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic |  | Malfunction | Yes | Yes | Key or Button Unresponsive/not Working | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | Us complainant reported that the patient with unknown race was placed onto impella 5.5 support due to cardiomyopathy. while on support, the automated impella controller (aic) experienced kbd/rotary knob issues with no resolution specified. patient ultimately expired . there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported issue has been completed. data analysis: as can be seen in (imc ic1258.pdf), at 12:37 on 1/5/23, the clinician changed the p-level from 6 to 5. the clinician then proceeded to open the p-level drop down multiple times but kept the p-level at 5 for the remainder of the procedure. later in the procedure, the clinician opened up other menus but only went in one direction on the gray knob. device analysis: when the device was returned for investigation, the failure mode was unable to be reproduced. despite attempts to generate emi or excessive heat, the gray knob was fully functional. the cause of the issue was not determined since it could not be reproduced. this report is being filed as part of a retrospective review of historical records.
15244 | 1220648-2024-22685 | 20644643 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20644643&pc=OZD | 07/08/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 07/08/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1351637 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) had a controller error. this aic was replaced with another aic, no further issues were reported. there was no patient harm or injury reported. | [5 empty] | The investigation into the controller error/placement issue has been completed. the impella automated controller (aic) was returned for investigation. data analysis: imc logs confirmed the reported failure mode (see attachments section). a placement signal unreliable alarm and eventually controller error alarm (opm comm crc error) was triggered on the complaint date when a demo pump was connected. initially, the demo pump was read by the console with all zero optical gauge factors, which most likely resulted in the immediate unexpected placement signal unreliable alarm. controller error alarm (opm comm crc error) was triggered a while after connection of the pump. see imc logs from the complaint date in the attachments section. upon further analysis of the rt logs, the opm-related signals went invalid. see rt logs in the attachments section. device analysis: the console was tested on an engineering bench by connecting a known-good pump/ purge and power cycling in between. no issues were observed during the testing. the unexpected controller error alarm (opm comm crc error) is known pattern failure, which is characterized by a temporary loss of placement signal and triggering of a controller error alarm. this failure has a low probability of reoccurrence and does not affect hemodynamic support. the cause of the optical placement signal issues (transient loss of opm-related signals) was unable to be determined because the issue was unable to be reproduced. this report is being filed as part of a retrospective review of historical records.
15245 | 1220648-2024-22717 | 20644651 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20644651&pc=OZD | 12/08/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 12/08/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1383042 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had an automated impella controller (aic) with power failure at the medical facility. the self-test failed to start, and the device exhibited a failure code to shut down and utilize another device. unknown how this issue was resolved, no report of patient harm or injury. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the aic (automated impella controller) issue was contamination. this report is being filed as part of a retrospective review of historical records.
15246 | 1220648-2024-22723 | 20644731 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20644731&pc=OZD | 12/22/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 12/22/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1087351 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The us complainant had an automated impella controller (aic) failure during support that prompted the team to replace the aic. no harm or injury reported to the patient. | [5 empty] | The investigation for the reported aic - controller failure (red alarm) has been completed. the product was returned and the aic - controller failure (red alarm) was confirmed. the root cause of the aic - controller failure (red alarm) was most likely due to a malfunctioning pud pca. this report is being filed as part of a retrospective review of historical records.
15247 | 1220648-2024-22701 | 20644752 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20644752&pc=OZD | 05/07/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 05/07/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1450938 |  | 0042-0040-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller's (aic) purge clip failed. no report of patient involvement, harm, and injury. | [5 empty] | D9: the date of the devices return for evaluation is unknown. e1, e3, e4: the name of the initial reporter and occupation is unknown. the investigation into purge disc/flag issues has been completed. the impella automated controller was returned for investigation. the returned product was received, and a visual inspection noted the purge flag was broken. the cause of the purge issue is due to a broken purge flag. the cause of the failed purge flag was likely due to the wear and tear of the plastic component as well as additional unnecessary force when inserting the purge disc into the purge disc retainer during prior cases. this report is being filed as part of a retrospective review of historical records.
15248 | 1220648-2024-22727 | 20644983 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20644983&pc=OZD | 10/20/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 10/20/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1251201 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Malfunction | Yes | Yes | Application Program Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was on an automated impella controller (aic) for mechanical circulatory support. during support, the placement signal waveform appeared with jagged waveforms and high-pressure readings, attempted re-plugging the driveline and checked the driveline for kinks. troubleshooting was unsuccessful and the device was replaced with another aic. the procedure was successful, with no report of harm attributed to this malfunction. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The date of patient death is unknown. the device was returned for evaluation. console logs from the reported day of event are consistent with complaint and show large number of error messages related to invalid placement signal samples due to low snr, followed by a ¿placement signal not reliable¿ alarm. device analysis: it¿s been determined that battery 2 was defective due to cell imbalance that caused the battery to shut down. the battery failure was unrelated to any optical placement signal issues. inspection of the console revealed large amount of contamination on the optical fiber of pump connector, which could not be cleaned, and resulted in an air gap between fibers, causing low snr and placement signal issues. after the optical pump connector (aka blue receptacle) was replaced, the console¿s 5s parameters were within specification. this report is being filed as part of a retrospective review of historical records.
15249 | 1220648-2024-22739 | 20645048 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20645048&pc=OZD | 02/15/2023 | 11/08/2024 |  | 11/08/2024 |  |  | 02/15/2023 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1620355 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Malfunction | Yes | Yes | Application Program Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  | No | Initial submission | [9 empty] | Us complainant reported that the patient was placed onto impella cp support due to acute myocardial infarction / cardiogenic shock. while on support, the aic experienced a controller error yellow alarm that was resolved by transferring to another console. the patient ultimately expired, but there was no allegation against the aic or impella related to the patient¿s death. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | Initially, the logs were incomplete; retrieved the imc, rt, and lt logs through telnet. console logs from the reported day of the event are consistent with a complaint and show that at some point during case start, optical bench data acquisition was stopped which resulted in controller error alarm and perceived placement signal issues. the controller error alarm as due to the software issue that the optical bench stops sampling, so no data interpreted from the bench. the placement signal is available, but the controller is not reading due to a software glitch. the console was tested in an engineering lab. the root cause of the controller error is not determined due to the inability to reproduce. the optical system operated within specification . further analysis of console data, performed later by sme on the software team, revealed a software defect, where a process erroneously stops optical bench data acquisition. this issue manifests if the pump is unplugged unexpectedly at a certain time during case start without a prior prompt from the case start wizard. this report is being filed as part of a retrospective review of historical records.
15250 | 1220648-2024-22741 | 20645115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20645115&pc=OZD | 12/14/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 12/14/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, JP |  |  | ABIOMED, INC. |  |  | US |  | 1453610 |  | 0042-0000-JP |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. during support, the automated impella controller (aic) became partially blank on the purge history screen. the device was replaced with another aic, and there was no reported patient harm. | [5 empty] | D9: no device return date available. the device was returned and evaluated, and the investigation for the reported display issue has been completed. data analysis: aic logs did not contain anything that highlighted the intermittent/ momentarily missing purge history. device analysis: this purge history was correctly presented when the console was inspected upon return in field service (fs). fs tested the functionality of the purge history with this console by running a test over time, entering the purge information (when prompted), and checking whether the history is populated in a timely manner. the software mechanism that populates this table was most likely delayed which was due to the aic software (internal). this report is being filed as part of a retrospective review of historical records.
15251 | 1220648-2024-22682 | 20645184 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20645184&pc=OZD | 07/01/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 07/01/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1468305 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced an optical signal issue and the console indicated the placement signal was unreliable. this aic was replaced with a new aic, which reportedly resolved the issue. there was no patient harm or injury reported. | [5 empty] | The investigation into the optical signal issue has been completed. data analysis: imc logs confirmed the reported failure and revealed a placement signal unreliable alarm (opm signal invalid ¿ 1036) was issued on the complaint date. rt logs showed the calculated placement signal railed to the max value of 3000 once the mc became pulsatile as the pump was positioned correctly. the snr value also slowly decreased to an abnormally low value (<150). see the attachments section for imc and rt logs. device analysis: the console was tested thoroughly on an engineering bench. running a known good optical pump and purge did not reproduce the failure. as received, the blue pump receptacle had contamination with a hair-line fracture or scratch which was not in the region of the optical fiber and most likely did not contribute to the failure. once cleaned, the 5s parameters and optical power was within specification according to the fc procedure (0042-7316). however, the snr was marginal (2673) and close to the threshold of 2000. the root cause of the placement signal issues was most likely due to a defective optical bench pca. this report is being filed as part of a retrospective review of historical records.
15252 | 1220648-2024-22700 | 20645214 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20645214&pc=OZD | 05/02/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 05/03/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1114580 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. while on support, the aic experienced a purge issue that was resolved by replacing the console. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported purge issue has been completed. the product was returned, and the issue was confirmed. the console logs from the reported day of event are consistent with complaint and show numerous errors related to purge operation: piston blocked, sau_purge too busy..., home not found, purge fluid not detected, and eventually purge system failure alarm (#417). this resulted in deairing step not being able to complete. the console's purge system was tested in danvers, and although no severe malfunctions were observed, there were recurring purge-related errors ("imc_ahp error: 0x90...") present in imc logs, especially after homing the motor/cassette. as a troubleshooting step the pud pca was temporarily replaced with a known-good one, and test repeated. this time there were no purge-related errors recorded in logs, and purge system operated within specification. per the results of the clinical review and investigation, the root cause was determined to be a defective pud pca (board). the cause of pud pca defect was not determined. this report is being filed as part of a retrospective review of historical records.
15253 | 1220648-2024-22693 | 20645256 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20645256&pc=OZD | 02/12/2023 | 11/08/2024 |  | 11/08/2024 |  |  | 02/12/2023 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 2023206963 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Application Program Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. the staff was not able to turn off the aic. the white cable was unplugged, however the aic continued to say to unplug the white cable. it was reported that the procedure was successful, and this issue occurred after pump removal. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: the date of the patient's death is unknown. d4: the serial number and udi are unknown. the cause of the issue was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records.
15254 | 1220648-2024-22709 | 20645290 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20645290&pc=OZD | 04/21/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 04/21/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1314857 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Connection Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that automated impella controller (aic) failed to recognize white cable plugged in during setup, causing the software not to advance to the next setup steps. a new aic was used, and setup was successful. it was also noted that while attempting to download logs, a download error received. the download was attempted on the back side of the aic usb port, which the error occurred at 2% of the download, then the download was attempted from the side port where the connector was removed from the usb port, which the error code occurred at 45% progress. multiple usb sticks were used as well. no reported harm or injury to the patient. | [5 empty] | The investigation for the reported issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the issue was due to bugs or defects in software (excludes firmware). this report is being filed as part of a retrospective review of historical records.
15255 | 1220648-2024-22698 | 20645313 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20645313&pc=OZD | 07/02/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 07/02/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1147985 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) did not recognize the purge cassette during startup. this aic was replaced with another aic and there were no further issues. there was no patient harm reported. | [5 empty] | D9: the date of the devices return for evaluation is unknown. e1, e3, e4: the name of the initial reporter and occupation is unknown. the investigation into purge disc/flag issues has been completed. the impella automated controller was returned for investigation. the returned device was tested and the issue with properly detecting the purge pressure disc was reproduced, and the purge flag was confirmed to be broken. the cause of the issue was due to the broken purge flag. this report is being filed as part of a retrospective review of historical records.
15256 | 1220648-2024-22714 | 20645337 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20645337&pc=OZD | 06/09/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 06/09/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1153712 |  | 0042-0000-EU |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) that during maintenance, the battery failure alarm is not beeping at all. there was no patient involvement. | [5 empty] | D9 - date device returned to manufacturer: unknown. the investigation into the buzzer / alarm malfunction issue has been completed. troubleshooting the issue led to replacement of the pbm assembly and fixing the console fully functional, reference the attached fs report. the pbm buzzer was confirmed to be out of specification. this report is being filed as part of a retrospective review of historical records.
15257 | 1220648-2024-22708 | 20645339 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20645339&pc=OZD | 06/10/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 06/10/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1360581 |  | 0042-0010-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) that when switched on in ac-mode the controller displays a battery failure alarm. batteries at 0% and not charging. there was no report of patient harm or injury. | [5 empty] | The investigation into the a/c power issue has been completed. the automated impella controller (aic) was returned for investigation. imc logs show battery failure alarms upon bootup. there are also battery comm. alarms (events #352 and 353). failure mode was reproduced in fs upon receiving the console. the console battery switch was found to be disengaged. once the battery switch was engaged, console boots up with 0% charge (on aic display) but no battery failure or comm. alarm. this seems like a 'misuse' case where either the battery switch was disengaged while in the field, or the console batteries were depleted. it cannot be confirmed because the console had a software update in the field, and there are no logs from before the update. when plugged in, the battery display on the batteries show that they are charging. the console was plugged in overnight to allow the batteries to charge completely. the console batteries were completely discharged and then charged back up to verify battery health. the cause of the battery failure alarm was due to depleted batteries this report is being filed as part of a retrospective review of historical records.
15258 | 1220648-2024-22745 | 20645387 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20645387&pc=OZD | 12/12/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 12/13/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1619563 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) developed a screen malfunction. there was no reported patient involvement. | [5 empty] | The device was returned and evaluated, and the investigation for the reported display issue has been completed. data analysis: aic logs did not contain any information related to the display failure. device analysis: over long period of time, the console was run, while the display was manually monitored. occasionally, the display would exhibit artifacts, specifically horizontal darker sections. due to the intermittent nature of the failure mode, a picture could not be taken that captured this failure in full detail. after replacing the lcd display, no issues were observed. the root cause of the display issues was most likely due to the lcd display. however, the root cause is unable to be determined. this report is being filed as part of a retrospective review of historical records.
15259 | 1220648-2024-22696 | 20645399 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20645399&pc=OZD | 12/15/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 12/15/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1098585 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Unintended Application Program Shut Down | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that their automated impella controller ( aic) rebooted unexpectedly during patient support. following the reboot, the console returned to the previously set p level and an unexpected controller shutdown alarm appeared. as a result of the failure, the aic was replaced with another console and patient support continued. there was no patient alleged injury resulting from the noted issue; however, care was withdrawn and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported issue has been completed. imc logs are consistent with the complaint. console was on patient support steadily, suddenly console process ossiopsens skipped heartbeat during multiple checks. imc logs indicating 2 processes inactive, and pmd reset, the root cause of the unexpected controller shutdown was due to a software (sw)process failure. the root cause of the sw process failure was not determined. however, the console passed pm and other standard testing after proper repairs. this report is being filed as part of a retrospective review of historical records.
15260 | 1220648-2024-22691 | 20645411 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20645411&pc=OZD | 03/28/2023 | 11/08/2024 |  | 11/08/2024 |  |  | 03/28/2023 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1692374 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Decrease in Pressure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission |  |  |  |  | 03/28/2023 |  |  |  |  | The user facility reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the automated impella controller (aic) exhibited purge volume critically low alarm. this occurred despite resetting the volume. the resolution for this issue is unknown. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: date of patient death is unknown the device was returned for evaluation. the root cause of console mistakenly issuing purge volume alarm was not determined, because the failure mode was not reproduced. this report is being filed as part of a retrospective review of historical records.
15261 | 1220648-2024-22725 | 20645425 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20645425&pc=OZD | 12/30/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 12/30/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1548487 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that the complainant reported an automated impella controller was failing to hold a charge. no report of patient harm or injury. | [5 empty] | E1, e3, e4: the name of the initial reporter and occupation is unknown. the investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
15262 | 1220648-2024-22694 | 20645459 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20645459&pc=OZD | 10/16/2024 | 12/13/2024 |  | 11/08/2024 |  |  | 10/16/2024 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025518494 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that that patient on impella 5.5 support was brought to the interventional radiology (ir) for tunneled central line placement. upon transfer to the table, a hematoma immediately formed on right axillary cutdown site. impella white cord and catheter carefully managed, and the nurse stated at no point did the cable get pulled. impella position verified unchanged. manual pressure held. the patient was brought back to the room and sandbag was placed over site. vital signs stable. the patient was taken back to the operating room for pocket evacuation. the patient survived the event. | [5 empty] | The investigation into access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding was determined to be patient movement because patient developed the sudden hematoma developed after transfer. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-22694. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
15263 | 1220648-2024-22699 | 20645482 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20645482&pc=OZD | 05/27/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 05/27/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1313442 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) shutdown during patient support and attempts to reboot did not resolve as it just kept circling. additional information was provided to the investigator of the complaint indicating that a backup console was used and there was no patient harm related to this event. | [5 empty] | The investigation into the shutdown issue has been completed. the impella automated controller was returned for investigation. the representative supplied the investigator with the following information regarding the event indicating the issue may have been due to the software. after the console was set aside, the console would not boot successfully (repetitively rebooting), which explained the phrase ¿kept circling¿. there was mention of this failure mode occurring after the software (sw) had been upgraded, but the console had been utilized without any similar complaints since the sw upgrade (preventative maintenance) to v8.4.1 in january 2021. there was one complaint on january 29th, 2021, but the failure mode/ root cause was unrelated to this investigation¿s failure mode. the account¿s complaint history had also been reviewed, and there was only one complaint associated with a different console related to this failure mode. the data logs revealed the device stopped abruptly on the complaint date. video of the aic display from the impella connect servers revealed that the console was on the placement signal screen and then suddenly rebooted. the returned console as received would not boot up. the logs could not be retrieved. there were issues accessing the /userflash directory. after swapping with known-good compact flash cards (cfcs), the console would boot up successfully. the root cause of the console repetitively rebooting was due to the one of the original compact flash cards was corrupted, affecting normal boot up of the aic. this failure was most likely due to a 4-in-1 malfunction or power to 4-in-1 from the power battery manager (pbm). the component-level root cause of the product malfunction is unknown. the unexpected controller shutdown was most likely due to a 4-in-1 malfunction or power to 4-in-1 from the pbm. the component-level root cause of the product malfunction is unknown. this report is being filed as part of a retrospective review of historical records.
15264 | 1220648-2024-22731 | 20645487 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20645487&pc=OZD | 02/14/2022 | 12/26/2024 |  | 11/08/2024 |  |  | 02/14/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1219724 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported a console failure when attempting to download the logs for a sensor complaint that occurred on this device. rebooted several times and still had the same issue. in addition, when setting up the catheter, a calibrating sensor screen appeared. additional information was provided that noted the patient expired. | [5 empty] | Automated impella controller (console) logs from the reported day of event are consistent with complaint and show an alarm "placement signal not reliable" (#1036) 2h43m after pump (b)(4) was started. the operator(s) disabled placement monitoring, and support continued for another 10 hours. real time log data confirms that raw optical placement signal from that moment forward had value of -32768 and snr was reported as 0. when the console was powered-up on the next day, the optical bench malfunction was manifested upon boot-up with alarms generated. there is also evidence of an unsuccessful attempt to download console logs via usb. after console was brought to danvers, the placement signal issue was immediately reproduced, and identified as optical bench failure, resulting in increased current draw, causing pbm to reduce supply voltage which, in turn, causes the bench to be unable to communicate and power the lamp. after the optical bench was replaced, the issue was resolved. inability to download data on the usb stick was not reproduced during testing (using either usb port on the console's i/o panel or the usb port on the back of rlm), and was most likely caused by a defective usb stick used by operator(s). the cause of the console issue was a defective optical bench. this report is being filed as part of a retrospective review of historical records. | The aic (automated impella controller) device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records.
15265 | 1220648-2024-22748 | 20645533 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20645533&pc=OZD | 04/24/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 04/24/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1215712 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, automated impella controller (aic) placement signal alarm occurred, then pump stopped after startup. aic was switched and alarm was resolved. additional information noted that the care was withdrawn and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise | [5 empty] | The investigation for the reported optical signal issue and aic - pump stop has been completed. the product was returned and the reported issues were not confirmed. the cause of the optical signal issue was use related, most likely due to the pump was not plugged in fully. the cause of the aic - pump stop issue was improper positioning of the device. this report is being filed as part of a retrospective review of historical records.
15266 | 1220648-2024-22710 | 20645676 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20645676&pc=OZD | 04/14/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 04/14/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1175563 |  | 0042-0000 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that after the white cable was connected, the catheter not identified. automated impella controller (aic) was restarted, and the problem persisted. the console was switched with no issues with the new console. there was no reported harm or injury to the patient. | [5 empty] | The investigation for the reported issues has been completed. the device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the cause of the issue was likely due to bugs or defects in software (excludes firmware). this report is being filed as part of a retrospective review of historical records.
15267 | 1220648-2024-22715 | 20645732 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20645732&pc=OZD | 12/17/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 12/17/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1327996 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | Unintended Application Program Shut Down | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a controller error message. it was reported that this aic was replaced with a new console. there was no patient harm reported. | [5 empty] | The investigation into the controller error has been completed. the impella automated controller was returned for investigation. the data analysis of the logs for the reported day of event were analyzed but an issue matching the exact description of a controller error could not be identified. during aic testing of the returned device, the console operated within specification, and no unexpected shutdowns or other irregularities were observed. inspection of the console however, revealed a loose metal screw inside the console housing, which likely had shorted portion of the power battery manager circuitry, causing momentary loss of power and subsequently aic shutdown/restart. the most likely cause of the aic shut down issue found during the investigation was a loose fastener. the controller error was not confirmed as the device logs did not indicate there was a controller error. this report is being filed as part of a retrospective review of historical records.
15268 | 1220648-2024-22697 | 20645745 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20645745&pc=OZD | 07/07/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 07/07/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1095636 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the automated impella controller (aic) did not recognize the purge disc. this aic was replaced with another aic and had no further issues. there was no patient harm reported. | [5 empty] | The investigation into purge disc/flag issues has been completed. the impella automated controller was returned for investigation. the data analysis of the logs was not necessary as the returned device testing reproduced the issue, and purge flag was confirmed to be broken (missing at blue disc retainer). the root cause of the purge cassette not being recognized was due to the broken purge flag. this report is being filed as part of a retrospective review of historical records.
15269 | 1220648-2024-22712 | 20645833 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20645833&pc=OZD | 04/12/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 04/12/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1102537 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during an automated impella controller (aic) case start, the yellow warning alarm prompted the staff that the optical sensor was not supported. no patient harm or injury noted. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the pmd hw revision issue (causing optical sensor not supported during case start) was unable to be determined but most likely occurred during the software upgrade. this report is being filed as part of a retrospective review of historical records.
15270 | 1220648-2024-22721 | 20645855 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20645855&pc=OZD | 03/15/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 03/15/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1116344 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Application Program Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that after an hour of support, the aic (automated impella controller) alarmed for a failure and was replaced. the ifu was followed, and a backup controller was used. additional information was provided that noted the patient expired. | [5 empty] | The investigation for the reported aic - controller error (yellow alarm) issue has been completed. the product was returned; however, the root cause of the issue could not be determined since the issue could not be reproduced. console logs from the reported day of event are consistent with complaint and show that 45 minutes after pump 363146 was started, a controller error alarm due to optical bench communication crc error was presented. this was preceded by "opm crc error" event 2 minutes earlier and "invalid sample due to ambient pressure out of range" event (events, unlike alarms, are not visible to the operator). in about 1 minute since the alarm, the issue self-resolved, and did not reoccur for the rest of the support. rtlog data associated with this event confirmed this to be a known pattern failure "transient loss of optical data". abiomed is aware of the issue and is working on appropriate corrective action. there¿s no need for repairs if the condition self-corrects. repair: perform functional check of the console. this report is being filed as part of a retrospective review of historical records.
15271 | 1220648-2024-22705 | 20645888 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20645888&pc=OZD | 04/27/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 04/27/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1203420 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the patient was placed onto impella cp support for high-risk surgery. while on support, the automated impella controller (aic) experienced a controller failure red alarm that was resolved by transfer to another console. there was no reported harm or injury to the patient. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the impella defective alarm was most likely due to intermittent glitch of the impellatronic's epm circuitry. the issue could not be reproduced during testing. this report is being filed as part of a retrospective review of historical records.
15272 | 1220648-2024-22729 | 20645983 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20645983&pc=OZD | 10/12/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 10/12/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1263210 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Failure to Read Input Signal | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. it was reported that during preparation for the procedure the controller not recognizing purge disk. the device was replaced with another aic and the procedure was successful. with no reported harm due to this malfunction. the procedural outcome was patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: the date of patient death is unknown. the device was returned for evaluation. console purge assembly was inspected. upon opening the purge assembly door, it was confirmed the purge disc retainer had broken/snapped off of the purge assembly. this report is being filed as part of a retrospective review of historical records.
15273 | 1220648-2024-22711 | 20646020 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20646020&pc=OZD | 04/18/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 04/18/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1194666 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the patient was placed onto impella cp support due to acute myocardial infarction /cardiogenic shock. while on support, the automated impella controller (aic) experienced a controller error yellow alarm that was resolved by transfer to backup aic. there was no reported harm or injury to the patient. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the communication loss is undetermined. this report is being filed as part of a retrospective review of historical records.
15274 | 1220648-2024-22703 | 20646060 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20646060&pc=OZD | 05/05/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 05/05/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED. INC. |  |  | US |  | 1142691 |  | 0042-0000-CA |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | [9 empty] | The complainant reported that upon receipt of the automated impella controller (aic), the devices display was blank. no patient involvement. | [5 empty] | The device was returned for evaluation. device analysis shows that one of the pins of the lvds connector on lcd was bent and the lvds connector cable was also damaged. it is possible that excessive force may have damaged the components. this report is being filed as part of a retrospective review of historical records.
15275 | 1220648-2024-22702 | 20646072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20646072&pc=OZD | 05/08/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 05/08/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1142507 |  | 0042-0000-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. the biomedical engineer reported that after each purge fluid change through the purge fluid bag, the menu prompts the aic alarm purge disc not detected. there was no patient harm or injury reported. | [5 empty] | The investigation into purge disc/flag issues has been completed. the impella automated controller was returned for investigation. a visual inspection noted a cracked purge sensor disk retainer, preventing normal operation of the flag. the console was repaired. the cause of the disk unable to detect was a cracked purge disk retainer. this report is being filed as part of a retrospective review of historical records.
15276 | 1220648-2024-22718 | 20646130 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20646130&pc=OZD | 05/23/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 06/04/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1095654 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) when beginning case and hit ¿case start¿ on the aic. the purge cassette was placed, and the bag of dextrose spiked. startup wouldn¿t progress beyond this point and an error message would state no fluid detected. restarted process several times and ultimately turned on another aic to attempt the process. the second aic was successful. there was no report of patient harm or injury. | [5 empty] | The investigation into the pump not detected issue has been completed. the impella automated controller (aic) was returned for investigation. data analysis: imc logs confirmed the reported complaint, revealing that a purge fluid not detected screen was issued. the rt logs showed that purge pressure would increase but not build when the stepper motor was ticking. device analysis: the failure mode was not reproduced throughout testing with a known-good pump and purge cassette/ disc. the cause of the low purge pressure during the case start procedure was not determined. his report is being filed as part of a retrospective review of historical records.
15277 | 1220648-2024-22720 | 20646179 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20646179&pc=OZD | 05/31/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 06/07/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1124694 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Noise, Audible | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) that when inserting the purge cassette, a grinding noise could be heard. a second purge assembly was used and had the same issue. the purge cassette was placed into another aic console with no issues. there was no report of patient harm or injury. | [5 empty] | The investigation into purge disc/flag issues has been completed. the impella automated controller was returned for investigation. there were no entries in console logs on the reported day of event (2021-06-03), but logs on 2021-05-31 indicate numerous purge-related issues: piston blocked errors, purge fluid not detected messages, and communication errors with sau_purge / pud (see attached logs). at some point, alarm #501 (pbm voltage out-of-spec) was presented, implicating bad vpurge voltages. when console was brought to danvers for testing, the issue was not immediately reproduced, but couple minutes after purge start, an unusual noise was heard, and alarm #417 was triggered (purge system failure - piston blocked). inspection of the stepper motor assembly did not reveal any dextrose contamination, but the motor shaft was much more difficult to spin by hand, compared to other consoles (with console powered off). once the stepper motor was removed and disassembled, it's been observed that grinding noise may have been related to the relative alignment of parts of motor housing, that were held by 4. if parts of stepper motor were not well aligned (coaxial with shaft), the shaft was difficult to rotate by hand. if this situation was occurring during console operation, increased resistance of stepper motor would cause increased load, and pbm would regulate power and provide more current, while lowering vpurge voltage. as far as numerous sau_purge communication issues observed in logs - those were most likely caused by pud pca malfunction. the root cause of the grinding noise was defective stepper motor of purge assembly. the cause of stepper motor defect was not determined. this report is being filed as part of a retrospective review of historical records.
15278 | 1220648-2024-22722 | 20646181 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20646181&pc=OZD | 12/05/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 12/05/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1123073 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Read Input Signal | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. during support, the aic was unable to recognize new purge disc during routine purge cassette change. the device was replaced with another aic without any further issues. it was reported that the procedure was successful. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: date of patient death is unknown the device was returned for evaluation. device logs from the complaint date revealed that the console issued the purge disc not detected alarm (event set # 402) several times through a long time. re-insertion of pc/ purge disc was observed, yet the console still had problem in detecting purge disc. device evaluation: the console was tested after arriving in danvers. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken and contaminated by glucose residue. this report is being filed as part of a retrospective review of historical records.
15279 | 1220648-2024-22735 | 20646196 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20646196&pc=OZD | 02/27/2023 | 11/08/2024 |  | 11/08/2024 |  |  | 02/27/2023 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1276151 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Hispanic |  | Death | Yes | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  | No | Initial submission | [9 empty] | Us complainant reported that the patient in acute myocardial infarction/cardiogenic shock was placed onto impella cp support due to acute myocardial infarction / cardiogenic shock. while on support, the automated impella controller (aic) experienced a placement signal issue that was resolved by switching to another console. survival outcome at explant noted the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported placement signal issue has been completed. the product was returned, and the placement signal issue was not confirmed. console logs from the reported day of event were consistent with complaint and indicated that the reported issue is a known pattern failure caused by a temporary loss of optical placement signal. console was inspected in danvers, and its optical system was found to be operating within specification. this report is being filed as part of a retrospective review of historical records.
15280 | 1220648-2024-22740 | 20646211 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20646211&pc=OZD | 01/22/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 01/22/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1577697 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Application Program Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was needing an impella cp device for mechanical circulatory support. the complainant reported an alarm emitting from the automated impella controller (aic). the initial aic was replaced, and the case continued. additional information was provided that noted the patient expired. | [5 empty] | The investigation for the reported aic - controller failure (red alarm) issue has been completed. the product was returned; however, the root cause of the issue could not be determined since the issue could not be reproduced. this report is being filed as part of a retrospective review of historical records.
15281 | 1220648-2024-22742 | 20646234 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20646234&pc=OZD | 10/10/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 10/10/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2023240278 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Failure to Read Input Signal | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. upon changing purge cassette tubing, aic alarmed ¿purge disc not detected¿. the device was replaced with another aic. the procedure was successful with no reported harm to the patient due to this malfunction. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: the date of the patient death is unknown. the device was returned for evaluation. console purge assembly was inspected. it was observed that the purge disc retainer was broken off of the purge assembly. this report is being filed as part of a retrospective review of historical records.
15282 | 1220648-2024-22707 | 20646236 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20646236&pc=OZD | 04/22/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 04/22/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1450701 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the patient was placed onto impella 5.5 support due to cardiomyopathy. while on support, the automated impella controller (aic) experienced a controller failure red alarm for which the customer was advised to switch to a backup controller. there was no reported harm or injury to the patient. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the issues running the pump was unable to be determined because the console and data were not returned for review. this report is being filed as part of a retrospective review of historical records.
15283 | 1220648-2024-22753 | 20646293 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20646293&pc=OZD | 04/18/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 04/18/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1143011 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was noted when setting up the impella catheter the optical sensor did not detect initially. staff unplugged and tried again. the catheter was recognized and started, but no aortic valve placement signal was displayed. issue was resolved by swapping aic. patient expired on support. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported optical signal issue has been completed. the product was returned and the optical signal issue was not confirmed. the cause of the loss of opm comm was unable to be determined due to the inability to reproduce. this report is being filed as part of a retrospective review of historical records.
15284 | 1220648-2024-22713 | 20646306 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20646306&pc=OZD | 02/07/2023 | 11/08/2024 |  | 11/08/2024 |  |  | 02/07/2023 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1079579 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  |  |  | Death | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. during support, the aic alarmed indicating that there was air in the purge. the nurse attempted to de-air the purge using the aic guidance but all soft buttons and rotary knob on the aic were non-responsive. several attempts were made; however, the aic was locked up and did not respond to any buttons pushed. the device was replaced with another aic and the issue was resolved with no patient harm. the procedural outcome noted that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation into aic issues are complete. the device was received for evaluation. the device logs are consistent with the complaint. air in purge system alarm was triggered at the end of case start wizard, prompting the user to enter de-air purge wizard. despite the de-air wizard performed, the purge pressure was unable to build up. when the de-air wizard finished, the air in purge system alarm was triggered again. however, this time multiple of keyboard error was observed in logs instead of user pressing the corresponding keys to enter de-air purge wizard. device evaluation: the root cause of the kbd non-responsive was highly likely due to defective kbd assembly. the root cause of air in the purge alarm was unable to determine, as the condition was not reproduced. this report is being filed as part of a retrospective review of historical records.
15285 | 1220648-2024-22728 | 20646308 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20646308&pc=OZD | 10/05/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 10/05/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1480337 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  |  | Death | Yes | Yes | Application Program Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. it was reported that the patient coded, and the staff turned the aic power on, they noticed it was taking an abnormally long time to complete the boot up process and be ready for use. the staff and physician stated that the boot up time was between 4-5 minutes before it was ready to use. it was reported that the procedure was successful, however the procedural outcome is patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: the date of the patient death is unknown the device was returned for evaluation. the logs show that there was a skipped boot on (b)(6) 2022. this would have caused minute delay in powering up the console which could have been misinterpreted as ¿long boot up time¿. skipped boot does not likely point to malfunctioning of the console. the console was further power cycled for 150 times, out of which it had only 4 skipped boots (refer attachment) followed by a normal boot sequence each. this was compared with a known good console which had 4 skipped boots in 150 power cycles. device analysis: the console was tested in the lab, the reported issue of ¿long boot up time¿ was not determined due to inability to reproduce. this report is being filed as part of a retrospective review of historical records.
15286 | 1220648-2024-22730 | 20646317 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20646317&pc=OZD | 03/03/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 03/03/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1203422 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Application Program Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. there was an aic failure during support. the pump was switched to a new controller and support continued. additional information was provided that noted the patient expired. | [5 empty] | The investigation for the reported aic - controller failure (red alarm) issue has been completed. the product was returned, and the issue was confirmed. the cause of the aic issue was a defective impellatronics board. this report is being filed as part of a retrospective review of historical records.
15287 | 1220648-2024-22738 | 20646349 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20646349&pc=OZD | 12/09/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 12/09/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1545370 |  | 0042-0010-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that there was an automated impella controller (aic) with a battery charging issue. no report of patient involvement, harm, or injury. | [5 empty] | The investigation for the reported issue has been completed. the device was returned for evaluation. the issue was reproduced, and the console did not charge the batteries. when it was plugged into ac mains, the led above the rotating knob stayed off. after switching the controller on, the power plug icon was crossed out indicating no connection to power. the cause of the console's inability to charge batteries was failure of the internal power supply (ps). the cause of ps failure was not determined. this report is being filed as part of a retrospective review of historical records.
15288 | 1220648-2024-22750 | 20646350 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20646350&pc=OZD | 12/09/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 12/09/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1148981 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) exhibited a battery failure. no patient involvement. | [5 empty] | The investigation into the battery power issue has been completed. the automated impella controller (aic) was returned for investigation. the cause of the battery failure alarm was due to a defective battery (rapid decline in cell voltage). the exact cause of the defective battery was undetermined. this report is being filed as part of a retrospective review of historical records.
15289 | 1220648-2024-22746 | 20646421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20646421&pc=OZD | 05/28/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 05/28/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1137669 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. an attempt to transfer from one aic to another. the white cable was plugged into ic2553 but the screen that normally pops up with restart pump option never did. patient became unstable before being able to attempt again. additional information noted that the care was withdrawn and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported optical signal issue has been completed. the product was returned, and the optical signal issue was confirmed. the cause of the automated impella controller (aic) issue was a defective impellatronics board. the pump detection issue was most likely due to the rarely occurring epm crystal issue. the impellatronic pca does not have permanent damage, but this intermittent pump detection failure occurs infrequently and already has a sw mitigation in place. this report is being filed as part of a retrospective review of historical records.
15290 | 1220648-2024-22719 | 20646460 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20646460&pc=OZD | 09/27/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 09/27/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | ABIOMED, INC. |  |  | GM |  | 1240269 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Female |  |  |  | Malfunction | Yes | No | Pumping Stopped | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | Eu complainant reported patient was placed on impella 5.5 with smart assist for hemodynamic support. while on support the aic (automated impella controller) experienced multiple pumps stops. issue was resolved by switching aic. | [5 empty] | Data analysis: console logs from the reported day of event are consistent with complaint and show that after pump 364317 was successfully started and ran without any issues for about 6 hours, a sudden pump stop occurred, as evidenced by alarm #6 (pump stop due to motor driver guard error or its supply voltage out of spec), preceded by pe_cpld error messages. pump was auto-restarted successfully, but same alarm and pump stop symptom occurred again soon. the auto-restarts failed to permanently resolve the pump stop issue. rtlog data shows that there were no motor current irregularities before the first pump stop. please see attached imclogs ic4362.pdf and screenshot, rtlogs.png. device analysis: after cosnole ic4362 was brought to fs aachen for testing, upon console booted up, there was a permanent beeping sound from the impellatronic and a controller error yellow alarm (loss of comm (cpld)), indicating internal communication issue (between pmd and cpld). after impellatronic assembly was temporarily replaced with a known-good one, the issue was resolved. the cause of impellatronic communication issue involving cpld was not investigated further. the root cause of the pump stop that was not unsuccessfully self-resolved was due to the defective impellatronic. repair: replace impellatronic assembly (0042-1115), and perform functional check, and preventive maintenance.
15291 | 1220648-2024-22732 | 20646466 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20646466&pc=OZD | 12/26/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 12/26/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1205390( |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Moisture Damage | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) soft keys would not work after the nurse spilled fluid on the console during changing of the purge fluid bag. this aic was replaced with another aic and the procedure was successfully completed. there was no patient harm reported. | [5 empty] | The investigation into the damaged softkey keyboard has been completed. the impella automated controller was returned for investigation. the data analysis of the logs were consistent with the complaint. the returned console was inspected, and no signs of fluid were found on the console exterior or interior. the console keyboard ribbon cable was checked and properly reseated. the console softkeys were tested through various menus with no issues. the console was run on a known good pump and purge system with no issues. in conclusion the lack of softkey response is most likely due to a wet console screen and/or wet staff hands. this is evident by the keyboard electrostatic discharge present at the time of the purge bag change. the console is functioning properly after completely drying. the root cause of the lack of softkey response is most likely due to a wet console screen and/or wet staff hands. this report is being filed as part of a retrospective review of historical records.
15292 | 1220648-2024-22704 | 20646488 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20646488&pc=OZD | 06/08/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 06/08/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1454057 |  | 0042-0010-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on an automated impella controller (aic) for mechanical circulatory support. while on support, there was a battery failure red alarm. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported battery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the battery issue and its relation to the impella automated controller device was not determined since the product, data logs, and sufficient clinical information were not provided. this report is being filed as part of a retrospective review of historical records.
15293 | 1220648-2024-22706 | 20646492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20646492&pc=OZD | 04/23/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 04/23/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1194878 |  | 0042-0000-US |  | No |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | Insufficient Information | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a controller failure red alarm with no resolution specified. there was no patient involvement. | [5 empty] | The investigation into the battery failure (red alarm) power issue has been completed. the automated impella controller (aic) was not returned for investigation. the cause of the battery failure alarm was due to depleted batteries. this report is being filed as part of a retrospective review of historical records.
15294 | 1220648-2024-22751 | 20646542 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20646542&pc=OZD | 12/07/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 12/07/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1182294 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that the battery failed upon automated impella controller (aic) start up. a loaner was sent to the customer. no report of patient involvement, harm, or injury. | [5 empty] | E1, e3, e4: the name and occupation of the initial reporter is unknown. the investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the battery failure was due to a defective battery (rapid decline in cell voltage). the exact cause of the defective battery was unable to be determined. this report is being filed as part of a retrospective review of historical records.
15295 | 1220648-2024-22752 | 20646553 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20646553&pc=OZD | 11/08/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 11/08/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1409295 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that there was a battery fault and is inoperable, a loaner automated impella controller (aic) was sent, and there was no mention of patient involvement. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the battery failure alarm was due to a defective battery (rapid decline in cell voltage). the exact cause of the defective battery was utd. this report is being filed as part of a retrospective review of historical records.
15296 | 1220648-2024-22726 | 20646563 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20646563&pc=OZD | 12/29/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 12/29/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1261094 |  | 0042-0000-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported failure of the device to charge. no report of patient harm or injury. | [5 empty] | D9: the date of the devices return for evaluation is unknown. e1, e3, e4: the name and occupation of the initial reporter is unknown. the investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the root cause of the battery charging issues was due to a defective pbm (charging circuitry). this report is being filed as part of a retrospective review of historical records.
15297 | 1220648-2024-22716 | 20646569 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20646569&pc=OZD | 08/04/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 08/04/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1635221 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) is not working and will not turn on. this issue occurred during training, no patient involvement. | [5 empty] | The investigation for the reported blank screen issue has been completed. the product was returned, and the issue was confirmed. data analysis: aic logs revealed purge unit driver (pud) communication issues. device analysis: boot up issues were reproduced. console would not pass system self-check. the root cause was determined to be a defective pud. this report is being filed as part of a retrospective review of historical records.
15298 | 1220648-2024-22755 | 20646571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20646571&pc=OZD | 04/19/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 04/19/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1253201 |  | 0042-0000-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic |  | Death | Yes | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an (aic) impella cp device for mechanical circulatory support. the us complainant had an automated impella controller failure with the cp. the aic was replaced and a backup controller was used. additional information noted that the care was withdrawn, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | D9: device return date is unknown. the investigation for the reported aic - controller failure (red alarm) has been completed. the product was returned and the aic - controller failure (red alarm) was confirmed. console was tested in (b)(6), and although the exact issue was not reproduced, results of repeated power-cycling testing revealed one failed boot-up sequence due to pmd (sau_motor_1) communication failure. this implicates impellatronic assembly, as both subcircuitries - epm and pmd - are located on this board. this report is being filed as part of a retrospective review of historical records.
15299 | 1220648-2024-22734 | 20646593 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20646593&pc=OZD | 09/23/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 09/23/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1107437 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic |  | Death | Yes | Yes | Unintended Application Program Shut Down | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the automated impella controller (aic) exhibited "unexpected controller shutdown switch to back up controller". the device was replaced with another aic. it was reported that the procedure was successful with no harm to the patient due to the device malfunction. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: the date of patient death is unknown. the device was returned for evaluation. the console was tested on an engineering lab bench. the failure modes were not reproduced while running an optical pump simulator. the root cause of the detection issue of pump disconnection was due to an aic software issue. the cause of console shutdown was due to power cord not plugged in ac, aic shutdown due to battery depleted. this report is being filed as part of a retrospective review of historical records.
15300 | 1220648-2024-22754 | 20646704 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20646704&pc=OZD | 04/20/2022 | 11/08/2024 |  | 11/08/2024 |  |  | 04/20/2022 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1169678 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. automated impella controller (aic) failed to detect the purge cassette during use of the device. a new cassette and new aic were used. additional information noted that the care was withdrawn and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the product was returned and the aic - purge disc/flag issues was confirmed. the console was tested after arriving in fs. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. this report is being filed as part of a retrospective review of historical records.
15301 | 1220648-2024-22756 | 20646705 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20646705&pc=OZD | 12/06/2021 | 11/08/2024 |  | 11/08/2024 |  |  | 12/06/2021 | 11/08/2024 |  |  |  |  | 11/08/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1551790 |  | 0042-0010-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that there was a battery pack failure, tried some troubleshooting but that did not fix the problem. the instructions for use were followed and a backup controller was used. this was discovered during prep, and there was no harm to the patient. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the root cause of the battery failure was due to a defective battery (rapid decline in cell voltage). the exact cause of the defective battery could not be determined. this report is being filed as part of a retrospective review of historical records.
15302 | 1220648-2024-22274 | 20626308 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20626308&pc=OZD | 07/26/2022 | 11/06/2024 |  | 11/07/2024 |  |  | 07/26/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1339213 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Crack | Failure to Read Input Signal | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported purge disc slide notch is broken on automated impella controller (aic) .console returned for investigation / repair. no patient is involved . | [5 empty] | The investigation into the purge disc/flag issue has been completed. the automated impella controller (aic) was returned for investigation. as per the data logs reported complaint date is (b)(6)2022, however the issues related to the purge disc actually occurred on (b)(6) 2022. on (b)(6) 2022, staff began case start but several "purge disc not detected" alarms triggered. console arrived in danvers. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. this report is being filed as part of a retrospective review of historical records.
15303 | 1220648-2024-22251 | 20626344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20626344&pc=OZD | 09/23/2022 | 11/06/2024 |  | 11/07/2024 |  |  | 09/23/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1272375 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported that the automated impella controller (aic) console displayed the error message "failure optical sensor". this aic was replaced with another aic resulting in no issues with the replacement aic. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported optical bench issue has been completed. the product was returned, and the issue was confirmed. console logs from the reported day of event are consistent with complaint and show that during case start placement signal was not reliable, resulting in event 1036, case start wizard error message, and eventually alarm #1036. log data showed very low snr values. console was tested at aachen fs, as per instructions provided by the complaint investigator. inspection of the pump plug connector revealed dextrose residue and contamination of the optical fiber connector, which were deemed as most likely contributors to the reported malfunction. after blue receptacle replacement, ¿5s¿ parameters were initially passing, but during repeated testing, the contrast parameter was too low and optical bench had to be replaced. per the results of the clinical review and investigation, the root cause was determined to be contamination. this report is being filed as part of a retrospective review of historical records.
15304 | 1220648-2024-22275 | 20626364 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20626364&pc=OZD | 05/17/2021 | 11/06/2024 |  | 11/07/2024 |  |  | 05/17/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1446166 |  | 0042-0000-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller experienced a battery failure red alarm and battery error yellow alarm with no specified resolution. there was no patient involvement. | [5 empty] | D.9 exact date of when the console was returned for evaluation is unknown. e.1, e.2 and e.3 initial reporter name and health professional/occupation is unknown. the investigation for the reported issue has been completed. while prepping for a case, the console gave both a red and yellow alarm. battery status was only 50%. the clinical staff did not proceed with the case on this console. the logs showed a battery 1 comm. failure alarms and battery failure alarms. it was noted from the logs that the console was not turned on for about four months. it is possible the console was not plugged in and batteries depleted. the failure mode was reproduced in field service upon arrival. only battery 2 showed sufficient charge on the battery display. battery 1 had a blank display. batttest showed 'smb' (smbus error) on battery 1. swapping the battery terminals showed that the failure followed the battery and power nattery manager circuit board was not at fault. replacing the battery set fixed the failure mode. the cause of the battery failure alarm was most likely due to battery 1 malfunction. it is unable to determine as to why only battery one malfunctioned. this report is being filed as part of a retrospective review of historical records.
15305 | 1220648-2024-22289 | 20626413 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20626413&pc=OZD | 07/26/2021 | 11/06/2024 |  | 11/07/2024 |  |  | 07/26/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1288547 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the automated impella controller experienced a controller error yellow alarm in training. no clinical details regarding resolution or patient involvement/condition were provided. no patient was involved. | [5 empty] | The investigation has been completed. the console (ic5059) was sent back for further investigation. when inspecting internal components, it was observed that there was a marking on the top of the power battery manager (pbm). the pbm was removed and another marking was seen on the bottom of the board, mirroring the location from the top of the board. the failure to boot in the field was due to a failed pbm. this report is being filed as part of a retrospective review of historical records.
15306 | 1220648-2024-22290 | 20626418 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20626418&pc=OZD | 07/22/2021 | 11/06/2024 |  | 11/07/2024 |  |  | 07/22/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1556637 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant's automated impella controller (aic) was found to have a battery failure. there was no patient use and no harm or injury possible. | [5 empty] | E1: name of initial reporter is unknown. the investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. battery comm failure alarm is due to one malfunctioning battery. the exact root cause of the malfunctioning battery is unknown. this report is being filed as part of a retrospective review of historical records.
15307 | 1220648-2024-22291 | 20626438 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20626438&pc=OZD | 07/19/2021 | 11/06/2024 |  | 11/07/2024 |  |  | 07/19/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1100209 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  |  | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported an automated impella controller (aic) failure alarm occurred during patient care. there was no adverse impact to the patient. the instructions for use was followed, and the backup controller was used. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the controller error (opm comm crc error ¿ event set 1045) was unable to be determined because the issue was unable to be reproduced. this report is being filed as part of a retrospective review of historical records.
15308 | 1220648-2024-22293 | 20626445 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20626445&pc=OZD | 07/14/2021 | 11/06/2024 |  | 11/07/2024 |  |  | 07/14/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, EU |  |  | ABIOMED, INC. |  |  | US |  | 2022083595 |  | 0042-0010-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  |  | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. during use, the automated impella controller (aic) produced an "error in console - switch to standby controller" alarm. the aic was replaced with a backup unit, and there was no patient harm reported. | [5 empty] | D9: the date of device return is unknown. the investigation for the reported controller error (yellow alarm) has been completed. the device logs were reviewed and confirmed that the console issued a controller error, and simultaneously, the placement signal disappeared. this occurred with the pump plugged in, and after the pump was physically disconnected from the console, the console would not shutdown normally because the disconnection was not recognized due to the lack of communication to the optical bench. the console was tested at abiomed aachen in collaboration with the post market engineering team. an impella cp optical pump was run with purge, but the failure mode was not reproduced. therefore, the root cause could not be determined. this known pattern failure, which is characterized by a temporary loss of placement signal and triggering of an optical bench-related controller error alarm, has a low probability of reoccurrence. this report is being filed as part of a retrospective review of historical records.
15309 | 1220648-2024-22292 | 20626656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20626656&pc=OZD | 07/15/2021 | 11/06/2024 |  | 11/07/2024 |  |  | 07/15/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | ABIOMED, INC. |  |  | US |  | 1156294 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) produced an "aic error / switch to backup controller" message during a routine check. there was no reported patient involvement. | [5 empty] | The device was not returned, but the investigation for the reported controller error (yellow alarm) has been completed with data provided. console logs from the reported day of event are consistent with complaint and show controller error alarm #24 due to loss of communication with kbd. the issue was resolved after power-cycling the console. when console was received at the lab and inspected, a damaged connector on the kbd assembly was discovered. the cause of the controller error alarm was communication glitch/loss between kbd and 4-in-1 due to compromised connector on the kbd assembly. the cause of connector damage was not determined. this report is being filed as part of a retrospective review of historical records.
15310 | 1220648-2024-22261 | 20626657 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20626657&pc=OZD | 07/07/2022 | 05/19/2025 |  | 11/07/2024 |  |  | 07/07/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, JP |  |  | ABIOMED, INC. |  |  | US |  | 1483134 |  | 0042-0000 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant in japan had an aic (automated impella controller) in at the office, not in a patient use when it was found to have display issues. there was no harm or injury from the display/vga issue. should an aic fail in patient support the ifu states to use a backup console. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Data analysis: the imc and rlm logs were reviewed and deemed irrelevant for this failure mode. device analysis: the console was returned to japan fs. upon arrival, the aic was powered on and connected to an external monitor via vga from the rlm. the issue of no signal was reproduced. japan fs replaced the rlm module and tested the vga output of the rlm. the vga connection to the external monitor behaved as expected with no issues. this report is being filed as part of a retrospective review of historical records.
15311 | 1220648-2024-22280 | 20626665 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20626665&pc=OZD | 05/17/2021 | 11/06/2024 |  | 11/07/2024 |  |  | 05/17/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1383039 |  | 0042-0000-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. during support, the automated impella controller experienced a controller error yellow alarm that was resolved by replacing the console. there was no reported patient harm. | [5 empty] | The investigation for the reported controller error (yellow alarm) has been completed. the impella device was not received from the customer. the console logs from the reported day of event are consistent with the complaint and showed an alarm (opm communication stopped). the cause of the opm communication corruption was unable to be determined because the issue was unable to be reproduced. this report is being filed as part of a retrospective review of historical records.
15312 | 1220648-2024-22303 | 20626684 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20626684&pc=OZD | 06/17/2021 | 11/06/2024 |  | 11/07/2024 |  |  | 06/17/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1134037 |  | 0042-0000-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced a controller error yellow alarm that occurred during preventive maintenance. there was no reported patient involvement. | [5 empty] | D9: date of the device return is unknown. the device was evaluated at a field service center, and the investigation for the reported controller error (yellow alarm) has been completed. while being served under its annual field service maintenance procedure, the device failed to meet the spec for optical bench functionality and issued a controller error alarm (defective optical sensor lamp - #1039). replacing the optical bench gained the console its normal functionality. the root cause of the controller error alarm was due to a defective optical bench pca. this report is being filed as part of a retrospective review of historical records.
15313 | 1220648-2024-22307 | 20626704 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20626704&pc=OZD | 06/11/2021 | 11/06/2024 |  | 11/07/2024 |  |  | 06/11/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1311843 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  |  | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the patient was placed onto impella cp support due to cardiomyopathy. while on support, the automated impella controller (aic) experienced a controller error alarm yellow alarm, the pump position waveform became flat and disappeared 10 minutes later. it was reported that the device exhibited the same alarm (date unknown), and it resolved in 2 minutes. no further alarms were reported. no harm or injury to the patient reported. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the controller alarms was invalid optical placement signal due to undetermined communication issue with optical bench. the root cause of the communication issue was unable to be determined because the issue was unable to be reproduced. this report is being filed as part of a retrospective review of historical records.
15314 | 1220648-2024-22308 | 20626770 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20626770&pc=OZD | 06/03/2021 | 11/06/2024 |  | 11/07/2024 |  |  | 06/03/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1451015 |  | 0042-0000-US |  | No |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was on an automated impella controller (aic) for mechanical circulatory support. while on support, the aic experienced a battery failure with no resolution specified. no reported harm or injury to the patient. | [5 empty] | The investigation into the a/c power issue has been completed. the automated impella controller (aic) was not returned for investigation. according to the biomed contact, the issue has since resolved, and the account would not be returning the console. they have not experienced any further issues. the cause of the battery temperature high alarm was unable to be determined because the console and data were not returned for review. this report is being filed as part of a retrospective review of historical records.
15315 | 1220648-2024-22310 | 20626806 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20626806&pc=OZD | 05/03/2021 | 11/06/2024 |  | 11/07/2024 |  |  | 05/21/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1313571 |  | 0042-0000-UK |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) that the aic experienced a battery failure red alarm, which was not plugged in by the site for many months due to inactivity. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported battery failure issue has been completed. the product was returned (unknown return date), and the issue was confirmed. data review: the console logs revealed that it had not been powered on for 8 months and there were battery failure (transport connection blown/missing) alarms, confirming the complaint. also, the console was not powered on for more than 7 months, but it could not be determined whether the console¿s batteries were charged at all during that period. device analysis: fs confirmed the issue and swapped with known good batteries and ran sections 18, 19, and 20 from the pm procedure (0042-7309) to validate the pbm would properly charge the batteries. the console charged the newly installed batteries within the specifications defined in the pm procedure. per the results of the clinical review and investigation, the root cause was determined to be due to depleted batteries, which were not routinely charged. this report is being filed as part of a retrospective review of historical records.
15316 | 1220648-2024-22249 | 20626830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20626830&pc=OZD | 09/23/2022 | 11/06/2024 |  | 11/07/2024 |  |  | 09/23/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1111980 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) alarmed ¿controller error¿. the placement signals would intermittently turn to dashes and then reappear. then the ¿controller error¿ alarm appeared. the patient remained stable and the team switched to a new aic. alarms resolved after aic transfer. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported controller error issue has been completed. the product was returned, and the issue was confirmed. console logs from the reported day of event are consistent with complaint and show controller alarm due to optical bench communication crc error (#1045) as well as multiple errors ¿invalid_bad_pressure_sample_mask¿ falsely indicating ambient pressure out of range. rtlog data showed that optical placement signal, as well as all data from optical bench, were intermittent, and toggling between valid measurements and fixed values of -16384. this is consistent with a particular ¿flavor¿ (manifestation) of known pattern failure of transient loss of optical data. ltlog data from backup console ic6217 used later with the same pump showed no issues with optical bench data. the console was extensively tested in danvers, but the malfunction was not reproduced. no communication issues were observed, and optical bench ¿5s¿ parameters were within specification. it is most likely that some unspecified malfunction (affecting serial communication) of the 4-in-1 assembly contributed to this issue. per the results of the clinical review and investigation, the root cause was determined to be most likely a defective 4 in 1 board. this report is being filed as part of a retrospective review of historical records.
15317 | 1220648-2024-22259 | 20626843 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20626843&pc=OZD | 09/13/2022 | 11/06/2024 |  | 11/07/2024 |  |  | 09/13/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1502180 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported that the automated impella controller (aic) had a failure, the reporter did not indicate what type of failure occurred. the aic was replaced with another aic. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported controller failure issue has been completed. the product was returned, and the issue was not confirmed. data analysis: aic logs revealed that an impella 5.5 was disconnected from the console and then reconnected with eeprom reading crc errors and ultimately an impella defective alarm. the software (sw) watchdog then detected the calculator process was not responding, which led to a sw watchdog reboot. despite the console having rebooted, the pump could not be read successfully due to crc errors, resulting in an impella defective alarm. device analysis: the console was tested by aachen fs in collaboration with engineering. the pump eeprom reading mechanism was tested multiple times with no issues reproduced. the complete system could not be tested for any intermittent read/write issues that would result in corruption of the i2c data, which is transmitting eeprom pump data. the eeprom read issues were followed by a failure of the calculator process resulting in a sw watchdog partial/ full reset. no direct link could be made between the eeprom read issues and the reboot due to the inability to reproduce the failure. per the results of the clinical review and investigation, the impella defective alarm was unable to be reproduced. the suspected root cause was i2c communication issues during read of the impella 5.5 eeprom, which involves the impellatronics epm circuitry. this report is being filed as part of a retrospective review of historical records.
15318 | 1220648-2024-22317 | 20626865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20626865&pc=OZD | 04/22/2021 | 11/06/2024 |  | 11/07/2024 |  |  | 05/10/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1462815 |  | 0042-0010-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) that the aic does not detect the impella. the aic was turned off and on and still did not work. the impella was then connected to a new aic, and it worked without issues. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported pump not detected issue has been completed. the product was returned (exact return date unknown), and the issue was confirmed. the unrecognized pumps is known with this software revision and could be fixed by repowering the console. the customer stated they re-powered, but there is no record of second boot-up in the log file. this console is also due to its sw upgrade to v8.5 which has an implemented fix addressing the unrecognized connected pump - gid-dft-1506015: pump not detected (multiple external complaints). in-house troubleshooting was not able to replicate unrecognized pump ruling out new failure mode. per the results of the clinical review and investigation, the root cause was determined to be a known issue with the crystal of the epm on the impellatronic board. this report is being filed as part of a retrospective review of historical records.
15319 | 1220648-2024-22318 | 20626876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20626876&pc=OZD | 05/09/2021 | 11/06/2024 |  | 11/07/2024 |  |  | 05/09/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1090690 |  | 0042-0000-EU |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) that the aic experienced a controller failure red alarm. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported controller failure issue has been completed. the product was returned (no exact return date available), and the issue was confirmed. data analysis: complaint was confirmed via the console logs. imc logs revealed that the console triggered a controller failure (loss of communication to pud) and would also not pass system self-check few days later. device analysis: the failure was reproduced in fs aachen. per the results of the clinical review and investigation, the root cause was determined to be a defective pud board. the cause of the defective pud board was unable to be determined. this report is being filed as part of a retrospective review of historical records.
15320 | 1220648-2024-22253 | 20626888 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20626888&pc=OZD | 05/22/2023 | 11/06/2024 |  | 11/07/2024 |  |  | 05/22/2023 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1662880 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The us complainant reported that an automated impella controller (aic) had controller failure red alarm resulting in pump stop. plugging the pump to another aic resolved the issue. no patient harm was reported. | [5 empty] | The investigation for the reported aic - controller failure (red alarm) has been completed. the product was returned, and the complaint was confirmed. aic and rt logs align with the pattern failure mode, transient loss of optical data. aic logs confirmed the reported failure mode given there was a controller error alarm triggered during the clinical procedure. impella stopped, air in purge system and retrograde flow alarms were triggered on the complaint day but were false alarms due to this loss of optical data/ inability to detect pump disconnection. real time (rt) logs confirmed that this was a pattern failure mode in which all signals received from optical bench (placement, snr, and contrast) are lastly reported as 16384 values and then no more samples were recorded. the absence of snr samples impacted the console¿s ability to recognize the pump disconnection and required the user to perform a hard shutdown. the console was tested on an engineering lab bench. the failure mode was not reproduced while running an optical pump simulator. there was no issue with the console hardware. the root cause of the controller error was due to an aic software issue. this report is being filed as part of a retrospective review of historical records.
15321 | 1220648-2024-22319 | 20626890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20626890&pc=OZD | 05/08/2021 | 11/06/2024 |  | 11/07/2024 |  |  | 05/08/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1204375 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 22 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) that during changing the purge cassette, the aic was stuck on the change the box screen. multiple unsuccessful attempts were made to re-insert the cassette. additionally, the grounding plug on the aic reportedly was broken off. this aic was replaced with another aic. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported flag issue has been completed. the product was returned, and the issue was confirmed. device analysis: after console was received in house, it was observed that the purge flag was broken. per the results of the investigation, the root cause of the failed purge flag was likely due to the wear and tear of the plastic component as well as additional unnecessary force when inserting the purge disc into the purge disc retainer. during prior cases. this report is being filed as part of a retrospective review of historical records.
15322 | 1220648-2024-22309 | 20626895 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20626895&pc=OZD | 05/28/2021 | 11/06/2024 |  | 11/07/2024 |  |  | 05/29/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1082389 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) that there was a controller failure during console-to-console transfer from outside the hospital. white the power cable transfer was successful; the purge transfer was unsuccessful. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported controller failure has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible, and the clinical details provided were not sufficient to determine a root cause. this report is being filed as part of a retrospective review of historical records.
15323 | 1220648-2024-22287 | 20626898 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20626898&pc=OZD | 08/08/2021 | 11/06/2024 |  | 11/07/2024 |  |  | 08/08/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1203510 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported an automated impella controller (aic) failure, aic - controller failure (red alarm), that occurred during preparation of the device. a new aic was used without any harm to the patient. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the failure mode was most likely use related (not following case start wizard prompts). this report is being filed as part of a retrospective review of historical records.
15324 | 1220648-2024-22297 | 20626964 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20626964&pc=OZD | 07/07/2021 | 11/06/2024 |  | 11/07/2024 |  |  | 07/07/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1428325 |  | 0042-0010-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had an automated impella controller (aic) battery failure observed by the impella connect team at installation failure. there was no patient use and no harm or injury. the team utilized a backup console. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the controller alarms was battery failure (open fuse). the cause of battery failure was not investigated. this report is being filed as part of a retrospective review of historical records.
15325 | 1220648-2024-22285 | 20626993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20626993&pc=OZD | 08/16/2021 | 11/06/2024 |  | 11/07/2024 |  |  | 08/16/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1181672 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant's automated impella controller (aic) had a battery failure (yellow alarm). on boot up battery failure controller showed 50% power, the resolution for this issue is unknown. no report of patient involvement. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of battery communication error alarm was a defective power battery manager board (pbm). the cause of the defective pbm is unknown. this report is being filed as part of a retrospective review of historical records.
15326 | 1220648-2024-22316 | 20626995 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20626995&pc=OZD | 05/11/2021 | 11/06/2024 |  | 11/07/2024 |  |  | 05/11/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1251615 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary insertion was on an automated impella controller (aic) for mechanical circulatory support. on initial startup, they were unable to prime. switched purge cassette but still unable. switched to new aic and primed without problem. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported purge issue has been completed. the product was returned, and the issue was confirmed. console logs from the reported day of event are consistent with complaint and show multiple purge-related errors: piston blocked, sau_purge communication issues, "home not found" (unable to home purge motor), and eventually a "purge fluid not detected" error during case start wizard. also values of purge pressure, flow, and purge ticks recorded in log for the pump were observed to be abnormal and/or invalid. after the console was brought to danvers for troubleshooting, the issue was initially not reproduced, but after 10 minutes of running with known-good purge cassette and pump simulator, "piston blocked" events started occurring. after temporarily replacing pud pca with a know-good board, the purge system was tested, and no further purge errors were reported and measured values (pressure, purge ticks) were within specification. per the results of the clinical review and investigation, the root cause was determined to be a defective pud pca. the cause of the board defect was unable to be determined. this report is being filed as part of a retrospective review of historical records.
15327 | 1220648-2024-22298 | 20626999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20626999&pc=OZD | 07/02/2021 | 11/06/2024 |  | 11/07/2024 |  |  | 07/02/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1468360 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 25 YR |  | Female |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was placed onto impella 5.5 support due to cardiomyopathy. while on support, the automated impella controller (aic) experienced battery charging / other that was resolved with no further detail provided. | [5 empty] | The investigation for the reported aic - battery charging/other issues" has been completed. the product was returned and the "aic - battery charging/other issues" was confirmed. complaint cause was most likely due to the power battery manager (pbm) incorrectly reporting battery temperatures when ac power was connected because at the moment ac power was connected, battery data was sampled. this reading falsely reports the battery temp as something oor triggering the battery alarm. when the next sample was taken 10 seconds later, the pbm read a correct battery data sample. this could be seen in the logs when the battery temp high alarm was triggered, and exactly ten seconds later the alarm cleared. this false reporting will not require a replacement of the pbm; however, batteries will be replaced as they were "unsealed" in order to read their highest recorded temperature. this report is being filed as part of a retrospective review of historical records.
15328 | 1220648-2024-22299 | 20627010 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20627010&pc=OZD | 06/25/2021 | 11/06/2024 |  | 11/07/2024 |  |  | 06/25/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1229946 |  | 0042-0000-EU |  | Yes |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the automated impella controller (aic) exhibited an error message. the console was replaced. no report of patient or injury. | [5 empty] | D9: the date the device was returned is unknown. the device was returned for evaluation. the logs confirmed the reported alarms on the complaint date. an impella defective (error reading eeprom) alarm was issued after just booting up (without pump connected). the device was returned for evaluation. the console was tested with a known good pump, and the failure mode was reproduced. replacing the impellatronics pca resolved the issue. the cause of the impella defective alarm was due to a malfunction of the impellatronics epm circuitry. this report is being filed as part of a retrospective review of historical records.
15329 | 1220648-2024-22311 | 20627051 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20627051&pc=OZD | 05/21/2021 | 11/06/2024 |  | 11/07/2024 |  |  | 05/21/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1156150 |  | 0042-0000-US |  | No |  |  | No |  | No |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) exhibited a battery communication failure yellow alarm during preventative maintenance. the console was boot cycled and when the console completely powered on, the console buzzer began to alarm. it was noted that the console was on battery power. no patient involvement. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of beeping power battery manager board (pbm) was a defective pbm. the cause of the defective pbm is unknown. this report is being filed as part of a retrospective review of historical records.
15330 | 1220648-2024-22301 | 20627061 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20627061&pc=OZD | 06/18/2021 | 11/06/2024 |  | 11/07/2024 |  |  | 06/18/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1351867 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced battery charging issues. no clinical details regarding resolution or patient involvement / condition were provided. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the battery failure alarm was depleted batteries. the console was left unplugged for over eight months. this report is being filed as part of a retrospective review of historical records.
15331 | 1220648-2024-22313 | 20627074 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20627074&pc=OZD | 05/14/2021 | 11/06/2024 |  | 11/07/2024 |  |  | 05/14/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1184609 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a battery communication failure yellow alarm. the device was losing communication with battery intermittently and won't charge. no report of patient involvement and no report of patient harm or injury. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of battery comm. failure was a defective pbm. the cause of the defective pbm is unknown. this report is being filed as part of a retrospective review of historical records.
15332 | 1220648-2024-22314 | 20627088 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20627088&pc=OZD | 05/18/2021 | 11/06/2024 |  | 11/07/2024 |  |  | 05/18/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1173723 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | Battery Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) console failed self-check during power up, the customer cycled the battery, however it still failed. the device was removed from use, no report of patient involvement, injury, or harm. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the device logs revealed that there were issues booting up due to communication issues. when the device was received it failed to successfully boot. the power battery manager board (pbm) installed had corrosion. the corroded pbm was replaced with a known-good pbm, and the system check, and communication issues were resolved. the cause of boot up issue was corrosion near the c69 and c70 supercapacitors on the pbm. this report is being filed as part of a retrospective review of historical records.
15333 | 1220648-2024-22286 | 20627097 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20627097&pc=OZD | 08/10/2021 | 11/06/2024 |  | 11/07/2024 |  |  | 08/10/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1096324 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a failure of the device automated impella controller (aic) to hold a charge. there was no indication of patient involvement. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. log analysis of the battery data reveals that beginning before the complaint date of occurrence, cell 3 voltage of battery 1 had been declining before the complaint date. failure mode was confirmed, aic is unable to charge because battery 1 has failed. due to low/ flatlined voltage, battery 1 data is unavailable through ata or bq software. when console was booted for the first time in danvers, console alarms are consistent with what was seen in the field. when visually inspecting the batteries and pbm, it was observed that the status leds of battery 1 were completely blank. failed battery was checked, ocv was measured to be ~v2.9. compared to the other battery where the measured voltage was ~16v. the decline in cell 3 voltage of the failed battery is undetermined. this report is being filed as part of a retrospective review of historical records.
15334 | 1220648-2024-22271 | 20627107 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20627107&pc=OZD | 08/02/2022 | 11/06/2024 |  | 11/07/2024 |  |  | 08/02/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1462815 |  |  |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that while on support an automatic impella controller (aic) failure shut down occurred. this issue was resolved by switching consoles. the procedure was successful and no patient harm was reported. | [5 empty] | E1: the initial reporters name is unknown. the investigation for the reported aic - shutdown or restart issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. aic performed a series of soft resets (perceived as shutdowns) in response to failure of a software process (calculator) associated with eeprom reading issues. the cause of this issue was most likely a defective impellatronics board (its epm circuitry is responsible for eeprom reading). this report is being filed as part of a retrospective review of historical records.
15335 | 1220648-2024-22304 | 20627146 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20627146&pc=OZD | 06/04/2021 | 11/06/2024 |  | 11/07/2024 |  |  | 06/04/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1125342 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Hispanic | American Indian Or Alaskan Native | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was on an automated impella controller (aic) for mechanical circulatory support. while on support, a new staff member felt that the system did not purge through, however there were no alarms. on the side of caution, the decision was made to use another console. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported purge issue has been completed. the product was returned, and the issue was not confirmed. console logs from the reported day of event are consistent with complaint and show some irregularities during case start, related to de-airing, purge pressure drop, and multiple attempts to enter purge fluid info. however, no errors or alarms were present, and no evidence supporting claim of purge system malfunction. testing of the console and its purge system, performed in danvers, showed that console operated within specification. visual inspection of the purge system revealed slightly kinked ribbon cable of the purge pressure transducer, but no relation to the complaint issue could be established. per the results of the clinical review and investigation, the root cause was most likely use related. this report is being filed as part of a retrospective review of historical records.
15336 | 1220648-2024-22312 | 20627232 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20627232&pc=OZD | 05/20/2021 | 11/06/2024 |  | 11/07/2024 |  |  | 05/20/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1302712 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | Incomplete or Inadequate Priming | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary insertion was on an automated impella controller (aic) for mechanical circulatory support. while on support, the aic would not recognize when the impella was connected, also could not prime pump. attempted several times to connect and re-connect white cable without success. a back-up aic was used with successful pump connection and priming. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported pump not detected issue has been completed. the device was not received from the customer. the clinical details provided were sufficient to determine a root cause. in the imc logs, we see that during console boot-up the optical bench state machine (ossiopsens) gets stuck in post_start_state and does not reach post_ok_state. this prevents the software from reading pump eeprom. the root cause of the issue and its relation to the impella automated controller device bugs or defects in software (excludes firmware). this report is being filed as part of a retrospective review of historical records.
15337 | 1220648-2024-22300 | 20627243 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20627243&pc=OZD | 06/17/2021 | 11/06/2024 |  | 11/07/2024 |  |  | 06/17/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1353540 |  | 0042-0000-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a battery communication failure yellow alarm during preventive maintenance. the console was received with the battery level at zero. after the batteries were replaced, the batteries were at zero the next day with a controller alarm. no details regarding resolution were provided. no patient involvement. | [5 empty] | D9: the date of the device return is unknown. the device was returned for evaluation and the investigation for the reported issue has been completed. a review of the device logs was conducted and confirmed the battery failure alarms. during the evaluation of the device, one of the battery fuses was observed to be blown. battery connections were swapped on the power battery manager board (pbm) and another test was performed, showing the blown fuse was moving batteries/ positions. this observation led to investigation of the pbm. a known good pbm was placed into the console and the issue resolved. both batteries had shown intact fuses and passing charging currents. the complaint cause is due to a malfunctioning pbm, however the root cause of the issue with the pbm is unable to be determined. this report is being filed as part of a retrospective review of historical records.
15338 | 1220648-2024-22321 | 20627283 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20627283&pc=OZD | 04/16/2021 | 11/06/2024 |  | 11/07/2024 |  |  | 04/20/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1452267 |  | 0042-0010-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | Connection Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. during support, the aic asked for pump connection. the pump was plugged in, however there was no identification of the pump. the customer had another aic console and plugged in the same pump, the pump was identified and worked without an issue. no harm or issue with the patient in this situation, delay of implant was approximately 1-2min. | [5 empty] | D9: the date of the device return for evaluation is unknown. the device was returned for evaluation. the cause of the pump not detected failure was a defective impellatronic board. the cause of the defective impellatronic board is unknown. this report is being filed as part of a retrospective review of historical records.
15339 | 1220648-2024-22315 | 20627288 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20627288&pc=OZD | 05/11/2021 | 11/06/2024 |  | 11/07/2024 |  |  | 05/18/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1093343 |  | 0042-0000-EU |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  |  | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) that during prep, there was a aic alarm "failure in purgesystem"; no cause to be determined and alarm resolved after a short time; about 35 hours later, aic alarms with "air in purgesystem"; the nurse said, that there was no "low filling purgefluid alarm" before. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported purge issue has been completed. the product was returned, and the issue was confirmed. imc logs show purge piston block during the case start. this triggers the purge system failure alarm (#417). the staff replaced the bad purge cassette with a new one and continued with the case without any more piston block errors. once the pump is plugged in, we see a crc error from sau_sens_eep_1 when reading the eeprom structure "pump limits data". once the pump motor is started, we see sau_motor_1 errors consistently. these errors point to the impellatronics board or its ribbon cable. later we see an air in purge system alarm (#404) which clears itself. logs show a forward purge tick but no gain in purge pressure. soon after this alarm, the staff performs a purge bag change which deairs the purge system and clears the #404 error. about 14 hours later we get a low purge pressure alarm. logs show that before the low pressure alarm, the purge pressure was actually going high while the purge flow was decreasing. this indicates biomaterial buildup at the pump end of purge line which is blocking the purge flow, hence increasing the purge pressure. towards the end of the case, we see purge volume low alarms. this is calculated from the purge bag volume (entered during purge bag change) and time (since last bag change). it has nothing to do with the purge assembly. no purge failures were reproduced in field service. the console was tested with pump simulator and purge for over 24 hours without any issues. the impellatronics ribbon cable was inspected for any misalignment, kinks or discoloration. a kink was observed in the cable. since the errors were not reproduced, it is possible the ribbon cable may have moved during shipping and fixed the misalignment. per the results of the clinical review and investigation, the root cause was determined to be due to purge cassette failing to prime because of the blocked piston. this report is being filed as part of a retrospective review of historical records.
15340 | 1220648-2024-22306 | 20627325 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20627325&pc=OZD | 06/11/2021 | 11/06/2024 |  | 11/07/2024 |  |  | 06/11/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1133541 |  | 0042-0000-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Female |  |  | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. while on support, the aic experienced a controller failure red alarm. it was recommended to the physician to switch to another aic. the physician decided to explant the impella instead. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported controller failure issue has been completed. the product was returned (unk date of return), and the issue was confirmed. console logs from the reported day of event are consistent with complaint and show multiple kbd communication errors (communication error, corrupt register), possible esd events (kbd: high guard level), and controller error alarms #24 "loss of communication with kbd". after console was send to aachen fs, it was tested but issue was not reproduced. inspection of the console - and specifically kbd assembly and cable, did not reveal any anomalies, defects, or evidence of corrosion or fluid ingress. per the results of the clinical review and investigation, the root cause of the controller alarm was corrupt communication between 4-in-1 and kbd, most likely due to 4-in-1 carrier malfunction. this report is being filed as part of a retrospective review of historical records.
15341 | 1220648-2024-22320 | 20627444 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20627444&pc=OZD | 05/06/2021 | 11/06/2024 |  | 11/07/2024 |  |  | 05/07/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1401385 |  | 0042-0010-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  | White | Malfunction | No | Yes | No Audible Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) white console cable became disconnected from the red hub. the screen showed flat placement and motor signals, and ¿pump stopped - alarm condition resolved¿ at the top. it was noted that the controller did not alarm. abiomed staff were present and resolved the issue by troubleshooting. no patient harm reported. | [5 empty] | The device was not returned for evaluation. the device logs revealed that an impella stopped (current mismatch alarm ¿ event set 119) was triggered when the pump was accidentally unplugged, but the alarm was resolved the same second it was triggered. this likely occurred because the console then recognized the pump disconnection, which resulted in the aic resolving the impella stopped alarm. the console was not returned for this investigation. the failure mode was reproduced on an engineering test console. once the console recognized the pump has been disconnected, the impella stopped alarm resolved. the cause of the impella stopped alarm resolving immediately was due sw¿s alarm criteria (classified as a sw defect in this case), which resolves alarm upon disconnection of the stopped pump. this report is being filed as part of a retrospective review of historical records.
15342 | 1220648-2024-22322 | 20627455 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20627455&pc=OZD | 04/16/2021 | 11/06/2024 |  | 11/07/2024 |  |  | 04/16/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1502180 |  | 0042-0010-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. while on support, the automated impella controller (aic) experienced controller failure. the customer had a replacement aic onsite; the device was replaced with further reported issue. no patient harm reported. | [5 empty] | D9: the date of the device return for evaluation is unknown. the cause of the pump stop events was a defective impellatronic board. the cause of the defective impellatronic board is unknown. this report is being filed as part of a retrospective review of historical records.
15343 | 1220648-2024-22323 | 20628750 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20628750&pc=OZD | 04/28/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 04/29/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1075485 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock needed an impella for mechanical circulatory support. while on support, the automated impella controller experienced purge issues that were resolved by changing out to a backup console. there was no report of patient harm or injury. | [5 empty] | The investigation into the purge issue has been completed. the automated impella controller (aic) was returned for investigation. console logs from the reported day of event were consistent with the complaint and showed multiple "purge disk not detected" alarms that started during purge cassette change attempt. the console was observed that the purge disk detect flag was missing (broken). after the part was replaced, the console was able to correctly detect the purge disk. this report is being filed as part of a retrospective review of historical records.
15344 | 1220648-2024-22324 | 20628837 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20628837&pc=OZD | 04/06/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 04/06/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1560022 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  | White | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that at the end of a case, the connector cable was removed but the automated impella controller (aic) did not detect its removal. the aic would not shut down. there was no mention of patient harm. | [5 empty] | The investigation for the reported issue has been completed. the console logs from the reported day of event were consistent with the complaint and showed the console presenting with a message "shutdown requested but pump is connected" after the pump was stopped and purge cassette was removed, but no entries in the logs that would indicate the pump being disconnected. however, careful analysis of real time log data allowed to reconstruct the sequence of events which indicated that pump cable was actually disconnected at 13:58:39 on (b)(6) 2021, based on values of placement signal, signal-to-noise ratio (snr) and gain. real time logs also showed that the console registered optical placement signal glitch (value of -32768) shortly after the case start, but not long enough to cause a controller alarm. all other optical bench signals were valid. the logs from the entire case contained some isolated epm communication errors and some pmd communication issues. the console was troubleshooted but the issue was not reproduced during testing - either with optical or non-optical test pumps. the fact that the pump disconnection was not registered despite snr values below 150 thresholds, suggests that the failure had to be electrical in nature. the console was tested by an optical pump connector was being pulled out just enough to disconnect the optical fibers but keep contact pins engaged to maintain electrical connection, but the resulting symptoms were inconsistent with the malfunction observed in the field. because the console had functional check and software upgrade performed prior to being received by engineering, it could not be proved (or disproved) that a firmware corruption (epm or pmd) caused the malfunction during the case. after the console was opened and visually inspected, a minimally misaligned flat flex cable (between impellatronic and 4-in-1) was observed, so it was realigned and some testing was repeated. there were no more pmd comm. issues, but epm comm. issues reoccurred. in addition, it's been observed that each time a non-optical pump is plugged in, console logs contain errors. although there is no evidence directly relating these errors to the pump disconnection mechanism failure and the complaint issue was not reproduced, the evidence collected throughout testing indicates that the impellatronic assembly contributed to the investigated malfunction. after impellatronic assembly was replaced, error did not reoccur. the cause of the console not detecting pump being disconnected was a defective impellatronic board. the cause of the defective impellatronic board is unknown. this report is being filed as part of a retrospective review of historical records.
15345 | 1220648-2024-22325 | 20628959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20628959&pc=OZD | 05/04/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 05/04/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1180595 |  | 0042-0000-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a failure of the automated impella controller that it could not start. this did not occur during use with a patient. | [5 empty] | D.9 the exact date for when the device was returned for evaluation is unknown. e.1, e.2 and e.3 the first and last name of initial reporter and health professional/occupation are unknown. the investigation into the issue has been completed. console logs from the reported day of event were consistent with the complaint and showed that when the console was started while plugged into alternating current (ac), an alarm would show indicating battery failure. this was consistent with the inability to start the console when the ac cord was disconnected. the console was returned to germany field service and was tested. the issue was reproduced. further inspection of the batteries indicated that they were fully depleted and could not be charged. additional data collected, including internal battery status and power battery manager (pbm) charger status confirmed that batteries were fully depleted to the point that its internal electronics disabled charging or discharging of batteries but the communication between pbm and batteries was still possible. secondary analysis of console logs prior to the reported day of event showed that the console has not been used since its preventative maintenance one year prior, and evidently was placed in long-term storage not plugged into ac (which is inconsistent with the console instructions for use which instructs to keep the console plugged in, while in storage). after the batteries were replaced, it has been confirmed that the pbm was able to charge them. the cause of the inability to start the console was the fact that the batteries were fully depleted. this was due to storing the console for one year without ac connection. this report is being filed as part of a retrospective review of historical records.
15346 | 1220648-2024-22326 | 20629278 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20629278&pc=OZD | 10/13/2024 | 12/12/2024 |  | 11/07/2024 |  |  | 10/13/2024 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025536117 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported during impella cp support, the 88-year-old male patient experienced ventricular fibrillation (vf), arrested, and cardio-pulmonary resuscitation (cpr) was initiated. the patient coded for 15 minutes. post case support was provided in cath lab (ccl). patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation for the reported ventricular fibrillation/ventricular tachycardia (vf/vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf/vt could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
15347 | 1220648-2024-22327 | 20629536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20629536&pc=OZD | 10/13/2024 | 12/16/2024 |  | 11/07/2024 |  |  | 10/13/2024 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025548427 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient had an attempted impella 5.5 however it was unsuccessful. the staff could not make the turn at the aorta. multiple attempts were made with multiple wires with no success and the decision was made to abort the insertion. there was reported patient harm. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, the impella was unable to pass turn at aorta as per the clinical description provided. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
15348 | 1220648-2024-22328 | 20629855 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20629855&pc=OZD | 10/14/2024 | 12/12/2024 |  | 11/07/2024 |  |  | 10/14/2024 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025537957 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp support initiated for a 57-year-old male patient admitted in an acute myocardial infarction / cardiogenic shock (ami/cgs). the patient developed limb ischemia in the impella accessed limb and had thrombectomy performed. the pump was explanted after just over 4 1/2 hours. the team evaluated the decision to place an impella 5.5 pump after the impella cp explanted. the team looked to the flow and determined the ischemia to be acute on chronic vascular disease. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
15349 | 1220648-2024-22329 | 20629951 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20629951&pc=OZD | 10/01/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 10/01/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1337492 |  | 0042-0010-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 2.5 for mechanical circulatory support. while on support, the automated impella controller (aic) experienced a purge disc issue, which was resolved by changing the console. no patient harm reported. | [5 empty] | The investigation has been completed. the purge disc detection failure was confirmed by the disassembly of the purge system. purge disc detection flag was visibly broken. the root cause of the purge disc not detected was confirmed to be a broken flag. this report is being filed as part of a retrospective review of historical records.
15350 | 1220648-2024-22330 | 20630021 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20630021&pc=OZD | 09/20/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 09/20/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1158010 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) experienced a high battery temperature alarm. no clinical details regarding resolution or patient involvement / condition were provided. | [5 empty] | The investigation has been completed for battery charging. the console ((b)(6)) was sent back for further investigation. it was confirmed that persistent battery failure (transport connection blown/missing) was not due to a disengaged battery switch. the root cause of the false alarm for high battery temperature is due to the power battery manager misreading battery data during a single sample. this report is being filed as part of a retrospective review of historical records.
15351 | 1220648-2024-22331 | 20630088 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20630088&pc=OZD | 10/14/2024 | 12/10/2024 |  | 11/07/2024 |  |  | 10/14/2024 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025528034 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 83-year-old male patient in an acute myocardial infarction/cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. during impella support, the patient developed a brain bleed and decision made to stop heparin and explant impella. the patient is noted to be stable. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system: section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." section: warnings, cautions & precautions: "to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers." | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email.
15352 | 1220648-2024-22332 | 20630229 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20630229&pc=OZD | 10/12/2024 | 12/11/2024 |  | 11/07/2024 |  |  | 10/12/2024 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025522453 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Tachycardia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the a patient was on impella cp support. while on support, echocardiogram showed the impella was very shallow. the patient has ventricular tachycardias and does not flow on the impella. additionally, the lower leg extremity pulses were absent. the impella cp was weaned and explanted. a clot was noted in leg post closure and it was evacuated without issue. the patient survived the expalnt. | [5 empty] | The investigation for the thrombus, limb ischemia, and ventricular tachycardia (vt) has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus, limb ischemia, and ventricular tachycardia (vt) could not be determined as the device was not returned nor sufficient clinical details. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
15353 | 1220648-2024-22333 | 20630432 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20630432&pc=OZD | 10/13/2024 | 12/16/2024 |  | 11/07/2024 |  |  | 10/13/2024 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025552370 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, a massive hematoma to the right groin was noted after peelaway was removed. manual pressure was held. eventually femstop was placed and the decision was made to remove the impella due to excessive bleeding from the site. the patient survived the explant. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
15354 | 1220648-2024-22334 | 20630480 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20630480&pc=OZD | 07/28/2023 | 11/07/2024 |  | 11/07/2024 |  |  | 07/28/2023 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1438016 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) experienced optical signal issues. the console was replaced to address the issue. no patient harm reported. | [5 empty] | The investigation for the reported optical signal issue has been completed. the product was returned and the optical signal issue was confirmed. the root cause of the optical signal issue was determined to be a software issue. this report is being filed as part of a retrospective review of historical records.
15355 | 1220648-2024-22335 | 20630540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20630540&pc=OZD | 02/22/2023 | 11/07/2024 |  | 11/07/2024 |  |  | 02/22/2023 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | AUTOMATED IMPELLA CONTROLLER |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1205052 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced an issue in which the rotary knob did not scroll through the options during a patient management course. this was resolved by rebooting the controller, and there was no patient involvement. | [5 empty] | The investigation for the reported "aic - kbd/rotary knob issues" has been completed. the product was returned, and the reported issue was confirmed. data analysis: imc logs were reviewed, but not relevant to the failure mode. device analysis: during console testing, the failure mode was reproduced when a varistor on the impellatronic board was shorted. the reported intermittent issue of rpb malfunction was highly likely due to defective impellatronic board. this report is being filed as part of a retrospective review of historical records.
15356 | 1220648-2024-22336 | 20630597 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20630597&pc=OZD | 02/17/2023 | 11/07/2024 |  | 11/07/2024 |  |  | 02/17/2023 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1292317 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | Complainant reported that there was an automated impella controller ( aic) system self check failure. the aic was returned for investigation. no patient was involved. | [5 empty] | The investigation for the reported issue has been completed. the device and data logs were returned for analysis data analysis showed the failure originally occurred after an attempt to install the new 12.1 fw. the device outputted a system self check failure when the fw did not match what the console expected. the console reproduced the "system self check failure" in the lab. the console was attempted to revert the console back to 11.1.1 which succeeded for every subcomponent except for pbm (fw check.png). then the specific msc was attempted to be programmed but was still unsuccessful. upon analyzing the pbm, there was residue around the super cap which is indicative of the failure mode super cap corrosion. the cause of the aic issue was contamination of a communication line on the pbm.
15357 | 1220648-2024-22338 | 20630648 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20630648&pc=OZD | 01/04/2023 | 11/07/2024 |  | 11/07/2024 |  |  | 01/04/2023 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1637920 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) had a display issue/flickering. there was no reported patient involvement. | [5 empty] | The device was returned and evaluated, and the investigation for the reported display issue has been completed. data analysis: there was nothing in the logs that indicate the screen flickering. device analysis: the device failure was reproduced and confirmed during the investigation and through isolative testing can be isolated to the screen and not the 4-in-1. root cause: the cause of the aic issue was a flickering kbd/display. this report is being filed as part of a retrospective review of historical records.
15358 | 1220648-2024-22337 | 20630683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20630683&pc=OZD | 10/14/2024 | 12/10/2024 |  | 11/07/2024 |  |  | 10/14/2024 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025519536 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support, the 54-year-old female patient's labs were hemolyzed. the impella was repositioned but ultimately the team withdrew care, and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system: section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis." "patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis." | The investigation into hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-22337.
15359 | 1220648-2024-22339 | 20630686 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20630686&pc=OZD | 11/17/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 11/17/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1142691 |  | 0042-0000-CA |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Other | Foreign | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | Key or Button Unresponsive/not Working | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The biomedical engineer reported that the automated impella controller (aic) rotary knob broke. a loaner was sent to the customer to initiate return of the device. there was no patient harm reported. | [5 empty] | The investigation into the keyboard rotary knob issue has been completed. the impella automated controller (aic) was returned for investigation. the data analysis of the logs reviewed, but few suspicious log entries were found that might reveal the failure mode around the complained date. the user attempted to shut down the console by pressing the button on the side of the aic, but they tried a few times until the successful shutdown of the console, likely due to unable to confirm via the push button. no abnormalities were found during inspection within the console. the broken rotary knob issue was confirmed during testing and the intermittent rpb assembly replaced, the root cause was unable to be determined. this report is being filed as part of a retrospective review of historical records.
15360 | 1220648-2024-22341 | 20630770 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20630770&pc=OZD | 10/27/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 10/27/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1420045 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Nurse | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's nurse reported that during patient (no patient details provided) support, the automated impella controller (aic) screen was flickering. the aic was replaced with a backup unit, and no patient harm was reported. | [5 empty] | The device was returned and evaluated, and the investigation for the reported display issue has been completed. data analysis: imc logs were reviewed, but not relevant to the failure mode. device analysis: the flickering screen failure mode was initially not reproduced during idling. however, once the console glass keys were pressed hard, the screen dimmed. console was disassembled and the internal cables were inspected without any abnormalities. and after the console was re-assembled, the failure mode was no longer reproduced. the root cause of intermittent screen flickering failure mode was not determined. this report is being filed as part of a retrospective review of historical records.
15361 | 1220648-2024-22340 | 20630802 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20630802&pc=OZD | 11/11/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 11/11/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1164189 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported that when the automated impella controller (aic) was powered on, it displayed a blue screen with a message stating ¿system self check failed, change console immediately.¿ a backup aic was used for the upcoming case, and no patient harm was reported. | [5 empty] | The device was returned and evaluated, and the investigation for the reported system self-check error has been completed. failure mode was reproduced. upon visual inspection, the power battery manager (pbm) super cap shows signs of leaking causing a damaged communication line. the cause of the syscall error was contamination from the pbm cap's corrosion. this report is being filed as part of a retrospective review of historical records.
15362 | 1220648-2024-22342 | 20630834 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20630834&pc=OZD | 10/21/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 10/21/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1146850 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the purge disc was detected during training. the initial automated impella controller (aic) was exchanged for a new device. no patient was involved. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the product was returned and the aic - purge disc/flag issues was confirmed. aic logs were looked at but were found not relevant to the complaint. the console was tested after arriving in fs. the purge disc retainer was confirmed to be broken (broken blue disc retainer). this report is being filed as part of a retrospective review of historical records.
15363 | 1220648-2024-22343 | 20630876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20630876&pc=PYX | 06/19/2024 | 12/12/2024 |  | 11/07/2024 |  |  | 10/14/2024 | 11/07/2024 |  |  | 02/28/2024 |  | 11/07/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023204373 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding, the complainant reported a 62-year-old male patient with right heart failure was implanted with an impella rp device for mechanical circulatory support. the impella rp was supporting for 2 weeks when the pump was explanted. the team reported low flows with possible clot entrainment. there was no noted treatment for the low flows despite explant. | [5 empty] | The investigation into low pump flow has been completed. the impella device was not returned for evaluation. the cause of the low pump flow issue was most likely biomaterial, based on data log review. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-22343. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed.
15364 | 1220648-2024-22344 | 20630928 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20630928&pc=OZD | 08/31/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 08/31/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, EU |  |  | ABIOMED, INC. |  |  | US |  | 1479095 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) had screen flickering. there was no reported patient involvement. | [5 empty] | E1: name of initial reporter is unknown. the device was returned and evaluated, and the investigation for the reported display issue has been completed. when console was brought to field service, the issue was reproduced. inspection of the screen assembly revealed a misaligned connector on the lvds port of the tft array, that was unable to be fully latched, due to bent pins inside the connector. the issue presumably originated during aic manufacturing process, but has not been observed before. after the screen assembly and cable were replaced, the issue was resolved and there was no further flicker observed. the cause of the aic issue was a defective lvds port on the display assembly, which was most likely caused during manufacturing process. this report is being filed as part of a retrospective review of historical records.
15365 | 1220648-2024-22346 | 20631015 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20631015&pc=OZD | 06/21/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 06/21/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1311572 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) that the purge disc holder was broken off the aic console. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported purge disc/flag issue has been completed. the product was returned, and the issue was confirmed. data analysis: aic logs were not relevant to this failure mode. device analysis: the console was inspected on an engineering lab bench. the purge disc retainer was confirmed to be broken (broken blue disc retainer). per the results of the clinical review and investigation, the root cause occurred during normal use. this report is being filed as part of a retrospective review of historical records.
15366 | 1220648-2024-22347 | 20631028 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20631028&pc=OZD | 06/21/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 06/21/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1108091 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) that the console powered on correctly, but the soft keys were not producing results. unable to move past start case screen with the soft keys. attempted to turn the console off and back on without result. the console was switched out. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported kbd/rotary knob issue has been completed. the product was returned, and the issue was not confirmed. data analysis: imc logs from the complaint occurred date (6/21/2022) did not reveal any pressing on the kbd, which could also be due to kbd not responsive/ not producing any results. power cycles were observed, but case start wizard was never entered in all power cycles on 6/21. however, imc logs on 6/23/2022 revealed that console could successfully enter case start wizard with soft key selection. device analysis: the failure mode keyboard not responsive was not reproduced during several power cycles. the console could enter case start via pressing soft keys. inspection of internal connections did not find any abnormalities. per the results of the clinical review and investigation, the root cause was not determined since the failure mode was not reproduced. this report is being filed as part of a retrospective review of historical records.
15367 | 1220648-2024-22348 | 20631045 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20631045&pc=OZD | 10/13/2024 | 12/11/2024 |  | 11/07/2024 |  |  | 10/14/2024 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025546215 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had impella 5.5 pump support ongoing for a patient admitted in with non-st segment elevation myocardial infarction. the pump was placed in escalation from an impella cp. after six days, the team diagnosed a extensive intracranial hemorrhage and severe mass effect with hydrocephalus. the team stopped the heparin and an external ventricular drain was placed by neurology. the family decided to withdraw care and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock; section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿. | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email.
15368 | 1220648-2024-22349 | 20631055 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20631055&pc=OZD | 06/20/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 06/20/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1126035 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) a controller error alarm during use of the device. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported controller error issue has been completed. the product was returned, and the issue was not confirmed. data analysis: according to imc logs this is a known pattern failure in which all signals received from optical bench (placement, snr, contrast, lamp, gain) are interpreted as -16384 values. this lasts between 2 to 10 minutes, after which all signals go back to normal. device analysis: the console was placed on engineering lab bench running a pump, while monitoring optical communication lines using standard lab equipment. console was allowed to run overnight to reproduce failure mode. failure mode was reproduced, exhibiting the same symptoms reported by the field. loss of placement signal for a brief period of time, as well as a controller error. per the results of the clinical review and investigation, the root cause could not be determined. this report is being filed as part of a retrospective review of historical records.
15369 | 1220648-2024-22350 | 20631062 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20631062&pc=OZD | 06/10/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 06/10/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1076062 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility¿s biomedical engineer reported for an automated impella controller (aic) a broken blue purge clip. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported purge disc/flag issue has been completed. the product was returned, and the issue was confirmed. data analysis: imc logs are not relevant for this failure mode. device analysis: console purge assembly was inspected. upon opening the purge assembly door, it was confirmed the purge disc retainer had broken/snapped off of the purge assembly. per the results of the clinical review and investigation, the root cause was determined to be a defective component. this report is being filed as part of a retrospective review of historical records.
15370 | 1220648-2024-22345 | 20631098 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20631098&pc=OZD | 08/06/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 08/06/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1486033 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. during support, the automated impella controller (aic) screen began "shifting" and disappeared. the console was swapped with a backup unit, and there was no patient harm reported. | [5 empty] | The device was returned and evaluated, and the investigation for the reported display issue has been completed. device analysis: console was inspected and tested on an engineering lab bench. the lcd display was monitored for more than two hours with a camera to observe any artifacts related to the reported "screen shifting". this test was repeated when console was power by batteries and ac mains. no issues were observed. it is very unlikely that the lcd was mechanically shifting given no issues were observed. the lvds and backlight cables were secured properly as received. cable strain did not produce any artifacts on the lcd screen. the display issue was not reproduced throughout testing, and the cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
15371 | 1220648-2024-22351 | 20631277 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20631277&pc=OZD | 10/14/2024 | 12/12/2024 |  | 11/07/2024 |  |  | 10/14/2024 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025466314 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was on impella cp support and had some hemolysis showing up in the urine. additionally, the impella started getting an intermittent position alarm after being repositioned in bed. an echocardiogram showed the pump still in the left ventricle by one centimenters. the medical doctor did not feel comfortable advancing the impella cp back into position so the impella cp was removed. there was no harm to the patient. | [5 empty] | The investigation for the hemolysis and the positioning issues has been completed. data logs were reviewed which confirmed pump was in improper position corresponded with clinical description that triggered alarms impella position in aorta. no other issues were found on the logs. product was not returned for review. based on clinical description and data analysis, the root cause of positioning issue was most likely use related. the root cause of hemolysis was most likely due to pump was not in the proper position. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
15372 | 1220648-2024-22360 | 20631324 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20631324&pc=OZD | 10/14/2024 | 12/12/2024 |  | 11/07/2024 |  |  | 10/14/2024 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025519973 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the 71-year-old male patient was attempted to be implanted with impella 5.5, but procedure was aborted due to patient anatomy. no known patient harm or injury. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because implantation was aborted due to patient's anatomy. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.
15373 | 1220648-2024-22355 | 20631331 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20631331&pc=OZD | 06/06/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 06/06/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1080191 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) that it was not detecting the purge disc. discovered during training with demo equipment. there was no patient involvement. | [5 empty] | The investigation for the reported purge disc/flag issue has been completed. the product was returned, and the issue was confirmed. data analysis: imc logs from the complaint date revealed that the console issued the purge disc not detected alarm (event set #402) several times. device analysis: the issue with properly detecting the purge pressure disc was reproduced, and the purge flag was confirmed to be broken. per the results of the clinical review and investigation, the root cause was determined to be a defective component. this report is being filed as part of a retrospective review of historical records.
15374 | 1220648-2024-22358 | 20631335 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20631335&pc=OZD | 06/07/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 06/07/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1175561 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) the blue clip that holds the purge transducer was broken to the point that it came off the aic from behind the purge cassette door. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported purge disc/flag issue has been completed. the product was returned, and the issue was confirmed. device analysis: the console was tested after arriving in fs. the purge disc retainer was confirmed to be broken. per the results of the clinical review and investigation, the root cause was likely a combination of normal use and material weakness over repetitive use. this report is being filed as part of a retrospective review of historical records.
15375 | 1220648-2024-22364 | 20631339 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20631339&pc=OZD | 04/25/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 04/25/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1205609 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) experienced a error alarm. the console was replaced to address the issue. no patient harm reported. | [5 empty] | The investigation for the reported aic - controller error (yellow alarm) has been completed. the product was returned and the aic - controller error (yellow alarm) was confirmed. the root cause of the aic - controller error (yellow alarm) was determined to be defective impellatronics board. this report is being filed as part of a retrospective review of historical records.
15376 | 1220648-2024-22356 | 20631343 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20631343&pc=OZD | 10/12/2024 | 12/13/2024 |  | 11/07/2024 |  |  | 10/12/2024 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025544943 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Sepsis | Tachycardia | Thrombocytopenia | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support. while on support, patient had runs of atrial fibrillation (afib) and was given digoxin and add precedex. the patient was cardioverted and also had runs of ventricular tachycardia (vt) and supraventricular tachycardia (svt). additionally, the patient became septic and was given two antibiotics. the patient also developed thrombocytopenia and pressers were increased. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation into vf/vt/af/at and infection/sepsis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly on manufacturer device report 1220648-2024-22356. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
15377 | 1220648-2024-22352 | 20631378 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20631378&pc=OZD | 06/16/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 06/16/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1087350 |  | 0042-0000-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary insertion was on an automated impella controller (aic) for mechanical circulatory support. while on support, there was a pump stop that occurred during transport. there was no report of patient harm or injury. | [5 empty] | Data analysis: case began on (b)(6) 2022 with impella 5.5 sn: (b)(6) on (b)(6). no issues occurred with pump and console until (b)(6) 2022 where a brief (~2 second) pump stop alarm (impella stopped) occurred. without user interaction, pump function restarted automatically by design, resuming previous p-level. at that point in time, staff decided to swap to ic7004 out of an abundance of safety. case continued and a successful explant occurred on (b)(6) 2022. no patient harm had occurred due to the pump stop event. on 6/16/22, imc logs reported the following: 16/06/22 15:16:06 sau_sens_eep_1 sau name already set 16/06/22 15:16:06 sau_sens_eep_1 shared memory:sh_eeprom_1 already exist.nothing to do. 16/06/22 15:16:06 sau_sens_eep_1 reinitialisation of impt device 16/06/22 15:16:06 sau_sens_eep_1 stm: wait before rst 16/06/22 15:16:06 scumain:unexpected impella pump connected following this unexpected reinitialization, the impella stopped alarm is triggered: 16/06/22 15:16:07 imcgui: event set: #115 from 9 (0) - impella stopped (mechanical/ electrical issue) after this event, the pmd is power cycled, and pump operation is resumed: 16/06/22 15:16:09 systemrtm:pmd reset complete... 16/06/22 15:16:09 imcgui: event clear: #115 from 9 (1) - impella stopped (mechanical/ electrical issue) cleared however pump reinitialization occurs again in imc logs, but no impella stopped alarm is issued and no pmd reset occurs. pump flow is uninterrupted by this event occurring at 6/16/22 15:44:16. shortly after this, staff remove the pump from (b)(6) and perform an aic-aic transfer to ic7004. see imc logs "sau_sens_eep reinitializing and impella stopped" attached. rt log analysis of the event shows the sudden loss of motor current and pump flow coinciding with the unexpected reinitialization of the pump through the sau_sens_eep channel. see rt logs "(b)(6) motor current and pump flow" attached. no issues were reported with impella 5.5 sn: 372270 occurred during operation with ic7004. imc logs for ic7004 were not available via constellation or salesforce complaint intake. ic7004 was not returned to danvers. device analysis: console was returned to danvers. (no exct date of return available) console interior was inspected; all connections were verified for connection integrity. console was ran with a optical pump simulator and purge for an extended period of time with no issues occurring during pump operation. (b)(6) was ran with an impella ld and purge setup for an extended period of time, during which the console was tested with a phase cut simulator/ short circuit simulator to rule out electrical failure of the pump/console. no issues occurred during these tests, console behaved as expected. conclusion/ repair task: the unexpected resets/ reinitialization of the pump via sau_sens_eep and its potential correlation to the pump stop are undetermined. out of an abundance of safety, fs will be instructed to: - replace impellatronics pca (0042-1116p) - perform phase cut simulation/ short circuit simulation per pm (0042-7309) - perform fc (0042-7316) the root cause of the pump stop (2 second pause - restarted automatically) was most likely due to a malfunction with the impellatronics board. the exact root cause could not be determined due to the inability to reproduce. this report is being filed as part of a retrospective review of historical records.
15378 | 1220648-2024-22353 | 20631385 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20631385&pc=OZD | 06/13/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 06/13/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1302108 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) had no vga output to a separate display monitor. there was no reported patient involvement. | [5 empty] | The device was returned and evaluated, and the investigation for the reported display issue has been completed. device analysis: the console was tested on an engineering lab bench. the failure mode was reproduced. no vga signal was available at the rlm vga output. the issue was isolated to the rlm. the root cause of the display issue was due to a defective rlm. this report is being filed as part of a retrospective review of historical records.
15379 | 1220648-2024-22367 | 20631389 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20631389&pc=OZD | 10/17/2024 | 12/24/2024 |  | 11/07/2024 |  |  | 10/17/2024 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025558316 |  | 0048-0044 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella cp support. while on support, oozing was noted at the access site. heparin was withheld due to bleeding. additionally, impella was identified to be incorrectly placed - trapped under a mitral leaflet. impella was then removed and washed in sterile water, wire access in ventricle redone, and then impella was placed correctly. the patient survived the event. | [5 empty] | The investigation for the positioning issues and access site bleed has been completed. the pump was not returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the positioning issue most likely was use related due to improper technique. the cause of the access site bleeding - major most likely was anticoagulation management due to improper anticoagulation: high acts. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
15380 | 1220648-2024-22354 | 20631396 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20631396&pc=PYX | 10/12/2024 | 12/11/2024 |  | 11/07/2024 |  |  | 10/12/2024 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | IMPELLA RPPYXTEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025515322 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Failure to Pump | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported during implant procedure of impella rp for 77-year-old male patient, the physician stated that he pulled the rp out because it wouldn¿t start. the physician then rinsed the impella in a basin with sterile water and flushed the inlet. the physician determined the rp was working and reinserted it via the right femoral artery (rfv). the patient coded and expired after approximately 5 hours of support. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ | The investigation into pump did not start has been completed. the impella device was not returned for evaluation. the root cause of the pump not starting was most likely biomaterial. the following have been reported incorrectly on manufacturer device report 1220648-2024-22354. g1 reporting contact email.
15381 | 1220648-2024-22361 | 20631466 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20631466&pc=OZD | 10/14/2024 | 12/16/2024 |  | 11/07/2024 |  |  | 10/14/2024 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025513172 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient had an attempted impella 5.5 implant and the pump was unable to be inserted completely as it appeared to get stuck as the motor housing approached the innominate drop-in. the surgeon had significant resistance causing kinking where the catheter meets the motor housing. this further reduced their drivability of the pump for insertion. the surgeons plan was to take the pump out and potentially rewire, but upon taking the pump back out of the body, he noticed a significant amount of tissue/intima in and around the inlet. this concern for dissection and/or stroke led to an aborted insertion. no dissection or stoke was noted and the patient required no further intervention at that time. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation into delivery issues and thrombus has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition due to the clotting and pump was unable to be delivered. as the necessary information was not provided, the root cause of the thrombus was not determined. the following have been reported incorrectly on manufacturer device report 1220648-2024-22361.
15382 | 1220648-2024-22357 | 20631472 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20631472&pc=OZD | 10/13/2024 | 12/27/2024 |  | 11/07/2024 |  |  | 10/13/2024 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025510113 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Death | Yes | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | Insufficient Information | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support, there was high purge pressure. tissue plasma activator was added to the purge. eventually, a pump stop occurred. several attempts were made to restart the pump with no success therefore the pump was removed. it was further reported that after the pump was removed the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation for the high purge pressure and the pump stop has been completed. impella 5.5 was returned for evaluation. no data logs were returned for evaluation. electrical testing was performed and all three motor cable lines were found to have open lines. upon further inspection in red handle, all three motor cable lines were found be broken and necking. additionally, ground line was found to be broken, and monkey fist glue joint was pulled out of place. clinical details noted that pump stop occurred and attempts to restart pump were unsuccessful. the root cause of the pump stop was most likely due to an open electrical lines. biomaterial was present in the purge gap and on impeller blade. pump was connected to console and high purge pressure alarms were triggered after approximately 20 minutes. no blood was present in the purge line just proximal of the motor housing. clinical details noted that tissue plasminogen activator (tpa) was already running in purge when purge pressure high alarms were triggered. the root cause of the high purge pressure was most likely due to biomaterial. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock: section: direct aortic insertion: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
15383 | 1220648-2024-22368 | 20631501 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20631501&pc=OZD | 07/05/2023 | 11/07/2024 |  | 11/07/2024 |  |  | 07/05/2023 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1381389 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) lost the placement signal. it was reported that the yellow controller error message was generated three separate times. a loaner was supplied to the customer by the local onsite team. there was no report of patient harm or injury. | [5 empty] | The investigation into the controller error issue has been completed. the complaint date was 7/5/2023. the console was returned over one year after the failure date. the data analysis of the logs was not possible as there were no logs available on the device for that complaint date. constellation and the logs were searched but there were no logs in those locations as well. due to the lack of logs, the nature of the complaint described is unknown. the returned device investigation did not find any issues with the console, the console operated as expected. there were no issues found during the case since the logs were unavailable and clinical information was insufficient to assist with the cause of the alleged issue. the console operated as expected when returned for investigation. there was not enough information to determine the cause, the product did not fail specifications when testing. this report is being filed as part of a retrospective review of historical records.
15384 | 1220648-2024-22365 | 20631515 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20631515&pc=OZD | 10/15/2024 | 12/11/2024 |  | 11/07/2024 |  |  | 10/15/2024 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025485777 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Increase in Pressure | [6 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 pump implanted to support 70-year-old male patient with post-cardiotomy cardiogenic shock (pccs) from coronary artery bypass graft (cabg) and valve surgery. the pump was placed but on day 11 there was thrombosis on the cannula. the team gave tissue plasminogen activator, (tpa) for purge system blocked alarms. the tpa did not remedy the issue and the following day the pump was explanted. the patient survived the explant and pocket closure. | [5 empty] | Data logs were not available. returned product was investigated and the cannula had a large amount of biomaterial in the cannula and outflow cage, even after soaking. the pump was connected to a console in lab and high purge pressure reproduced. then, a window was cut in the catheter just distal to the motor housing, and it was confirmed that no blood had ingressed into the purge line. the catheter was then cut in that location and the high purge pressure immediately resolved. clinical details state they attempted the tissue plasminogen activator (tpa) protocol, however, it is possible they didn't allow enough time for it to take effect, or there simply was too much buildup. based on clinical details and returned product, the root cause of the high purge pressure was determined to most likely be biomaterial. thrombus: since a thrombus was found to be the root cause of the high purge pressure, the investigation of that complication covers both reported complications. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-22365 b1 should have been product problem only e4 should have been left blank g1 reporting contact email h6 code 4644 and 4440 were reported incorrectly. new code was added to health effect - clinical code. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smart assist: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
15385 | 1220648-2024-22359 | 20631521 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20631521&pc=OZD | 10/13/2024 | 12/10/2024 |  | 11/07/2024 |  |  | 10/13/2024 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025527828 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient had an impella cp implant and once explanted, the impella groin access site was bleeding. heparin was stopped and a pseudoaneurysm was discovered. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation into access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was most likely patient condition since the bleeding started occurring after the impella was explanted. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-22359. e4 should have been left blank. g1 reporting contact email. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
15386 | 1220648-2024-22363 | 20631540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20631540&pc=OZD | 10/14/2024 | 12/11/2024 |  | 11/07/2024 |  |  | 10/14/2024 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025505028 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 65-year-old male patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 was being repositioned after echo showed inlet pointing towards mitral valve. during repositioning, the patient experienced ventricular tachycardia and was shocked. the patient was shocked 7 times since yesterday and started on lidocaine for ventricular tachycardia. patient remains on support. | [5 empty] | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
15387 | 1220648-2024-22362 | 20631596 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20631596&pc=PYX | 10/14/2024 | 12/10/2024 |  | 11/07/2024 |  |  | 10/14/2024 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025551604 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella rp flex support and during support, they had black tarry stool, significant bleeding in the esophagus, and an additional 1-2 liters of bright red blood was found in the stomach. heparin was in the purge of the rp flex. the family decided to withdraw care and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter: use of the repositioning sheath and the 23 fr peel-away introducer: ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound .¿ | The investigation into vascular damage - peripheral vessel has been completed. due to limited clinical details and no product returned, the root cause of the vascular damage - peripheral vessels cannot be determined. the following have been reported incorrectly on manufacturer device report 1220648-2024-22362. e4 should have been left blank. g1 reporting contact email.
15388 | 1220648-2024-22366 | 20631621 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20631621&pc=OZD | 10/16/2024 | 12/03/2025 |  | 11/07/2024 |  |  | 10/16/2024 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1497915 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant in japan reported that the automated impella controller used in a clinical setting had a frozen screen. the buttons were pressed to see if the frozen screen would clear, and the issue was resolved. there was no harm to the patient. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
15389 | 1220648-2024-22370 | 20631627 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20631627&pc=OZD | 01/30/2023 | 11/07/2024 |  | 11/07/2024 |  |  | 01/30/2023 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1103516 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a controller error. a loaner was sent to the customer to initiate return of this device. there was no report of patient harm or injury. | [5 empty] | The investigation into the controller error/optical bench issues have been completed. the impella automated controller was returned for investigation. the data analysis of the logs revealed on january 30th, when ic 1634 was turned on, the console twice checked 5 signal (5s) measurements with lamp voltages of 0.46v and 0.45v. after twice failing to pass 5s measurement thresholds, the console outputted a "controller error (defective optical sensor lamp)." upon turning on the returned console, the error was reproduced with 5s measurements similar to what was seen in the field. the lamp was unplugged resulting in similar results. after confirming voltage into the lamp was within spec, the original optical bench was replaced with a known-good bench which passed its 5s parameters with no errors. isolating the failure to the optical bench, the original bench was plugged back in which passed 5s measurements pointing to intermittent connectivity issues causing the failure mode. the cause of the aic issue was most likely a defective optical bench. this report is being filed as part of a retrospective review of historical records.
15390 | 1220648-2024-22369 | 20631662 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20631662&pc=OZD | 10/11/2024 | 12/16/2024 |  | 11/07/2024 |  |  | 10/15/2024 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025528419 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 77-year-old female noted as high risk coronary intervention was implanted with an impella cp device for mechanical circulatory support. during impella support, the patient developed upper and lower limb ischemia after the insertion of the retention sheath. the physician stated that blood pressure was not measured when the device was inserted under the right clavicle, and the color of the hand was pale. the impella cp was explanted and the issue resolved. | [5 empty] | The investigation into limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-22369. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system: section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation."
15391 | 1220648-2024-22371 | 20631667 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20631667&pc=OZD | 10/14/2024 | 12/12/2024 |  | 11/07/2024 |  |  | 10/14/2024 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025551224 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Hemodynamic instability | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and after the emergency implant during cardiopulmonary resuscitation, the pump was deep in the left ventricle. upon attempting to pull the impella cp back, it bounced back into the aorta and the cannula bent inwards on itself near the outlet cage. impella cp flows were severely reduced, and the impella cp was unable to be re-advanced. the patient coded during this time and eventually expired after 1.58 hours of impella cp support and after the doctor decided to withdraw care. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation for the reported positioning issue has been completed. product was not returned for review. the root cause of the positioning issue was most likely use related to improper technique during the advancement of the repositioning unit resulting in the deep pump position. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
15392 | 1220648-2024-22376 | 20631770 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20631770&pc=OZD | 10/15/2024 | 12/10/2024 |  | 11/07/2024 |  |  | 10/15/2024 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025516619 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention |  |  |  |  | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient was bleeding at the impella access site and out of mouth, post trans-thoracic echocardiography probe insertion. two units of packed red blood cells and a six pack of platelets were given in the operation room. fibrillar inside pocket, bio-glue to anastomosis site packed with fibrillar and pressure dressing was applied. it was further reported that the patient expired on impella support after care was withdrawn. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of the access site bleeding was most likely patient condition since the patient was bleeding from multiple sites. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrections to the initial mfr report: g1 mdr reporting contact email.
15393 | 1220648-2024-22372 | 20631807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20631807&pc=OZD | 04/04/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 04/04/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1223939 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The clinical consultant reported that the automated impella controller (aic) red alarm battery failure generated when the aic was turned on. there was no patient harm reported. | [5 empty] | The investigation into the battery failure red alarm has been completed. the impella automated controller was returned for investigation. the data analysis of the logs from the reported day of event are consistent with the complaint and show battery failure alarm due to "transport connection blown/missing" (#360), as well as battery 2 communication failure (#353). the logs also show that alarm #360 occurred for the first time on 2022-04-01, and prior to that console was operated last on 2021-08-31 (with no battery-related alarms that day). the long-term log power data indicates state of charge for both batteries around 65%-70% on 2021-08-31 and 0% the next time console was powered on (2022-04-01). the returned console and its batteries were inspected, and batttest diagnostic tool reveled very low battery voltage (2.9v and 4.4v inspead of nominal 16v), confirming that both batteries were fully depleted (one of batteries was depleted to the point that it wouldn't support communication). it is most likely that it was caused by the console not being plugged into ac power while in storage for 7 months. after batteries were replaced, the power battery manager was verified, and the console operated within specification. the cause of the aic issue was a depleted battery. this report is being filed as part of a retrospective review of historical records.
15394 | 1220648-2024-22374 | 20631966 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20631966&pc=OZD | 01/06/2023 | 11/07/2024 |  | 11/07/2024 |  |  | 01/06/2023 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1105809 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility¿s biomed reported for an automated impella controller (aic) generated a placement signal not reliable alarm which then switched to a controller error. this aic was replaced with another aic. there was no report of patient harm or injury. | [5 empty] | The investigation into the controller error/placement signal issue has been completed. the impella automated controller (aic) was returned for investigation. the data analysis of the logs from the reported day of event are consistent with the complaint and show several instances of unreliable placement signal, as well as alarms ¿placement signal unreliable¿ (due to invalid opm signal, #1036) and ¿controller error¿ (due to opm communication stopped, #1043). the long term(lt) log data and partial real time(rt) log data (some rt log files for pump 416992 were deleted by aic to reclaim disk space) allowed to identify corresponding failure modes as ¿unreliable placement signal with oscillating contrast¿ and ¿transient loss of optical data¿ respectively. the log data analysis also indicated prior malfunction of the software serial/universal serial bus (usb) driver (resulting in possible usb communication malfunction). during testing of the returned device neither of the placement signal issues were reproduced, however the rlm was found to be unable to boot-up, as evidenced by continues flashing of both light-emitting diodes (led) indicators. the "oscillating contrast" pattern failure is known to be caused by either defective light bulb (lamp assembly) or malfunctioning optical bench electronics. inspection of the console confirmed malfunction of the remote lan management (rlm) diagnostic (defective electronics preventing the rlm from booting-up even with a known-good som card and a known-good backstrap cable), and also revealed disconnected speaker wire (resulting in no audible alarms), and contaminated fiber on the blue receptacle - the last two issues were previously unreported but were unrelated to current complaint. since the presence of core dump file might indicate malfunction of the software serial/usb driver (resulting in possible usb communication malfunction), it is recommended to replace 4-in-1 and cf cards as the most likely components related to the issue. the cause of the aic issue was a defective impella connect module. this report is being filed as part of a retrospective review of historical records.
15395 | 1220648-2024-22375 | 20631973 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20631973&pc=OZD | 02/01/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 02/01/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1541587 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) displayed incorrect video colors in ethernet and wifi modes. the engineer replaced several components, and the problem was resolved with the replacement of the carrier/epc (4 in 1) component. there was no reported patient involvement. | [5 empty] | The investigation for the reported display issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, investigation details provided were sufficient to determine a root cause. after the suspect component (4-in-1) was replaced with a new one, the issue was resolved. the cause of 4-in-1 malfunction was not investigated. the cause of the aic issue was a defective 4-in-1 board. this report is being filed as part of a retrospective review of historical records.
15396 | 1220648-2024-22378 | 20632080 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20632080&pc=OZD | 10/16/2024 | 12/10/2024 |  | 11/07/2024 |  |  | 10/16/2024 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025543781 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, left ventricle thrombus found on a computed tomography scan. impella remained implanted. the patient survived the event. | [5 empty] | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation; section: impella cp with smart assist catheter insertion; section: axillary insertion of the impella cp with smart assist catheter; section: use of impella with transcatheter aortic valves; ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.".
15397 | 1220648-2024-22383 | 20632149 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20632149&pc=OZD | 12/27/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 12/27/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1314432 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a battery failure alarm, despite having automated impella controller/console plugged in for an extended period of time. no patient involvement or impact noted. | [5 empty] | The investigation for the reported aic - battery failure (red alarm) issues has been completed. the device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the battery failure was due to a defective battery 1. the root cause of the defective battery 1 was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
15398 | 1220648-2024-22377 | 20632186 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20632186&pc=OZD | 02/02/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 02/02/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1071329 |  | 0042-0000-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) was showing white/blue video during maintenance. the engineer replaced several components to resolve the issue, and there was no reported patient involvement. | [5 empty] | D8/9: device returned to field service, date unknown. the device was returned and evaluated, and the investigation for the reported display issue has been completed. device analysis: troubleshooting was completed in collaboration with field service. the 4-in-1 was swapped to resolve the video issues. the root cause of the off-color video on impellaconnect.com was due to a defective 4-in-1. this report is being filed as part of a retrospective review of historical records.
15399 | 1220648-2024-22384 | 20632232 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20632232&pc=OZD | 12/26/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 12/26/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1207018 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) experienced a purge disc not detected alarm, which was resolved by changing the console. no patient harm reported. | [5 empty] | The investigation has been completed. the issue with properly detecting the purge pressure disc was reproduced, and blue housing for purge pressure sensor was confirmed to be damaged. the root cause of the purge disc detection issue was the broken blue housing. this report is being filed as part of a retrospective review of historical records.
15400 | 1220648-2024-22380 | 20632263 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20632263&pc=OZD | 12/03/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 12/05/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1464036 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) that after explanting an impella device and removing the purge line, the purge disc retainer broke off of the console. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. data analysis: imc logs are not relevant for this failure mode. device analysis: console arrived in (b)(6). a physical inspection of the purge assembly revealed that the purge disc retainer had broken off. clinical details were sufficient to determine the root cause. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15401 | 1220648-2024-22385 | 20632267 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20632267&pc=OZD | 12/10/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 12/21/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1417603 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Unknown |  | Non Hispanic |  | Malfunction | No | Yes | No Audible Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support, the automated impella controller (aic) experienced fluid was coming from impella purge cassette. impella never alarmed. no patient harm reported. | [5 empty] | The investigation has been completed. the product was not returned. the root cause of the no alarm was unable to be determined because the console was not returned for review. this report is being filed as part of a retrospective review of historical records.
15402 | 1220648-2024-22386 | 20632333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20632333&pc=OZD | 11/10/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 11/10/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1380109 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) that it had a broken purge clip. the console was removed from service and a replacement was provided to the customer for support as needed. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. data analysis: imc logs are not relevant for this failure mode. device analysis: console arrived in danvers. the console purge assembly was inspected. upon opening the purge assembly door, it was confirmed the purge disc retainer had broken/snapped off of the purge assembly. clinical details were sufficient to determine the root cause. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15403 | 1220648-2024-22382 | 20632391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20632391&pc=OZD | 12/29/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 12/29/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1168791 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automatic impella controller (aic) had a battery failure alarm. there was no mention of patient involvement. no additional information provided. | [5 empty] | The investigation for the reported aic - battery failure (red alarm) issues has been completed. the device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the cause of the battery failure alarm was due to a defective battery (rapid decline in cell voltage). the exact cause of the defective battery was unable to be determined. this report is being filed as part of a retrospective review of historical records.
15404 | 1220648-2024-22398 | 20632422 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20632422&pc=OZD | 06/16/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 06/16/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1103719 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) that there was a broken purge clip there was no report of patient harm or injury. | [5 empty] | The investigation for the reported issue has been completed. the product was returned, and the issue was confirmed. data analysis: imc logs are not relevant for this failure mode. device analysis: console arrived in danvers. console purge assembly was inspected. upon opening the purge assembly door, it was confirmed the purge disc retainer had broken/snapped off of the purge assembly. per the results of the clinical review and investigation, the root cause was determined to be a defective component. this report is being filed as part of a retrospective review of historical records.
15405 | 1220648-2024-22405 | 20632450 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20632450&pc=OZD | 06/17/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 06/17/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1128485 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) that during preventive maintenance, a red line through the display appeared suddenly during testing. the console was rebooted, and the red line still appeared. reinstall software and the issue was still present. there was no patient involvement. | [5 empty] | The investigation for the reported issue has been completed. the product was returned, and the issue was confirmed. device analysis: failure mode was reproduced. lvds cables were secured properly to associated connectors. failure was isolated to the lcd display by replacing the cables and 4-in-1 with known good and observing the same symptoms. the root cause of the display issue was due to a defective lcd display. this report is being filed as part of a retrospective review of historical records.
15406 | 1220648-2024-22408 | 20632471 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20632471&pc=OZD | 10/16/2024 | 12/17/2024 |  | 11/07/2024 |  |  | 10/16/2024 | 11/07/2024 |  |  | 11/30/2025 |  | 11/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024435424 |  | 0550-0002 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Female |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella 5.5 support due to acutely worsened left ventricle function. while on support, purge flow low and purge pressure high were present in the pump two different times. high purge pressure/purge system blocked alarm occurred. the patient survived the explant. | [5 empty] | The investigation into high purge pressure has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-22408. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
15407 | 1220648-2024-22435 | 20632473 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20632473&pc=OZD | 01/03/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 01/03/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1308463 |  | 0042-0010-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a battery failure red alarm that prompted replacement of the automatic impella controller (aic). no report of patient involvement. | [5 empty] | D9: the date of the devices return for evaluation is unknown. e4: the name of the initial reporter and occupation is unknown. the investigation for the reported aic - battery failure (red alarm) issues has been completed. the device was returned for evaluation. the cause of the battery failure alarm was due to a defective battery (rapid decline in cell voltage). the battery failure was due to a defective battery 1. the root cause of the defective battery 1 was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
15408 | 1220648-2024-22429 | 20632480 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20632480&pc=OZD | 06/02/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 06/02/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1103723 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the aic experienced a purge disc/flag issues. when inserting purge cassette, the console was not reading purge disc while disc was in place and inserted correctly. the same purge cassette/disc was inserted in another aic with no issue. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported issue has been completed. the product was returned, and the issue was confirmed. data analysis: imc logs from the complaint date revealed that the console issued the purge disc not detected alarm (event set #402) several times. device analysis: the console was tested after arriving in fs. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. the root cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15409 | 1220648-2024-22407 | 20632539 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20632539&pc=OZD | 12/05/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 12/05/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1422863 |  | 0042-0000-JP |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the automated impella controller (aic) exhibited a purge cassette stopped alarm. they also received the please replace the console alarm. neither the pump could be started, nor the purge cassette could be replaced, and when the console was replaced, the pump could be started. no report of patient injury or harm. | [5 empty] | The investigation for the reported aic - purge issue - other issues has been completed. the device was returned for evaluation (exact date of return is unknown). the cause of the controller alarm was malfunction of the purge pressure sensor assembly. this report is being filed as part of a retrospective review of historical records.
15410 | 1220648-2024-22388 | 20632602 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20632602&pc=OZD | 12/16/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 12/16/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1204376 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a battery failure with an automatic impella controller (aic) console. there was no impact to the patient noted. no additional information provided. | [5 empty] | The investigation for the reported aic - battery failure (red alarm) issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the battery failure (red alarm) was due to a defective battery (rapid decline in cell voltage). the exact cause of the defective battery was utd. this report is being filed as part of a retrospective review of historical records.
15411 | 1220648-2024-22397 | 20632644 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20632644&pc=OZD | 12/10/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 12/10/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1432818 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) that a battery failure occurred. a loaner was shipped to initiate the return of the aic for investigation/repair. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported battery failure issue has been completed. the product was returned, and the issue was confirmed. data analysis: imc logs at the complaint occurred date revealed battery 2 communication failure (smbus comm loss or sda short) ¿ event # 353 alarm. attempts of reboot aic failed in resolving this alarm. device analysis: batt test was performed in engineering bench. the appearance of ¿smb¿ in channel 2 is consistent with the battery 2 communication failure alarm. swapping battery 1 and battery 2 did not change the ¿smb¿ appearance in channel 2 of pbm. battery 2 was temporarily replaced with a known-good battery, but the communication failure persisted. batt test confirmed the battery communication failure resulted from defective channel 2 of pbm, the root cause of pbm malfunction was utd. per the results of the clinical review and investigation, the root cause was determined to be a defective pbm. the cause of the defective pbm is unknown. this report is being filed as part of a retrospective review of historical records.
15412 | 1220648-2024-22400 | 20632698 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20632698&pc=OZD | 12/10/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 12/10/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1460757 |  | 0042-0000-UK |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with high-risk percutaneous coronary intervention was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) experienced a purge disc issue, which was resolved by changing the console. no patient harm reported. | [5 empty] | The investigation has been completed. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. the root cause of the broken purge flag is due to repeated force and wear and tear of the plastic component. this report is being filed as part of a retrospective review of historical records.
15413 | 1220648-2024-22390 | 20632703 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20632703&pc=OZD | 12/01/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 12/01/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1100890 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that there was always an "impella defective" error message displayed no matter what kind of pump was connected to the receptible optical or non-optical. the same pumps worked fine in other consoles. no additional information was provided. | [5 empty] | The investigation for the reported aic - controller failure (red alarm) issues has been completed. the device has been returned for evaluation (exact return date unk). clinical details were sufficient to determine the root cause. the cause of the pump not being recognized was malfunction of the epm circuitry on impellatronic resulting in inability to read pump eeprom. this report is being filed as part of a retrospective review of historical records.
15414 | 1220648-2024-22422 | 20632734 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20632734&pc=OZD | 10/15/2024 | 12/12/2024 |  | 11/07/2024 |  |  | 10/17/2024 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025534224 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 10 YR |  | Female |  | Non Hispanic | White | Malfunction | Yes | Yes | Detachment of Device or Device Component | [6 empty] | Foreign Body In Patient | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and upon removing the pump, the surgeon opted to use a side clamp to remove the device. this lead to a cut proximal and a cut distal to the side clamp resulting in three pieces of the device to be removed. the distal and proximal portions were removed without incident. however, when the surgeon went to remove the piece inside the side clamp they were unable to locate it. via x-ray, a four centimeter portion of the catheter was found to be in the right femoral artery. it was decided by interventional cardiology that the patient would go to the cardiac catheterization lab for removal of this migrated portion of the catheter via contralateral access and snare. the entire four centimeter portion was successfully removed without incident. the patient survived the event. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the product damage has been completed. the pump was not returned for investigation. data log review was not required for this case. as per the clinician, the doctor cut the cannula into three pieces during removal. one piece of the catheter was found in the femoral artery, and it needed a surgical approach to be removed from the body. product damage issue failure mode was changed to removal issue as investigated failure mode. the cause of the removal issue was user-related, as the doctor lost the third piece of the pump during the cross-clamping cut approach.
15415 | 1220648-2024-22415 | 20632885 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20632885&pc=OZD | 10/16/2024 | 12/12/2024 |  | 11/07/2024 |  |  | 10/16/2024 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025547730 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Pumping Stopped | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient arrived in the intensive care unit after impella cp insertion with a slight oozing at the access site. activated clotting time was 244. brilinta was given. the site was re-sutured and re-dressed. several days later, the impella had a high motor current with an impella stopped alarm. the pump stopped and restarted at previous p-level. the impella was running with elevated motor current and aorta/left ventricle waveforms inverted. it was further reported that the patient expired while on venous-arterial extra-corporeal membrane oxygenation and impella cp support. | [5 empty] | The investigation into access site bleeding and pump stop has been completed. pump was not returned for investigation. the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical information was provided. the pump was restarted at the same p-level automatically, but saturated flows were observed following pump stop alarm. therefore, the root cause of the pump stop issue was sudden biomaterial ingestion. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-22415. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
15416 | 1220648-2024-22404 | 20632970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20632970&pc=OZD | 12/06/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 12/06/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, JP |  |  | ABIOMED, INC. |  |  | US |  | 2022100381 |  | 0042-0000-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Component Misassembled | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) cart and caster caps are connected upside down. there was no patient use and no harm or injury. this is was discovered at the time of delivery. | [5 empty] | The investigation for the reported cart issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the japan field service confirmed the aic cart caster was upside down and had sent a brand new replacement aic cart for customer. the root cause of the cart and caster cap connection upside down is undetermined. this report is being filed as part of a retrospective review of historical records.
15417 | 1220648-2024-22432 | 20633019 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20633019&pc=OZD | 05/19/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 05/19/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1261272 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automated impella controller experienced a purge disc/flag issue. the staff reported that during case start, it was observed that the purge retainer had broken off of the purge assembly. there was no reported harm or injury to the patient. | [5 empty] | The investigation has been completed. the console was returned for evaluation. the console purge assembly was inspected. upon opening the purge assembly door, it was confirmed the purge disc retainer had broken/snapped off of the purge assembly. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15418 | 1220648-2024-22410 | 20633020 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20633020&pc=OZD | 11/26/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 11/26/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1146156 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) that after impella insertion, a controller error alert was displayed. the aic also would not acknowledge the purge disc and was absent of placement signal. the aic was changed and the case went smoothly after changing consoles. it was unclear if a restart was attempted with the unit or not there was no report of patient harm or injury. | [5 empty] | The investigation for the reported purge disc/flag issue has been completed. the product was returned, and the issue was confirmed. data review: on the complaint reported day, a purge disc not detected (event #402) alarm occurred and disappeared instantaneously. according to the time stamps from imc logs, the alarm was triggered because of the missing purge disc after 10 seconds since insertion of pc. the alarm condition was cleared as insertion of purge disc. console logs from the reported day of event are consistent with complaint about controller error (yellow alarm). console initially reported a placement signal not reliable (opm invalid signal, optical pump connected) ¿ event set #1036 alarm, indicating the pump sensor signal is invalid as optical bench reporting. alarm #1036 revolved itself in 20 seconds, but there issued another alarm event set #1043, controller error, in few minutes. the alarm #1043 indicated the loss of communication to the optical bench. also, it was noted that 10 days prior to complaint date, the optical bench reported bad optical parameters in imc logs. the optical bench failed to reset, hence it triggered a controller error (defective optical sensor lamp) ¿ event set #1039 alarm. rtlogs at the complaint date and time revealed the optical placement signal and other optical parameters spiked into +16384 which is not reasonable. device analysis: after the console arrived in danvers, it was investigated for the two reported failure modes. few pairs of pc and purge disc were utilized, no purge disk detection problem was found. it was speculated that the alarm reported for purge disk detection was due to no insertion of purge disc on time based on the testing and imc log analysis. it is not a product malfunction. as for the controller error and placement signal unreliable, the alarms were not reproduced in engineering bench. during overnight testing in burn-in room with simulator pump and pc, the console functioned well without any failure. no abnormality was found when checking the ribbon cable connections for optical bench. the optical relating alarms were not reproduceable. per the results of the clinical review and investigation, the root cause was most likely due to defective optical bench.
15419 | 1220648-2024-22373 | 20633024 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20633024&pc=OZD | 02/02/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 02/02/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1097384 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Application Program Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Life Threatening | Required Intervention | [23 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. during support, the automated impella controller (aic) experienced a controller failure and was swapped with another device. while the console was being exchanged, the patient went into ventricular fibrillation. the patient coded for ten minutes, and a return of spontaneous circulation was achieved. no further patient harm was reported. | [5 empty] | The device was returned and evaluated, and the investigation for the reported controller failure and ventricular fibrillation has been completed. console logs from the reported day of event are consistent with complaint and show communication errors between 4-in-1 and optical bench: few seconds after pump was started, there was no data coming from optical bench (event 1043: communication stopped) followed by alarm #1043 (controller error due to opm communication stop), two minutes later. this was followed by a series of pump-related alarms (#115, impella stopped due to mechanical/electrical issue), however at that moment pump was already disconnected but console did not register that action. this, in turn, resulted in inability to shut down the console due to erroneous pump detection. after unplugging ac, the batteries were disengaged and console shut down. the issue was not reproduced in danvers during testing, however the symptoms are consistent with a known pattern failure where all signals from optical bench disappear, after a single sample of +16384. the only way issue is resolved is by power-cycling the console. the cause of the issue was not determined because issue could not be reproduced this report is being filed as part of a retrospective review of historical records.
15420 | 1220648-2024-22421 | 20633056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20633056&pc=OZD | 10/19/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 10/19/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1108441 |  | 0042-0000-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had an automated impella controller (aic) fail to power up and the battery was discharged. there was no patient involved. aic was sent for investigation. | [5 empty] | Device was returned to regional engineers and evaluated and the return date is unknown. e1: the name of the initial reporter is unknown. the investigation for the reported issue has been completed. the impella device was received from the customer. aachen fs noted that when console arrived, the console could not be powered on. console was plugged in and allowed to charge for a brief period of time before attempting to boot. console booted with low battery charge. aachen fs additionally noted there were no alarms on the console. the cause of the console unable to boot was due to depleted batteries, not routinely charged by connecting console to ac. this report is being filed as part of a retrospective review of historical records.
15421 | 1220648-2024-22416 | 20633126 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20633126&pc=OZD | 11/14/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 11/14/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1561776 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 for mechanical circulatory support. while on support, the automated impella controller (aic) experienced a purge disc issue, which was resolved by changing the console. no patient harm reported. | [5 empty] | The investigation has been completed. the root cause of the broken purge disk retainer was not determined. this report is being filed as part of a retrospective review of historical records.
15422 | 1220648-2024-22406 | 20633196 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20633196&pc=OZD | 12/01/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 12/01/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1267116 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | Connection Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) was unable to recognize the pump, which was resolved through new console. no patient was involved. | [5 empty] | The investigation has been completed for pump not detected. the console (ic4769) was sent back for further investigation. the console was tested thoroughly on an engineering bench. the inability to detect the purge fluid during the case start procedure was likely due to an open line in the purge tubing but could not be confirmed due to the lack of information from the clinical team. this report is being filed as part of a retrospective review of historical records.
15423 | 1220648-2024-22411 | 20633538 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20633538&pc=OZD | 11/17/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 11/17/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1426097 |  | 0048-0000-JP |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary insertion was on an automated impella controller (aic) for mechanical circulatory support. while on support, displayed an abnormal alarm on different occasions. the aic console cable could not be connected/disconnected. the initial aic was replaced, and the case continued. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported controller error issue has been completed. the product was returned (exact return date is unknown), and the issue was not confirmed. data analysis: imc logs confirmed the reported failure mode given there were controller error alarms (opm comm crc error : event set # 1045, and opm signal invalid event set # 360) triggered during the clinical procedure. real time (rt) logs confirmed that this was a pattern failure mode in which all signals received from optical bench (placement, snr, contrast, lamp, gain) are interpreted as -16384 values. device analysis: console was sent back to japan fs for further investigation. this known pattern failure, which is characterized by a temporary loss of optical data (placement, snr, contrast, lamp, gain) and triggering of a controller error alarm, has a low probability of reoccurrence. if needed, patient hemodynamics and impella position can be monitored with imaging. per the results of the clinical review and investigation, the root cause unable to be determined because the issue was unable to be reproduced. this report is being filed as part of a retrospective review of historical records.
15424 | 1220648-2024-22396 | 20633564 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20633564&pc=OZD | 10/17/2024 | 12/12/2024 |  | 11/07/2024 |  |  | 10/17/2024 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025561198 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support. while on support, the wire was removed, and the physician was manipulating the impella in an attempt to advance the impella deeper into the left ventricle (lv). during the process the impella perforated the backside of the lv. the physician was able to repair the perforation and successfully place the impella. the patient survived the event. | [5 empty] | The investigation for the ventricle perforation has been completed. pump was not returned for investigation. data log review was not required for this case run. as per the clinical team, the lv was perforated during manipulation of the catheter in attempted to advanced deeper in lv. doctor was able to repair successfully. no information was provided about the patient's condition or past medical history related to the heart. the cause of the ventricle per issue was not determined since product was not returned and sufficient clinical information was not provided. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
15425 | 1220648-2024-22395 | 20633741 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20633741&pc=OZD | 12/16/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 12/16/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1461907 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) would not recognize the purge disc, this occurred prior to the start of the procedure. the customer had an alternate aic device onsite and the device was placed without further issue. no report of patient harm or injury. | [5 empty] | The device was returned for evaluation. the console was tested and the issue with properly detecting the purge pressure disc was reproduced, the purge flag was confirmed to be broken. the cause of the issues properly detecting the purge disc was due to a broken purge flag. this report is being filed as part of a retrospective review of historical records.
15426 | 1220648-2024-22426 | 20633752 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20633752&pc=OZD | 09/18/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 09/18/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER EU |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER EU |  |  | ABIOMED, INC. |  |  | US |  | 1234873 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Crack | Shipping Damage or Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The field service engineer reported that the automated impella controller (aic) loaner console was damaged during shipping. it was reported that the customer refused shipment of the console from fedex as the box was damaged and wet. there was no patient involvement. | [5 empty] | The investigation into the damage has been completed. the impella automated controller was returned for investigation. the returned console was inspected the console was confirmed to have sustained damage due to being mishandled during transport. despite adequate packaging, the force that aic was subject to caused the purge assembly to break and detach itself from the housing. also, the front housing was cracked, and further inspection revealed a loose battery bracket inside the console, prompting battery replacement as a precaution. after replacement of damaged components, a functional check was performed and aic passed all testing, however analysis of console logs revealed some irregularities, most likely due to a keyboard communication glitch. after repeated inspection and testing in engineering, the issue was not reproduced. the console was ultimately deemed safe to use as a loaner. the root cause of the damage to the aic was a result of damage during transport to the customer. this report is being filed as part of a retrospective review of historical records.
15427 | 1220648-2024-22401 | 20633761 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20633761&pc=OZD | 12/09/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 12/09/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1349165 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automatic impella controller (aic) console was turned on and a "system self-check failure" alarm was received. unknown how this issue was resolved, no reported harm or injury to the patient. | [5 empty] | The investigation for the reported aic - system self-check error issues has been completed. the device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of system self check failure was corrosion near the c69 and c70 supercapacitors on the pbm. this report is being filed as part of a retrospective review of historical records.
15428 | 1220648-2024-22393 | 20633810 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20633810&pc=OZD | 12/19/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 12/19/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1249747 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support, the automated impella controller (aic) experienced a purge disc issue, which was resolved by changing the console. no patient harm reported. | [5 empty] | The investigation has been completed. the console ((b)(6)) was sent back for further investigation. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. the root cause of the purge disc not recognized was due to the broken flag. this report is being filed as part of a retrospective review of historical records.
15429 | 1220648-2024-22409 | 20633813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20633813&pc=OZD | 11/30/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 11/30/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED. INC. |  |  | US |  | 1088057 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during preventative maintenance the automated impella controller (aic) exhibited controller failure. a loaner was shipped to the customer. no patient involvement. | [5 empty] | E4: the name of the initial reporter and occupation is unknown. the device was returned for evaluation. the root cause of the opm communication loss is unable to be determined due to inability to reproduce.
15430 | 1220648-2024-22430 | 20633912 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20633912&pc=OZD | 05/20/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 05/20/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1194663 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported an automated impella controller (aic) system check failure. the aic was removed from service and the instructions for use were followed and a backup controller was utilized. there was no report of patient harm or injury. | [5 empty] | The investigation into the automated impella controller - system self check error issue has been completed. there were no log entries during the reported complaint date, but logs prior to the complaint date were consistent with the complaint description. the logs were filled with repetitive error messages causing the reported system self-check failure and console boot-up failure. the automated impella controller (aic) was returned for investigation. the failure mode was reproduced and a blue screen with ¿system self-check failure¿ presented. the console was opened for visual inspection which revealed multiple severe burnt marks on the power battery manager (pbm) and neighbor areas (batteries, fan cables, etc). a component was found with severe damage affecting the output of the converter. furthermore, impellatronic board received inappropriate power supply from the pbm, which manifested as the system self-check failure. the defective pbm was temporarily replaced with a known-good one with which console could boot up and operate as normal. the cause of the aic issue was a defective pbm board. this report is being filed as part of a retrospective review of historical records.
15431 | 1220648-2024-22379 | 20633914 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20633914&pc=OZD | 01/18/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 01/18/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1548396 |  | 0042-0010-UK |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that there is a loose fitting on the automated impella controller (aic). the customer noted that there was no patient involvement. | [5 empty] | D9: the date of the devices return for evaluation is unknown. e1, e3, e4: the initial reporters name and occupation is unknown. the device was returned for evaluation. it was determined that the "loose fitting" was an inline ac fuse holder not having been fully secured, which would have prevented the console from being powered on from ac. the fuse holder was tightened, and issue was resolved. the cause of the aic issue was a loose component (fuse holder in rear console housing). this report is being filed as part of a retrospective review of historical records.
15432 | 1220648-2024-22392 | 20633972 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20633972&pc=OZD | 12/20/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 12/20/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1181074 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced a missing purge disc connection. a new console was used without any issues. no patient was involved. | [5 empty] | The investigation has been completed. the console ((b)(6)) was sent back for further investigation. console was inspected on an engineering bench and console¿s purge pressure sensor assembly sustained mechanical damage and the disk retainer is missing. the root cause of the aic missing purge disk connection was due to disk holder missing. this report is being filed as part of a retrospective review of historical records.
15433 | 1220648-2024-22387 | 20634041 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20634041&pc=OZD | 10/16/2024 | 11/11/2024 |  | 11/07/2024 |  |  | 10/16/2024 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025550678 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the impella cp was being primed and before the pump was inserted into the patient, an impella stopped alarm was noticed. the doctor decided to open and use a new catheter. the patient remains on impella support. | [5 empty] | E.1 country should be united states and not puerto rico on the initial manufacturer device report number 1220648-2024-22387. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
15434 | 1220648-2024-22413 | 20634135 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20634135&pc=OZD | 11/04/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 11/04/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1077381 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced multiple sensor errors. the console was replaced to address the issue. no patient was involved. | [5 empty] | The investigation into the reported issue has been completed. console was ran on an optical pump for 1 day with no issues. console case start was attempted with an optical pump several times however the problem experienced in the field was unable to be reproduced. the cause of the placement signal not reliable alarms is undetermined. this report is being filed as part of a retrospective review of historical records.
15435 | 1220648-2024-22381 | 20634144 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20634144&pc=OZD | 01/04/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 01/04/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1187999 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a battery failure alarm that could not be resolved, despite attempting to charge for 24 hours. no additional information provided. no harm or injury noted. | [5 empty] | The investigation for the reported aic - battery failure (red alarm) issues has been completed. the device has been returned for evaluation. clinical details were sufficient to determine the root cause. the root cause of the battery failure alarm is likely due to a disengaged battery switch. this report is being filed as part of a retrospective review of historical records.
15436 | 1220648-2024-22389 | 20634178 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20634178&pc=OZD | 11/09/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 11/09/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1146764 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a controller error. the device was removed from use and a loaner was sent to the customer to initiate return of this device. there was no report of patient harm or injury. | [5 empty] | The investigation into the controller error has been completed. the impella automated controller was returned for investigation. data analysis: as can be seen in the imc log, shutdown was requested by the user on 22/08/22 with the pump still connected then emergency shutdown was initiated when the console rejected the initial user shutdown. the rt logs additionally show there is still purge pressure and pump was running at the time of shutdown. then when the console was next booted up on 08/11/22 the console showed an alarm from previous improper shutdown. power cycling removed the alarm as expected. device analysis: the same alarm as the field was triggered trying to shutdown with the pump running and then getting an error message after forcing shutdown. the cause of the issue was use related due to improper technique by the end user. this report is being filed as part of a retrospective review of historical records.
15437 | 1220648-2024-22419 | 20634202 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20634202&pc=OZD | 10/25/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 10/25/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER EU |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER EU |  |  | ABIOMED, INC. |  |  | US |  | 1303853 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The field service engineer reported that the automated impella controller (aic) was sent for preventive maintenance and upon return of the console there was a big hole in the rear housing. there was no patient involvement. | [5 empty] | E1: name of the initial reporter is unknown. the investigation into the damage to the rear housing has been completed. the impella automated controller was received from the customer. the field service engineer (fse) confirmed extensive damage to the bottom of the rear housing, the housing was cracked. the fse inspected the interior of the console and noted the power battery manager (pbm) had also been damaged. the fse replaced the console rear housing and pbm. the cause of the damage to the console rear housing and pbm was likely due to mishandling of the console. this report is being filed as part of a retrospective review of historical records.
15438 | 1220648-2024-22414 | 20634237 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20634237&pc=OZD | 06/15/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 06/15/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1202889 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) that there was a broken purge clip. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported issue has been completed. the product was returned, and the issue was confirmed. data analysis: aic logs were non-existent on the reported complaint date. 15 days before, there was a pump and purge connected which was unplugged shortly after due to a purge pressure low alarm, which likely had relation to the failure mode. device analysis: the console was tested after arriving in fs. the purge disc retainer was confirmed to be broken. the root cause of the issue was a broken purge retainer. this report is being filed as part of a retrospective review of historical records.
15439 | 1220648-2024-22433 | 20634269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20634269&pc=OZD | 10/20/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 10/20/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1618123 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a battery failure alarm on the automated impella controller (aic). the initial aic was replaced with a new device and the case continued. no patient harm was reported. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. when visually inspecting the batteries and power battery manager (pbm), it was observed that one of the batteries statuses leds were completely blank. the battery failure alarm was due to a defective battery 2 (decline of individual cell voltage). the root cause of the battery cell failure was undetermined. this report is being filed as part of a retrospective review of historical records.
15440 | 1220648-2024-22402 | 20634271 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20634271&pc=OZD | 10/17/2024 | 12/16/2024 |  | 11/07/2024 |  |  | 10/17/2024 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025567115 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient had oozing from the impella access site. it was treated with purse-string stitch and light femstop pressure. hemoglobin and hematocrit had decreased from the oozing from insertion site. fem stop pressure device was reapplied overnight. impella cp ultimately explanted due to bleeding. the patient remained stable and survived the event. | [5 empty] | The investigation into access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-22402. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
15441 | 1220648-2024-22425 | 20634309 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20634309&pc=OZD | 09/27/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 09/27/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1164186 |  | 0042-0000-US |  | Yes |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that during preventive maintenance, no data would stream on the automated impella controller (aic). troubleshooting was performed, and it was reported that the 'relay board (rlm) was faulty. the rlm was replaced and the device functioned as intended. no patient is involved. | [5 empty] | D9 : the date when console was sent for preventive maintenance is unknown. the investigation for the reported display issues has been completed. the aic device was analysed. replacing the module resolved the data streaming issue, and the console was sent out with a new module. the module was investigated further. the module¿s fpga fw was found to be out-of-date, not 0042-96344-a fpga fw, which is needed for data streaming enabled consoles. without this fw, the console would not be able to establish initial communication/ handshake with the module, which was observed on the complaint date. this report is being filed as part of a retrospective review of historical records.
15442 | 1220648-2024-22456 | 20634367 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20634367&pc=OZD | 01/30/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 01/30/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1540253 |  | 0042-0010-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an automated impella controller for mechanical circulatory support. while on support, there was a "controller error" alarm. after two minutes, the error fixed without replacement. the physician decided not to exchange. there was no report of patient harm or injury. | [5 empty] | D.9 - the exact date the device was returned is unknown. the investigation for the reported issue has been completed. the data logs showed that the complaint date was consistent with complaint intake. the console was returned to field service in germany. the console was power cycled several times and tested with a pump and purge system with no issues. the console was disassembled and interior connections were inspected. it was observed that the ribbon cable from the 4-in-1 to the impellatronics board was slightly misaligned, as well as the ribbon cable from the 4-in-1 to the optical bench. it is most likely that the misaligned connections could have caused intermittent communication issues. the cause of the controller error and momentary interruption of communication was most likely due to loose connectors. this report is being filed as part of a retrospective review of historical records.
15443 | 1220648-2024-22394 | 20634385 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20634385&pc=OZD | 10/16/2024 | 12/24/2024 |  | 11/07/2024 |  |  | 10/16/2024 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025538438 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Female |  |  | White | Malfunction | Yes | Yes | Difficult to Remove | [6 empty] | Unspecified Vascular Problem | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp placed to support a patient admitted in with acute myocardial infarction/cardiogenic shock. the pump supported for 20 hours and was explanted. the explant occurred but there were issues with the removal, and the leg/iliac artery became occluded. the team performed an angioplasty and closed by surgical methods. the perclose and angioseal had failed to achieve appropriate hemostasis. the patient survived the explant. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the removal issue has been completed. the pump was not returned for investigation. data log review was not required for this case. as per the clinical details, the closure device failed, and a surgical approach was taken to close the puncture site (vascular damage). ct also revealed no blood flow to the leg during the closure device failure. the cause of the vascular damage issue was not determined since product was not returned and sufficient clinical information was not provided. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
15444 | 1220648-2024-22438 | 20634428 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20634428&pc=OZD | 05/19/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 05/19/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1103520 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted with high risk percutaneous coronary intervention was needing an impella for mechanical circulatory support. it was reported that the automated impella controller (aic) failed to recognize the pump as the hospital staff was preparing the pump for insertion. after the white cable was connected, the aic prompted the staff to connect ¿gray to gray¿ as if the pump was not an optical device. the staff was able to eventually bypass the screen and continue with setup. there was no report of patient harm or injury. | [5 empty] | The investigation into the pump not detected issue has been completed. the automated impella controller logs revealed the following: the window which was instructing the clinical team to check grey-to-grey connectors occurred but was then bypassed once the epm was power cycled, leading to the proper detection of the impella pump. the console not returned for further investigation. the pump detection issue was due to the rarely occurring epm crystal issue. the impellatronic pca does not have permanent damage, but this intermittent pump detection failure occurs infrequently and already has a software mitigation in place. this report is being filed as part of a retrospective review of historical records.
15445 | 1220648-2024-22428 | 20634437 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20634437&pc=OZD | 09/09/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 09/09/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1100025 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the purge disc was not being recognized by the automated impella controller (aic) console. there was no report of patient harm or injury. | [5 empty] | The console (aic) was returned for evaluation. data analysis: imc logs from the complaint date revealed that the console issued the purge disc not detected alarm (event set #402) several times. device analysis: the console was tested after arriving in fs. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken (missing at blue disc retainer.) the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15446 | 1220648-2024-22431 | 20634524 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20634524&pc=OZD | 05/31/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 06/01/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1199622 |  | 0042-0000-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. the complainant had an aic with 2 hours of data logs missing from the console. the team noted no harm or injury from this and the red alarm is captured in the history of the aic. | [5 empty] | D9 - date device returned to manufacturer: unknown. the investigation for the reported issue has been completed. the product was returned, and the issue was confirmed. according to the report from staff, a case notification email was received via impellaconnect.com at 01:19 for the occurrence of a red banner alarm. staff attempted to view video history via impellaconnect.com but noticed that a gap in the video history spanning from ~00:43 to 02:43 had occurred. staff were able to confirm the red alarm through console alarm history. data analysis: red alarm - impella in ventricle/impella stopped: per imc logs, two red banner alarms occurred from the lost of network connectivity at ~00:43 to when a brief connection was made at 01:19. 31/05/2022 00:52:07 imcgui: event set: #139 from 10 (0) - impella in ventricle. 31/05/2022 00:59:50 imcgui: event set: #18 from 2 (0) - impella stopped. an email notification from impellaconnect.com was sent at 01:19 in response to the red alarm, however staff were unable to view the alarm due to the missing video history. due to discrepancies in imc log time, rlm log time, impellaconnect.com time, and the hospital site time, it is most likely that one of these two alarms is the triggering cause for the email notification generated. however, after imc log review, it has been determined that it is unlikely these two alarms have lead or were caused by a console malfunction. gap in impella connect video history: the gap/missing video history in question was reviewed through different means. data pulled from the impella connect server showing different connection parameters was reviewed at the time of the event. it can be seen from this data that the last good connection was valid from 5/31/2022 at 00:45 to 00:52. continuing forward, overall device connectivity was very poor from 5/31/2022 00:52 until 03:30. with the impella stopped/impella in ventricle alarms triggering between 00:52 and 00:59, no email was generated at the alarm time due to unstable network connectivity. however, a brief connection (~ 4 secs) was made at 01:19, which allowed for the impellaconnect.com email notification to be made and sent. due to the procedure of how video history is structured via video stitching, proper video history could not be 'stitched' due to lack of cloud connectivity from ~00:52 to ~02:45. at the time of this case, the other cases were underway at the same site. impella connect server data was reviewed from these three other cases occurring. it was observed that there were no connectivity issues with these consoles. device analysis: console returned to (b)(6) on 8/30/2022 for its annual preventative maintenance. per pm, rlm rssi is checked, where the rlm rssi was found to be within spec. conclusion: missing impellaconnect.com video history is most likely due to poor network strength in relation to the aic/rlm location. this resulted in the rlm consistently attempting to connect and disconnect from the ssid at the hospital site, creating an unstable connection. the root cause of rlm connectivity issue was incorrectly configured or inadequate site network infrastructure (includes low wifi signal strength). this report is being filed as part of a retrospective review of historical records.
15447 | 1220648-2024-22499 | 20634559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20634559&pc=OZD | 12/08/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 12/08/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1122734 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported the automated impella controller (aic) experienced battery charging/other issues. no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the failure mode was reproduced on an engineering bench. the battery 2 lcd indicator was blank (fully discharged). batt test was recorded for the single cell decline. the appearance of ¿smb¿ and missing of ¿fu¿ indicated battery 2 was either not providing any voltage or not communicating at all. see the attached screenshots batttest.png. the reported symptom ¿not charging¿ was associated with the battery 2 lockup, not accepting any charge. the battery failure was due to a defective battery 2. the root cause of the defective battery 2 was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
15448 | 1220648-2024-22502 | 20634594 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20634594&pc=OZD | 11/03/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 11/03/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1117651 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) experienced a purge disc issue, which was resolved by changing the console. no patient harm reported. | [5 empty] | The investigation has been completed. the console ((b)(4)) was sent back for further investigation. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. the root cause of the broken purge flag is due to repeated force and wear and tear of the plastic component. this report is being filed as part of a retrospective review of historical records.
15449 | 1220648-2024-22441 | 20634688 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20634688&pc=OZD | 12/16/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 12/16/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1569441 |  | 0042-0040-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automatic impella controller (aic) had a battery failure alarm when the aic was turned on. another aic console was used. there was no mention of patient involvement. | [5 empty] | The investigation for the reported aic - battery failure (red alarm) issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the cause of the battery failure alarm was due to a defective battery (rapid decline in cell voltage). the exact cause of the defective battery was unable to be determined. this report is being filed as part of a retrospective review of historical records.
15450 | 1220648-2024-22503 | 20634724 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20634724&pc=OZD | 11/01/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 11/01/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1471015 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) experienced an "placement signal not reliable" alarm once connected. it was recommended that the aic be disconnected and restarted. however, multiple efforts to reconnect resulted with the same alarm being received. the console was replaced to address the issue. no patient harm reported. | [5 empty] | The investigation has been completed for optical signal issue. the console (ic8037) was sent back for further investigation. the console was tested thoroughly on an engineering lab bench. the 5s parameters were within specification as received. the console¿s ferrule was not contaminated or damaged. the blue receptacle pump connector was slightly sticky with purge fluid residue. the root cause of the placement signal alarm was most likely due to an air gap at the optical connector. this report is being filed as part of a retrospective review of historical records.
15451 | 1220648-2024-22427 | 20634782 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20634782&pc=OZD | 09/15/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 09/17/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1112882 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The biomedical engineer reported that the automated impella controller (aic) did not recognize the cassette during use. this aic was replaced with another aic resulting in no issues with the replacement aic. there was no patient harm reported. | [5 empty] | The investigation into purge disc/flag issues has been completed. the impella automated controller was returned for investigation. the data analysis of the logs indicated while in the field, the console failed to detect a pump during case start. the were no issues detecting the cassette, the logs show that the console was able to detect pump connection but failed to read pump eeprom. this is a pattern failure mode only found in aic_v8.4 and aic_v8.4.1 and has been documented as a software (sw) defect. a mitigation for the issue has been added to aic_v8.5. the cause of the failure to read the pump eeprom was due to a software defect (in-house). this report is being filed as part of a retrospective review of historical records.
15452 | 1220648-2024-22391 | 20634807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20634807&pc=OZD | 12/20/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 12/20/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1376304 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) experienced an "optical sensor error" alarm was displayed which prompted continuation without sensor placement. the console was replaced to address the issue. no patient harm reported. | [5 empty] | The investigation has been completed. the console ((b)(4)) was sent back for further investigation. the root cause of the optical sensor error is undetermined due to the inability to reproduce. this report is being filed as part of a retrospective review of historical records.
15453 | 1220648-2024-22399 | 20634818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20634818&pc=OZD | 12/13/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 12/13/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1336757 |  | 0042-0000-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automatic impella controller (aic) had a battery failure alarm during preventive maintenance. there was no harm and no patient was involved. | [5 empty] | The investigation for the reported battery failure issue has been completed. the product was returned exact return date is unknown), and the issue was confirmed. data analysis: aic logs confirmed the reported failure. there was a battery failure alarm (transport connection blown/ missing ¿ event set #360). the ltpowerdata from the complaint date showed a cell voltage decline in only one of battery 1¿s cells which occurred as the battery failure alarm triggered. device analysis: the console batteries were replaced with batteries from known good batch and charging/ discharging was validated. the battery failure was due to a defective battery 1. per the results of the clinical review and investigation, the root cause was determined to be a defective battery (rapid decline in cell voltage). the exact cause of the defective battery was utd. this report is being filed as part of a retrospective review of historical records.
15454 | 1220648-2024-22477 | 20634884 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20634884&pc=OZD | 07/25/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 07/25/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1090444 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 16 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported a controller failure alarm that resolved on its own. they did not switch to a backup console. there was no reported harm to the patient. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the root cause of the transient loss of opm-related signals could not be determined due to the inability to reproduce. this report is being filed as part of a retrospective review of historical records.
15455 | 1220648-2024-22443 | 20634898 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20634898&pc=OZD | 12/15/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 12/15/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1147695 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during support the team observed a broken bar on the back of the automated impella controller (aic). no harm or injury from the malfunction. a backup console utilized for support. | [5 empty] | The investigation for the reported aic - controller failure (red alarm) issues has been completed. the device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the housing's support bar damage was most likely mishandling. this report is being filed as part of a retrospective review of historical records.
15456 | 1220648-2024-22469 | 20634959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20634959&pc=OZD | 01/29/2023 | 11/07/2024 |  | 11/07/2024 |  |  | 01/29/2023 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1438016 |  | 0042-0010 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The us complainant in the prep stage of support had a red battery failure observed on the aic (automated impella controller). there was no harm or injury. | [5 empty] | The console (aic) was returned for evaluation. upon evaluation of the console and logs, aic logs confirmed the reported failure modes. internal software, revealed cell2 of battery1 to be at a significantly lower voltage than all the other cells in battery1 causing cell lockout due to cell imbalance. the lcd screen on battery 1 was blank indicating cell lockout. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15457 | 1220648-2024-22440 | 20635019 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20635019&pc=OZD | 05/15/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 05/18/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1169580 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was placed on an impella cp for hemodynamic support. while on support, the automated impella controller (aic) experienced battery purge disc/flag issues. the clinical staff reported that the aic would not detect the purge disc when it was in place. the console was switched out with the same purge cassette and purge disc without further issue. there was no reported harm or injury to the patient. | [5 empty] | The investigation has been completed. the logs from the complaint date revealed that the console issued the purge disc not detected alarm during case start wizard. the alarm was noticed and console was switched out. the console was returned and tested. the issue with properly detecting the purge pressure disc was reproduced and the purge flag was confirmed to be broken (missing at blue disc retainer). the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15458 | 1220648-2024-22445 | 20635137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20635137&pc=OZD | 03/10/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 03/10/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1306044 |  | 0042-0000-JP |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) failed the self check when being turned on for inspection. restarting did not resolve the issue, so the engineer checked the firmware version of each board. since pbm1 was displayed as "na", it was presumed that the power battery manager (pbm) board was malfunctioning. this part was replaced resolving the issue. there was no reported patient involvement. | [5 empty] | D.9 the exact date the device was returned is unknown. e.1 the initial report name is unknown. the investigation for the reported automated impella controller (aic) - system self-check error has been completed. aic logs revealed that this console had clear communication issues with sau powermod 1 which resulted in failure of system self-check. the console was tested in collaboration with japan field service. fwcheck showed n.a under sau powermod 1 section, which indicated communication issues that prevent the fw/hw version from being read. to resolve the communication/ start up issues, the power battery manager (pbm) was replaced. the defective board was not inspected for any damage or corrosion. the cause of the syscall error was a defective pbm board. this report is being filed as part of a retrospective review of historical records.
15459 | 1220648-2024-22418 | 20635222 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20635222&pc=OZD | 10/26/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 10/26/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1351867 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) had multiple controller failure alarms. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported controller failure issue has been completed. the product was returned, and the issue was confirmed. data analysis: complaint intake is consistent with what was found in the logs. logs from the date of the complaint show the following: 26/10/21 10:51:38 imcgui: event set: #1029 from 40 (0) - controller failure (epm1 sw corruption). device analysis: console booted alarming for controller failure. when reviewing logs again, the controller failure was the same as the alarm exhibited in the field. a known good impellatronics board was swapped and resolved the controller failure on next boot. per the results of the clinical review and investigation, the root cause was determined to be corruption of the epm firmware on the impellatronic board. this report is being filed as part of a retrospective review of historical records.
15460 | 1220648-2024-22424 | 20635303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20635303&pc=OZD | 10/01/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 10/01/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1458318 |  | 0042-0010-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 42 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during set up, the insertion of purge disk was not detected by the automatic impella controller (aic). after several attempts the aic was switched for another. there was no harm to the patient. | [5 empty] | D9: device was returned to regional engineers and evaluated and the return date is unknown. the investigation for the reported aic - purge disc/flag issues has been completed. the console was tested after arriving in fs aachen. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. the root cause of the inability to detect the purge disc was due to a broken purge flag. this report is being filed as part of a retrospective review of historical records.
15461 | 1220648-2024-22521 | 20635315 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20635315&pc=OZD | 10/08/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 10/08/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1092251 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced a purge disc issue. no clinical details regarding resolution or patient condition were provided. no patient was involved. | [5 empty] | The investigation has been completed. the console (b)(4) was sent back for further investigation. the purge flag was intact, but the blue pressure disc retainer had glucose build up and was cracked. although the purge pressure sensor was visibly clean upon inspection, fluid ingress on the purge pressure sensor¿s optical sensor could have momentarily impacted the purge pressure disc detection mechanism. the root cause of the inability to detect the purge disc was due to a broken purge pressure disc retainer/ fluid ingress in this particular component. this report is being filed as part of a retrospective review of historical records.
15462 | 1220648-2024-22495 | 20635330 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20635330&pc=OZD | 06/13/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 06/13/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER EU |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER EU |  |  | ABIOMED, INC. |  |  | US |  | 1287216 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The field service engineer reported that the automated impella controller (aic) front housing was broken. there was no patient involvement. | [5 empty] | E1: name of the initial reporter is unknown. the investigation into the damage has been completed. the impella automated controller was not received from the customer. the data analysis of the logs did not reveal any malfunctions. the field service technician identified cracks on the front housing and glass keyboard and replaced the front housing, then performed preventive maintenance. the cause of the aic damage to the housing was unable to be determined. this report is being filed as part of a retrospective review of historical records.
15463 | 1220648-2024-22501 | 20635375 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20635375&pc=OZD | 10/04/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 10/04/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1339213 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported an automated impella controller (aic) that was experiencing battery issues. when plugged in overnight, it was verified that the battery on switch was lit. the battery on switch verified to be on. the device was removed from service because of the noted issue. there was no patient harm as a result of the reported failure. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. console interior connections were checked and verified. when visually inspecting the batteries and pbm, it was observed that one of the batteries status leds were completely blank. the battery failure alarm was due to a defective battery 2 (decline of individual cell voltage). the root cause of the battery cell failure was undetermined. this report is being filed as part of a retrospective review of historical records.
15464 | 1220648-2024-22455 | 20635418 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20635418&pc=OZD | 08/30/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 08/30/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1371809 |  | 0042-0000-JP |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) that immediately after launching automated impella controller (aic), an abnormal noise was generated and the screen blacked out. the power could not be turned off and it was forcibly terminated. the issue occurred immediately after the startup. since there was a case of impella standby, the event was discovered when the power was turned on. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported power on issues/blank screen issue has been completed. the product was returned (exact product return date unk), and the issue was confirmed. data analysis: aic logs confirmed that the console was powered on reported complaint date but did not give any insight on the failure cause. device analysis: fs japan received the console in in their lab and found a burn mark on a capacitor on 4-in-1 single board computer (sbc). per the results of the clinical review and investigation, the root cause was determined to be due to a defective 4-in-1 assembly. this report is being filed as part of a retrospective review of historical records.
15465 | 1220648-2024-22478 | 20635433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20635433&pc=OZD | 10/29/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 10/29/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1359283 |  | 0042-0000-JP |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automatic impella controller (aic) was started up for a software update, but the message "self-diagnosis failed. please replace the control device" was displayed. the same message appeared when the aic was forced to close and started up again, and it could not be started. therefore, a backup machine was prepared. there was no report of patient harm or injury. | [5 empty] | D9: date device returned to manufacturer: unknown. the investigation for the reported system self-check error issue has been completed. the product was returned, and the issue was confirmed. console logs from the reported day of event are consistent with complaint and show multiple failed boot-ups, and persistent powermod_1 (pbm1) communication failures. when console was opened and inspected by (b)(6), some corrosion was observed on the pbm, near capacitors c69 and c70. after pbm assembly was replaced the issue was resolved and defective pbm was sent to danvers for further analysis. in danvers, the issue was confirmed to be consistent with a known pattern failure of pbm corrosion due to electrolyte leak, that was in the past addressed by fcn 71. in this case, however, serial number and lot number of affected pbm placed it outside known range of affected material, making this the first observed and documented occurrence of this failure mode affecting material 0042-1125 (so far only a subset of 0042-1025 was affected). during visual inspection of the pbm some component damage on the underside of pbm was observed, but based on log analysis and recognized symptoms, the original issue was caused by corrosion, and the secondary damage sustained most likely during removal of already defective pbm from aic, therefore unrelated to the original complaint. per the results of the clinical review and investigation, the root cause was determined to be corrosion near the c69 and c70 supercapacitors on the pbm. this report is being filed as part of a retrospective review of historical records.
15466 | 1220648-2024-22417 | 20635519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20635519&pc=OZD | 11/13/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 11/13/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1168789 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced a purge disc issue. no clinical details regarding resolution or patient condition were provided. no patient was involved. | [5 empty] | The investigation has been completed. the console (ic3384) was sent back for further investigation. the blue disc retainer was broken off of the purge assembly. the root cause of the purge disc not detected alarms is due to a broken purge disc retainer. this report is being filed as part of a retrospective review of historical records.
15467 | 1220648-2024-22403 | 20635579 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20635579&pc=OZD | 12/07/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 12/07/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1148749 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) was returned from preventative maintenance. it was reported that when the console was initiated, a noise was observed from the cooling fan. it was stated that the console was potentially damaged during shipping. however, no visual damage to the shipping container was noted. the customer has another device onsite to utilize. no report of patient involvement, harm, or injury. | [5 empty] | The impella device has been returned for evaluation. it was found that the speaker wires were out of position and slipped into the cooling fan enclosure. it was verified that the complained noise was caused from the cooling fan leaves hit the speaker wires (blue and red) during spinning. the cause of this issue is unknown. this report is being filed as part of a retrospective review of historical records.
15468 | 1220648-2024-22423 | 20635603 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20635603&pc=OZD | 09/29/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 09/29/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1446476 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | Device Fell | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The biomedical engineer reported that the automated impella controller (aic) requested the console be checked out as the console had fallen and broke. a loaner was sent to the customer to initiate return of this device. there was no patient involved. | [5 empty] | The investigation into the damage has been completed. the impella automated controller was returned for investigation. the returned console was inspected, and it was noted the front panel optical connector cover was broken off, and the front panel optical connector was loose and freely moving. the console interior was inspected, and no damage was observed. the console was opened, and the interior components were checked for damage, no other components were damaged from the fall. the failure mode is due to damage from the console fall. the root cause of the console housing damage is due to the mishandling of the console by the end user. this report is being filed as part of a retrospective review of historical records.
15469 | 1220648-2024-22489 | 20635624 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20635624&pc=OZD | 10/16/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 10/16/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1276515 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. it was reported that the automated impella controller (aic) could not be switched off. despite the fact that the white connector cable is disconnected from the aic, the display shows the message that the aic cannot be switched off because the connector cable is connected. this issue was discovered after the removal of the device. no report of patient harm or injury. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the aic issue was most likely a defective impellatronics board, that resulted in misread gauge factors and snr threshold during pump eeprom read. this report is being filed as part of a retrospective review of historical records.
15470 | 1220648-2024-22496 | 20635634 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20635634&pc=OZD | 06/14/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 06/14/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1108441 |  | 0042-0000-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient of unknown race/ethnicity noted as high-risk percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. during use, the automated impella controller (aic) produced a "system check failed, replace console immediately." the console was replaced, and there was no reported patient harm. | [5 empty] | Device was returned to field service, date unknown the device was returned and evaluated, and the investigation for the reported system self-check error has been completed. imc logs from the complaint date were consistent with the complaint description. the failure mode was reproduced and a blue screen with ¿system self-check failure¿ presented. internal connection and components were visually inspected, finding burnt mark on pbm, underneath the super capacitors. the defective power battery manager (pbm) was temporarily replaced with a known-good one, with which console could boot up and operate as normal. the root cause of system self-check failure was due to a defective pbm. the damage to the pbm was due to leaking electrolyte from the neighboring super capacitors which impacted i2c communication. this report is being filed as part of a retrospective review of historical records.
15471 | 1220648-2024-22467 | 20635765 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20635765&pc=OZD | 08/06/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 08/06/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1119398 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the automated impella controller (aic) was started then approximately three minutes later the alarm "placement signal not reliable" started. this aic was replaced with another aic resulting in no alarming with the replacement aic. there was no patient harm reported. | [5 empty] | The investigation into the optical/placement signal issue has been completed. the impella automated controller was returned for investigation. the data analysis of the logs confirmed that a placement signal not reliable (opm signal invalid ¿ event set #1036) alarm was issued on the complaint date. the real time logs revealed the signal not reliable (snr) value was very low and placement signal was spiking to the max value of 3000. the returned console was tested thoroughly on an engineering lab bench, and the failure mode was reproduced while running a known-good optical pump and purge. spiking had occurred and eventually a placement signal not reliable (opm signal invalid ¿ event set #1036) alarm was issued. with all original components, the 5 signal parameters were within specification, but snr (measured at 4294) was relatively low compared to other functional consoles. the ferrule at the blue receptacle was confirmed to be clean. reseating the optical connectors at the bulkhead did not improve the snr value. replacing the optical bench pca not only improved the snr value but also resolved the intermittent spiking. after swapping the optical bench, a known-good pump and purge was run with no issues observed. the cause of the placement signal issues was due to a defective optical bench. the exact root cause of the defective optical bench was not determined. this report is being filed as part of a retrospective review of historical records.
15472 | 1220648-2024-22442 | 20635766 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20635766&pc=OZD | 04/07/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 04/07/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1579730 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) there was a controller failure (red alarm). there was no report of patient harm or injury. | [5 empty] | The investigation into the controller failure (red alarm) issue has been completed. the automated impella controller (aic) was returned for investigation. there were no entries in console logs on the reported day of event, 2022-04-07. based on all available logs between 2022-03-29 and 2022-05-04 we were unable to identify any issues that would match exactly the problem description "a few times, after turning on the console, it will say controller failure. it will resolve, but it keeps happening". events on 2022-04-05, specifically a single "piston blocked" event, several purge fluid not detected error messages, as well as purge system open alarms (#407) and air in purge system alarms (#404) during purge changes, may have been what the complaint referred to as "purge issues". on the same day (during same case), there were also two controller failure alarms (#414) triggered by "purge pressure sensor defective" condition, that self-resolved within seconds, which presumably prompted the complaint narrative. alarms #414 were most likely due to purge line leak or use issue rather than malfunction of the purge pressure transducer, as evidenced by valid values of purge pressure signal during the event. this alarm condition could have been triggered by a small leak which resulted in measured pressure not increasing during forward stroke of purge cassette piston - a condition that triggers alarm #414 as per design (0042-9702). same purge cassette was removed and re-inserted a few times during the case, apparently as troubleshooting attempts. alarms #407 and #404 were most likely caused by an open line or disconnected luers (use issue). neither the pump or purge cassette were returned for investigation. no irregularities were discovered during testing of the console in danvers, and the console operated within specification - including the purge system that underwent additional thorough testing as per sec. the cause of the issue was most likely use related. this report is being filed as part of a retrospective review of historical records.
15473 | 1220648-2024-22471 | 20635770 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20635770&pc=OZD | 12/18/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 12/18/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1087949 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | No Audible Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that suction and positioning alarms during impella placement did not produce audible sounds when controller buttons were pressed or settings were changed. attempts to disable and reenable the alarms were unsuccessful. it was reported that a review of the automated impella controller (aic) would be initiated to resolve the issue. the initial aic was replaced and the case was continued. | [5 empty] | The investigation for the reported automated impella controller (aic) - buzzer/ alarm inaudible has been completed. the device has been returned for evaluation. the speaker assembly cable was not connected to epc. the cause of the no audible alarms was a loose speaker cable. this report is being filed as part of a retrospective review of historical records.
15474 | 1220648-2024-22462 | 20635796 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20635796&pc=OZD | 08/08/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 08/08/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1273696 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The clinical consultant reported that the automated impella controller (aic) stopped recognizing the catheter. this aic was replaced with another aic resulting in no issues with the replacement aic. there was no patient harm reported. | [5 empty] | The investigation into the pump not detected/controller failure issue has been completed. the impella automated controller (aic) was returned for investigation. the data analysis of the logs did not confirm the reported failure to detect the impella pump. during the preparation for the clinical procedure, the pump was detected, and eeprom was read. there were eventually issues with the placement signal preceding optical bench communication issues, which resulted in a loss of optical data and an issue detecting pump disconnection. as a result, the console triggered a controller error (opm communication loss - event set 1043) which did not resolve until the console was power cycled. the returned console (ic4825) was tested thoroughly by the post market engineering team, and the failure was not reproduced. this known pattern failure, which is characterized by a temporary loss of optical data and triggering of a controller error alarm, has a low probability of recurrence. the cause of the loss of optical data was unable to be determined because the issue was unable to be reproduced. this report is being filed as part of a retrospective review of historical records.
15475 | 1220648-2024-22475 | 20635927 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20635927&pc=OZD | 07/30/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 07/30/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1137163 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. the complainant reported a console failure alarm during support. the device was replaced with no reported harm or injury to the patient. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the optical signal loss was unable to be determined because the issue was unable to be reproduced. this report is being filed as part of a retrospective review of historical records.
15476 | 1220648-2024-22450 | 20635986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20635986&pc=OZD | 01/10/2022 | 01/02/2025 |  | 11/07/2024 |  |  | 01/10/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility's biomedical engineer reported a vulnerability that could impact the automated impella controller. no clinical details regarding resolution. no patient was involved. | [5 empty] | This is a corrected mdr, submitted to replace the original mfr (b)(4) which was filed in error. as there was no impact to any abiomed product or service for a reported potential log4j v1.2.x vulnerability there is no reportable incident that occurred. | Investigation into the reported issue has been completed. no device was received for evaluation. product security impact assessment concluded that none of abiomed products or services is affected by this vulnerability. there was no impact on any abiomed product or service. this report is being filed as part of a retrospective review of historical records.
15477 | 1220648-2024-22472 | 20636034 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636034&pc=OZD | 11/08/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 11/08/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1424180 |  | 0042-0010-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Read Input Signal | Connection Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella 5.0 for mechanical circulatory support. the automated impella controller (aic) was unable to recognize the pump. both cable connections were disconnected and reconnected without success; impella cable changed without success. the console was replaced to address the issue. no patient harm reported. | [5 empty] | The investigation has been completed for pump not detected. the console was not returned for further investigation. there were no apparent issues with the pump detection mechanism after the complaint date. the root cause of the pump detection issues was most likely due to the epm crystal issue, which did not resolve after the epm sau reset (mitigation). this report is being filed as part of a retrospective review of historical records.
15478 | 1220648-2024-22492 | 20636058 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636058&pc=OZD | 04/05/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 04/05/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1338086 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The field service engineer was performing preventative maintenance, it was discovered that the fan on the automated impella controller (aic) console was inoperable. it was discovered that a fuse was open at the power manager pca. no patient involvement. | [5 empty] | E1, e3, e4: the name of the initial reporter and occupation is unknown. the investigation for the reported issue has been completed. field service evaluated the device, the cause of the one inoperable fan was an open fuse on the pbm. fuse f11 was found open, root cause is defective component. this report is being filed as part of a retrospective review of historical records.
15479 | 1220648-2024-22474 | 20636072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636072&pc=OZD | 08/02/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 08/02/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1150347 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported that during preventative maintenance the automated impella controller experienced low bulb power, the field engineer resolved this issue by replacing the bulb. no patient involvement. | [5 empty] | The investigation has been completed. the root cause of the low bulb output was due to a defective optical bench. the cause of the defective optical bench was not determined. this report is being filed as part of a retrospective review of historical records.
15480 | 1220648-2024-22484 | 20636088 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636088&pc=OZD | 03/13/2023 | 11/07/2024 |  | 11/07/2024 |  |  | 03/13/2023 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1606415 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automated impella controller (aic) experienced a battery failure red alarm. no patient was involved. | [5 empty] | The investigation has been completed. the console (ic9143) was sent back for further investigation. aic logs confirmed the reported failure. there was a battery failure alarm (transport connection blown/ missing ¿ event set #360) due to low pack voltage. the cause of the aic issue was a defective battery due to a single cell imbalance. this report is being filed as part of a retrospective review of historical records.
15481 | 1220648-2024-22444 | 20636096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636096&pc=OZD | 05/13/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 05/13/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1324632 |  | 0042-0010-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced power on issues/blank screen. no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | D.9 the exact date the device was returned is unknown. e.1, e.2 and e.3 the first and last name of the initial reporter and health professional/occupation are unknown. the investigation for the reported issue has been completed. logs are consistent with complaint intake. subsequent boots following the complaint date trigger these alarms. long term log analysis of the battery data revealed that beginning before the complaint date of occurrence, cell 1 voltage of battery 1 had been declining for an extended period of time. the console was returned to field service in germany. it was confirmed that persistent battery failure (transport connection blown/missing) was not due to a disengaged battery switch. when console was booted for the first time, the console alarms were consistent with what was seen in the field. the console interior connections were checked and verified. when visually inspecting the batteries and power battery manager, it was observed that one of the batteries status leds were completely blank. the battery failure alarm was due to a defective battery 1 (decline of individual cell voltage). the root cause of the battery cell failure was undetermined. this report is being filed as part of a retrospective review of historical records.
15482 | 1220648-2024-22458 | 20636098 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636098&pc=OZD | 08/24/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 08/24/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1100968 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | Insufficient Information | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. the customer plugged the console in for several hours and alarm battery failure / battery communication. it is unknown as to how the issue was managed. failure alarms were still present and there was no report of patient involvement or harm. | [5 empty] | The investigation into the a/c power issue has been completed. the automated impella controller (aic) was returned for investigation. the battery failure alarm was due to a defective battery 2 (decline of individual cell voltage). the root cause of the battery cell failure was undetermined. before returning to customer battery was replaced. this report is being filed as part of a retrospective review of historical records.
15483 | 1220648-2024-22412 | 20636114 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636114&pc=OZD | 11/16/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 11/16/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1138128 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) experienced a purge disc issue, which was resolved by changing the console. no patient harm reported. | [5 empty] | The investigation has been completed. the purge disc detection failure was confirmed by the disassembly of the purge system. purge disc detection flag was visibly broken. therefore, the root cause of the purge disc detection issue was a broken cantilever (purge disc detection flag) in purge assembly. the root cause of the purge disc not recognized was a broken purge flag. this report is being filed as part of a retrospective review of historical records.
15484 | 1220648-2024-22527 | 20636118 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636118&pc=OZD | 05/14/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 06/21/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1234571 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  | Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) experienced false placement signal alarms were noted. no clinical details regarding resolution or patient involvement / condition were provided. | [5 empty] | The investigation has been completed. the console (b)(6) was sent back for further investigation. the cause of the aic issue was most likely due to a defective blue receptacle. this report is being filed as part of a retrospective review of historical records.
15485 | 1220648-2024-22508 | 20636186 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636186&pc=OZD | 06/22/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 06/22/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1550078 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility biomedical engineer reported for an automated impella controller (aic) that the blue purge clip was broken. a loaner sent to initiate the return for investigation/repair. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported purge disc/flag issue has been completed. the product was returned, and the issue was confirmed. data analysis: imc logs are not relevant for this failure mode. device analysis: console purge assembly was inspected. upon opening the purge assembly door, it was confirmed the purge disc retainer had broken/snapped off of the purge assembly. per the results of the clinical review and investigation, the root cause was determined to be a defective component. this report is being filed as part of a retrospective review of historical records.
15486 | 1220648-2024-22448 | 20636206 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636206&pc=OZD | 12/29/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 12/29/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1126792 |  | 042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 18 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support, the automated impella controller (aic) did not recognize the purge cassette or disc, which was resolved by changing the console. no patient harm reported. | [5 empty] | The investigation has been completed. the console (ic2196) was sent back for further investigation. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. the root cause of the broken purge flag is most likely due to repeated force and wear and tear of the plastic component. this report is being filed as part of a retrospective review of historical records.
15487 | 1220648-2024-22483 | 20636209 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636209&pc=OZD | 11/07/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 11/07/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1561365 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella rp for mechanical circulatory support. while on support, the automated impella controller (aic) experienced an unexpected shutdown. the aic rebooted and returned to the prior pump speed level. the patient was off of support for approximately twenty seconds, and the patient was supported by titration of epinephrine during the event. no patient harm. | [5 empty] | The investigation has been completed for aic - shutdown or restart issue. the console was not returned. the root cause of the unexpected shutdown was unable to be determined because the console and data were not returned for review. this report is being filed as part of a retrospective review of historical records.
15488 | 1220648-2024-22494 | 20636217 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636217&pc=OZD | 03/31/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 03/31/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1122787 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Nurse | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the console alarmed "system self-check failed/change console immediately", this occurred upon console startup. no further information is available. no report of patient involvement, harm, or injury. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the system self-check failure was due to a failed pbm because of corrosion near the c69 and c70 supercapacitors. this report is being filed as part of a retrospective review of historical records.
15489 | 1220648-2024-22493 | 20636250 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636250&pc=OZD | 08/31/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 08/31/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER US |  |  | ABIOMED, INC. |  |  | US |  | 1129486 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 27 YR |  | Female |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an automated impella controller (aic) for mechanical circulatory support. while on support, the aic experienced a/c power issues that were resolved by changing out the console. no reported harm to patient. | [5 empty] | The investigation into the a/c power issue has been completed. the automated impella controller (aic) was returned for investigation. when the console arrived, the battery switch was engaged and ac power was connected, however the console still would not recharge. console was opened and interior connection integrity was verified. the ac power supply was temporarily replaced by the known good one, and the failure mode was resolved, and the batteries began to properly charge. the root cause of the aic not recognizing being plugged was due to a defective ac power supply. this report is being filed as part of a retrospective review of historical records.
15490 | 1220648-2024-22498 | 20636264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636264&pc=OZD | 06/08/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 06/08/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1420044 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a jammed front door on the automatic impella controller (aic). the device was removed, and a loaner was shipped. no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | The investigation for the reported aic - purge door won't close issues has been completed. the impella device has been returned for evaluation. inspection on purge door found the purge door push button unable to bound back after being pushed. the jammed purge door push button was due to glucose residue build up around the button. after the residue was cleaned, the purge door push button could work as expected. the cause of the aic issue was contamination. this report is being filed as part of a retrospective review of historical records.
15491 | 1220648-2024-22509 | 20636269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636269&pc=OZD | 06/17/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 06/17/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1327996 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced with an electrical issue in the procedure room. no clinical details regarding resolution or patient condition were provided. no patient was involved. | [5 empty] | The investigation has been completed. the console (b)(6) was sent back for further investigation. console logs from the reported day of event were not analyzed as the console had not been powered on at that time. after console was brought to danvers for troubleshooting, visual inspection of the ac cord revealed sings of arcing, confirming reports of "electrical issue / short circuit". also, one prong was observed to be bent, and the outside screw on the ac plug had visible oxidation, suggesting possible exposure to moisture or liquids. subsequent testing of aic's power entry module and power supply did not reveal any irregularities. both inline fuses in console rear housing were in good condition. after ac cord set replacement console (b)(6) passed electrical safety testing. it is most likely that the reported electrical issue originated at the ac wall socket, but there's no evidence pointing to the aic's hardware as the cause. the root cause of the issue was not determined due to insufficient evidence. this report is being filed as part of a retrospective review of historical records.
15492 | 1220648-2024-22452 | 20636332 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636332&pc=OZD | 01/04/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 01/04/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 119474 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced a broken purge disc. a new console was used without any issues. no patient was involved. | [5 empty] | The investigation has been completed. the device was received for evaluation. logs revealed the purge disk being inserted (purger status: sensor = 1), but suddenly the status changed to 0, which represented the blue housing (purge disk retainer) came apart and failed to hold the purge disk. the root cause of defective blue retainer was unable to be determined. this report is being filed as part of a retrospective review of historical records.
15493 | 1220648-2024-22447 | 20636353 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636353&pc=OZD | 01/17/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 01/17/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1133108 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced a purge disc issue. no clinical details regarding resolution or patient condition were provided. no patient was involved. | [5 empty] | The investigation has been completed. the console (ic2379) was sent back for further investigation. console was ran with pump and purge to test the stepper motor priming. it was observed that the stepper motor was stalling attempting to prime the pump. the root cause of the failure mode was contamination. this report is being filed as part of a retrospective review of historical records.
15494 | 1220648-2024-22485 | 20636366 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636366&pc=OZD | 05/03/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 05/03/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1167442 |  | 0042-0000-UK |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Fell | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The field service engineer reported that the automated impella controller (aic) fell on the floor and broke. there was no patient harm reported. | [5 empty] | E1: name of the initial reporter is unknown. the investigation into the damaged console has been completed. the impella automated controller (aic) was not received from the customer and therefore, an evaluation of the device was not possible. the console fell on the floor and broke. multiple cracks and scuffs on the housing. according to the pictures, the console housing (front and rear) was damaged. it is possible that the optical pump connector was also damaged and may be replaced on a precautionary basis. the cause of the cracked housing was user mishandling of the console where the console fell to the floor. this report is being filed as part of a retrospective review of historical records.
15495 | 1220648-2024-22464 | 20636373 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636373&pc=OZD | 08/10/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 08/10/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1210871 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced a system self check failure. the resolution for this issue is unknown. no report of patient involvement, harm, or injury. | [5 empty] | The device was returned for evaluation. failure mode was reproduced on first boot, console booted into a system self-check error. impellatronics board was inspected and both power leds were observed and lit. both supply lines from the pbm were at expected voltages of ~5v and ~20v. a known good impellatronics board was swapped into console. console booted without issue or alarms. this report is being filed as part of a retrospective review of historical records.
15496 | 1220648-2024-22482 | 20636376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636376&pc=OZD | 06/01/2023 | 11/07/2024 |  | 11/07/2024 |  |  | 06/01/2023 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1314428 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) failed the system self-check. there was no reported patient harm. | [5 empty] | The device was returned and evaluated, and the investigation for the reported system self-check error has been completed. imc logs revealed persistent unsuccessful boots on complaint date. in each boot, console was unable to retrieve aic software revision from memory. communication initialization with other subcomponents was successful. at field service, console was unable to boot up as reported. however, firmware check did not reveal any abnormality on peripheral subcomponents. console¿s compact flash cards were temporarily replaced, and it could boot up normally. the root cause of the syscall error was due to defective compact flash cards. this report is being filed as part of a retrospective review of historical records.
15497 | 1220648-2024-22486 | 20636396 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636396&pc=OZD | 11/07/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 12/20/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1325616 |  | 0042-0000-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 42 YR |  | Female |  |  | Asian | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) experienced a purge disc issue. no clinical details regarding resolution. no patient harm reported. | [5 empty] | The investigation has been completed. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. the root cause of the inability to progress through the purge bag change procedure is due to a missing/broken purge flag. this report is being filed as part of a retrospective review of historical records.
15498 | 1220648-2024-22470 | 20636423 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636423&pc=OZD | 08/03/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 08/03/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1303853 |  | 0042-0000-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) was failing to start. it was reported that when switching the console on the screen, it remains black. soft keys were beeping, but not illuminated. rotary knob not beeping. power light-emitting diode (led) green. telnet connection not possible. the resolution for this issue is unknown. no report of patient involvement, harm, or injury. | [5 empty] | The date of the devices return for evaluation is unknown. the name of the initial reporter and occupation is unknown. the investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the root cause of the consoles intermittent failure to boot is due to the pbm. the pbm malfunction is undetermined. this report is being filed as part of a retrospective review of historical records.
15499 | 1220648-2024-22460 | 20636452 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636452&pc=OZD | 08/24/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 08/24/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1314856 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The biomedical engineer reported that the automated impella controller (aic) when setting up for a case start the surgeon was unable to insert the pump cable into the console pump connector. a loaner was sent to the customer. there was no patient harm reported. | [5 empty] | The investigation into the case start issue has been completed. the impella automated controller (aic) was returned for investigation. the data analysis of the logs revealed that the case start procedure was attempted a few days before the reported complaint date, but the console would not detect a pump at the screen which prompts for pump connection. the console was tested thoroughly on an engineering bench. the failure mode was not reproduced as there were no issues properly detecting the pump in between power cycles. this is a pattern failure mode called the epm crystal issue, which affects the ability of the impellatronic¿s epm circuitry to detect the pump. software release 8.3 had brought the feature to reset the epm circuitry, which was intended to be a solution for this issue. however, even with a reset of the (epm) circuitry, the console still had issues detecting the pump. the root cause of the issues properly detecting the pump was due to the impellatronic¿s epm crystal issue. the cause of the failure of the pump detection mechanism was an issue with the crystal of the epm on the impellatronic board. this report is being filed as part of a retrospective review of historical records.
15500 | 1220648-2024-22473 | 20636473 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636473&pc=OZD | 08/03/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 08/03/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1272366 |  | 0042-0000-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported that when checking the software version, the automated impella controller experienced persistent beeping. the console was replaced to address the issue. no patient involvement. | [5 empty] | D9: the date of the devices return for evaluation is unknown. the device was returned for evaluation. the failure mode was reproduced after shutting the console down. the cause of the failure mode is due to a malfunction of the pbm. the exact root cause of the failure is undetermined. this report is being filed as part of a retrospective review of historical records.
15501 | 1220648-2024-22479 | 20636596 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636596&pc=OZD | 10/29/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 10/29/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1194666 |  | 0042-0000-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) experienced a system self-check failure after the battery was placed. the device logs observed that the 'pbm' needed to be replaced. the 'pbm' was replaced, and the aic functioned as intended. there was no report of patient harm or injury. | [5 empty] | D9 - date device returned to manufacturer: unknown. the investigation for the reported system self-check error issue has been completed. the product was returned, and the issue was confirmed. data analysis: imc logs show on boot up, post fails and powers console into a system self check failure. device analysis: console was retrieved by engineering. console was powered on multiple times where the system self check error was exhibited. a visual inspection of the console interior was performed, with specific attention to the pbm, no visible damage was observed. a known good pbm was replaced into (b)(6) and the console was booted multiple times without issue. per the results of the clinical review and investigation, the root cause was determined to be due to powermod_2 communication errors. the exact root cause of the powermod_2 errors is undetermined. this report is being filed as part of a retrospective review of historical records.
15502 | 1220648-2024-22457 | 20636610 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636610&pc=OZD | 01/03/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 01/04/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1115051 |  | 0042-0000-EU |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. during preparation of the automated impella controller (aic), the purge disc was not recognized, which was resolved by changing the console. no patient harm reported. | [5 empty] | The investigation has been completed. logs from the complaint date revealed that the console issued the purge disc not detected alarm several times. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. before returning this console back to the customer, field service was instructed to replace the purge flag and perform a full functional check on the console and optical system. the root cause of the inability detect the purge disc was due to a broken flag. this report is being filed as part of a retrospective review of historical records.
15503 | 1220648-2024-22524 | 20636635 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636635&pc=OZD | 08/25/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 08/25/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1365450 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced a purge disc/flag issues. a new console was used without any issues. no patient was involved. | [5 empty] | The investigation has been completed. the console ((b)(6)) was sent back for further investigation. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. the root cause of the broken purge flag is due to repeated force and wear and tear of the plastic component. this report is being filed as part of a retrospective review of historical records.
15504 | 1220648-2024-22497 | 20636649 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636649&pc=OZD | 11/04/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 11/05/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1284913 |  | 0042-0010-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced a controller error yellow alarm, which was resolved by changing the console. no patient involvement reported. | [5 empty] | The investigation has been completed. no device was returned for evaluation, therefore a confirmed root cause can not be determined. after review of data logs, the root cause of the failure mode was most likely a software issue. this report is being filed as part of a retrospective review of historical records.
15505 | 1220648-2024-22465 | 20636653 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636653&pc=OZD | 01/24/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 01/24/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1319822 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported an automated impella controller was powered on during preparation for a case. three alarms were noted on the screen. they were the following: battery failure, battery level low, and battery comm. failure. a call was made for instructions on to resolve the issue but all attempts to reset the controller were unsuccessful. the alarms remained. there was no reported harm or injury to the patient. | [5 empty] | The investigation for the reported automated impella controller (aic) - battery failure (red alarm) has been completed. logs were consistent with complaint intake on the complaint date. subsequent boots following the complaint date trigger these alarms. long term log analysis of the battery data revealed that beginning before the complaint date of occurrence, cell 3 voltage of battery 2 had been declining for an extended period of time. the console was returned. it was confirmed that persistent battery failure (transport connection blown/missing) was not due to a disengaged battery switch. when console was booted for the first time, the console alarms were consistent with what was seen in the field. console interior connections were checked and verified. when visually inspecting the batteries and power battery manager, it was observed that one of the batteries status leds were completely blank. the battery failure alarm was due to a defective battery 2 (decline of individual cell voltage). the root cause of the battery cell failure was undetermined. the cause of the battery failure and battery comm error alarms is due to the decline of a single cell within battery 2. this report is being filed as part of a retrospective review of historical records.
15506 | 1220648-2024-22466 | 20636661 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636661&pc=OZD | 12/13/2021 | 11/07/2024 |  | 11/07/2024 |  |  | 12/13/2021 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1574757 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automatic impella controller (aic) had a battery failure alarm whenever it was turned on. no additional information provided, no report of patient involvement, injury, or harm. | [5 empty] | The investigation for the reported aic - battery failure (red alarm) issues has been completed. the device has been returned for evaluation. the cause of the battery failure alarm was due to a defective battery (rapid decline in cell voltage). the exact cause of the defective battery was unable to be determined. this report is being filed as part of a retrospective review of historical records.
15507 | 1220648-2024-22511 | 20636675 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636675&pc=OZD | 05/05/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 05/05/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1129883 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Audible Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automatic impella controller (aic) had no audible alarm when the yellow or red alarm was active. no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | The investigation for the reported aic - buzzer/ alarm inaudible issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the aic issue (no audible alarm) was an improperly seated speaker cable, most likely due to work instruction / technician error. this report is being filed as part of a retrospective review of historical records.
15508 | 1220648-2024-22490 | 20636700 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20636700&pc=OZD | 09/29/2022 | 11/07/2024 |  | 11/07/2024 |  |  | 09/29/2022 | 11/07/2024 |  |  |  |  | 11/07/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1117233 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) exhibited a purge controller error, and a loaner device was sent to the customer. no report of patient harm or injury. | [5 empty] | The investigation for the reported purge controller error has been completed. the device logs were reviewed revealed that the console issued the purge disc not detected alarm. the product was returned, and the display issue was confirmed. per the results of the investigation: the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. this report is being filed as part of a retrospective review of historical records.
15509 | 1220648-2024-22121 | 20616101 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20616101&pc=OZD | 04/16/2021 | 11/05/2024 |  | 11/06/2024 |  |  | 04/16/2021 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1126092 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella for mechanical circulatory support. during support, the automated impella controller (aic) experienced a controller failure red alarm that was resolved by switching to the backup aic. there was no reported patient harm. | [5 empty] | The investigation for the reported controller failure (red alarm) has been completed. console logs from the reported day of event are not consistent with complaint, and do not show any evidence of console failure or malfunction. the only significant alarms, pump stopped due to mechanical or electrical issue and running pump disconnected presented towards the end of case and were caused by the pump being disconnected while running. this is consistent with the automated impella controller software design. the console was returned for investigation. the console successfully passed all tests associated with preventive maintenance procedure and operated within specification. the root cause was unable to be determined. this report is being filed as part of a retrospective review of historical records.
15510 | 1220648-2024-22100 | 20616111 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20616111&pc=OZD | 08/24/2022 | 11/05/2024 |  | 11/06/2024 |  |  | 08/24/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1545370 |  | 0042-0010 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that after power up, aic (automated impella controller) shows battery fault without any pump connected. field service was contacted and console must be sent in for repair. successfully switched to new console. no harm done to the patient. | [5 empty] | The investigation for the reported aic - battery failure (red alarm) issues has been completed. the controller device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the aic charging issue was a defective pbm board. this report is being filed as part of a retrospective review of historical records.
15511 | 1220648-2024-22123 | 20616138 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20616138&pc=OZD | 02/09/2022 | 11/05/2024 |  | 11/06/2024 |  |  | 02/09/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1570900 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported an automated impella controller (aic) battery failure during operation. the initial device was replaced, and the case continued. there was no reported patient harm. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the root cause of the battery failure is due to the decline of a single cell voltage within battery one. this report is being filed as part of a retrospective review of historical records.
15512 | 1220648-2024-22114 | 20616148 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20616148&pc=OZD | 07/13/2021 | 11/05/2024 |  | 11/06/2024 |  |  | 07/13/2021 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1218087 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | Connection Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced emitting an "impella defective" alarm prior to connecting any impella catheter. staff tried to then power off aic and was unable to. no clinical details regarding resolution were provided. no patient was involved. | [5 empty] | The investigation into the reported issue has been completed. logs confirmed the reported alarm triggering issue. an impella defective alarm was issued with no pump connected. there were also issues properly shutting the console down as a result. see the attachments section for the imc logs from the complaint date. the failure mode was reproduced by power cycling the console. after swapping the impellatronic pca, the failure was resolved. the root cause of the erroneous impella defective alarm was due to a defective impellatronic pca, specifically the epm circuitry. this report is being filed as part of a retrospective review of historical records.
15513 | 1220648-2024-22173 | 20616152 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20616152&pc=OZD | 02/14/2022 | 11/05/2024 |  | 11/06/2024 |  |  | 02/14/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, UK |  |  | ABIOMED, INC. |  |  | US |  | 1475782 |  | 0042-0010-UK |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was on an automated impella controller (aic) for mechanical circulatory support. it was reported that the device was connected to an ac power source but kept alarming the "ac power disconnected" and could not be charged. the staff tried other ac sources without success. no report of patient harm. | [5 empty] | Device return date unknown. the investigation into the a/c power issue has been completed. the automated impella controller (aic) was not returned for investigation. the cause of the issue was not determined since the failure mode was not reproduced after thorough testing and inspection. this report is being filed as part of a retrospective review of historical records.
15514 | 1220648-2024-22156 | 20616201 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20616201&pc=OZD | 09/07/2021 | 11/05/2024 |  | 11/06/2024 |  |  | 09/07/2021 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1172893 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced a display issue. there was no resolution specified and no reported patient involvement. | [5 empty] | The device was returned and evaluated, and the investigation for the reported display issue has been completed. device analysis: console arrived in danvers and on first boot issue was reproduced. console screen was blank, softkeys were responsive, rpb was not responsive. console was opened and connections were inspected. it was observed that the j13 connector of the pbm was not seated properly. j13 carries the pc_on signal responsible for powering the 4-in-1 when console power button is pressed. the console was power-cycled several times with the connector as is to reproduce the issue. the pc_on signal was checked during console boot and gave an expected high voltage value of ~5v. j13 connector was reseated and the console booted properly. failure mode is due to lack of power on signal from the pbm to the 4-in-1. the root cause is due to a loose cable/technician error during annual service. this report is being filed as part of a retrospective review of historical records.
15515 | 1220648-2024-22094 | 20616221 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20616221&pc=OZD | 04/11/2022 | 11/05/2024 |  | 11/06/2024 |  |  | 04/11/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED. INC. |  |  | US |  | 1084300 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during preparation the automated impella controller's (aic) purge disk reader was completely broken off. no further information available, no reported harm or injury to patient. | [5 empty] | The device was returned for evaluation, visual inspection of the purge assembly area confirmed a ripped/ missing blue retainer. the root cause is defective component. this report is being filed as part of a retrospective review of historical records.
15516 | 1220648-2024-22124 | 20616251 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20616251&pc=OZD | 02/02/2022 | 11/05/2024 |  | 11/06/2024 |  |  | 02/02/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1184608( |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported an automated impella controller (aic) battery failure alarm during operation. it was reported that the issue did not resolve despite the aic being charged. there was no reported patient involvement. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the aic issue was a defective pbm board, more specific, capacitor c62 corrosion. this report is being filed as part of a retrospective review of historical records.
15517 | 1220648-2024-22110 | 20616299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20616299&pc=OZD | 12/21/2021 | 11/05/2024 |  | 11/06/2024 |  |  | 12/21/2021 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1299913 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported automatic impella controller (aic) red alarms was not producing correct sequence of beeps or audible tone. there was no harm or injury to the patient. no additional information provided. | [5 empty] | The investigation for the reported aic - battery failure (red alarm) issues has been completed. the device has not been returned for evaluation. the cause of the aic alarm issue was unable to be determined because the issue was unable to be reproduced. this report is being filed as part of a retrospective review of historical records.
15518 | 1220648-2024-22142 | 20616320 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20616320&pc=OZD | 03/03/2022 | 11/05/2024 |  | 11/06/2024 |  |  | 03/03/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1323483 |  | 0042-0000-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) produced a "system self check failure" alarm during startup. the engineer checked the firmware and found that powermod 1 displayed "na" for revision and version. a visual inspection found damage at the power battery manager (pbm) consistent with leakage of the super caps. the engineer then replaced the pbm, which resolved the issue. this occurred during preventative maintenance, and there was no patient involvement. | [5 empty] | D8/9: device returned to field service, date unknown. the device was returned and evaluated, and the investigation for the reported system self-check error has been completed. data analysis: aic logs from the complaint date showed sau powermod 1 communication issues which resulted in failure of system self-check. device analysis: the power battery manager (pbm) was confirmed to have corrosion on the communication traces next to the super caps. the root cause of syscall error was due to a defective pbm. the root cause of the defective pbm was corrosion due to leaking super capacitors. this report is being filed as part of a retrospective review of historical records.
15519 | 1220648-2024-22145 | 20616336 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20616336&pc=OZD | 04/03/2022 | 11/05/2024 |  | 11/06/2024 |  |  | 04/03/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1095657 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. a yellow screen ¿optical sensor system error¿ was noted during pump setup. registered nurse switched consoles and was able to setup the same pump without issue. no patient harm was reported. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, no error/ failure was reproduced during long term simulation with test pump and purge. the optical bench was also tested without problem in parameters. given the testing and data analysis for the console, it was concluded that the optical sensor error was highly likely due to pump not fully connected to blue receptacle to make enough optical connection when the event occurred. before returning this console back to the customer, a functional check on the console performed. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15520 | 1220648-2024-22098 | 20616367 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20616367&pc=OZD | 06/15/2022 | 11/05/2024 |  | 11/06/2024 |  |  | 06/15/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | ABIOMED, INC. |  |  | US |  | 1446166 |  | 0042-0000 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant in the eu had a red controller aic (automated impella controller) alarm. the team noted no harm or injury. | [5 empty] | Console logs from the reported day of event are consistent with complaint and show controller failure alarm due to epm software corruption (#1029) as well as sporadic epm communication errors. after console was brought to aachen fs, the issue was reproduced and controller failure alarms (#1029) were persistently presented after aic boot-up. visual inspection of the impellatronic assembly (where the epm sub-circuitry is located) and its communication cable did not reveal any irregularities (no kinks or misalignment of the cable). after the impellatronic assembly was replaced, the issue was resolved, and console operated within specification. repair: replace impellatronic assembly (0042-1115), and perform functional check.
15521 | 1220648-2024-22165 | 20616392 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20616392&pc=OZD | 03/14/2022 | 11/05/2024 |  | 11/06/2024 |  |  | 03/14/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1614404 |  | 0042-4201 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) exhibited a "system self check failed, change console immediately" message and they were unable to get to the home screen. the resolution for this issue is unknown. no report of patient involvement, harm, or injury. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, the issue was reproduced, and console was unable to boot-up correctly. the self-check process was temporarily disabled and manually running the script (checkversions), which allowed us to identify the cause as file corruption. re-installing aic sw fixed corrupt files, and although the cause of file/disk corruption could not be identified, it is most likely that it was due to cf card defect. therefore, root cause of the system self-check error was due to a defective component. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15522 | 1220648-2024-22127 | 20616394 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20616394&pc=OZD | 02/05/2022 | 11/05/2024 |  | 11/06/2024 |  |  | 02/05/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1465870 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported an automated impella controller (aic) battery failure during operation. there was no reported patient harm. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the root cause of the battery alarms was due to gradual depletion of the batteries which were not routinely charged. this report is being filed as part of a retrospective review of historical records.
15523 | 1220648-2024-22157 | 20616472 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20616472&pc=OZD | 03/18/2022 | 11/05/2024 |  | 11/06/2024 |  |  | 03/18/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1082388 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) that the pole on the console stand has broken off. a backup controller was utilized for support. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported cart issue has been completed. the product was not received from the customer and therefore, an evaluation of the device was not possible. the root cause of the issue was not determined since the product was not returned. this report is being filed as part of a retrospective review of historical records.
15524 | 1220648-2024-22154 | 20616524 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20616524&pc=OZD | 04/10/2021 | 11/05/2024 |  | 11/06/2024 |  |  | 04/10/2021 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1098594 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) produced an error message asking to switch out consoles. the team tried to reboot several times but received the same error message. the aic was then replaced with a backup console and no patient harm was reported. | [5 empty] | The investigation for the reported controller error (yellow alarm) has been completed. while supporting the patient, the console displayed a yellow grade controller alarm. the staff switched to a backup console successfully. they saw the yellow controller alarm after performing a power cycle as well. the logs confirmed controller alarm (optical bench ambient pressure out of range) was present intermittently. towards the end, alarm (opm communication stopped) was seen. there are sau_motor_1 errors present after ten minutes of when the pump was started. upon first boot in field service, there was no light seen on the optical fiber at pump connector. the 5s values were out of spec. replacing the optical bench fixed the issue. console was run on a pump simulator for two hours and sau_motor_1 errors were not reproduced. no sau_motor_1 errors were seen. the cause of the controller error alarm was a malfunctioning optical bench. replacing the optical bench fixed the controller error. this report is being filed as part of a retrospective review of historical records.
15525 | 1220648-2024-22120 | 20616539 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20616539&pc=OZD | 02/10/2022 | 11/05/2024 |  | 11/06/2024 |  |  | 02/10/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1361559 |  | 0042-0010-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Female |  |  | White | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. while on support, the aic showed a blue screen and restarted automatically. the console was replaced, and there were no reports of patient harm or injury. | [5 empty] | The investigation for the reported restart issue has been completed. the product was returned, and the issue was confirmed. data analysis: complaint date is consistent with complaint intake, after a few days of support, the aic shutdown unexpectedly. imc logs show no adverse or irregular log lines prior to console crash. after the console reboots successfully, the "unexpected controller shutdown" alarm triggers as intended. however, there are no core dumps associated with the crash. console had rebooted and resumed pump, function however staff stopped pump to swap the console out. device analysis: the console was returned to aachen. upon arrival, a physical inspection of the console was performed with no issues found. multiple electrical power connection point were tested and verified from the 4-in-1 and pbm, no issues were found. a 24 hour test was performed, running the console on pump and purge. no issues occurred and the failure mode was not reproduced. per the results of the clinical review and investigation, the root cause could not be determined due to the inability to reproduce. this report is being filed as part of a retrospective review of historical records.
15526 | 1220648-2024-22163 | 20616570 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20616570&pc=OZD | 07/28/2021 | 11/05/2024 |  | 11/06/2024 |  |  | 08/04/2021 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER EU |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER EU |  |  | ABIOMED, INC. |  |  | US |  | 1360581 |  | 0042-0010-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) did not recognize the purge disc during preparation. this aic was replaced with another aic and the preparation was successfully completed. there was no reported patient involvement. | [5 empty] | The investigation into purge disc/flag issues has been completed. the impella automated controller was not received from the customer and therefore, an evaluation of the device was not possible. the data analysis of the logs show that console was not used on the reported day of event, however intermittent purge disk detection recorded on 2021-07-28 is consistent with the complaint. several attempts were made to go through a case with start wizard, but each time the purge disk was not being detected the wizard would not allow it to go to the next step. the field service engineer observed that purge disk detect flag was broken, which explained the malfunction of the detection mechanism. after the broken purge disk detect flag was replaced, the console operated within specification. the cause of the console's inability to properly detect purge pressure disk was a broken purge disk detect flag. this report is being filed as part of a retrospective review of historical records.
15527 | 1220648-2024-22164 | 20616652 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20616652&pc=OZD | 03/15/2022 | 11/05/2024 |  | 11/06/2024 |  |  | 03/15/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1228178 |  | 0042-0000-EU |  | Yes |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during a procedure the automated impella controller (aic) the purge cassette would not work properly, this occurred prior to pump insertion. the customer replaced the device with another aic they had onsite. no report of patient harm or injury. | [5 empty] | D9: the date of the device return for evaluation. e1, e2, e3: the initial reporters name and occupation is unknown. the console (aic) was returned for evaluation. upon inspection of the console, it was discovered that purge motor shaft was immobilized by dried-up glucose (dextrose). after the motor shaft gasket and purge insert were cleaned, the issue was resolved, and the purge system operated within specification. therefore, the root cause of the purge cassette was due to contamination. the purge motor area was cleaned, and a functional check was performed. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15528 | 1220648-2024-22167 | 20616669 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20616669&pc=OZD | 08/03/2021 | 11/05/2024 |  | 11/06/2024 |  |  | 08/03/2021 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER JP |  |  | ABIOMED, INC. |  |  | US |  | 1471970 |  | 0048-0024-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) generated a control device error alarm. it was reported that the alarm disappeared when plugged in and unplugged was performed. during the alarm appearance, the waveform display of the aortic position waveform and left ventricle position waveform disappeared. there was no problem with the operating condition of the pump, the alarm appeared 4 times in approximately 24 hours. the clinician recommended the replacement of the control device, but at the discretion of the doctor, it was decided to keep watching without replacing the console, a backup was prepared in case it was needed, which it was note. the patient remained on impella support and was successfully weaned. | [5 empty] | The investigation into the controller error/placement issue has been completed. the impella automated controller was not received from the customer and therefore, an evaluation of the device was not possible. the data analysis of the logs confirmed the reported failure. on the day of the complaint, the console issued a controller error (opm communication crc error- event set #1045). real time logs revealed a transient loss of valid optical data. the console (ic7982) was tested thoroughly by the field service team, and the failure was not reproduced. this known pattern failure, which is characterized by a temporary loss of optical data and triggering of a controller error alarm, has a low probability of recurrence. the cause of the transient loss of optical data was unable to be determined because the issue was unable to be reproduced. this report is being filed as part of a retrospective review of historical records.
15529 | 1220648-2024-22169 | 20616707 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20616707&pc=OZD | 08/01/2021 | 11/05/2024 |  | 11/06/2024 |  |  | 08/02/2021 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1255438 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) had a controller failure and was unable to recognize the impella catheter. it was reported the system was running software version 8.4.1. there was no patient harm reported. | [5 empty] | The investigation into the pump not detected/controller failure issue has been completed. the impella automated controller (aic) was returned for investigation. the data analysis of the logs revealed that while in the field, the console failed to detect a pump during case start. the logs show that the console was able to detect the pump connection but failed to read pump eeprom. this is a pattern failure mode only found in aic_v8.4 and aic_v8.4.1 and has been documented as a software (sw) defect. the returned console booted-up normally, but the optical bench state machine (ossiopsens) gets stuck in post_start_state and did not reach post_ok_state. this prevented the sw from reading pump eeprom, a mitigation for the issue has been added to aic_v8.5. there was also an unexpected controller shutdown right after this reported issue. the clinical staff did not clarify whether this unreported issue was due to toggling the battery transport switch/ unplugging ac. the console did not immediately reboot according to the logs, but the boot count on the epc went from 97 to 99 which may indicate an unsuccessful reboot attempt. the cause of the failure to recognize the pump was due to a software defect (in-house). the console was tested thoroughly in an attempt to reproduce any behavior linked to the unexpected controller shutdown. while power cycling, system self check had failed (on 20th boot cycle) due to failure checkversions, but an unexpected controller shutdown was not reproduced. to resolve the intermittent failure of system self check, the 4-in-1 and cfcs were swapped for known-good parts, and the console was power cycled 200 times with no related issues. the root cause of the unexpected controller shutdown was most likely due to a defective 4-in-1. the 4-in-1 was replaced because of the failure of system self check during testing. the root cause of the unexpected controller shutdown was most likely due to a defective 4-in-1. this report is being filed as part of a retrospective review of historical records.
15530 | 1220648-2024-22170 | 20616709 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20616709&pc=OZD | 02/20/2022 | 11/05/2024 |  | 11/06/2024 |  |  | 02/20/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2022099212 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported a purge system stopped alarm resulting in the need to switch to a backup controller. there was no adverse impact to the patient. | [5 empty] | The investigation into the purge disc/flag issues issue has been completed. the automated impella controller (aic) was returned for investigation. imc logs show that several controller failure (purge pressure sensor defective) alarms were triggering, as well as purge system stopped alarms. these were preceded by a purge bag change attempt. several log entries occurred for purge rfid errors, as well as the inability for the purge cassette to reset to "home" position. lt logs from the complaint date show valid purge pressure and purge flow tick rate, however a spike in purge pressure occurs, rising from ~500 mmhg to ~1100 mmhg. purge assembly was inspected, no issues were found. console was tested on pump and purge with no issues. out of an abundance of safety, fs was instructed to replace the purge assembly. this report is being filed as part of a retrospective review of historical records.
15531 | 1220648-2024-22171 | 20616759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20616759&pc=OZD | 02/15/2022 | 11/05/2024 |  | 11/06/2024 |  |  | 02/15/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 210122 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported an automated impella controller (aic) battery failure persisting, despite having the device plugged in to charge for over two hours. there was no patient involved or impacted. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the aic issue was a defective battery. this report is being filed as part of a retrospective review of historical records.
15532 | 1220648-2024-22135 | 20616772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20616772&pc=OZD | 07/09/2022 | 11/05/2024 |  | 11/06/2024 |  |  | 07/09/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1228182 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 48 YR |  | Female |  |  |  | Malfunction | No | Yes | Defective Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support, the automated impella controller (aic) experienced a short alarm "unexpected console shutdown", which was resolved by itself. no patient harm reported. | [5 empty] | The investigation has been completed. the console (ic4232) was sent back for further investigation. the root cause of the crash of epc processes could not be determined due to the inability to reproduce. this report is being filed as part of a retrospective review of historical records.
15533 | 1220648-2024-22175 | 20616778 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20616778&pc=OZD | 01/21/2022 | 11/05/2024 |  | 11/06/2024 |  |  | 01/21/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1542729 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced intermittent controller error alarm during a patient case. no clinical details regarding resolution or patient condition were provided. no patient harm reported. | [5 empty] | The device was returned for evaluation. this known pattern failure, which is characterized by a temporary loss of optical data (placement, snr, contrast, lamp, gain) and triggering of a controller error alarm, has a low probability of reoccurrence. if needed, patient hemodynamics and impella position can be monitored with imaging. unable to determine the root cause. this report is being filed as part of a retrospective review of historical records.
15534 | 1220648-2024-22093 | 20616781 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20616781&pc=OZD | 03/03/2022 | 11/05/2024 |  | 11/06/2024 |  |  | 03/03/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1323491 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant notified the clinical manager that the automated impella controller (aic) generated a battery failure alarm. it was reported the alarm was present when the aic was plugged into the ac outlet, the outlet was also changed and did not resolve the alarm. there was no patient harm reported. | [5 empty] | The investigation of the reported battery failure has been completed. the impella automated controller was returned for investigation. the data log review from the reported day of event are consistent with complaint and show multiple battery-related alarms: battery failure (#360, #350), battery 1 communication failure (#352), battery level low (#23). although these alarms started appearing on (b)(6) 2022. the long-term log power data shows that a mismatch between cell voltages of battery 1 started around (b)(6) 2022, but did not result in any alarms or warnings that would be visible to operators. this is indicative of cell imbalance (pattern failure) that ultimately led to battery 1 failure. during troubleshooting the returned product, battery 1 was found to be internally shut down, with no output voltage, and unable to communicate with either the power battery manager circuit board (pbm) or diagnostic equipment (ata interface cable). after the console batteries were replaced, the issue was resolved, and no further battery alarms were present. the cause of the aic issue was a defective battery. this report is being filed as part of a retrospective review of historical records.
15535 | 1220648-2024-22172 | 20616832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20616832&pc=OZD | 02/09/2022 | 11/05/2024 |  | 11/06/2024 |  |  | 02/09/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1455140 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a battery failure, the customer was sent a loaner device. no reported patient involvement. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the root cause of the incorrect battery level was not determined but could be due to a glitch in pbm communication. this report is being filed as part of a retrospective review of historical records.
15536 | 1220648-2024-22126 | 20616834 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20616834&pc=OZD | 07/19/2022 | 11/05/2024 |  | 11/06/2024 |  |  | 07/19/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1558681 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced battery communication failure alarm and a battery failure alarm. no clinical details regarding resolution. no patient was involved. | [5 empty] | The investigation has been completed. the console ((b)(6)) was sent back for further investigation. the cause of the aic issue was a defective battery. this report is being filed as part of a retrospective review of historical records.
15537 | 1220648-2024-22174 | 20616863 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20616863&pc=OZD | 02/11/2022 | 11/05/2024 |  | 11/06/2024 |  |  | 02/11/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1442220 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic), that upon start up with the console that there was a blank screen for a long period of time. a loaner was sent to initiate the return of the device for investigation/repair and preventive maintenance service. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported power on/blank screen issue has been completed. the product was returned, and the issue was confirmed. data analysis: imc logs are not related to this failure mode, no indication of the blank screen. device analysis: console was tested on an engineering bench and was power cycled 100 times, however, the blank screen symptom was not reproduced. per the results of the clinical review and investigation, the root cause was not determined since the failure mode was not reproduced. this report is being filed as part of a retrospective review of historical records.
15538 | 1220648-2024-22151 | 20616864 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20616864&pc=OZD | 03/25/2022 | 11/05/2024 |  | 11/06/2024 |  |  | 03/25/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1095654 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) was booted up to the start screen, however the soft keys would not work. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported kbd/rotary knob issues issue has been completed. the product was returned, and the issue was confirmed. data analysis: imc logs on event occurred time were consistent with the complaint. at the beginning of one power cycle, imc indicated that kbd functioned as normal. then, repetitive entries of kbd error occurred, indicating kbd was not responsive momentarily (due to esd guard). however, imc logs show that kbd was back to working condition later without kbd error. device analysis: the console was investigated and tested on an engineering bench. however, the kbd not responsible issue was not reproduced. the console was opened, visual inspection on kbd revealed a contamination on capacitor c4, but it was unlikely the root cause of kbd malfunction. per the results of the clinical review and investigation, the root cause of the issue was not determined since the failure mode was not reproduced. this report is being filed as part of a retrospective review of historical records.
15539 | 1220648-2024-22133 | 20616889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20616889&pc=OZD | 05/16/2022 | 11/05/2024 |  | 11/06/2024 |  |  | 05/16/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1086075 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) experienced a controller error yellow alarm that was resolved by unspecified troubleshooting. no patient harm reported. | [5 empty] | The investigation has been completed. there was no issue found during the case. the device functioned/operated to specification. the controller error was triggered according to sw specifications. this report is being filed as part of a retrospective review of historical records.
15540 | 1220648-2024-22176 | 20616892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20616892&pc=OZD | 02/03/2022 | 11/05/2024 |  | 11/06/2024 |  |  | 02/03/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1095116 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with high-risk percutaneous coronary intervention was implanted with an impella 5.5 for mechanical circulatory support. while on support, the automated impella controller (aic) experienced a controller error yellow alarm that self-resolved, reoccurred, and was subsequently resolved by switching to another console. no patient harm reported. | [5 empty] | The device was returned for evaluation. complaint cause is a known pattern failure characterized by the temporary loss of placement signal. no repair will be necessary as the issue has a low probability of reoccurrence, lasts only up to 10 minutes. if needed, patient hemodynamics and impella position can be monitored with imaging. the root cause of this issue is unable to be determined. this report is being filed as part of a retrospective review of historical records.
15541 | 1220648-2024-22158 | 20616926 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20616926&pc=OZD | 09/02/2021 | 11/05/2024 |  | 11/06/2024 |  |  | 09/09/2021 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1542729 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  |  |  | Malfunction | No | Yes | Material Fragmentation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The biomedical engineer reported that the automated impella controller (aic) screen was becoming detached on one side and requested a loaner while having repairs completed. there was no patient harm reported. | [5 empty] | E1: name of the initial reporter is unknown the investigation into the damage has been completed. the impella automated controller was returned for investigation. the returned console was inspected and the reported detachment of one corner of the glass keyboard assembly was measured to be minuscule (<0.0295") and therefore negligible and won't affect console operation. there was no console damage - more like "cosmetic imperfections" that do not affect console operation. no corrective action required. the cause was not determined. this report is being filed as part of a retrospective review of historical records.
15542 | 1220648-2024-22161 | 20616940 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20616940&pc=OZD | 09/01/2021 | 11/05/2024 |  | 11/06/2024 |  |  | 09/01/2021 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER, JP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | 0042-0000-JP |  | Yes |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | Insufficient Information | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a control device battery failure alarm and a battery status unknown alarm. there was no report of patient harm or injury. | [5 empty] | The investigation into the battery power issue has been completed. the automated impella controller (aic) was returned for investigation. the battery failure alarm was due to a defective battery 1 (decline and flatline of individual cell voltage). the root cause of the battery cell failure was undetermined. this report is being filed as part of a retrospective review of historical records.
15543 | 1220648-2024-22162 | 20616952 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20616952&pc=OZD | 03/14/2022 | 11/05/2024 |  | 11/06/2024 |  |  | 03/14/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1218087 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported that an automated impella controller (aic) had a controller error. a back up aic was used. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported controller failure (red alarm) issue has been completed. the product was returned, and the issue was confirmed. data analysis: complaint date is consistent with complaint intake. imc logs from the complaint date show the following: controller failure (purge pressure sensor defective). device analysis: the controller failure was not reproduced on boot up after receiving console from field service. the console was run on pump and purge for extended period of time with no issues. per the results of the clinical review and investigation, the root cause could not be determined due to the inability to reproduce. this report is being filed as part of a retrospective review of historical records.
15544 | 1220648-2024-22177 | 20616958 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20616958&pc=OZD | 01/31/2022 | 11/05/2024 |  | 11/06/2024 |  |  | 01/31/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1088167 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. during support the fiber optic signal was lost then console gave out " console error". unknown how this issue was resolved, no report of patient harm or injury. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, the reported issue was not reproduced during troubleshooting. the console behavior is consistent with a known pattern failure "transient loss of optical data", and abiomed is aware of the issue and working on appropriate corrective action. there is no need for repairs, as the issue self-corrects. after functional check, the console will be returned to the customer. therefore, the root cause of the optical signal issues was unable to be determined because the issue was unable to be reproduced. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15545 | 1220648-2024-22178 | 20616979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20616979&pc=OZD | 01/31/2022 | 11/05/2024 |  | 11/06/2024 |  |  | 01/31/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1094061 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support, the automated impella controller (aic) experienced failure alarm. the console was replaced to address the issue. no report of patient harm attributed to this issue. | [5 empty] | The device was returned for evaluation. complaint cause is a known pattern failure characterized by temporary loss of placement signal. no repair will be necessary as the issue has a low probability of reoccurrence, lasts only up to 10 minutes. if needed, patient hemodynamics and impella position can be monitored with imaging. the root cause of the transient loss of optical data could not be determined due to the inability to reproduce. this report is being filed as part of a retrospective review of historical records.
15546 | 1220648-2024-22182 | 20617021 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20617021&pc=OZD | 01/19/2022 | 11/05/2024 |  | 11/06/2024 |  |  | 01/19/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1218082 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during a demonstration with no patient involvement, the automated impella controller (aic) exhibited battery failure. | [5 empty] | The automated impella controller (aic) was returned for evaluation. upon evaluation of the console, the reported failure mode was duplicated on an engineering bench. the battery failure was due to a defective battery 1. the root cause of the defective battery 1 was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
15547 | 1220648-2024-22185 | 20617025 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20617025&pc=OZD | 01/18/2022 | 11/05/2024 |  | 11/06/2024 |  |  | 01/18/2022 | 11/06/2024 |  |  | 09/27/2013 |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1142855 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) was damaged during shipping. the side was cracked. no patient involvement. | [5 empty] | The console (aic) was returned for evaluation. upon visual inspection of the console, cracks were identified around aic io panel. no internal damage was found. the cause of the damage to the console rear housing was likely due to mishandling of the console. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15548 | 1220648-2024-22166 | 20617031 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20617031&pc=OZD | 03/11/2022 | 11/05/2024 |  | 11/06/2024 |  |  | 03/11/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1185327 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) exhibited a system self-check error, no additional information is available. no report of patient involvement, injury, or harm. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, the failure mode was reproduced on an engineering bench. the console was opened for inspection, revealing corrosion around c261, near the voltage supply to the pmd board. therefore, the root cause of the system self-check error was due to a defective pbm board. before returning this console back to the customer, the pbm pca was replaced, and the sw was upgraded while a preventive maintenance and functional check was performed. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15549 | 1220648-2024-22189 | 20617041 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20617041&pc=OZD | 01/10/2022 | 11/05/2024 |  | 11/06/2024 |  |  | 01/10/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1317757 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had to replace an automated impella controller (aic) (per ifu recommendations) as there were battery alarms, battery communication failure (yellow alarm). no patient was involved. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the battery failure was due to a defective battery. the root cause of the defective battery was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
15550 | 1220648-2024-22180 | 20617066 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20617066&pc=OZD | 01/26/2022 | 11/05/2024 |  | 11/06/2024 |  |  | 01/27/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1502180 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was on an automated impella controller (aic) for mechanical circulatory support. during setting up/doing case start for an impella cp, the aic malfunctioned and didn't recognize the purge fluid. a backup aic was used for support. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported purge issue has been completed. the product was returned, and the issue was confirmed. data analysis: aic logs confirmed the reported issue since the error was triggered during the case start procedure. however, a purge system failure alarm (event set #417 ¿ piston blocked) was also issued because the available range of the stepper motor (piston) was 0%. device analysis: the console was inspected for any signs of glucose build that would explain the stepper motor issues driving the piston to build purge pressure. glucose build-up was found on/ around the purge assembly. see the attachments section for image. per the results of the clinical review and investigation, the root cause was build-up of glucose on the stepper motor gasket. this report is being filed as part of a retrospective review of historical records.
15551 | 1220648-2024-22184 | 20617080 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20617080&pc=OZD | 01/25/2022 | 11/05/2024 |  | 11/06/2024 |  |  | 01/25/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1199223 |  | 0042-0000-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the automated impella controller (aic) exhibited the error message "system failure optical sensor occurred". the device was replaced with another aic that the customer had onsite. no report of patient harm or injury. | [5 empty] | D9: the date of the device return for evaluation is unknown. the console (aic) was returned for evaluation. upon inspection of the console, there were no error/failure was reproduced during long term simulation with test pump and purge. the optical bench was also tested without problem in parameters. given the testing and data analysis for the console, it was concluded that the system failure optical sensor error was likely due to a defective pump. however, the pump was discarded on the field, the root cause could not be confirmed. before returning this console back to the customer, the optical bench was replaced and functional check on the console was performed. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15552 | 1220648-2024-22159 | 20617099 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20617099&pc=OZD | 09/01/2021 | 11/05/2024 |  | 11/06/2024 |  |  | 09/01/2021 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2022079876 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Nurse | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The nurse reported that the automated impella controller (aic) soft menu buttons would not respond to touches for running the de-air process. there was no patient harm reported. | [5 empty] | E1: name of the initial reporter is unknown. the investigation into the keyboard rotary knob (kbd) issue has been completed. the impella automated controller (aic) was returned for investigation. the data analysis of the logs from the reported day of event are consistent with complaint and show multiple "high guard level" errors, indicating possible electrostatic discharge (esd) events, which are known to be causing keyboards to be temporarily unresponsive. after 4 minutes of the kbd being unresponsive (during which all other systems and components were operating within specification) the console was shut down. during the next 90 days, when console was being used in the field, and before it was shipped to the manufacturer for troubleshooting, there were no esd events shown in the logs and no evidence (or reports) of non-responsive kbd. the issue was not reproduced, and the console was found to be within specification and passing all testing. it is most likely that some undetermined esd event occurred, however the aic did not receive any permanent damage. the cause of the kbd being unresponsive was unable to be determined because the issue could not be reproduced. this report is being filed as part of a retrospective review of historical records.
15553 | 1220648-2024-22183 | 20617115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20617115&pc=OZD | 01/24/2022 | 11/05/2024 |  | 11/06/2024 |  |  | 01/24/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1129486 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) experienced an error alarm with loss of placement signal. the console was replaced to address the issue. no patient harm reported. | [5 empty] | The device was returned for evaluation. complaint cause is a known pattern failure characterized by temporary loss of placement signal. no repair will be necessary as the issue has a low probability of reoccurrence, lasts only up to 10 minutes. if needed, patient hemodynamics and impella position can be monitored with imaging. the root cause is unable to be determined. this report is being filed as part of a retrospective review of historical records.
15554 | 1220648-2024-22190 | 20617150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20617150&pc=OZD | 01/11/2022 | 11/05/2024 |  | 11/06/2024 |  |  | 01/11/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1316939 |  | 0042-0000-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had a battery failure, battery comm. failure (yellow alarm), observed at the setup for a case. there was no harm or injury noted from any delay in the swapping to a backup console. the instructions for use was followed with a backup console presented to the patient. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the battery failure was due to a defective battery (rapid decline in cell voltage). the root cause of the cell failure was undetermined. this report is being filed as part of a retrospective review of historical records.
15555 | 1220648-2024-22193 | 20617151 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20617151&pc=OZD | 12/07/2021 | 11/05/2024 |  | 11/06/2024 |  |  | 12/07/2021 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1313452 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported having a battery failure with an automated impella controller (aic). the battery is backup for a/c power and for transport. the instructions for use was followed and aic removed and backup console utilized for patient use. | [5 empty] | E1: name of initial reporter is unknown. e2/e3: unknown. the investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the battery failure alarm was due to a defective battery (rapid decline in cell voltage). the exact cause of the defective battery was utd. this report is being filed as part of a retrospective review of historical records.
15556 | 1220648-2024-22191 | 20617172 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20617172&pc=OZD | 12/17/2021 | 11/05/2024 |  | 11/06/2024 |  |  | 12/17/2021 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1438016 |  | 0042-0010-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) had a syscall error after a software update. there was no reported patient involvement. | [5 empty] | The investigation for the reported syscall error issue has been completed. the product was returned, and the issue was confirmed. (unknown return date). data analysis: aic logs were not available for review. device analysis: console troubleshooting was conducted in collaboration with (b)(4). fwcheck (telnet command) revealed that the epm sau¿s hw and sw versions were listed as ¿n.a.¿ indicating the epm¿s microcontroller either had incompatible fw or could not communicate. this was a result of a bad sw upgrade. however, after manually reprogramming the microcontroller, the sw upgrade was still not successful (failed epm). replacing the impellatronic resolved the sw upgrade issues. per the results of the clinical review and investigation, the root cause of the sw upgrade issues was due to a malfunctioning impellatronic pca (epm circuitry). this report is being filed as part of a retrospective review of historical records.
15557 | 1220648-2024-22188 | 20617193 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20617193&pc=OZD | 01/24/2022 | 11/05/2024 |  | 11/06/2024 |  |  | 01/24/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1146037 |  | 0042-0000-EU |  | Yes |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automated impella controller (aic) failed to charge. the aic is plugged in, but the battery is not charging, and no green power indicator is lit. the staff tried several outlets and check to make sure the outlet was functioning. the transport fuse is off and on again, the aic still had no power and battery wasn't charging. it was reported that the device was not in use, no patient involvement. | [5 empty] | D9: the date of the devices return for evaluation is unknown. e1, e2, e3: the initial reporters name and occupation is unknown. the investigation for the reported issue has been completed. the device was returned for evaluation, upon visual inspection, it was observed the 2 fuses in line with the ac input had blown. the root cause is a defective component. this report is being filed as part of a retrospective review of historical records.
15558 | 1220648-2024-22179 | 20617194 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20617194&pc=OZD | 01/27/2022 | 11/05/2024 |  | 11/06/2024 |  |  | 01/27/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1314157 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) exhibited a controller error. no resolution was specified, no patient involvement reported. | [5 empty] | The automated impella controller (aic) was returned for evaluation. upon evaluation of the console, it was confirmed persistent battery failure (transport connection blown/missing) was not due to a disengaged battery switch. when the console was booted, the alarms were consistent with what was seen in the field. therefore, when visually inspecting the batteries and pbm, it was observed that one of the batteries status leds were completely blank. the cause of the failure mode was a defective battery. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15559 | 1220648-2024-22194 | 20617272 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20617272&pc=OZD | 12/03/2021 | 11/05/2024 |  | 11/06/2024 |  |  | 12/03/2021 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1075891 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had a biomed observed a battery communication failure error message. the instructions for use was followed, and the backup controller was used. no patient use and no harm or injury. | [5 empty] | E1: name of initial reporter is unknown. e2/e3: unknown. the investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the battery communication failure was due to a defective battery. the exact cause of the defective battery was most likely cell imbalance. this report is being filed as part of a retrospective review of historical records.
15560 | 1220648-2024-22187 | 20617293 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20617293&pc=OZD | 01/21/2022 | 11/05/2024 |  | 11/06/2024 |  |  | 01/21/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2022139888 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 91 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. during support, the automated impella controller (aic) exhibited red alarm back up battery failure. the device was replaced with another console without any further reported issues. no report of patient harm or injury. | [5 empty] | The console (aic) was returned for evaluation. upon evaluation of the console, the reported failure mode was duplicated on an engineering bench. the root cause for the battery failure was due to a defective battery 2. the root cause of the defective battery 2 was due to single cell failure. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15561 | 1220648-2024-22186 | 20617321 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20617321&pc=OZD | 01/23/2022 | 11/05/2024 |  | 11/06/2024 |  |  | 01/23/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1370139 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the automated impella controller (aic) exhibited an alarm controller error alarm followed by pump stop. the aic was replaced with another aic device and there were no further reported issues. no report of injury or harm to the patient. | [5 empty] | The automated impella controller (aic) was returned for evaluation. console logs from the reported day of event are partially consistent with complaint and show controller error alarm due to invalid crc in optical bench communication, however, there's no evidence that pump stopped on its own. the next two alarms, "impella stopped due to currents mismatch" and "impella disconnected / running pump disconnected" were the result of pump being purposefully unplugged from the console, after the first controller error alarm. the physical device was evaluated, and the reported problem could not be reproduced, and root cause is unable to be determined. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15562 | 1220648-2024-22192 | 20617370 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20617370&pc=OZD | 12/16/2021 | 11/05/2024 |  | 11/06/2024 |  |  | 12/16/2021 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1152333 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant had an automated impella controller (aic) battery failure. the team did all troubleshooting, but the battery failure persisted. the aic can be replaced per the instructions for use. there was no harm or injury. | [5 empty] | E1: name of initial reporter is unknown. e2/e3: unknown. the investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the aic battery failure alarm was due to a defective battery (rapid decline in cell voltage). the exact cause of the defective battery was utd. this report is being filed as part of a retrospective review of historical records.
15563 | 1220648-2024-22195 | 20618977 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20618977&pc=OZD | 05/06/2021 | 11/06/2024 |  | 11/06/2024 |  |  | 05/06/2021 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1109479 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) would not read the purge disc. it was reported that the purge disk was inserted and reinserted multiple times. however, the aic continued to alarm no purge disk detected. the aic was replaced with another aic. there was no patient harm reported. | [5 empty] | The investigation into purge disc/flag issues has been completed. the automated impella controller was returned for investigation. the data analysis of the logs for this failure mode was not necessary. the console was observed and showed that the purge flag within the purge disc retainer housing had been broken. the complaint cause is likely due to natural wear and tear of the plastic componentry on top of unintentionally aggressive use when inserting purge discs. this report is being filed as part of a retrospective review of historical records.
15564 | 1220648-2024-22196 | 20619103 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20619103&pc=OZD | 04/30/2021 | 11/06/2024 |  | 11/06/2024 |  |  | 04/30/2021 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1134036 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an automated impella controller was being used and that while preparing for a primary percutaneous coronary intervention, the cradle/holder for the purge pressure sensor was missing/broken off. the staff switched to a backup console successfully. it was reported that the complainant did not know when it occurred. there was no patient harm reported. | [5 empty] | The investigation has been completed. the automated impella controller was returned for investigation. clinical details reported that while in the field, when preparing for a primary percutaneous coronary intervention it was observed that the purge disc holder was missing. the staff switched to a backup console successfully. the cause of the broken purge disc holder was unable to be determined. this report is being filed as part of a retrospective review of historical records.
15565 | 1220648-2024-22197 | 20619114 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20619114&pc=OZD | 05/03/2021 | 11/06/2024 |  | 11/06/2024 |  |  | 05/03/2021 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1292759 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) purge cassette slot was broken. the aic was replaced with another aic without any issues. there was no patient harm reported. | [5 empty] | The investigation into purge disc/flag issues has been completed. the automated impella controller was returned for investigation. the data analysis of the logs for this failure mode was not necessary. the console booted without issue dueing investigation. the console purge assembly was inspected and it was observed that blue purge clip retainer had snapped. the complaint cause likely was due to excessive force when inserting past purge discs, and the natural wear and tear of the plastic component. this resulted in the plastic component failing and ultimately breaking. this report is being filed as part of a retrospective review of historical records.
15566 | 1220648-2024-22198 | 20619452 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20619452&pc=OZD | 10/15/2024 | 12/16/2024 |  | 11/06/2024 |  |  | 10/15/2024 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025547752 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 being placed for support of patient undergoing high risk surgery. the patient had multiple comorbidities. the impella 5.5 failed to deliver and was removed. the team made plans to instead place an impella cp via the access. the patient was calcified. no patient harm has been reported. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, unsuccessful 5.5 placement due to inability of 5.5 to pass through subclavian due to calcification as per the clinical description provided. corrections to the initial mfr report: g1 mdr reporting contact email | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist: section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
15567 | 1220648-2024-22199 | 20619504 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20619504&pc=PYX | 10/14/2024 | 01/08/2025 |  | 11/06/2024 |  |  | 10/14/2024 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025534048 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella rp flex support and after dilating, the doctor noticed the 23 french sheath's hemostatic valve was bleeding quite a bit. after trying to assess what the issue was, the doctor decided to open a new sheath. the patient survived the explant. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smartassist instructions for use and clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 french peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was determined to be most likely lack of vessel recoil because the bleeding from the hemostatic valve of the introducer sheath. g.1 revised reporting contact email since the initial manufacturer device report 1220648-2024-22199 was submitted in accordance with updated procedures.
15568 | 1220648-2024-22200 | 20619654 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20619654&pc=OZD | 10/13/2024 | 12/04/2024 |  | 11/06/2024 |  |  | 10/13/2024 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025544200 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 58-year-old male was implanted with an impella cp device for mechanical circulatory support. during impella support, there was a pump stop that occurred. there was flow saturation prior to the pump stop. the impella was therefore explanted. | [5 empty] | The investigation for the saturated flow and the pump stop has been completed. no product was returned for analysis. the data logs confirm pump stop alarms and motor current spikes with speed deviations before pump stop. there was a rise in purge pressure after spike and a drop in purge flow, characteristic of a possible biomaterial ingestion. there was no motor current spike before flow saturation. the pump flow was trending upwards with flows higher than p6 range (2.5-2.9l/min). the root cause of the pump stop was most likely biomaterial. the root cause of the saturated flow was not determined as no product was returned for analysis. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
15569 | 1220648-2024-22202 | 20619876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20619876&pc=OZD | 10/09/2024 | 12/23/2024 |  | 11/06/2024 |  |  | 10/09/2024 | 11/06/2024 |  |  | 09/30/2025 |  | 11/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025459460 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Female |  | Non Hispanic | Asian | Malfunction | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Low Blood Pressure/ Hypotension | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient had both lower limbs and both axillary/subclavian arteries could not be used for impella cp insertion. an approach was attempted from the arm blood vessels, but there was no suitable blood vessels. another approach was made by cutting down the common femoral artery and a 16fr sheath was inserted. the cp was then advanced, but it became stuck around the brachiocephalic artery, making insertion difficult. during this process, blood pressure decreased, so the cp was removed and venoarterial extracorporeal membrane oxygenation was set up through the same sheath of the common femoral artery. next, the left subclavian artery was widely exposed, and the cp was successfully inserted. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, the impella was unable to pass through the vascular system as per the clinical description provided. corrections to the initial mfr report: g1 mdr reporting contact. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
15570 | 1220648-2024-22206 | 20619880 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20619880&pc=OZD | 10/11/2024 | 12/12/2024 |  | 11/06/2024 |  |  | 10/11/2024 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025545514 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 57-year-old male patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. during impella support, the patient developed a gastrointestinal bleed. heparin was removed from the purge line in response to the bleed. support remained ongoing with the implanted pump. | [5 empty] | The investigation for the vascular damage has been completed. pump was not returned for investigation. clinical information is very limited and important information regarding the damage is not available. therefore, the root cause of the vascular damage issue was not determined since product was not returned and insufficient clinical information as provided. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
15571 | 1220648-2024-22210 | 20619890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20619890&pc=OZD | 10/14/2024 | 09/30/2025 |  | 11/06/2024 |  |  | 10/14/2024 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025542535 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | Us complainant reported the patient was on impella cp support. while on support, placement signal low alarm occurred. additionally, a nurse called for assistance with the pump stop with many unsuccessful attempts to restart impella. signs of flow saturation were observed prior to the pump stop. the pump was explanted due to the pump stop. the patient survived the event. | [5 empty] | The investigation for the pump stop has been completed. data logs were reviewed and real time (rt) log from field shows motor current spike with simultaneous speed deviation. purge pressure steps up with a drop in purge flow at same time as motor current (mc) spike. pump flow also becomes completely saturated and is able to regain some pulsatility after approximately 30 minutes. next rt log from case shows high motor current pump stop with a mc spike, speed deviation, and a spike in ps right before alarm. pump was able to be restarted and two more mc spikes with speed and pump flow deviations were noted after restart. second lt log from case shows pump was connected to new console and unable to be started with immediate pump stop alarms occurring. clinical details noted that pump flow showed signs of saturation prior to pump stop occurring and that changing aic did not resolve issue. the pump was used beyond the impella cp 8 day design life. the root cause of the pump stop was most likely due to biomaterial. corrections to the initial mfr report: g1 mdr reporting contact email | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. no product was returned for evaluation. the failure mode of "placement signal issue" is for issues related to the dp sensor and not the optical sensor. the failure mode was changed from "placement signal issue" to "optical signal issue" as the impella cp only has an optical sensor and does not have an dp sensor. data logs were reviewed and optical parameters do not show any hard failures of the optical sensor, all parameters withing a normal range. the root cause of the optical signal issue cannot be definitively determined as limited clinical details were provided. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-22210. e4 erroneously entered as "no."
15572 | 1220648-2024-22207 | 20619974 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20619974&pc=OZD | 10/11/2024 | 12/10/2024 |  | 11/06/2024 |  |  | 10/12/2024 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025534686 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Seroma | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support. while on support a seroma had to be surgically evacuated from the axillary pocket. a few days later hematoma had to be surgically evacuated as well. the patient remains on support therefore explant date and patient outcome are unknown. | [5 empty] | The investigation into access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly on manufacturer device report 1220648-2024-22207. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
15573 | 1220648-2024-22203 | 20619997 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20619997&pc=OZD | 10/10/2024 | 11/26/2024 |  | 11/06/2024 |  |  | 10/10/2024 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025553174 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Material Split, Cut or Torn | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Cardiovascular Dysfunction/Insufficiency | [10 empty] | Death | [5 empty] | EXTRA CORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, the patient was bleeding from surgical site overnight. suction was noted. the patient was taken back to the operating room for repair. unable to get flow above p1 without suction. the patient was given volume. the sterile sleeve was ripped at the junction of the tuohy lock. provider continued to reposition and the sleeve was secured. additionally, the patient required bilateral calf fasciotomies. the patient continued to have decreased pulses. the patient eventually expired after the family decided to withdraw care. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the low pump flow, limb ischemia, access site bleed, and product damage has been completed. the pump was not returned for evaluation. without additional clinical details, the root cause of the access site bleeding - major cannot be determined due to lack of clinical information and product returned. the root cause of the product damage - sterile sleeve damage is most likely due to use as it became ripped when the physician was attempting to reposition the device. the data logs were not returned. based on the clinical description provided, the root cause of the low pump flow is most likely due to the patients condition. the root cause of the limb ischemia could not be determined without additional clinical details. corrections to the initial mfr report: d4 model number.
15574 | 1220648-2024-22209 | 20620000 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20620000&pc=OZD | 10/13/2024 | 12/30/2024 |  | 11/06/2024 |  |  | 10/13/2024 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025534090 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Female |  |  | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Cardiac Perforation | Pericardial Effusion | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the 81-year-old female patient was on impella cp support and during the implant, a pericardial effusion occurred. the cause of the effusion is unknown. additionally, when the impella was repositioned and retracted, it dislocated and slipped completely back into the aorta. the pump was explanted. | [5 empty] | The investigation for the positioning and perforation issues has been completed. no product or data logs were returned for evaluation. clinical details noted that a pericardial effusion occurred during pump implant. additionally, after patient received cpr, pump was too deep in ventricle and when repositioning, pump became dislocated and slipped back completely into aorta. the pump was unable to re-cross valve and was explanted. the root cause of the ventricle perforation cannot be definitively determined as limited clinical details were provided. the root cause of the positioning issues was most likely a use related issue as the pump came out of position when cpr was being performed. corrections to the initial mfr report: d4 updated udi number g1 mdr reporting contact email. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation, section: direct aortic insertion, section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
15575 | 1220648-2024-22205 | 20620012 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20620012&pc=OZD | 10/11/2024 | 12/19/2024 |  | 11/06/2024 |  |  | 10/11/2024 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025523731 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | Yes | Pumping Stopped | Infusion or Flow Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [5 empty] | EXTRA CORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Eu complainant reported the patient was on impella cp support when there was a sudden pump stop. before the pump stop, there was a sudden increase in the motor current from about 850ma to over 1400ma. the pump was explanted. the patient expired after 65.54 hours of impella support. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation into pump stop has been completed. the root cause of the pump stop was not determined since product was not returned and insufficient clinical detail was provided. the impella cp was used beyond the 8 day design life per design trace matrix the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-22205. e4 should have been left blank. g1 reporting contact email revised in accordance with updated procedures.
15576 | 1220648-2024-22211 | 20620023 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20620023&pc=OZD | 10/14/2024 | 12/11/2024 |  | 11/06/2024 |  |  | 10/14/2024 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025518542 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the 72-year-old male patient was attempted to be implanted with an impella 5.5 but the pump was unable to be advanced through the axillary. the implant was aborted. there was no patient harm reported. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was returned for evaluation and upon inspection, there were abnormalities found on the pump. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since md unable to advance impella 5.5 via axillary artery due to notable calcium at distal end of axillary artery. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.
15577 | 1220648-2024-22208 | 20620047 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20620047&pc=OZD | 10/09/2024 | 11/14/2025 |  | 11/06/2024 |  |  | 10/12/2024 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1849988 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5. | [18 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | It was reported that the patient on impella 5.5 support knocked over the automated impella controller (aic) resulting in damage to the impella catheter plug. the white connector cable plug pins were broken in the aic. the aic was removed and returned for evaluation. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation has been completed. data analysis: no evidence was found in the logs to confirm the reported failure mode. device analysis: this console was tested after arriving in field service, and it was visually confirmed that the broken part of the white catheter plug with the pins remains inside the front panel optical connector. this confirms the reported failure mode. also found the rotating knob was missing as received. the most likely knob could also be knocked off of the aic. root cause: the root cause of the damaged front panel optical connector was most likely the use issue. there was no evidence in the logs indicating a "pump stop due to aic failure." it is likely that damage to the front-panel optical connector was misreported as "pump stop due to aic failure.¿ d2 common device name updated. d10 concomitant medical products and therapy dates updated. e4 should have been left blank on the initial report.
15578 | 1220648-2024-22213 | 20620049 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20620049&pc=OZD | 06/25/2022 | 11/06/2024 |  | 11/06/2024 |  |  | 06/27/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1137157 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) produced a "system self-check failed" alarm during maintenance. restarting the device did not resolve the issue. there was no reported patient involvement. | [5 empty] | The device was returned and evaluated, and the investigation for the reported system self-check error has been completed. imc logs from the complaint date showed sau powermod 1 communication issues which resulted in failure of system self-check. the power battery manager (pbm) was confirmed to have corrosion on the communication traces next to the super caps. the root cause of syscall error was due to a defective pbm. the root cause of the defective pbm was corrosion due to leaking super capacitors. this report is being filed as part of a retrospective review of historical records.
15579 | 1220648-2024-22201 | 20620074 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20620074&pc=OZD | 10/11/2024 | 12/05/2024 |  | 11/06/2024 |  |  | 10/11/2024 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563295 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  |  | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 65-year-old female noted as high risk for coronary intervention was implanted with an impella cp device for mechanical circulatory support. during impella support, there were positioning issues that led to hemolysis. hemolysis was evident with red urine. the device was repositioned however, it moved after the patient was coughing. the doctor did not think that they could cross the pump along the valve. the patient needed longer and more support, so the pump was explanted and an impella 5.5 was placed. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation into positioning issues and hemolysis has been completed. data logs confirmed pump was in improper position corresponded with clinical description that triggered alarms impella position unknown/aorta. no other issues were found on the logs. the product was not returned for review. the cause of the positioning issue was most likely wireless re-access. the impella device is not indicated for wireless re-access. the root cause of hemolysis was most likely due to pump was not in the proper position.
15580 | 1220648-2024-22204 | 20620079 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20620079&pc=OZD | 10/11/2024 | 12/12/2024 |  | 11/06/2024 |  |  | 10/11/2024 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025530473 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | Infusion or Flow Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 64-year-old male patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella 5.5 device for mechanical circulatory support. during impella support, the impella 5.5 had high purge pressure system blocked alarm. the motor current and flows increased. tissue plasminogen activator (tpa) was started but impella stopped then restarted multiple times, ultimately not being able to be restarted. the impella was removed. | [5 empty] | The investigation into saturated flow and pump stop has been completed. the impella device was not returned for evaluation. data logs were reviewed and the logs shows gradual increase in purge pressure over approximately 10 hours before ¿high purge pressure¿ alarms are triggered. motor current started to rise 2 hours before pump stop occurred. pump flow also becomes saturated at during motor current rise with purge pressure high. pump flows became fully saturated after a 30-minute rise in pump flow and a pump stop occurred 2 hours after saturated flows were noted. clinical details noted that after high purge pressure alarms occurred, motor current increased and pump flows increased to 5.2l/min at p-7. the root cause of the saturated flow was most likely due to biomaterial. the root cause of the pump stop was most likely due to biomaterial build-up leading to cascading events of high purge pressure then motor current rise before pump stop occurred. instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-22204. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist. section: using the automated impella controller with the impella catheter. ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿. section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿. section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high-risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿.
15581 | 1220648-2024-22212 | 20620178 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20620178&pc=OZD | 10/12/2024 | 09/22/2025 |  | 11/06/2024 |  |  | 10/12/2024 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025544042 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Defective Device | Device Sensing Problem | Material Split, Cut or Torn |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a 52-year-old male patient with cardiomyopathy and other systemic comorbidities was implanted with an impella 5.5 device for mechanical circulatory support. during impella support, the sheath for the impella 5.5 was exposed. there was a tear when repositioning. the sheath was covered with sterile dressing. additionally, the impella sleeve was compromised. non-alcohol based clean solution for sleeve and catheter was used when redressing. it was further reported that placement signal alarms occurred. the alarm was disabled and the staff will be monitoring the motor current and flows. patient was successfully weaned from the device after 14 days of support. | [5 empty] | The investigation into product damage (sterile sleeve damage) and introducer damage ¿ mechanical has been completed. clinical says that during repositioning of the pump, the sterile sleeve was torn and was wrapped with sterile cover later. therefore, the root cause of the product damage - sterile sleeve issue was most likely use related to excessive force. introducer damage failure mode was deleted since it was incorrectly added due to confusing clinical information. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-22212. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation for the optical signal issue has been completed. pump was not returned for investigation. data logs were investigated. the placement signal returns after 40 hours and 5s parameters return to normal. this signs are similar to that of an air gap. therefore, the root cause of the optical signal failure was an air gap from between the console and the patient cable plug or within the middle of the red handle ferrule. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2024-22212.
15582 | 1220648-2024-22214 | 20620261 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20620261&pc=OZD | 06/08/2022 | 11/06/2024 |  | 11/06/2024 |  |  | 06/08/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1153139 |  | 0042-0000-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the automated impella controller (aic) exhibited purge disc not reading. this occurred while changing the purge cassette. no reported harm to patient. | [5 empty] | D9: the date of the device return for evaluation is unknown. the console (aic) was returned for evaluation. after review of the console logs, the purge change wizard, alarm event #402, purge disc not detected was triggered after purge cassette was inserted. troubleshooting of pc removing, and re-inserting failed to resolve the alarm. another purge cassette was used, but same alarm was triggered for not detecting the purge disk. upon inspection of the console, the purge assembly area confirmed broken purge clip/ blue retainer. the purge disc retainer was replaced, and a functional check was performed; before the console was returned to the customer. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15583 | 1220648-2024-22218 | 20620287 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20620287&pc=OZD | 06/23/2022 | 11/06/2024 |  | 11/06/2024 |  |  | 06/23/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1191155 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 90 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) that the purge disc blue slot was completely removed. a new console was brought in and the case was completed. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported purge disc/flag issue has been completed. the product was returned, and the issue was confirmed. data analysis: imc logs are not relevant for this failure mode. device analysis: console purge assembly was inspected. upon opening the purge assembly door, it was confirmed the purge disc retainer had broken/snapped off of the purge assembly. per the results of the clinical review and investigation, the root cause was determined to be a defective component. this report is being filed as part of a retrospective review of historical records.
15584 | 1220648-2024-22215 | 20620302 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20620302&pc=OZD | 06/21/2022 | 11/06/2024 |  | 11/06/2024 |  |  | 06/21/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1598659 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The biomedical engineer reported that the automated impella controller (aic) connector cover was broken. a loaner was sent to the customer to initiate return of the device. there was no patient involvement. | [5 empty] | The investigation into the damage has been completed. the impella automated controller was returned for investigation. the returned console was inspected the optical pump connector dust covered was confirmed to be damaged. the damage was most likely due to use. the optical connector ferrule was cleaned and inspected but no damage was found. the cause of the damaged blue receptacle dust cover was most likely due to improper use. this report is being filed as part of a retrospective review of historical records.
15585 | 1220648-2024-22216 | 20620380 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20620380&pc=OZD | 06/21/2022 | 11/06/2024 |  | 11/06/2024 |  |  | 06/22/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1113656 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the burst pipe in the unit resulted in a flood which affected the console. no clinical details regarding resolution. no patient was involved. | [5 empty] | The investigation has been completed. the console (ic1876) was sent back for further investigation. the batteries were also likely exposed to water given their placement underneath the power battery manager, but there were no signs of visible damage. the root cause of the aic issue was contamination. this report is being filed as part of a retrospective review of historical records.
15586 | 1220648-2024-22219 | 20620393 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20620393&pc=OZD | 03/17/2023 | 11/06/2024 |  | 11/06/2024 |  |  | 03/20/2023 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1326727 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. while on support, the aic experienced a controller failure yellow alarm that was resolved by switching to another console. there was no reported patient harm. | [5 empty] | The investigation has been completed; the device was returned for evaluation. loss of optical data (placement, snr, contrast, lamp, gain) could not be reproduced and is unable to be determined. controller error alarm, with yellow alarm grade, the cause is most likely software related. this report is being filed as part of a retrospective review of historical records.
15587 | 1220648-2024-22217 | 20620441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20620441&pc=OZD | 10/13/2024 | 12/11/2024 |  | 11/06/2024 |  |  | 10/13/2024 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025524399 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Increase in Pressure | Pumping Stopped | [5 empty] | Ventricular Fibrillation | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and the purge flow was 3.9. there was a kink at the yellow leur. it was advised to secure the purge side arm to straighten out the tubing. support continued. furthermore, there was oozing at the access site. the physician was not concerned and no known intervention was performed. additionally, there was a pump stop. multiple attempts to restart the pump were made. there was a motor current spike with a controller failure at the time of the pump stop. the pump was explanted. the patient was on impella support for 146.54 hours before expiring. the patient experienced a ventricular fibrillation arrest before expiring. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation into high purge pressure and pump stop has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-22217.
15588 | 1220648-2024-22221 | 20620472 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20620472&pc=OZD | 10/13/2024 | 01/03/2025 |  | 11/06/2024 |  |  | 10/13/2024 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1095119 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | Low Cardiac Output | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported during impella cp support for 71-year-old male patient, there was a pump stop. there was a controller failure on the automated impella controller (aic) at the time of the pump stop and the aic was replaced in hopes that the pump would restart. the pump would not restart despite replacing the aic. the patient expired after 146.54 hours of impella support. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high-risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ | The investigation of automated impella controller (aic) - pump stop has been completed. the impella device was not returned for evaluation. the root cause for the pump stop was not related to any problems with the aic. no issues were found given the console exchange did not resolve issue. please refer to investigation for pump in manufacturer device report 1220648-2024-22217 for more details. g1 reporting contact email revised on manufacturer device report 1220648-2024-22221 in accordance with updated procedures.
15589 | 1220648-2024-22226 | 20620516 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20620516&pc=OZD | 10/14/2024 | 11/23/2024 |  | 11/06/2024 |  |  | 10/14/2024 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1242769 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had an automated impella controller being prepped for use and found to have a start alarm that prompted the team to follow recommendations / instructions for use and replace the console. no patient harm has been reported. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into automated impella controller (aic) - system self-check error has been completed. the logs from the console confirm that the system self check failure alarm was issued. the console booted up in system-self check failure. inspection under a digital microscope confirmed that the corrosion of copper traces in vicinity of super capacitors. the root cause of the system self check failure was contamination of pbm through supercapacitor electrolyte leakage leading to pbm corrosion.
15590 | 1220648-2024-22222 | 20620521 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20620521&pc=OZD | 10/12/2024 | 12/12/2024 |  | 11/06/2024 |  |  | 10/12/2024 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025503061 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 71-year-old male patient was on impella 5.5 support when there was left sided weakness on the patient, slurred speech, and pupillary changes. the patient was on heparin. magnetic resonance imaging was ordered. the pump was explanted. the patient survived the explant. | [5 empty] | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
15591 | 1220648-2024-22220 | 20620580 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20620580&pc=OZD | 06/09/2023 | 11/06/2024 |  | 11/06/2024 |  |  | 06/09/2023 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1668177 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the door on automated impella controller (aic) console would not open when pressing the side button. no reported harm to the patient. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the aic issue was contamination of purge door. this report is being filed as part of a retrospective review of historical records.
15592 | 1220648-2024-22223 | 20620741 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20620741&pc=OZD | 07/11/2023 | 11/06/2024 |  | 11/06/2024 |  |  | 07/11/2023 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1125041 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | Defective Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automated impella controller (aic) was dropped and parts of the aic were broken. there was a placement signal unreliable alarm that went off. the aic was replaced the issues were resolved. no known patient harm or injury. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the root cause for the failure mode 1 - placement signal not reliable was most likely due to a damaged blue receptacle as the user dropped the aic, which resulted in the loss of optical connection to the bench, though the pump was electrically connected. the root cause for the failure mode 2 - unexpected controller shutdown was most likely because the user dropped the aic, which caused the aic to lose power and resulted in an unexpected shutdown. this report is being filed as part of a retrospective review of historical records.
15593 | 1220648-2024-22228 | 20620796 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20620796&pc=OZD | 02/16/2023 | 11/06/2024 |  | 11/06/2024 |  |  | 02/16/2023 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1176696 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The field service engineer (fse) reported that the automated impella controller (aic) generated an alarm communicating that the controller needed to be replaced. the fse arrived at the site and checked the alarm history and did not find confirmation of an alarm in the history. a loaner was sent to the customer to initiate return of this device. there was no patient harm reported. | [5 empty] | E1: name of the initial reporter is unknown. the investigation into the alarm/optical signal issue has been completed. the impella automated controller (aic) was returned for investigation. the data analysis of the logs showed repeated 1036 errors resulting in unreliable placement signal controller error. the real time logs noted (unreliable placement signal) and showed repeated extended outputs of -3276.8 and placement signal of 3000. when the console was returned, an optical pump was unable to make good electrical connection. the connector was discovered to have a piece of an old pump lodged into the connector. this setup was able to reproduce the errors in the field. the cause of the aic issue was a defective blue receptacle. this report is being filed as part of a retrospective review of historical records.
15594 | 1220648-2024-22230 | 20620887 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20620887&pc=OZD | 10/12/2024 | 12/11/2024 |  | 11/06/2024 |  |  | 10/12/2024 | 11/06/2024 |  |  | 01/31/2025 |  | 11/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024327921 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, pallor and cold temperature was noted to the lower left extremity distal to the impella site. the staff increased inotropes and impella was weaned for removal. it was also reported that the patient was declared brain dead but hemodynamically stable and the procedural outcome was the patient survived the explant. | [5 empty] | The investigation for the reported limb ischemia and stroke has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of limb ischemia and stroke could not be determined as the device was not returned nor sufficient clinical details. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
15595 | 1220648-2024-22224 | 20620975 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20620975&pc=OZD | 04/21/2023 | 11/06/2024 |  | 11/06/2024 |  |  | 04/21/2023 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1383039 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced a failed optical bench in field service review. no clinical details regarding resolution or patient condition were provided. no patient was involved. | [5 empty] | The investigation has been completed. the console (ic6642) was sent back for further investigation. the root cause of the bulb power out of spec was due to a defective optical bench. this report is being filed as part of a retrospective review of historical records.
15596 | 1220648-2024-22227 | 20620978 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20620978&pc=PYX | 10/12/2024 | 12/12/2024 |  | 11/06/2024 |  |  | 10/12/2024 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025529649 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella rp flex had elevated placement signal and decrease in flows. this change occurred after removing the x-ray board from under the patient. there was a sharp increase in motor current. the impella rp flex was replaced and the patient survived the explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation into low pump flow has been completed. due to lack of product returned and no data logs provided, the root cause of the low pump flow could not be determined. g.1 revised reporting contact email since the initial manufacturer device report 1220648-2024-22227 was submitted.
15597 | 1220648-2024-22225 | 20621003 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20621003&pc=OZD | 03/10/2023 | 11/06/2024 |  | 11/06/2024 |  |  | 03/10/2023 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1586850 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Audible Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The biomedical engineer reported that the automated impella controller (aic) loaner had no audible alarms. a loaner was sent to the customer to initiate return of this device. there was no patient harm reported. | [5 empty] | E1: name of the initial reporter is unknown. the investigation into the inaudible alarm issue has been completed. the impella automated controller (aic) was returned for investigation. the data analysis of the logs indicated the purge pressure low alarm was triggered during the procedure, so the absence of an audible alarm would have been noticeable. the inaudible alarm was reproduced in the laboratory. the alarm cable was unplugged from the speaker and after plugging it in, the alarms returned. the console also was unable to select by depressing the gray knob during investigation. the voltage between tp101 and tp100 on the impellatronic stayed at 0v and did not rise to 5volts when the knob was pressed. the cause of the aic issue was a loose alarm connector. the cause of the aic issue was a defective output from u101.0 on the impellatronic board. this report is being filed as part of a retrospective review of historical records.
15598 | 1220648-2024-22233 | 20621055 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20621055&pc=OZD | 10/14/2024 | 12/16/2024 |  | 11/06/2024 |  |  | 10/14/2024 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025528177 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 60-year-old male patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. the patient developed a clot in the aorta. the pump remained for 6 days of support and was explanted and care escalated to the 5.5 pump. | [5 empty] | The investigation for the thrombus has been completed. impella cp returned for evaluation. upon visual inspection, no biomaterial was present on pump. dried dextrose present in outflow cage, pump was soaked in water for approximately 5 minutes and dextrose was removed. no additional biomaterial present on pump after soaking. clinical details noted that physician wanted pump inspected to rule out pump as potential cause of clot. no evidence of biomaterial ingestion from lt logs. the root cause of the thrombus could not be determined without additional clinical details. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
15599 | 1220648-2024-22231 | 20621059 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20621059&pc=OZD | 10/14/2024 | 12/10/2024 |  | 11/06/2024 |  |  | 10/14/2024 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025547793 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Delivered as Unsterile Product | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported a slash mark was discovered on the exterior of the impella 5.5 box. the sterile sleeves holding the 5.5 device were investigated and it was found that there was a compromise of the sterile sleeves progressing into the 5.5 compartment. the device was removed from the sterile field, a new device was opened and the case proceeded with new and separate impella 5.5. no patient harm has been reported. | [5 empty] | The investigation for the packaging issues has been completed. the device was not returned for evaluation. due to limited clinical details and lack of product returned, the root cause of the packaging issue - sterility cannot be determined. added h6 impact code 4629 corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
15600 | 1220648-2024-22229 | 20621158 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20621158&pc=OZD | 01/03/2023 | 11/06/2024 |  | 11/06/2024 |  |  | 01/03/2023 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1696215 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The biomedical engineer reported that the automated impella controller (aic) had an out of box failure. it was reported that the console was unboxed and plugged in and smoke was coming from the electrical cord. this aic was replaced by customer service with a brand new aic. there was no patient involvement. | [5 empty] | E1: name of the initial reporter is unknown. the investigation into the power issue/damage issue has been completed. the impella automated controller (aic) was returned for investigation. the data analysis of the logs determined the logs were irrelevant to the event. the returned console exterior power cable and connectors were checked without any finding of abnormalities. after opening the console, there was a burnt capacitor c175 found on the power battery manager (pbm). the burnt pbm had no visual damage to other nearby components. the root cause of the smoke was most likely due to a defective pbm. the root cause of the burnt pbm was not determined. this report is being filed as part of a retrospective review of historical records.
15601 | 1220648-2024-22232 | 20621277 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20621277&pc=OZD | 10/12/2024 | 12/12/2024 |  | 11/06/2024 |  |  | 10/12/2024 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025527424 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support, the 77-year-old male patient's access site was oozing. the patient had severely low platelets and diabetic ketoacidosis. heparin was withheld. the patient was on impella cp support for 21.78 hours before expiring. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation into access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was most likely patient condition related, oozing from groin site related to severely low platelets and diabetic ketoacidosis (dka) as per the clinical description provided. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-22232. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
15602 | 1220648-2024-22235 | 20621527 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20621527&pc=OZD | 10/10/2024 | 12/11/2024 |  | 11/06/2024 |  |  | 10/10/2024 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025546182 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Insufficient Information | [12 empty] | Death | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 had a purge pressure high alarm. tissue plasminogen activator was administered. it was further reported that the family made the decision to withdraw care and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist: section: using the automated impella controller with the impella catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation into high purge pressure has been completed. the impella device was not returned for evaluation. the root cause of high purge pressure was most likely due to biomaterial. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-22235.
15603 | 1220648-2024-22234 | 20621647 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20621647&pc=PYX | 10/11/2024 | 12/12/2024 |  | 11/06/2024 |  |  | 10/11/2024 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025476503 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 45-year-old female patient in an acute myocardial infarction/cardiogenic shock was implanted with an impella rp flex device for mechanical circulatory support. during impella support, there were high purge pressures reported. tissue plasminogen activator (tpa) was added to the purge. after one bag of tpa, the pressures returned to normal and they switched the purge back to bicarbonate. | [5 empty] | The investigation for the high purge pressure has been completed. data logs were reviewed and as soon as the pump was started up, the purge pressure began to trend upwards for 13 hours until the "purge pressure high" alarm triggered. leading up to the alarm, there are several potential motor current spikes that could indicate biomaterial impeding rotation of the impeller. one in particular coincided with a marked drop in motor speed as well. however, none of these motor current spikes correlate with an immediate affect on the purge pressure. this aligns with the fact that the pressure was gradually building over 13 hours, indicating a buildup over the purge gap. as reported through clinical details, purge pressures had returned to normal values roughly 7 hours after the alarm triggered, likely due to tpa being infused in the purge. there were no further issues for the remainder of the case. product was not returned for investigation. based on clinical details and data log review, the root cause of the high purge pressure was determined to most likely be biomaterial. the instructions for use for the impella rp with smartassist system states the following: section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ added h6 impact code 4644 corrections to the initial mfr report: g1 mdr reporting contact email | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
15604 | 1220648-2024-22238 | 20621689 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20621689&pc=OZD | 09/17/2023 | 12/13/2024 |  | 11/06/2024 |  |  | 10/15/2024 | 11/06/2024 |  |  | 12/31/2024 |  | 11/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024318722 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bacterial Infection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support. according to abiomed¿s long-term outcome and quality indicator (loqi) impella registry, a patient endured pseudomonas sepsis with a possible relationship to the impella device. the patient developed a spiked fever 103.1 and white count elevated to 9100. the patient was treated with antibiotics - vancomycin and zosyn. the decision was made to remove impella 5.5. the patient survived the event. | [5 empty] | The investigation for the sepsis has been completed. the impella device was not returned for evaluation. therefore, without sufficient clinical details, the root cause of the reported issues sepsis could not be determined. corrections to the initial mfr report:g1 mdr reporting contact email. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
15605 | 1220648-2024-22237 | 20621819 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20621819&pc=OZD | 10/09/2024 | 06/25/2025 |  | 11/06/2024 |  |  | 10/09/2024 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Consumer | [5 empty] | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Social media manager for johnson and johnson reported a patient expressed, "sadly the impella caused got stuck in my aorta causing almost death." no further information has been provided regarding the incident. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. a.1, a.2. a.3, a.4 and a.5 are unknown. b.1 exact brand name is unknown. d.4 exact model number, catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. e.1 facility name and address are unknown. h.4 device manufacture date is unknown. | No product or data logs returned for analysis. no information on case or complaint intake provided. complaint mentions impella got stuck in aorta. the root cause of the positioning issue was not determined as no product or data logs were returned for analysis and insufficient information related to failure was provided g1 reporting contact email revised on manufacturer device report 1220648-2024-22237 in accordance with updated procedures.
15606 | 1220648-2024-22240 | 20621992 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20621992&pc=OZD | 10/15/2024 | 12/12/2024 |  | 11/06/2024 |  |  | 10/15/2024 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025551011 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, oozing noted at the swan sheath site. the patient received two units of blood. additionally, the patient was explanted at bedside due to malposition. the patient survived the event. | [5 empty] | The investigation for the access site bleed and the positioning issue has been completed. the pump was not returned for investigation, and a data log review was not required for this case run. according to the clinical team, the patient received 2 units of blood after reported bleeding from the swan site. additionally, the impella was pulled back during repositioning. the cause of the positioning issue was most likely use related since pump was pulled out during repositioning. the cause of the access site bleeding issue was most likely patient condition related. the bleeding was unrelated to the impella, as it was reported at a non-impella access site. added h6 impact code 4627 corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
15607 | 1220648-2024-22239 | 20622001 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20622001&pc=OZD | 10/07/2024 | 12/20/2024 |  | 11/06/2024 |  |  | 10/07/2024 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025515648 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 85 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support the patient had compartment syndrome. while on support, the patient had an embolectomy and compartmental division. the clinical representative confirmed that the patient had the compartment syndrome prior to the impella implant. it was not thought to be impella related. furthermore, the patient had bleeding due to coagulation issues. the clinical representative confirmed the patient was bleeding from all sites. the intervention for the bleeding is unknown. additionally, after the surgery, the patient suddenly deteriorated. the impella had permanent suction with "position unknown" alarms. the left ventricle was filled. good pump position was confirmed and volume was given to the patient. the impella flows never returned. therapy was discontinued and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation low pump flow and access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of low flow was most likely patient condition related. the root cause of access site bleeding was most likely anticoagulation management related. note: removed failure mode thrombus due to no mention of a thrombus found in clinical description. g1 reporting contact email revised on manufacturer device report 1220648-2024-22239 in accordance with updated procedures.
15608 | 1220648-2024-22245 | 20623009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20623009&pc=OZD | 04/28/2021 | 11/06/2024 |  | 11/06/2024 |  |  | 04/28/2021 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1390016 |  | 0042-0000-JP |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported during inspection, smoke was noted to be emitted from the gap between the outer frame and the display. it was confirmed that the flex on the display board was inserted diagonally. no report of patient involvement, harm, or injury. | [5 empty] | D.9 the device return date for evaluation is unknown. the investigation has been completed. the console logs from the reported day of event were not available. the console was inspected by japan field service but no source of smoke or burning could be identified. the console inspection revealed a misaligned ribbon cable between kbd assembly and the 4-in-1, but no relation to any burning/smoke could be established. liquid crystal display module and power battery manager pca have been replaced as a precaution. no further troubleshooting was performed. the console was tested and confirmed to operate within specification. the cause of the smoke was unable to be determined because the source was unable to be identified. this report is being filed as part of a retrospective review of historical records.
15609 | 1220648-2024-22247 | 20623169 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20623169&pc=OZD | 12/02/2022 | 11/06/2024 |  | 11/06/2024 |  |  | 12/02/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1181833 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | Insufficient Information | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) that as the aic was turned on during case preparation. alarm console failure - switch to backup console¿ appeared. there was no pump connected at this time. the nurse switched the console for the case and successful pump placement, patient stable there was no report of patient harm or injury. | [5 empty] | The investigation for the reported controller failure issue has been completed. the product was returned, and the issue was confirmed. console logs from the reported day of event are consistent with complaint and show controller alarm due to defective optical sensor lamp (# 1039), preceded by the optical bench failing internal check of 5s parameters (and not recovering after optical bench is power-cycled). troubleshooting of the console performed in aachen revealed minor damage to the blue receptacle fiber, which was unrelated to this failure mode. it¿s been also discovered that there was no light output from the optical bench. further testing showed that the issue was not caused by the lamp assembly but rather optical bench defect resulting in intermittent power to the light bulb. per the results of the clinical review and investigation, the root cause was determined to be a defective optical bench. the cause of the optical bench defect was not determined. this report is being filed as part of a retrospective review of historical records.
15610 | 1220648-2024-22241 | 20623390 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20623390&pc=OZD | 12/18/2021 | 11/06/2024 |  | 11/06/2024 |  |  | 12/19/2021 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1345073 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella 2.5 for mechanical circulatory support. when attempting to change the purge bag, the aic stopped responding. neither the soft buttons nor the selection button react when pressed. aic was replaced and the pump runs with the new aic without problems. no patient harm reported. | [5 empty] | The investigation has been completed. console was powered on without any issues. console was ran overnight with a pump and purge setup with no issues occurring. softkeys were behaving as expected. the root cause of the non-responsive softkeys was likely due to a wet console screen and/or wet staff hands. this report is being filed as part of a retrospective review of historical records.
15611 | 1220648-2024-22242 | 20623681 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20623681&pc=OZD | 06/04/2021 | 11/06/2024 |  | 11/06/2024 |  |  | 06/04/2021 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1147480 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that the automated impella controller (aic) would not recognize the catheter. it was reported that when the impella was plugged into aic it was not recognized. the aic was replaced with another aic resulting in no issues with the replacement aic. there was no patient harm reported. | [5 empty] | The investigation into the pump not detected/case start issue has been completed. the automated impella controller (aic) was returned for investigation. while in the field, the console failed to detect a pump during case start. the logs showed that the console was able to detect the pump connection but failed to read the pump eeprom. this is a new failure mode only found in aic_v8.4 and aic_v8.4.1 and has been documented as a software defect. in the logs, it can be seen that during console boot-up, the optical bench state machine (ossiopsens) got stuck in post_start_state and did not reach post_ok_state. this preveneds the software from reading pump eeprom. the root cause is bugs or defects in software (excuding firmware). this report is being filed as part of a retrospective review of historical records.
15612 | 1220648-2024-22236 | 20623783 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20623783&pc=OZD | 10/15/2024 | 12/05/2024 |  | 11/06/2024 |  |  | 10/15/2024 | 11/06/2024 |  |  | 11/30/2025 |  | 11/06/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024430724 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Hematuria | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported shortly following impella cp implant, the patient began bleeding from their access site and developed hemolysis noted by hematuria. the pump was explanted that same day in response to the conditions. the patient survived the event. | [5 empty] | The investigation into access site bleeding and hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-22236 g1 reporting contact facility name missing. g1 reporting contact email. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
15613 | 1220648-2024-22244 | 20623883 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20623883&pc=OZD | 12/06/2022 | 11/06/2024 |  | 11/06/2024 |  |  | 12/06/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1433698 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The biomedical engineer reported that during preventive maintenance the automated impella controller (aic) purge pressure displayed 65 mmhg when 100 mmhg was applied. there was no patient involvement. | [5 empty] | The investigation for the reported purge issue has been completed. the product was returned, and the issue was confirmed. data analysis: aic logs were not relevant to this failure mode. device analysis: the pressure reading was below the 75 mm hg limit when pressure of 100 mm hg was applied. failure mode was isolated to the purge pressure sensor. this component was replaced so that the console passed this section of the procedure. per the results of the clinical review and investigation, the root cause was determined to be due to defective purge pressure sensor. this report is being filed as part of a retrospective review of historical records.
15614 | 1220648-2024-22277 | 20623915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20623915&pc=OZD | 07/21/2022 | 11/06/2024 |  | 11/06/2024 |  |  | 07/22/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1379222 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported during setup the automated impella controller ( aic) would not recognize the purge disc. staff attempted using a different/second purge cassette and encountered the same problem. swapped to a new aic and were able to proceed with case. no patient harm was reported. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, it was confirmed the purge disc flag had broken/snapped off of the purge assembly. this report is being filed as part of a retrospective review of historical records.
15615 | 1220648-2024-22284 | 20624034 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20624034&pc=OZD | 08/27/2021 | 11/06/2024 |  | 11/06/2024 |  |  | 08/30/2021 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | 0042-0000-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Female |  |  | Asian | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary insertion was on an automated impella controller (aic) for mechanical circulatory support. while on support, the pump was turned on, the numbers were displayed, but suddenly disappeared (lv waveforms are displayed). the console was replaced to address the issue. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported the lack of lv waveform data has been completed. the cause of the issue and its relation to the impella automated controller device was not determined since the product, data logs, and sufficient clinical information were not provided. per the results of the clinical review and investigation, the root cause could not be determined this report is being filed as part of a retrospective review of historical records.
15616 | 1220648-2024-22248 | 20624096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20624096&pc=OZD | 11/11/2022 | 11/06/2024 |  | 11/06/2024 |  |  | 11/11/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1371823 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported that upon turning on the automated impella controller (aic) console, it displayed "system self check failed, change console immediately.¿ the team used a backup console for the procedure. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported system self-check error issue has been completed. the product was returned, and the issue was confirmed. data analysis: when the console was booted several log lines triggered reporting powermod_1 communication errors. device analysis: console arrived in danvers. aic was booted and the system self-check failure was reproduced. the console was opened and the interior was inspected. corrosion on the pbm was found near super capacitors c69 and c70. per the results of the clinical review and investigation, the root cause was determined to be contamination. this report is being filed as part of a retrospective review of historical records.
15617 | 1220648-2024-22257 | 20624139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20624139&pc=OZD | 05/27/2021 | 11/06/2024 |  | 11/06/2024 |  |  | 05/27/2021 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1096322 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The biomedical engineer reported that the automated impella controller was in use and both the screen and knob froze. the engineer also reported that when the biomedical engineer received the console, it was booted up without issue and the screen and knob worked correctly. there was no patient harm reported. | [5 empty] | The investigation into the keyboard/rotary knob has been completed. the automated impella controller (aic) was returned for investigation. the clinical staff reported that the console became unresponsive as the screen and knob froze. the biomedical staff rebooted the console to resolve the issue. this complaint was related to a clinical procedure, but there were no adverse events related to this product malfunction. the logs revealed there were many keyboard guard errors a couple days before the complaint date. the console was tested manually to validate that the keyboard functioned properly and was power cycled looking for any signs of keyboard guard errors. the failure mode (keyboard unresponsive) could not be reproduced throughout testing, but the keyboard pca subcomponent was most likely responsible for triggering the errors in the clinical setting. the cause of the unresponsive keyboard and keyboard esd errors was unable to be determined because it could not be reproduced. the board-level root cause was most likely the keyboard pca. this report is being filed as part of a retrospective review of historical records.
15618 | 1220648-2024-22256 | 20624146 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20624146&pc=OZD | 07/18/2022 | 11/06/2024 |  | 11/06/2024 |  |  | 07/18/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1373062 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The us complainant had an aic (automated impella controller) replaced by field service due to a broken purge clip. there was no noted harm or injury from the aic replacement. ifu being followed and backup console will be used. | [5 empty] | The impella device was received from the customer on 01-aug-2022. data analysis: imc logs were reviewed and deemed irrelevant for the failure mode. device analysis: console arrived in danvers. console purge assembly was inspected. upon opening the purge assembly door, it was confirmed the purge disc retainer had broken/snapped off of the purge assembly. see photo "(b)(6) broken purge disc retainer" attached. repair task: before returning to customer, fs will be instructed to: - replace purge disc retainer and purge flag (5500-0222) - validate purge pressure sensor via pm (0042-7309) - perform fc (0042-7316).
15619 | 1220648-2024-22258 | 20624235 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20624235&pc=OZD | 07/15/2022 | 11/06/2024 |  | 11/06/2024 |  |  | 07/15/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1085774 |  |  |  | No |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The us complainant had a battery failure. the red alarm remained despite the team plugging in the console overnight. the ifu will be followed and a backup controller used. no harm or injury. | [5 empty] | Data analysis: imc logs are consistent with the complaint for the battery failure red alarm. power cycling on the console did not help in resolving the alarm. it was also noticed that console was plugged in ac source all the time due to no alarms issued for ac disconnected. (b)(4). device analysis: console was confirmed for battery switch not engaged on the bottom of aic by the account. the console issued battery failure red alarm with console able to boot up was due to disengaged battery switch with ac power plugged in. console was sent back to danvers for regular check on console and optical system.
15620 | 1220648-2024-22263 | 20624273 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20624273&pc=OZD | 04/25/2022 | 11/06/2024 |  | 11/06/2024 |  |  | 04/26/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced system self-check errors. no clinical details regarding resolution or patient condition were provided. no patient was involved. | [5 empty] | The investigation has been completed. cybersecurity; compromised availability of applicable system or product; networking resources or services rendered unavailable to its users. the cause of the security vulnerabilities were vulnerabilities discovered with yocta linux. this report is being filed as part of a retrospective review of historical records.
15621 | 1220648-2024-22252 | 20624274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20624274&pc=OZD | 05/26/2021 | 11/06/2024 |  | 11/06/2024 |  |  | 05/26/2021 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1308002 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was on an automated impella controller (aic) for mechanical circulatory support. during preparation for a procedure, the aic experienced an optical signal issue. the screen exhibited an "impella priming" message and it went to a screen "connect cable to pump". the customer had another aic device onsite and replaced the device without any further reported issues. no reported injury or harm to patient was reported. | [5 empty] | The investigation has been completed. clinical reviewed stated that the¿pump cable connected to aic, purge line attached also. while on screen "impella priming" it went to a screen " connect cable to pump but this was a cp with smart assist. disconnected and turned off aic and attempted 2 more times and went back to that screen. used a 2nd pump and new aic and worked correctly. i did use my demo with the first aic (the next day) to ensure it was not the aic and my demo worked correctly.¿ this complaint was related to a clinical procedure, and there were no adverse events related to this product malfunction. partial logs from the servers were reviewed and the failure mode was confirmed to the pump not recognized failure mode. this is a pattern failure mode only found in aic_v8.4 and aic_v8.4.1 and has been documented as a software defect. the console was not returned for the investigation. the cause of the failure to read the pump eeprom was due to a software defect (in-house). the field service team was informed to upgrade the console software to version aic_v8.5. the root cause is bugs or defects in software (excludes firmware). this report is being filed as part of a retrospective review of historical records.
15622 | 1220648-2024-22255 | 20624288 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20624288&pc=OZD | 07/19/2022 | 11/06/2024 |  | 11/06/2024 |  |  | 07/19/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1098593 |  | 0042-0000 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The us complainant had a console issue discovered in which the team had a replace aic (automated impella controller) alarm. the ifu will be followed and a backup controller used. no harm or injury possible as it occurred in a pm and not patient use. | [5 empty] | The impella device was received from the customer on 01-aug-2022. device analysis: console arrived at danvers for a full investigation and repair. fwcheck (telnet command) revealed that the epm sau¿s hw and sw versions were listed as ¿n.a.¿ indicating the epm¿s microcontroller either had incompatible fw or could not communicate. this was a result of system self check failure. however, after manually reprogramming the microcontroller, the sw upgrade was still not successful (failed epm). replacing the impellatronic resolved the sw upgrade issues. the root cause of the sw upgrade issues was due to a malfunctioning impellatronic pca (epm circuitry). before sending the console back to the customer, fs was instructed to: replace the impellatronic pca (0042-1115). upgrade aic sw.
15623 | 1220648-2024-22246 | 20624483 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20624483&pc=OZD | 12/02/2022 | 11/06/2024 |  | 11/06/2024 |  |  | 12/02/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1161398 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) had a self-check failure during preparation to implant. a backup console was used, and there was no report of patient harm or injury. | [5 empty] | The investigation for the reported system self-check error issue has been completed. the product was returned, and the issue was confirmed. console logs from the reported day of event are consistent with complaint and show that console failed initial boot-up (system check failure), and then restarted, however the issue - manifested as sau_powermod_2 communication loss - remained unresolved. analysis of the console logs prior to that date revealed that this issue first manifested itself already, immediately after the console came back from pm work. the corresponding fsr document confirms that console housing had to be opened to perform necessary work on the console (hardware upgrades). inspection of the console revealed that a communication cable between pbm and 4-in-1 carrier has not been fully inserted into j13 port on the pbm, which resulted in intermittent (or loss of) communication. per the results of the clinical review and investigation, the root cause was determined to be a loose connector. this report is being filed as part of a retrospective review of historical records.
15624 | 1220648-2024-22264 | 20624523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20624523&pc=OZD | 02/10/2022 | 11/06/2024 |  | 11/06/2024 |  |  | 02/10/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1418929 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 98 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with high-risk percutaneous coronary intervention was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) experienced optical signal issues, which resolved through console replacement. no patient harm reported. | [5 empty] | The investigation has been completed. the console ((b)(6)) was sent back for further investigation. there were no optical bench-related alarms triggered during more than 24 hours of testing with an optical pump simulator and purge. the root cause of the controller error (opm comm. loss) could not be determined because the failure mode was not reproduced during testing. this report is being filed as part of a retrospective review of historical records.
15625 | 1220648-2024-22243 | 20624726 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20624726&pc=OZD | 12/15/2022 | 11/06/2024 |  | 11/06/2024 |  |  | 12/15/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1389732 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) failed the self-test during startup. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported system self-check error issue has been completed. the product was returned, and the issue was confirmed. data analysis: aic logs showed continual powermod_1 communication errors and repeated unexpected controller shutdowns pointing to a potential issue with the power battery management (pbm). device analysis: upon visual inspection, the pbm super cap shows signs of leaking causing a damaged communication line. per the results of the clinical review and investigation, the root cause was determined to be contamination from the pbm super cap leaking into the communication lines. this report is being filed as part of a retrospective review of historical records.
15626 | 1220648-2024-22276 | 20624758 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20624758&pc=OZD | 07/27/2022 | 11/06/2024 |  | 11/06/2024 |  |  | 07/27/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1321781 |  | 0042-0000-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported an alarm was issued for an abnormality in the automated impella controller (aic) control device, and the position waveform became flat and disappeared. a few minutes later, the waveform appeared, but the control device abnormality alarm occurred again and the position waveform disappeared, so we suspected an abnormality in the control device and replaced it, after which the position waveform was displayed normally and no alarms were issued. there was no report of patient harm or injury. | [5 empty] | The investigation into the aic - optical signal issue has been completed. console logs from the reported day of event are 'consistent with complaint and show intermittent loss of data from optical bench, identified as a ¿transient loss of optical data. this is a known pattern failure caused by a temporary loss of optical placement signal. abiomed is aware of the issue, and is working on appropriate corrective action. there¿s no need for repairs if the condition self-corrects. this is a low probability event and lasts only a few minutes. if needed, monitor patient hemodynamics and impella position with imaging. this report is being filed as part of a retrospective review of historical records.
15627 | 1220648-2024-22279 | 20624856 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20624856&pc=OZD | 07/19/2022 | 11/06/2024 |  | 11/06/2024 |  |  | 07/19/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER EU |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER EU |  |  | ABIOMED, INC |  |  | US |  | 1343736 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The field service engineer reported that the automated impella controller (aic) was switched on for testing and displayed a yellow controller error and the soft keyboard buttons did not buzz when pressed. there was no patient involved. | [5 empty] | The investigation into the keyboard rotary knob issue has been completed. the impella automated controller (aic) was returned for investigation. the data analysis of the logs from the reported day of event are consistent with complaint and show persistent keyboard (kbd) communication errors and alarms upon boot-ups, on multiple days. the returned device visual inspection of the ribbon cable connecting the kbd assembly to the 4-in-1 carrier revealed that the cable was misaligned. the cause of the aic issue was a loose connector. this report is being filed as part of a retrospective review of historical records.
15628 | 1220648-2024-22281 | 20624919 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20624919&pc=OZD | 07/07/2022 | 11/06/2024 |  | 11/06/2024 |  |  | 07/07/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1362195 |  | 0042-0000-US |  | Yes |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary insertion was on an automated impella controller (aic) for mechanical circulatory support. while on support, a pump emitting a "pump failure" alarm when plugged into the aic controller. a new pump was used without any harm to the patient. console was found to be a possible contributing factor in this event, upon investigation. it was noted that the patient was not effected. | [5 empty] | The investigation for the reported pump not detected issue has been completed. the product was returned, (the return date is unknown) and the issue was confirmed. console logs from the reported day of event are consistent with complaint and show that pump was connected (detected) twice but its eeprom was unable to be read, due to an unspecified communication issue between epc and epm (impellatronic). a third connection was successful (after initial problems reading the first section of eeprom) and the pump was detected and started, and ran for 1 hour, before console was powered down. during testing in danvers we were able to reproduce eeprom reading issues, but of a different nature than those experienced in the field: our tests showed instances of data corruption during i2c transmission from pump eeprom to eep/epm. serial communication issues between epm and epc were not reproduced. nevertheless, the issue was most likely due to impellatronic malfunction. there was also a prior complaint with same symptoms. per the results of the clinical review and investigation, the root cause was most likely a defective impellatronics board. this report is being filed as part of a retrospective review of historical records.
15629 | 1220648-2024-22262 | 20625222 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20625222&pc=OZD | 09/01/2022 | 11/06/2024 |  | 11/06/2024 |  |  | 09/09/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1300878 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Fell | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The biomedical engineer reported that the automated impella controller (aic) fell off the table in the catheterization laboratory during a swap out and upon impact the knob came off. a follow up with the engineer indicated that the device was being returned to be checked over to assure it was safe for use. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported damage issue has been completed. the product was returned, and the issue was not confirmed. imc logs were reviewed but no abnormality was found. after the console arrived in danvers, it was fully tested, especially on rotary push button. no abnormality was found. per the results of the clinical review and investigation, there was no issue found during the case and the device functioned/operated to specification. this report is being filed as part of a retrospective review of historical records.
15630 | 1220648-2024-22265 | 20625265 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20625265&pc=OZD | 08/16/2022 | 11/06/2024 |  | 11/06/2024 |  |  | 08/16/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1670339 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  | White | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported that the automated impella controller (aic) was causing shocks during walks with the patient by the intensive care unit nurse. the nurse felt a shock coming from the aic and reported the location as the right upper above the grey selector knob. the physician elected to change the console at the bedside. this aic was replaced with another aic resulting in no issues with the replacement aic. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported cart issue has been completed. the product was returned, and the issue was not confirmed. data analysis: aic logs were looked into but they were not found to be relevant to the failure mode, device analysis: during the time of the complaint, the console was not plugged into ac and was one the cart. the console was tested in the lab, the failure did not reproduce during this process and passed the electrical safety test. the coating inside the console was equipotential to the grounding wires all throughout the console. the continuity for different pins of the connector cable between rotary and impellatronic was tested using a multimeter, and also compared with a known good connector cable, the connector cable was found to be good. while operating on battery, such a high voltage generation inside the console is not possible under normal operation. the charge source was most likely outside of console electrical circuits. console was being walked isolated from ground ¿ this was an ideal condition for the equipment and clothing to build electrostatic charge. theoretically, a clothing material can build 30kv+ in some cases. the aic was not the cause of the shock. per the results of the clinical review and investigation, the root cause was determined to be use related. the source of the "shock" was likely the electrostatic discharge from the user (charge build up on the clothes) this report is being filed as part of a retrospective review of historical records.
15631 | 1220648-2024-22278 | 20625341 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20625341&pc=OZD | 07/15/2022 | 11/06/2024 |  | 11/06/2024 |  |  | 07/15/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1504720 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) that it did not recognize the pump at the connector (covered by the swivel door). there was no report of patient harm or injury. | [5 empty] | The investigation for the reported pump not detected issue has been completed. the product was returned, and the issue was not confirmed. device analysis: failure mode was attempted to be reproduced however did not. the aic was tested through case start without issue. several different pumps were plugged into the aic repeatedly without issue of recognition. per the results of the clinical review and investigation, there was no issue found during the case. the device functioned/operated to specification. this report is being filed as part of a retrospective review of historical records.
15632 | 1220648-2024-22267 | 20625443 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20625443&pc=OZD | 08/15/2022 | 11/06/2024 |  | 11/06/2024 |  |  | 08/15/2022 | 11/06/2024 |  |  | 06/09/2014 |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1161035 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Fell | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The biomedical engineer reported that the automated impella controller (aic) was dropped or suffered a heavy impact resulting in housing damage. a loaner was onsite for patient support. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported damage issue has been completed. the product was returned, and the issue was confirmed. data analysis: imc logs around the complaint date do not reveal any malfunctions. device analysis: during inspection, the console housing was found chipped and broken. no abnormality was found for console internal modules and connections. per the results of the clinical review and investigation, the root cause was determined to be use related. this report is being filed as part of a retrospective review of historical records.
15633 | 1220648-2024-22269 | 20625697 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20625697&pc=OZD | 08/08/2022 | 11/06/2024 |  | 11/06/2024 |  |  | 08/08/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1420045 |  | 0042-0010-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) after the pump was inserted and the patient was ready to come off the table, it was noted that the power cord was frayed and was not charging the console. a new console was brought in, and the pump was swapped to the new one. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported a/c power issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the root cause of frayed power cord was unable to determine because neither product was returned nor a picture of the frayed cord is available. this report is being filed as part of a retrospective review of historical records.
15634 | 1220648-2024-22272 | 20625782 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20625782&pc=OZD | 05/24/2021 | 11/06/2024 |  | 11/06/2024 |  |  | 05/24/2021 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1352449 |  | 0042-0000-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during training, the automated impella controller exhibited transmitter improper installation error. no patient was involved. | [5 empty] | E4: the name and health professional/occupation of the initial reporter is unknown. the cause of the improper installation alarm was unable to be determined because the console and data were not returned for review. this report is being filed as part of a retrospective review of historical records.
15635 | 1220648-2024-22295 | 20625801 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20625801&pc=OZD | 07/14/2021 | 11/06/2024 |  | 11/06/2024 |  |  | 07/14/2021 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1148976 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | Battery Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a pump's placement signal failing during support. the white connector cable was reseated and the issue resolved. additionally, the device exhibited a battery communication failure. the console was replaced, no report of harm or injury to the patient. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the battery is inconsistently failing and appears to work at some points, and not others. charging and discharging correctly at times, and not at others, likely due to a malfunctioning battery. it was noted that the console did not receive its scheduled battery replacement on the three-year interval. the root cause of the battery communication failure is likely due to the fact that the battery set is out of compliance, however, root cause was unable to be determined. the device logs were reviewed for the placement signal loss and confirmed the loss of placement signal occurred for around three minutes, ending when the controller error alarm cleared. this is a known pattern failure in which all signals received from optical bench. this lasts between 2 to 10 minutes, after which all signals go back to normal. no repair will be necessary as the issue has a low probability of reoccurrence, lasts only up to 10 minutes, and does not affect hemodynamic support. this report is being filed as part of a retrospective review of historical records.
15636 | 1220648-2024-22268 | 20625848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20625848&pc=OZD | 05/27/2021 | 11/06/2024 |  | 11/06/2024 |  |  | 05/28/2021 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1327897 |  | 0042-0010-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The biomedical engineer reported that when inserting the purge cassette, the automated impella controller (aic) did not recognize the purge cassette and did not vent the purge system. it was reported that a second purge assembly was used with a replacement aic with no issues. there was no patient harm reported. | [5 empty] | D.9 the exact date the device was returned for evaluation is unknown. the investigation into purge disc/flag issues has been completed. the clinical staff reported that the console had issues recognizing the purge cassette during preparation for a clinical procedure. the console was sent back to germany field service for further investigation. the logs did not confirm the reported complaint. rather, it was discovered that there were issues driving the purge cassette. the logs indicated that the piston was blocked with 0% range. the user interface presented a screen that the purge fluid could not be detected. while at germany field service, the purge system was scrutinized but was found to be fully functional. upon final inspection of the purge stepper motor gasket, white traces were observed around the purge insert side of the gasket, presumably dried glucose build up. the gasket and the purge housing around the hole was cleaned and the motor was reinstalled. the root cause of the purge fluid not detected error was most likely due to dried glucose build up, which bonded the stepper motor gasket to the purge insert plastic subcomponent. this report is being filed as part of a retrospective review of historical records.
15637 | 1220648-2024-22273 | 20625853 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20625853&pc=OZD | 07/27/2022 | 11/06/2024 |  | 11/06/2024 |  |  | 07/27/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1456617 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Crack | Failure to Read Input Signal | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported that during the prep an automated impella controller (aic) purge disc slide notch is broken. console returned for investigation / repair. no patient involved. | [5 empty] | The investigation into the purge disc / flag issue has been completed. the automated impella controller (aic) was returned for investigation. console arrived in danvers. console purge assembly was inspected. upon opening the purge assembly door, it was confirmed the purge disc retainer had broken/snapped off of the purge assembly. this report is being filed as part of a retrospective review of historical records.
15638 | 1220648-2024-22305 | 20625854 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20625854&pc=OZD | 06/14/2021 | 11/06/2024 |  | 11/06/2024 |  |  | 06/14/2021 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1274733 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the automated impella controller (aic) experienced a controller error yellow alarm. there was no reported patient involvement. | [5 empty] | The device was returned and evaluated, and the investigation for the reported controller error (yellow alarm) has been completed. log lines indicate a serial communication error that could potentially be involved with an issue of the pbm. console was brought in house for service and upon first boot, complaint cause was reproduced. console booted with "system self check error." console was opened, and a visual inspection of the interior was performed. connections were checked and observed to be ok. console was dismantled and further inspection of components was performed. pbm was taken out of console and after a visual inspection, a visible burn mark on the bottom of the board was observed. this burn mark is located on a track leading to component l38. complaint cause is due to a malfunctioning pbm. the root cause of the pbm malfunction could be due to a short that caused the burn seen on the bottom of the board. however, without sufficient evidence, the root cause of the pbm malfunction is utd. this report is being filed as part of a retrospective review of historical records.
15639 | 1220648-2024-22266 | 20625906 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20625906&pc=OZD | 01/19/2022 | 11/06/2024 |  | 11/06/2024 |  |  | 01/19/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 117580 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced a purge disc issue and could not be detected. a new console was used without any issues. no patient was involved. | [5 empty] | The investigation has been completed. console logs from the reported day of event are consistent with complaint and show intermittent purge disk detection. the root cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15640 | 1220648-2024-22254 | 20625988 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20625988&pc=OZD | 07/27/2022 | 11/06/2024 |  | 11/06/2024 |  |  | 07/27/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1322215 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. the complainant in the eu had a 5.5 support ongoing when the aic alarmed for a controller failure. the instruction for use (ifu) was followed and a backup controller was used instead. no reported harm or injury. | [5 empty] | The investigation into the controller failure (red alarm) issue has been completed. the automated impella controller (aic) was returned for investigation. data analysis: imc logs on the complaint date are consistent with the description. after few days of patient support, imc logs suddenly started logging with repetitive entries of sau_powermode_1 communication error, and controller error yellow alarms. device analysis: console arrived at aachen field service. the controller error failure was reproduced in aachen. the power battery manager (pbm) was confirmed to have corrosion on the communication traces next to the super caps. the root cause of controller error (pbm1 loss of communication) was due to a defective pbm. the root cause of the defective pbm was corrosion due to leaking super capacitors. this report is being filed as part of a retrospective review of historical records.
15641 | 1220648-2024-22288 | 20626042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20626042&pc=OZD | 08/06/2021 | 11/06/2024 |  | 11/06/2024 |  |  | 08/06/2021 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1502288 |  | 0042-0040-JP |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | Insufficient Information | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the aic experienced purge issues. no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | D8/d9 date of return is unknown. the investigation for the reported purge issue has been completed. the automated impella controller (aic) was returned for investigation. console logs were not analyzed, as they did not contain any relevant information. after console was shipped to japan field service, the issue was not reproduced with any purge cassette used for testing, however, when the purge cassette alignment tool was used - the issue was reproduced. measurements of opening in the purge cassette insert confirmed that its vertical dimension was out of spec (measured: 36.49mm, vs. specification 36.6mm+0.3mm). the cause for the out-of spec dimension was not investigated. this is the first known occurrence of this failure mode. this report is being filed as part of a retrospective review of historical records.
15642 | 1220648-2024-22296 | 20626068 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20626068&pc=OZD | 07/12/2021 | 11/06/2024 |  | 11/06/2024 |  |  | 07/12/2021 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | TESTED OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1534633 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) failed multiple times during testing. a loaner device was sent to the customer. there was no reported patient involvement. | [5 empty] | The device was returned and evaluated, and the investigation for the reported controller error (yellow alarm) has been completed. in line with the complaint entry, it is most likely that the failure that was exhibited was a console system self-check failure. the console was boot cycled 100 times in an attempt to reproduce the system self-check failure of the console/ antagonize the impellatronics to fail again. however, this was unsuccessful, as the console booted successfully each time. a pump simulator was then run on the console with a multimeter probing the 20v pmd voltage cable. the multimeter was set to catch a minimum voltage drop while the console was running a pump at maximum flow for 1 hour. no failures occurred as the voltage was stable through the duration of the pump running. the complaint cause is likely due to a malfunctioning impellatronics board that was exhibited by the consoles system self-check failure. the cause of the failure can be pointed to the cpld circuitry of the impellatronics. this report is being filed as part of a retrospective review of historical records.
15643 | 1220648-2024-22282 | 20626074 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20626074&pc=OZD | 08/31/2021 | 11/06/2024 |  | 11/06/2024 |  |  | 08/31/2021 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2022058404 |  | 0042-0020-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) that there was a problem with the stand pole screw receiving side. the screws cannot be tightened properly, and the stand pole cannot be fixed. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported cart issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. japan field service confirmed the aic cart was defective and had sent a replacement aic cart pole. the damage to the threaded screw holes on the aic cart stand pole is undetermined. this report is being filed as part of a retrospective review of historical records.
15644 | 1220648-2024-22294 | 20626137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20626137&pc=OZD | 07/09/2021 | 11/06/2024 |  | 11/06/2024 |  |  | 07/09/2021 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1182671 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) experienced an error message during the case and the catheter was switched to another console. no patient harm was reported. | [5 empty] | The investigation for the reported optical signal issue has been completed. the product was returned, and the optical signal issue was not confirmed. the cause of the alarm during case start and inability to shut down console was loss of communication to optical bench. the root cause of opm communication loss could not be determined because issue was unable to be reproduced. this report is being filed as part of a retrospective review of historical records.
15645 | 1220648-2024-22283 | 20626154 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20626154&pc=OZD | 08/27/2021 | 11/06/2024 |  | 11/06/2024 |  |  | 08/27/2021 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1302706 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility biomedical engineer reported for an automated impella controller (aic) that it was exhibiting a controller error. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported controller error issue has been completed. the product was returned, and the issue was confirmed. data analysis: complaint date is consistent with reported occurrence date. this is a known pattern failure in which all signals received from optical bench (placement, snr, contrast, lamp, gain) are interpreted as -16384 values. this lasts between 2 to 10 minutes, after which all signals go back to normal. device analysis: complaint cause is a known pattern failure characterized by temporary loss of placement signal. no repair will be necessary as the issue has a low probability of reoccurrence, lasts only up to 10 minutes, and does not affect hemodynamic support. per the results of the clinical review and investigation, the root cause could not be determined due to the inability to reproduce. this report is being filed as part of a retrospective review of historical records.
15646 | 1220648-2024-22270 | 20626157 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20626157&pc=OZD | 08/02/2022 | 11/06/2024 |  | 11/06/2024 |  |  | 08/02/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1232830 |  | 0042-0000-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) had a self check failure on startup. restarting the console did not resolve the issue. the biomedical engineer opened the console to inspect the pbm and found what appears to be leakage from the super caps, and the pbm board had discolored through the solder point near the power mod. there was no reported patient involvement. | [5 empty] | The investigation for the reported system self-check error issue has been completed. the product was returned, (unknown return date) and the issue was confirmed. data analysis: imc logs for 08/03/2022 show repetitive ¿sau_ser10 received duplicate command seq: (b)(4). send ack eok ! ¿ which points towards error in serial communication with pbm vial channel 10. this issue resulted in ¿post: syscall10 checkfirmware failed¿ errors. device analysis: during pm, pbm was visually inspected and corroded connections were found near the super capacitor. the damage to the pbm was due to leaking electrolyte from the neighboring super capacitors which impacted i2c communication. per the results of the clinical review and investigation, the root cause was determined to be contamination. there was a electrolyte leakage from the supercapacitor which caused damage to the pbm serial connection. this report is being filed as part of a retrospective review of historical records.
15647 | 1220648-2024-22260 | 20626162 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20626162&pc=OZD | 07/13/2022 | 11/06/2024 |  | 11/06/2024 |  |  | 07/13/2022 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1205389 |  | 0042-0000 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The us complainant had a purge disc issue discovered in use at the hospital by the nurse. the aic will be sent for service and the ifu followed and a backup console used. no harm or injury. | [5 empty] | The impella device was received from the customer on 1-aug-2022. data analysis: aic logs from the complaint date showed purge related issues during the purge cassette change procedure. see ic3907 aic logs complaint date in the attachment section. device analysis: the console was tested after arriving in fs. the purge disc retainer was confirmed to be broken (broken blue disc retainer - see attached image). before returning this console back to the customer, fs was instructed to replace the purge disc retainer, validate sensor accuracy via section 12 of the pm procedure (0042-7309), and perform a full functional check on the console and optical system (0042-7316).
15648 | 1220648-2024-22250 | 20626166 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20626166&pc=OZD | 05/07/2021 | 11/06/2024 |  | 11/06/2024 |  |  | 05/07/2021 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1407754 |  | 0042-0010-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Break | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the aic grounding plug was broken off, and there was a micro crack on the purge luer proximal to the purge pressure reservoir. no report of patient harm or injury was noted. | [5 empty] | The investigation into the damage has been completed. the console was never returned for investigation. the automated impella controller was not received from the customer. the root cause is inadvertent console damage by the customer. this report is being filed as part of a retrospective review of historical records.
15649 | 1220648-2024-22302 | 20626189 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20626189&pc=OZD | 06/08/2021 | 11/06/2024 |  | 11/06/2024 |  |  | 06/08/2021 | 11/06/2024 |  |  |  |  | 11/06/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1319742 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility called customer support and reported that the automated impella controller (aic) experienced a controller failure alarm. during the customer support call the alarm resolved and the placement signal returned. customer support advised the customer to look for a backup console in case the alarm returned, there was no further information indicating additional issues. there was no report of patient harm or injury. | [5 empty] | Clinical review: clinical staff reported that ic5691 issued a controller error and had a ¿transient¿ loss of the optical placement signal. this complaint was not related to a clinical procedure, and there were no adverse events related to this product malfunction. after reviewing the imc logs from the impella connect servers, the customer was advised not to send the console back for further investigation since this is an intermittent pattern failure with a low risk of reoccurrence. data analysis: the controller error (crc error ¿ event set #1045) was confirmed in the imc logs (see attached). the cause of the placement signal issue (opm signals went invalid) was unable to be determined because the issue was unable to be reproduced. this report is being filed as part of a retrospective review of historical records.
15650 | 1220648-2024-21953 | 20605864 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20605864&pc=OZD | 08/03/2022 | 11/04/2024 |  | 11/05/2024 |  |  | 08/03/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | AUTOMATED IMPELLA CONTROLLER |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1283659 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a damaged purge disc holder in the automated impella controller (aic) . a loner aic was sent. there was no patient involvement noted. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, the purge flag and blue retainer were found to be loose at initial visual inspection (see attached image) as well as via logs (attached). the issue with properly detecting the purge pressure disc was reproduced. during this investigation, the purge flag completely broke off. therefore, the root of the loose purge clip not reading purge was due to a defective component. this report is being filed as part of a retrospective review of historical records.
15651 | 1220648-2024-21944 | 20605914 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20605914&pc=OZD | 06/14/2022 | 03/27/2025 |  | 11/05/2024 |  |  | 06/14/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1076062 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a battery failure with an aic (automated impella controller). the team contacted field service for aic replacement per ifu recommendations. there was no reported harm or injury to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Data analysis: event logs were consistent with complaint intake. subsequent boots following the complaint date triggered alarms. long term log analysis of the battery data reveals that beginning before the complaint date of occurrence, cell 4 voltage of battery 1 had been declining for an extended period of time. device analysis: console arrived in danvers. it was confirmed that persistent battery failure (transport connection blown/missing) was not due to a disengaged battery switch. when console was booted for the first time in danvers, console alarms are consistent with what was seen in the field. console interior connections were checked and verified. when visually inspecting the batteries and pbm, it was observed that one of the batteries status leds were completely blank. conclusion: the battery failure alarm was due to a defective battery 1(decline of individual cell voltage). the root cause of the battery cell failure was undetermined. this report is being filed as part of a retrospective review of historical records. d9: device available for evaluation? corrected to yes. device returned to manufacturer checked, and date device returned to manufacturer added.
15652 | 1220648-2024-21933 | 20605915 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20605915&pc=OZD | 08/15/2022 | 11/04/2024 |  | 11/05/2024 |  |  | 08/15/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1351112 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported that during preventive maintenance of a automated impella controller (aic) that the console displayed (controller failure-purge system has stopped). upon receipt, downloaded log files. updated sw to v11.1. upon power "on" the console displayed the controller failure. recycled the power three times with the same failure displayed. | [5 empty] | The investigation for the reported controller failure (red alarm) issue has been completed. the product was returned, and the issue was confirmed. data analysis: imc logs on the complaint date were consistent with the complaint. after software upgraded to v11.1, console presented a controller failure alarm that was not resolved by power cycling. device analysis: the console was tested with purge commends, which indicated that the purge pressure sensor was defective. after pps was temporarily replaced, the console would not alarm for controller failure (also indicating only defective pps was the root cause of failure). per the results of the clinical review and investigation: the root cause was determined to be a defective component. this report is being filed as part of a retrospective review of historical records.
15653 | 1220648-2024-21896 | 20605923 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20605923&pc=OZD | 10/05/2024 | 12/06/2024 |  | 11/05/2024 |  |  | 10/05/2024 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025512554 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Gastrointestinal Hemorrhage | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the 53-year-old male patient was on impella cp support due to st elevated myocardial infarction (stemi). while on support, a critical alarm noted per automated impella controller (aic): impella in aorta. the patient had ventricular tachycardia (vt) arrest and was defibrillated twice with return of spontaneous circulation (rosc). patient also had a run of supraventricular tachycardia (svt) overnight requiring 1 shock. amio drip started. patient had frequents run of non-sustained vt without losing pulsatility. additionally, kidney function continued to deteriorate. patient received 2 units of red blood cells after hh dropped post bloody bowel movement. patient continued to deteriorate with increased cr and no urine output overnight despite increased diuresis. patient required blood products. hh increased today. heparin was withheld due to gastrointestinal (gi) bleed. gi bleed remained active. the patient continued to deteriorate. care was withdrawn and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the gastrointestinal bleed (gi bleed) and ventricular tachycardia (vt) has been completed. the pump was not returned for investigation. the data log shows no signs related to improper positioning or any interaction with the device during the case. as per the given clinical information, the patient experienced vt and svt, requiring 1 shock. additionally, a gi bleed was reported, and blood products were given. improvement in hemoglobin was noted, though the gi bleed remained continuous. the patient was continued to deteriorate. care was withdrawn and the patient expired. the cause of the gi bleed issue was not determined since product was not returned. the relation between gi bleed and impella device was unknown. the root cause of the vt could not be determined without additional clinical details. added-b2 selected required intervention corrections to the initial mfr report: h6 clinical code 4476.
15654 | 1220648-2024-21903 | 20605936 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20605936&pc=OZD | 09/01/2022 | 11/04/2024 |  | 11/05/2024 |  |  | 09/01/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1122631 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was on an automated impella controller (aic) for mechanical circulatory support. while on support, the aic experienced a priming issue that was resolved by replacing the console. there was no report of patient harm or injury. | [5 empty] | The investigation into the purge disc/flag issues issue has been completed. the automated impella controller (aic) was returned for investigation. data analysis: impella defective alarm confirmed on console ic2046 with two different pumps. cloud data was checked for the usage of the second console (ic8192) with the pump sn 389815 (2nd pump), and there were no issues observed. device analysis: impella defective alarm was reproduced with console ic2046 with a known-good test pump. the impellatronic on console ic2046 was momentarily replaced to resolve the alarm/ issue. the root cause of the impella defective alarm was due to a defective impellatronic. this report is being filed as part of a retrospective review of historical records.
15655 | 1220648-2024-21951 | 20606036 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20606036&pc=OZD | 08/02/2022 | 11/04/2024 |  | 11/05/2024 |  |  | 08/02/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1133831 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a controller error. a loaner was sent to the customer to initiate return of this device. there was no patient harm reported. | [5 empty] | The investigation into the controller error data has been completed. the impella automated controller was returned for investigation. the data log reports revealed controller error (defective optical sensor lamp ¿ event set #1039) was triggered upon multiple power cycles in the clinical setting. the returned console was inspected and tested on an engineering laboratory bench. the failure mode was not reproduced in the original console condition. the optical system was thoroughly inspected and tested. the ferrule at the blue pump connector was contaminated enough to warrant replacement. theoretically, contamination on the pump connector ferrule can cause less light back reflection and trigger the controller error (1039) due to less than 0.75v of light 5s parameter. this was proven by intentionally contaminating the pump connector ferrule and triggering the controller error. the root cause was not determined due to the inability to reproduce the controller error in the original console condition. this report is being filed as part of a retrospective review of historical records.
15656 | 1220648-2024-21937 | 20606151 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20606151&pc=OZD | 08/14/2022 | 11/04/2024 |  | 11/05/2024 |  |  | 08/14/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1330532 |  | 0042-0000-JP |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 43 YR |  | Female |  |  | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient on an automated impella controller (aic) for mechanical circulatory support. upon starting the aic, it produced a self-check error upon boot up of this console. a backup console was used, and there was no report of patient harm or injury. | [5 empty] | The investigation for the reported system self-check error issue has been completed. the product was returned, and the issue was confirmed. data analysis: imc logs for 08/14/2022 show repetitive ¿sau_ser10 received duplicate command seq: 43560. send ack eok ! ¿ which points towards error in serial communication with pbm vial channel 10. (pdf attached). this issue resulted in ¿post: syscall10 checkfirmware failed¿ errors. device analysis: during pm, pbm was visually inspected and corroded connections were found near the super capacitor. the damage to the pbm was due to leaking electrolyte from the neighboring super capacitors which impacted i2c communication. the root cause of system self-check failure was due to a defective pbm. fcn 71 had been implemented to address other pbm boards manufactured between march and november 2013 that do not yet show signs of damage due to leaking capacitors. pbm super cap is a pattern failure mode and usually occurs with pbms that have a sn between (b)(6) (inclusive). this board's sn was outside of this date range. per the results of the clinical review and investigation, the root cause was determined to be a electrolyte leakage from the supercapacitor which caused damage to the pbm serial connection. this report is being filed as part of a retrospective review of historical records.
15657 | 1220648-2024-21923 | 20606238 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20606238&pc=OZD | 11/28/2022 | 11/04/2024 |  | 11/05/2024 |  |  | 11/28/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1233949 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the aic (automated impella controller) experienced purge disc/flag issues. no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15658 | 1220648-2024-21942 | 20606289 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20606289&pc=OZD | 08/11/2022 | 11/04/2024 |  | 11/05/2024 |  |  | 08/11/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1478138 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) is not registering the purge disc when changing cassettes or starting a new case. the staff tried replacing the purge cassette multiple times, however this did not resolve the problem. no report of patient harm or injury. | [5 empty] | The console (aic) was returned for evaluation. the device logs from the complaint date revealed that the console issued the purge disc not detected alarm several times. upon inspection of the console, issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. therefore, the root cause of the purge disc/flag issues was due to a defective component. purge flag was broken. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15659 | 1220648-2024-21894 | 20606493 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20606493&pc=OZD | 08/17/2022 | 11/04/2024 |  | 11/05/2024 |  |  | 08/17/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1213829 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic), while doing an education training, the aic had an alarm prior to plugging in the catheter. there was no patient involvement. | [5 empty] | The investigation for the reported controller failure (red alarm) issue has been completed. the product was returned, and the controller failure (red alarm) issue was confirmed. the console logs from the reported day of event are consistent with the complaint and show console seemingly detecting pump being connected, but unable to read its eeprom. as the result, the console could not be powered down, because it ¿thought¿ that the pump was still connected. power-cycling the console did not resolve the issue. after the console was received in danvers, the issue was eventually reproduced. considering the pump was observably not actually connected, this behavior is consistent with a known pattern failure where due to a defect of pump detection mechanism on impellatronic, pump is falsely detected. per the results of the clinical review and investigation: the root cause was determined to be a defective impellatronics board. this report is being filed as part of a retrospective review of historical records.
15660 | 1220648-2024-21901 | 20606678 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20606678&pc=OZD | 08/18/2022 | 11/04/2024 |  | 11/05/2024 |  |  | 08/18/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1193561 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during a demonstration the automated impella controller (aic) the purge cassette door was opened, and it was observed that the purge disk holder was broken. no patient involvement. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, it was confirmed the purge disc retainer had broken/snapped off of the purge assembly. therefore, the root cause of the purge disc/flag issues was due to a defective component. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15661 | 1220648-2024-21917 | 20606685 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20606685&pc=OZD | 12/20/2022 | 11/04/2024 |  | 11/05/2024 |  |  | 12/20/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1314428 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the automated impella controller (aic) experienced purge disc/flag issues that were resolved by replacing the aic. no patient harm or injury was reported. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. massive glucose residue was found between the blue retainer and purge unit. the cause of the issue was a broken purge flag. this report is being filed as part of a retrospective review of historical records.
15662 | 1220648-2024-21956 | 20608663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20608663&pc=OZD | 10/07/2024 | 12/10/2024 |  | 11/05/2024 |  |  | 10/07/2024 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025529553 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support when there was access site bleeding. two units of blood products were provided to the patient. angle matching and pressure dressing were performed and the bleeding subsided. it was further reported that care was withdrawn and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was most likely use related, the bleeding has stopped after angle matching and pressure dressing were done as per the clinical description provided. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-21956.
15663 | 1220648-2024-21957 | 20608825 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20608825&pc=OZD | 10/07/2024 | 12/09/2024 |  | 11/05/2024 |  |  | 10/07/2024 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025545556 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Life Threatening | Death | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump placed to support a patient admitted in with cardiogenic shock and cardiomyopathy. the left ventricle was seen to have a thrombus present and this prevented the impella cp removal and escalation to impella 5.5 pump. the patient continued to deteriorate after the initial pump was placed and the team was unable to stop the patient's decline. the pump remained on for support until the patient expired. the patient expired during transport to the intensive care unit. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can led to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can led to systemic embolization, serious injury, or death." | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details. correction to the initial mfr report: g1 mdr reporting contact email.
15664 | 1220648-2024-21958 | 20608998 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20608998&pc=OZD | 04/28/2024 | 12/02/2024 |  | 11/05/2024 |  |  | 10/07/2024 | 11/05/2024 |  |  | 09/30/2025 |  | 11/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025445675 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured supraventricular tachycardia with a possible relationship to the impella device used. the patient went into sustained supraventricular tachycardia with hypotension. the patient was cardioverted nine times and received aggressive medical management with digoxin, lidocaine, amiodarone, adenosine, and aggressive electrolyte management with potassium, magnesium, calcium and bicarbonate. the patient had some improvement in hemodynamics but remained tachycardic. the impella was adjusted at bedside with direct echocardiography with some improvement. the patient recovered with sequelae. it was further reported that the family elected to do-not-resuscitate and the patient expired seven minutes after the removal of the impella cp. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the reported ventricular fibrillation/ventricular tachycardia (vf/vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf/vt could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
15665 | 1220648-2024-21959 | 20609439 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20609439&pc=OZD | 11/09/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 11/09/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1170636 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the purge clip broke on the aic (automated impella controller). it was reported that the initial console was replaced, and the case continued. no harm or injury noted. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15666 | 1220648-2024-21960 | 20609441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20609441&pc=OZD | 11/04/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 11/04/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1417603 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a broken purge disc on the automated impella controller (aic). the initial controller was replaced with a new device and the case continued. no patient harm was reported. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15667 | 1220648-2024-21961 | 20609480 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20609480&pc=OZD | 11/03/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 11/03/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1302874 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during preparation, the purge cassette was not recognized. a replacement automated impella controller (aic) was required, and the case resumed. no patient harm was reported. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the issue was a defective component. the purge flag was broken. this report is being filed as part of a retrospective review of historical records.
15668 | 1220648-2024-21966 | 20609484 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20609484&pc=OZD | 10/10/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 10/10/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1443894 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a purge failure was observed on the automated impella controller (aic) as the purge disc was not being detected. the initial aic was replaced with a new device. no report of patient involvement, harm, or injury. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15669 | 1220648-2024-21962 | 20609488 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20609488&pc=OZD | 11/03/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 11/03/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 200402 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) had a broken purge disk. a replacement aic was required, and the case resumed. no patient harm was reported. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the case of the inability to detect the purge disc was due to a broken retainer. this report is being filed as part of a retrospective review of historical records.
15670 | 1220648-2024-21963 | 20609514 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20609514&pc=OZD | 10/25/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 10/25/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1245288 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a broken purge clip was observed on the automated impella controller (aic). the device was replaced with a new aic. no patient harm was reported. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15671 | 1220648-2024-21964 | 20609549 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20609549&pc=OZD | 10/19/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 10/19/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1170632 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported the purge disc compartment on the automated impella controller (aic) was damaged. no patient harm was reported. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15672 | 1220648-2024-21972 | 20609592 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20609592&pc=OZD | 10/08/2024 | 12/11/2024 |  | 11/05/2024 |  |  | 10/08/2024 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532115 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, the patient went into ventricular tachycardia and required defibrillation. additionally, oozing was noted from bilateral groin sites; femstop applied and two units of red blood cells were given. it was further reported that the patient expired while on support. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the ventricular tachycardia (vt) and the access site bleed has been completed. the impella device was not returned for evaluation. additionally, cinical details provided were not sufficient to determine the root cause. therefore the root cause of the ventricular tachycardia (vt) and the access site bleed could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email
15673 | 1220648-2024-21965 | 20609596 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20609596&pc=OZD | 10/08/2024 | 11/27/2024 |  | 11/05/2024 |  |  | 10/08/2024 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025501293 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump placed to support a patient admitted in acute myocardial infarction and cardiogenic shock and found to have multivessel coronary artery disease. it was reported that the patient required blood transfusion post impella cp placement however it is unknown the origin of the bleed. it was further reported that care was withdrawn and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
15674 | 1220648-2024-21967 | 20609611 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20609611&pc=OZD | 10/15/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 10/15/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1071329 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. the complainant reported a purge disc was not detected on the automated impella controller (aic). the initial aic was replaced with a new device, it was reported that the procedure was successful. no harm or injury to patient. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. the cause of the issue was a defective component. the purge flag was confirmed to be broken. this report is being filed as part of a retrospective review of historical records.
15675 | 1220648-2024-21969 | 20609641 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20609641&pc=OZD | 09/07/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 09/07/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1205400 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automated impella controller (aic) was just received back from preventative maintenance and would not recognize the purge cassette. the customer had a loaner onsite to utilize. no indication of patient involvement, harm, or injury. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. this report is being filed as part of a retrospective review of historical records.
15676 | 1220648-2024-21975 | 20609643 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20609643&pc=OZD | 10/05/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 10/05/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1399832 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automatic impella controller (aic) had a broken purge sensor. no report of patient involvement, injury, or harm. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. the console was tested after arriving in danvers. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. this report is being filed as part of a retrospective review of historical records.
15677 | 1220648-2024-21970 | 20609651 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20609651&pc=OZD | 10/07/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 10/07/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1356082 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automated impella controller (aic) would not recognize the purge cassette. the initial device was replaced, and the case continued. no patient harm was reported. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the issue was broken/ missing purge flag. this report is being filed as part of a retrospective review of historical records.
15678 | 1220648-2024-21971 | 20609678 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20609678&pc=OZD | 06/15/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 06/15/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1574757 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The us complainant had a cp support ongoing when the aic (automated impella controller) alarmed for battery failure. the ifu was followed and a backup controller used for support. there was no harm or injury reported. | [5 empty] | The investigation has been completed. the console was sent back for further investigation. console logs from the reported day of event are consistent with complaint and show that battery failure alarm (#360) and battery level low alarm (#23) presented 1h30m after aic was powered on (and 30 minutes after case start). running pump was disconnected and console was shut down. further analysis of power and battery data embedded in ltlogs revealed cell imbalance of battery 1, that resulted in alarm on 2022-06-15, however the imbalance was already evident as early as 2022-04-04, but did not result in any alarm. there was no imbalance recorded during last known use before 2022-04-04, which was on 2021-09-14. when console was tested in danvers, the issue was reproduced, and ic8835 presented with battery failure alarm upon boot-up. diagnostic data collected with ata interface cable confirmed battery 1 permanent failure due to cell imbalance. battest diagnostic data was consistent with these findings, showing cell 3 of battery 1 voltage being ~200mv lower than the other cells. this report is being filed as part of a retrospective review of historical records.
15679 | 1220648-2024-21968 | 20609685 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20609685&pc=OZD | 10/13/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 10/13/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1114578 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the purge clip on the automated impella controller (aic) was broken. the initial device was removed from service due to this issue. a loaner was sent to the customer. no indication of patient involvement, injury, or harm. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. console purge assembly was inspected. upon opening the purge assembly door, it was confirmed the purge disc retainer had broken/snapped off of the purge assembly. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15680 | 1220648-2024-21973 | 20609721 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20609721&pc=OZD | 10/06/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 10/06/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1206660 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had an automated impella controller (aic) with purge disc issues. the aic was having issues reading the purge disc. no reported injury or harm to patient. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken, and the purge pressure transducer retainer was also cracked. this report is being filed as part of a retrospective review of historical records.
15681 | 1220648-2024-21976 | 20609767 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20609767&pc=OZD | 10/10/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 10/10/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1565140 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. the complainant reported a purge disc not detected alarm from their automatic impella controller (aic). it was reported that the procedure was completed successfully with another device. no reported harm or injury to patient. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. the purge assembly of the console was inspected. it was observed that the purge disc retainer had broken off of the purge assembly. this report is being filed as part of a retrospective review of historical records.
15682 | 1220648-2024-21974 | 20609785 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20609785&pc=OZD | 10/12/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 10/12/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1468305 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a purge disc failure on an automated impella controller (aic). the purge disc sensor is broken, and the console is unusable. this occurred during an in-service and there was no patient involvement. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. the purge disc retainer was confirmed to be broken. this report is being filed as part of a retrospective review of historical records.
15683 | 1220648-2024-21984 | 20609789 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20609789&pc=OZD | 09/13/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 09/13/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1159467 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the automated impella controller (aic) console was not recognizing 2 consecutive purge cassettes. a new console was used successfully with no harm to the patient. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15684 | 1220648-2024-21978 | 20609795 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20609795&pc=OZD | 10/04/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 10/04/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1350280 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. while preparing for a procedure, the aic was unable to recognize purge disc, and the purge door was not fully closing. it was reported that another device was used and the procedure was successful. no reported harm or injury to the patient. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the device has been returned for evaluation. the purge assembly was inspected, with no issues occurring with purge door functionality. purge door was opened and closed multiple times with and without a purge cassette and disc installed. it was observed that the door detection mechanism was properly sensing whether the door was opened or closed. no issue with the purge assembly door functionality was found or reproduced. while inspecting the purge assembly area, it was observed that the purge disc flag was broken. this report is being filed as part of a retrospective review of historical records.
15685 | 1220648-2024-21979 | 20609812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20609812&pc=OZD | 10/08/2024 | 12/17/2024 |  | 11/05/2024 |  |  | 10/08/2024 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025534446 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Infusion or Flow Problem | [6 empty] | Insufficient Information | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp had a sudden increase in motor current while supporting a patient. impella flows were saturated. a structure can be seen sonographically at the inlet. the patient was taken to the catheterization lab where a transesophageal echocardiogram was performed. part of the papillary muscles was seen at the inlet. the pump was repositioned, which solved the problem for the time being. it was therefore decided to leave the pump in place and to continue monitoring. it was further reported that the patient expired while on impella support. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation into saturated flow has been completed. no product or data logs were returned for analysis. the root cause of the saturated flow was not determined g.1 revised mdr reporting contact email since the initial manufacturer device report number: 1220648-2024-21979 was submitted.
15686 | 1220648-2024-21985 | 20609816 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20609816&pc=OZD | 08/02/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 08/02/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1164961 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a controller alarm during power on with the automated impella controller . a new console was used to resolve the issue. no patient was involved. | [5 empty] | The investigation into the controller failure (red alarm) issue has been completed. the automated impella controller (aic) was returned for investigation. console logs from the reported day of event are consistent with complaint and show an alarm indicating epm software corruption (#1029), after each power-cycle, starting on 2022-08-02. analysis of the epm firmware extracted from the impellatronic revealed corruption within the code. after epm firmware was re-flashed, the issue was resolved. inspection of the console¿s internal components revealed some kinks on the impellatronic ribbon cable, and although it is not likely that it caused the firmware issue, it¿s recommended to be replaced as a precaution. this report is being filed as part of a retrospective review of historical records.
15687 | 1220648-2024-21980 | 20609818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20609818&pc=OZD | 10/03/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 10/03/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1232827 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the device exhibited a purge disc not detected alarm. the customer utilized a backup console and the procedure was successful. no reported harm or injury to the patient. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the device has been returned for evaluation. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. this report is being filed as part of a retrospective review of historical records.
15688 | 1220648-2024-21981 | 20609828 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20609828&pc=OZD | 09/23/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 09/23/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1117339 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant observed an automatic impeller controller (aic) with a purge disc failure. a loaner was already onsite. no report of patient involvement, injury, or harm. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the device has been returned for evaluation. the reported issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. the blue purge retainer also had a hairline fracture as seen under the microscope. this report is being filed as part of a retrospective review of historical records.
15689 | 1220648-2024-21977 | 20609830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20609830&pc=OZD | 10/06/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 10/06/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1108441 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automated impella controller (aic) did not detect purge cassette. the use of a backup console resolved the issue. no report of patient harm injury. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. the device logs revealed the reported purge disc not detected alarm triggered once around 8 seconds after purge cassette was inserted. however, the alarm was cleared in few seconds as the purge disc was inserted. it was noted that the console was not in the case start wizard, so the purge disc not detected alarm was supposed to be triggered when pc was connected without purge disc detected. the first alarm solely was not a direct indication of product malfunction given it was resolved moments later. visual inspection did not reveal any abnormalities in purge system. purge flag and purge retainer were intact. simulation of inserting pc and purge disc found the purge disc detection mechanism was functional. the root cause of purge disc not detection was more likely due to a use issue ¿ user did not insert the purge disc in time after insertion of purge cassette. this report is being filed as part of a retrospective review of historical records.
15690 | 1220648-2024-21982 | 20609862 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20609862&pc=OZD | 08/03/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 08/03/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | AUTOMATED IMPELLA CONTROLLER |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1646182 |  | 0042-4201 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a damaged purge disc clip in the automated impella controller (aic). there was no patient involved. a replacement loner was sent to address the issue. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, it was confirmed the purge disc retainer had broken/snapped off of the purge assembly. therefore, the root cause of the broken purge clip was due to a defective component. this report is being filed as part of a retrospective review of historical records.
15691 | 1220648-2024-21983 | 20609875 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20609875&pc=OZD | 09/22/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 09/22/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1254066 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Peeled/Delaminated | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) exhibited an alarm during set up and the gray piece on the device is loose. another device was utilized and there was no reported harm or injury to the patient. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. visual inspection of the purge assembly area confirmed a broken stand on blue purge retainer (the ¿grey piece¿ as reported) and a crack on retainer. the damage on the purge retainer resulted in the purge disc not detected issue. this report is being filed as part of a retrospective review of historical records.
15692 | 1220648-2024-21986 | 20609908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20609908&pc=OZD | 09/08/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 09/08/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1178829 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a broken purge disc on the automatic impella controller (aic). there was no harm or injury noted. the team utilized a backup controller for patient use. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15693 | 1220648-2024-21987 | 20609912 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20609912&pc=OZD | 09/08/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 09/08/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1350280 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a broken purge clip on the automatic impella controller (aic). the console was returned for investigation and repair. no patient was involved. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15694 | 1220648-2024-21988 | 20609992 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20609992&pc=OZD | 07/29/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 07/29/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1480337 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported an automated impella controller (aic) failing to recognize a purge disc. a new aic was used without any harm to the patient. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, it was confirmed the purge disc retainer had broken/snapped off of the purge assembly. therefore, the root cause of the purge disc/flag issues was due to a defective component. this report is being filed as part of a retrospective review of historical records.
15695 | 1220648-2024-21989 | 20609996 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20609996&pc=OZD | 07/27/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 07/27/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1308004 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported purge disc slide notch is broken on automated impella controller( aic). console returned for investigation / repair. no patient involved. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, it was confirmed the purge disc retainer had broken/snapped off of the purge assembly. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15696 | 1220648-2024-21990 | 20610121 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20610121&pc=OZD | 07/18/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 07/18/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER EU |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER EU |  |  | ABIOMED, INC. |  |  | US |  | 1226863 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a controller error, although the pump was not connected. there was no patient harm reported. | [5 empty] | E1: name of the initial reporter is unknown. the investigation into the controller error issue has been completed. the impella automated controller was returned for investigation. the data analysis of the logs revealed the impella defective alarm however the log was not fully consistent with the clinical review. the alarm was triggered every time the pump was plugged in even prior to the complaint date. the impella defective alarm would clear with pump disconnection. the returned product was tested and upon boot up, no alarm was issued. the failure mode, ¿impella defective¿ alarm, was eventually reproduced by connecting the demo pump. after the pump being disconnected, the console would not allow shutting down due to pump disconnection was not detected correctly by impellatronic. after swapping the impellatronic pca, the failure did not reoccur. the root cause of the impella defective alarm was due to a defective impellatronic pca, specifically the electrical power manager circuitry. this report is being filed as part of a retrospective review of historical records.
15697 | 1220648-2024-21991 | 20610197 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20610197&pc=OZD | 07/22/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 07/22/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER S |  |  | ABIOMED, INC. |  |  | US |  | 1414254 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Decrease in Pressure | Use of Device Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The clinical consultant reported that the automated impella controller (aic) was not detecting the purge flow and pressure accurately. it was reported that the grey holding device for the purge pressure transducer was loose. this aic was replaced with another aic. there was no patient harm reported. | [5 empty] | E1: name of the initial reporter is unknown the investigation into the purge issue/purge flag has been completed. the impella automated controller was returned for investigation. the data analysis of the logs from the complaint date revealed that the console issued the purge disc not detected alarm (event set # 402) several times. also, low purge pressure (event #431) can also be seen throughout the logs on 07/21/22) as well as purge_system_priming_fluid_not_detected. the returned console was tested and the issue of properly detecting the purge pressure disc was reproduced. the purge flag was confirmed to be broken/missing and the blue disc retainer was broken/missing. the cause of the issues was a defective component. this report is being filed as part of a retrospective review of historical records.
15698 | 1220648-2024-21993 | 20610211 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20610211&pc=OZD | 07/12/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 07/12/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1075485 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) exhibited a purge disc not detected error message. no report of patient involvement, harm, or injury. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, it was confirmed the purge disc retainer had broken/snapped off of the purge assembly. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15699 | 1220648-2024-21994 | 20610233 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20610233&pc=OZD | 07/14/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 07/14/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1207016 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Use of Device Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the automated impella controller (aic) experienced an air in purge system alarm, the alarm was ongoing after multiple attempts to de-air. the console was swapped out with resolution of the alarm issue and significant improvement in purge flow as well as improved waveforms and impella flow. there was no patient harm reported. | [5 empty] | The investigation into the purge issue has been completed. the impella automated controller was returned for investigation. the data analysis of the logs indicated the case began on 7/9/22 with the reported complaint occurring on 7/14/22, an air in purge system alarm triggered. several attempts to de-air the line were attempted through the purge menu, however these attempts were unsuccessful with clearing the alarm. the real time logs show that after the first occurrence of the air in purge system alarm, purge pressure was unstable and unable to recover despite de-air attempts. the staff additionally noted that after switching to ic7189, no issues occurred with the purge system. however, after reviewing the logs for ic7189, it was observed that a different purge cassette was used after the aic transfer. the returned console was placed on pump and purge and flow level was set to p-6 to mimic the scenario where air in purge system had occurred. the console was left running for an extended period of time with no alarms occurring. the root cause of the air in purge system alarms was not determined due to the inability to reproduce. this report is being filed as part of a retrospective review of historical records.
15700 | 1220648-2024-21992 | 20610334 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20610334&pc=OZD | 07/18/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 07/18/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER EU |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER EU |  |  | ABIOMED, INC. |  |  | US |  | 1327897 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) exhibited persistent yellow alarms during preventive maintenance. additionally, it was found that the impella connect module (icm) does not connect to the cloud. there was no patient involvement as the error occurred during maintenance. | [5 empty] | The investigation into the controller error and impella connect module connectivity issue identified during preventive maintenance (pm) has been completed. the impella automated controller was returned for investigation. the data analysis of the logs indicated a persistent controller error (defective optical sensor lamp ¿ event set #1039) was issued upon each startup. this alarm would occur immediately after the optical bench would reset due to the bad_lamp_recovery_state. the returned console was tested, and the controller error failure mode was observed multiple times during the pm. the investigator swapped the lamp assembly to resolve the controller error and retested the 5 signal parameters/ optical to confirm the bench was within the optical-related specifications per the preventative maintenance procedure (0042-7309). the controller error failure mode was due to a defective lamp assembly. the cause of the controller error issue was a defective component. issues relating to connectivity of impella connect module, icm unable to connected to network. the cause of console unable to connect to network was due to it reverted to mfg mode, however, the root cause of the problem was unable to be determined. this report is being filed as part of a retrospective review of historical records.
15701 | 1220648-2024-21996 | 20610371 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20610371&pc=PYX | 10/08/2024 | 12/06/2024 |  | 11/05/2024 |  |  | 10/08/2024 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025551075 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Defective Device | Patient-Device Incompatibility | [5 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an impella rp flex device placement was attempted but the physician encountered significant difficulty due to patient anatomy. the decision was made to abort ij approach and use the right femoral. additionally, a possible leak was noted from the introducer membrane, kink difficult insertion. the patient expired on support. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ impella rp flex with smartassist system with automated impella controller section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ impella rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ impella rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ | The investigation for the delivery issues and introducer damage has been completed. no product returned for analysis. the root cause of the delivery issue was most likely patient condition related as evidenced by clinical mention of complex anatomy. the root cause of the introducer damage - mechanical was not determined as no product or images were returned for analysis.
15702 | 1220648-2024-21995 | 20610396 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20610396&pc=OZD | 06/01/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 06/01/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1274736 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a controller error. there was no patient harm reported. | [5 empty] | E1: name of the initial reporter is unknown. the investigation into the controller error/syscall error issue has been completed. the impella automated controller was returned for investigation. the data analysis of the logs are consistent with complaint and show evidence of failed boot sequences and unexpected restarts, as early as (b)(6) 2022, but only reported after several more occurrences, on (b)(6) 2022. imc log entries indicate severe and persistent communication failure of powermod_1 component (pbm). the returned console troubleshooting revealed the issue was immediately reproduced, as (b)(6) was unable to complete the boot sequence due to communication failure. visual inspection of the console's components revealed corrosion on the pbm, caused by an electrolyte leak from supercaps c69, c70. the cause of the aic issue was pbm corrosion due to an electrolyte leak from supercaps c69,c70. this report is being filed as part of a retrospective review of historical records.
15703 | 1220648-2024-21997 | 20610421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20610421&pc=OZD | 10/10/2024 | 12/20/2024 |  | 11/05/2024 |  |  | 10/10/2024 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025548332 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 80-year-old male patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. while on support, bleeding was noted from the puncture site. extremely stiff vessels; 4-5 puncture attempts; bleeding reduction by means of manual compression and injection of adrenaline around the puncture site. hemostasis was not achieved. patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. corrections to the initial mfr report: d4 updated udi number. g1 mdr reporting contact email.
15704 | 1220648-2024-21998 | 20610496 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20610496&pc=OZD | 07/12/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 07/12/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1246897 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during a training the automated impella controller's (aic) purge disc holder broke off. no patient involvement. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, it was confirmed the purge disc retainer had broken/snapped off of the purge assembly. the purge disc retainer and purge flag were replaced, purge pressure sensor was validated, and a functional test was performed; before the console was returned to the customer. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15705 | 1220648-2024-22000 | 20610529 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20610529&pc=PYX | 10/12/2024 | 12/10/2024 |  | 11/05/2024 |  |  | 10/12/2024 | 11/05/2024 |  |  | 12/31/2024 |  | 11/05/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024320059 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella rp support. while on support, elevated pulmonary artery pressure (paps) was observed during a sedation vacation. automated impella controller (aic) showed motor current spike with simultaneous paps rise and decrease in flows. family decided to withdraw care and patient eventually expired. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation into low pump flow has been completed. the product was not returned. about 5 hours into the case there are motor current spikes with a correlating spike in placement signal. the cvp goes negative signaling the pump is in suction conditions. the flow decreases during this time. throughout the case there are more instances of suction and changes in flows. the root cause of the low pump flow is most likely due to biomaterial. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-22000. b1 should have been both adverse event and product problem. g1 reporting contact email. g1 reporting contact fax number missing.
15706 | 1220648-2024-21999 | 20610565 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20610565&pc=OZD | 10/10/2024 | 12/12/2024 |  | 11/05/2024 |  |  | 10/10/2024 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025533785 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, the patient had a slight ooze at the access site. the staff monitored the ooze. additionally, the patient had hemolysis while on support. intervention for the hemolysis is unknown. the family elected to withdraw patient care. the patient was on impella support for 47.11 hours before the patient expired. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The pump was not returned for investigation. data log shows the placement signal trending during hemolysis event, without observed motor current spikes, positioning issues, or suction alarms. according to the clinical team, the patient required a fluid bolus. additionally, a small groin bleed was reported, and heparin was paused for a few hours. the cause of the issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the hemolysis issue was most likely patient condition related, based on given clinical details and data log review. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-21999. g1 reporting contact email. h6 health effect - clinical code 1888 and health effect - impact code 4641 were missing.
15707 | 1220648-2024-22003 | 20610568 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20610568&pc=OZD | 10/10/2024 | 12/11/2024 |  | 11/05/2024 |  |  | 10/10/2024 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025509595 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 80-year-old male patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. during impella support and upon exchange of the peelaway and repositioning sheath, a large hematoma formed. the patient was taken to the operating room for femoral cutdown repair. additionally, the patient developed hemolysis. the patient lost pulsatility on impella and could not flow without any suction alarms, so, impella weaned to p2. code called for pea arrest. the patient received multiple rounds of cardio-pulmonary resuscitation (cpr). patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the hemolysis and the access site bleed has been completed. the pump was not returned for investigation. it was utilized for 8 hours. data logs show the trending placement signal and motor current, with observed low pump flow but no major suction or positioning alarms. according to the clinical team, vascular surgery successfully stopped the bleeding at the access site, which was reported during the placement of the repositioning sheath. also, regarding concerns over hemolysis, the doctor noted a good pump position with no alarms at the end of the case. the cause of the vascular damage and access site bleeding issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the hemolysis issue was most likely patient condition related, based on data log review and given clinical details. added h6 impact code 4624 corrections to the initial mfr report: g1 mdr reporting contact email.
15708 | 1220648-2024-22002 | 20610585 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20610585&pc=OZD | 10/08/2024 | 12/09/2024 |  | 11/05/2024 |  |  | 10/08/2024 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025551927 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | Tachycardia | [9 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 50-year-old male patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had decreased flow to right leg where impella was placed. external fem-fem bypass was placed to increase flow to distal extremity. patient developed hemolysis. echo showed impella was deep with patient¿s left ventricle being on the smaller side. doctor pulled back a couple centimeters, but did not want to pull back any further. additionally, the patient was bleeding from mouth, so, heparin was withheld. family made the decision to withdraw care. patient experienced a sustained run of ventricular tachycardia (vt) after the decision to withdraw care and the impella was turned off and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the ventricular tachycardia (vt), limb ischemia, hemolysis, and vascular damage has been completed. clinical details suggest that the pump was positioned deep in the lv, but this was due in part to the patient's anatomy being relatively small. data logs were reviewed and there were no positioning alarms. there were 3 suction alarms, but this is not consistent enough to suggest a volume issue. no other abnormalities were noted. product was not returned for investigation. based on data log review and clinical details, the root cause of the hemolysis was determined to most likely be improper positioning due to patient anatomy. the root cause of the limb ischemia, vt, and vascular damage could not be determined without additional details. added- h6 clinical code 1888 corrections to initial mfr report: g1-mdr reporting contact email.
15709 | 1220648-2024-22001 | 20610611 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20610611&pc=OZD | 10/07/2024 | 11/21/2024 |  | 11/05/2024 |  |  | 10/07/2024 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025551910 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 74-year-old female patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. during impella support, the patient had access site bleeding. the patient was transferred and upon arrival, the patient was not positioned correctly and was wrapped tightly. the site was redressed and angle matching was performed. two units of blood products were given to the patient. the patient expired after 23.03 hours of impella support. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation for the impella device has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the most likely root cause. the root cause of the access site bleeding issue was most likely use related, dressing change and angle matching done to achieve hemostasis as per the clinical description provided. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
15710 | 1220648-2024-22004 | 20610647 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20610647&pc=OZD | 10/11/2024 | 12/10/2024 |  | 11/05/2024 |  |  | 10/11/2024 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025524431 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and there were positioning issues. staff attempted to reposition but it was unsuccessful. the pump was explanted and there was no reported patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation into positioning issues has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-22004.
15711 | 1220648-2024-22005 | 20610676 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20610676&pc=OZD | 10/11/2024 | 12/12/2024 |  | 11/05/2024 |  |  | 10/11/2024 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025515054 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support. while on support, bleeding was noted at the access site. the jackson pratt drain had been emptied multiple times. the patient was given one unit of red blood cells. additionally, the patient suddenly became somnolent/unresponsive. a computed tomography scan of the head revealed multiple recent infarcts. per bedside nurse, patient had ¿mini strokes¿. the patient remains on support therefore explant date and patient outcome are unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the reported stroke and access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke access site bleed could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email.
15712 | 1220648-2024-22007 | 20610688 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20610688&pc=OZD | 10/13/2024 | 12/12/2024 |  | 11/05/2024 |  |  | 10/13/2024 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025533080 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemolysis | Hematuria | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, clot was removed from middle cerebral artery. patient showed signs of stroke. impella repositioned slightly bedside due to significant hematuria. waveforms and flows looked appropriate, but cardiologist believed patient has hemolysis. prognosis was poor. care was withdrawn and the patient eventually expired. | [5 empty] | The investigation for the hemolysis and stroke has been completed. the impella was not returned for evaluation. the cause of the hemolysis issue most likely patient condition based on the clinical and log review. data logs were reviewed finding no motor current (mc) spikes outside of p-level changes observed. no positioning issue observed. no suctions occurred. no indication for rc regarding the reported issue. the root cause of the stroke could not be determined without additional clinical details. corrections to the initial mfr report: b2-selected death and added date of death. g1 mdr reporting contact email. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
15713 | 1220648-2024-22006 | 20610740 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20610740&pc=OZD | 10/12/2024 | 12/27/2024 |  | 11/05/2024 |  |  | 10/12/2024 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025535854 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 87 YR |  | Female |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella cp support. while on support, bleeding was noted at the groin with hemorrhagic shock. the pump was explanted and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
15714 | 1220648-2024-22008 | 20610784 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20610784&pc=OZD | 10/14/2024 | 12/13/2024 |  | 11/05/2024 |  |  | 10/14/2024 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025555927 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support and during support, there was pain at the insertion site. it was swollen and slightly hard. additionally, there was a hematoma at the access site. heparin was stopped due to the concern of the hematoma. the patient remains on impella support. | [5 empty] | The investigation into access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-22008 e4 should have been left blank g1 reporting contact email g1 reporting contact fax number missing. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
15715 | 1220648-2024-22010 | 20610785 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20610785&pc=OZD | 06/13/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 06/13/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1469546 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a controller error. there was no patient harm reported. | [5 empty] | The investigation into the controller error/loss of opm data issue has been completed. the impella automated controller was returned for investigation. the data analysis of the logs confirmed the reported failure mode given there was a controller error alarms (opm comm crc error ¿ event set # 1045) triggered during the clinical procedure. real time (rt) logs confirmed that this was a pattern failure mode in which all signals received from optical bench (placement, snr, contrast, lamp) are interpreted as -16384 values. the returned console was tested and the known pattern failure, which is characterized by a temporary loss of optical data (placement, snr, contrast, lamp, gain) and triggering of a controller error alarm, has a low probability of reoccurrence. if needed, patient hemodynamics and impella position can be monitored with imaging. the cause of the transient loss of optical data was not determined due to the inability to reproduce. the device did not meet specifications.this report is being filed as part of a retrospective review of historical records.
15716 | 1220648-2024-22009 | 20610848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20610848&pc=OZD | 07/07/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 07/07/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1464036 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) exhibited an error code stating the latch for the purge cassette was not connected. a loaner device was sent to the customer. no report of patient involvement, harm, or injury. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, the reported damage on purge system was visually confirmed. the purge flag was confirmed to be missing at blue disc retainer and the stand for blue retainer was broken. the root cause of the reported error code was missing purge flag and broken blue retainer. this report is being filed as part of a retrospective review of historical records.
15717 | 1220648-2024-22011 | 20610871 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20610871&pc=OZD | 10/10/2024 | 12/13/2024 |  | 11/05/2024 |  |  | 10/11/2024 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025558142 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 72-year-old male patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. during impella support, the patient had limb ischemia. the patient limb was cold and mottled. a fasciotomy was performed, and the patient also received a femoral bypass. the patient¿s leg is now warm. | [5 empty] | The investigation into limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-22011. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
15718 | 1220648-2024-22012 | 20610929 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20610929&pc=OZD | 07/07/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 07/07/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2023194841 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported an aic controller emitting a "controller error" alarm during support. a backup console was brought into the room, however, issue resolved before the need to switch consoles. there was reported harm to the patient. | [5 empty] | The investigation into the controller failure (red alarm) issue has been completed. the automated impella controller (aic) was returned for investigation. console logs from the reported day of event are consistent with complaint and show two instances of controller failure alarms due to purge pressure sensor defect occurring during two out several purge changes. all other purge changes, throughout the case, were successful. data revealed, that during the alarm conditions, values of measured purge pressure were negative (approx. -15mmhg). during testing in (b)(6) (purge cassette and purge bag change) the issue was not reproduced, however observations of non-averaged values of measured purge pressure during purge changes showed small negative values (-1 to -2). same measurements conducted on a known-good engineering console never showed values lower than 0. although the purge pressure transducer was verified to be within specification as per 0042-7909 sec. 12, it¿s been decided that in abundance of caution the pud pca and the transducer should be replaced, as both components could have contributed to this malfunction. because the purge disk retainer was also found to be cracked (unrelated to this complaint), we requested for the entire purge assembly to be replaced. repair: replace purge assembly and perform functional check. this report is being filed as part of a retrospective review of historical records.
15719 | 1220648-2024-22015 | 20610981 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20610981&pc=OZD | 07/06/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 07/06/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1251612 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on an automated impella controller (aic) for mechanical circulatory support. during setup, optical sensor failure screen appeared, the white connector cable disengaged from the aic and optical light was observed. it was reported that the pump purged well and was functional. a replacement aic was not prepped, and the impella cp was placed without a placement signal. the pump did have the appropriate motor current. then the "controller error" alarm appeared with advisory to switch aics. the replacement aic was booted up and the switch was made without any complications. the procedure was successful, no reported harm or injury to the patient. | [5 empty] | The investigation into the aic - optical signal issue has been completed. the automated impella controller (aic) was returned for investigation. data analysis: the device logs from the complaint occurred date were consistent with reported content. during the initial setup, the pump was connected, however the optical sensor system error was alarmed because the pump did not push all the way. after the pump was reconnected and placed, the appropriate mc was observed but without ps. later, the reported alarm was triggered, controller error (opm comm stopped) [optical pump connected] - alarm issued. also, clinical staff decided to switch aic but the console was unable to be shut down because console determined that pump was still connected. device analysis: the reported failure mode was not reproduced with test pump and pc. blue receptacle for pump was ensured for no significant contamination. optical 5s parameters were measured to be within specifications. no problem was found on console. the reported controller error (event set #1043) is likely result of optical placement signal loss, due to opm communication corruption. this is a known pattern failure, abiomed is aware of the issue, and is working on appropriate corrective action. issue is resolved after aic is power cycled. in a clinical setting - if needed, monitor patient hemodynamics and impella position with imaging. the speculation matches the facts that: 1. no valid ps was observed on monitor after pump was reinserted and placed, 2. pump disconnection from console was unable to be detected . although the speculation could not be confirmed due to missing of rtlogs, the functionalities of console ic4526 was ensured by in-lab testing. this report is being filed as part of a retrospective review of historical records.
15720 | 1220648-2024-22016 | 20611003 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20611003&pc=OZD | 06/08/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 06/08/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1595838 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was on automated impella controller (aic) for mechanical circulatory support. while on support, the aic experienced a/c power issues that were resolved by replacing the console. there was no patient harm reported. | [5 empty] | The investigation into the a/c power issue has been completed. the automated impella controller (aic) was returned for investigation. the cause of ac power issues was due to damaged ac power cord. biomedical field team did not take image of damage and repaired it. cord was inspected, and console was tested for electrical safety. this report is being filed as part of a retrospective review of historical records.
15721 | 1220648-2024-22013 | 20611018 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20611018&pc=OZD | 10/10/2024 | 12/10/2024 |  | 11/05/2024 |  |  | 10/10/2024 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025545506 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Increase in Pressure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 76-year-old female patient in an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during impella support, there was high purge pressure reported. tissue plasminogen activator was added to the purge. the staff monitored. additionally, it was noted that the patient was bleeding from all sites. five units of blood products were given to the patient. furthermore, the patient had renal failure on support. it is unknown if the impella contributed or caused the renal failure. the patient was on impella support for 141.58 hours before the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the renal failure and the access site bleed has been completed. the pump was not returned for investigation. the pump was on bipella support. data logs show no signs related to improper positioning or biomaterial interaction. according to the clinical report, the patient underwent crrt after impella placement, and later, after 5 days, bleeding was reported from all sites. additionally, high purge pressure hpp was reported on the rp device, and it will be addressed under 24-37981 and removed as an investigated failure mode here. the cause of the access site bleeding issue was patient condition related since bleeding was reported from multiple access site. the cause of the renal failure issue was not determined since product was not returned and sufficient clinical information was not provided. corrections to the initial mfr report: g1 mdr reporting contact email.
15722 | 1220648-2024-22020 | 20611105 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20611105&pc=OZD | 06/06/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 06/06/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1145121 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant notified the clinical consultant that the automated impella controller (aic) generated a battery failure alarm. there was no patient harm reported. | [5 empty] | The investigation into the battery failure issue has been completed. the impella automated controller was returned for investigation. a review of the logs noted consistency with the complaint intake. the long-term log analysis of the battery data reveals that beginning before the complaint date of occurrence, cell 1 voltage of battery 2 had been declining for an extended period of time. it was confirmed that the persistent battery failure (transport connection blown/missing) was not due to a disengaged battery switch. when the console was booted for the first time during bench testing console alarms were consistent with what was seen in the field. the battery failure alarm was due to a defective battery 2 (decline of individual cell voltage). the root cause of the battery cell failure was undetermined. the cause of the aic issue was a defective battery. this report is being filed as part of a retrospective review of historical records.
15723 | 1220648-2024-22014 | 20611136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20611136&pc=OZD | 06/06/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 06/06/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1496398 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant notified the clinical consultant that the automated impella controller (aic) generated a battery failure red alarm. a loaner was sent to the customer to initiate return of this device. there was no patient harm reported. | [5 empty] | The investigation into the battery failure issue has been completed. the impella automated controller was returned for investigation. the long-term log was reviewed and noted that prior to the complaint date of occurrence, cell 1 voltage of battery 1 had been declining for an extended period of time. it was confirmed that the persistent battery failure (transport connection blown/missing) was not due to a disengaged battery switch. during bench testing when the console was booted for the first time, console alarms were consistent with what was seen in the field. the battery failure was due to a defective battery 1 (decline of individual cell voltage). the cause of the aic issue was a defective battery. this report is being filed as part of a retrospective review of historical records.
15724 | 1220648-2024-22018 | 20611138 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20611138&pc=OZD | 10/11/2024 | 12/11/2024 |  | 11/05/2024 |  |  | 10/11/2024 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025536831 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and there were bleeding concerns at the impella access site. the patient was taken back to the operating room to explore the pocket and clean out the site. the physician felt that pocket wasn¿t properly closed. support continued following the events. the patient remained stable until the successful wean and explant. | [5 empty] | The investigation into access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding is determined to be use issue because pocket wasn¿t properly closed by the fellow during the procedure yesterday evening. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-22018. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
15725 | 1220648-2024-22017 | 20611205 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20611205&pc=OZD | 05/27/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 05/27/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1117656 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant notified the field service logistics coordinator that the automated impella controller (aic) generated a battery failure alarm. a loaner was sent to the customer to initiate return of this device. there was no patient harm reported. | [5 empty] | The investigation into the battery failure issue has been completed. the impella automated controller (aic) was returned for investigation. the staff reported that a battery failure occurred on 5/27/22. the logs show no occurrence of a battery failure. however, on 5/21 and 5/24 separate instances of "battery level low" occurred. the log was reviewed and noted both instances of battery level low triggering were preceded by the same sau_powermod_1 communication errors. the long-term logs were reviewed at the time of the battery level low alarms, it could be seen that the battery remaining capacity parameter was well above the triggering threshold of "battery level low", except for one sample of data, occurring twice. it was noted upon return of the console when powered on, the console was nearly fully charged. the console was powered on, and a telnet batttest was performed, observing that the console batteries were appropriately discharging while the console was powered on idly. ic1968 was left with pump and purge intentionally disconnected from ac power and allowed to discharge down to 50% battery according to the aic screen. as expected, the console alarmed for "battery level low". the console was then plugged into ac power and allowed to charge. a batttest was performed during charging, where it was observed to have appropriate charging current. it is most likely that the cause of the two false battery level low alarms are due to a single sample of communication that was misinterpreted from the batteries by the pbm. the root cause of the miscommunication between the battery and pbm is undetermined. the console did not alarm for a battery failure. battery level low was what actually triggered. it most likely triggered due to the constant plugging and unplugging of ac power. never allowing the console to charge up fully. the root cause was not determined. this report is being filed as part of a retrospective review of historical records.
15726 | 1220648-2024-22042 | 20611255 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20611255&pc=OZD | 05/11/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 05/11/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1126094 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  | Non Hispanic | White | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported an aic turning off on its own. no harm to patient was reported. | [5 empty] | The investigation into the aic -shutdown or restart issue has been completed. the automated impella controller (aic) was returned for investigation. console logs from the reported day of event are consistent with complaint and show multiple resets of the system in response to calculator process crashes (as per design, aic software will perform soft-reset, and - if not effective - then restart the console when a software process is found unresponsive). after impella 5.5 pump 376627 was first plugged in, initialized, and started around 10:23am on (b)(6) 2022, it's been unexpectedly unplugged (while running) around 10:29am, despite valid placement signal and motor current. when pump was re-connected after 10 seconds, it was recognized but there was an issue while reading flow characteristic curves segment of pump eeprom. few seconds later, communication channel between scu (system control unit process) and calculator process started failing, followed by a crash of said process. in response, aic sw watchdog triggered partial reset, and soon after a restart of the console, which was perceived by the operators as "aic lost power". after the restart, the pump (that was still connected to aic) was detected and recognized but eeprom reading issues (calibration data segment) caused "impella defective" alarm (#3), which presumably prompted the operators to unplug the pump again for 3 seconds. after pump was reconnected, eeprom reading issues reoccurred (flow characteristic curves segment of pump eeprom), but pump was able to be started. couple seconds later, the calculator process crashed again, resulting in another soft-reset, followed by aic restart. when console restarted for the second time, although the pump was still connected, it had not been recognized as "previously running" (initialized). the pump had been unplugged and console shut down. there were no reports of any issues using pump 376627 on the backup console ic8353. after console ic2165 was received in (b)(6) and subject to testing, the issue was not reproduced. there were no observed eeprom reading issues, or calculator process (or any other processes) crashes. issues during eeprom reading and calculator process crashes could have been related, but the exact root cause couldn't not be determined. there's a number of possible failure mechanisms and components responsible for the malfunctions: software corruption (leading to crc errors) due to either cf cards corruption or 4-in-1 defect, hardware malfunction (undetermined 4-in-1 issue causing calculator process crashes), or epm circuitry malfunction on impellatronic. this report is being filed as part of a retrospective review of historical records.
15727 | 1220648-2024-22022 | 20611264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20611264&pc=OZD | 05/27/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 05/27/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1139666 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant notified clinical support that the automated impella controller (aic) generated a battery failure alarm, despite being plugged in and being fully charged. a loaner was sent to the customer to initiate return of this device. there was no patient harm reported. | [5 empty] | The investigation into the battery failure issue has been completed. the impella automated controller was returned for investigation. the data analysis was consistent with the reported symptom, upon boot-up of the console the red alarm generated. through the two on-site power cycles, no ¿ac power disconnected¿ alarm was issued, which indicated the console was connected to ac power. the returned console battery switch was engaged, and the console was powered on via batteries. no alarms occurred. a battest was performed, where no issues were found. both batteries were inspected and denied for potential battery failure. the console had expected behavior both on ac power and battery power with no issues. it is likely that the console did not have the battery switch engaged upon 1st boot up and the failure was resolved by engaging the battery switch on bottom of the console. the battery failure (red alarm) was confirmed via logs, but not reproduced. the battery failure red alarm was most likely due to disengagement of battery switch on the bottom of console. the cause of the issue was use related due to improper technique by the end user. this report is being filed as part of a retrospective review of historical records.
15728 | 1220648-2024-22023 | 20611297 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20611297&pc=OZD | 09/14/2024 | 12/12/2024 |  | 11/05/2024 |  |  | 10/13/2024 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025494457 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 58 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured right iliac artery stenosis with a possible relationship to the impella device used: ¿angiography post tavr which reported stenosis of right iliac artery as well as findings of possible external iliac thrombus versus dissection -s/p balloon inflation at the site with improvement of iliac artery stenosis, persistent lucency at external iliac artery -currently asymptomatic apart from pain/tenderness to groin si abi's & arterial duplex completed, reviewed by dr > no significant stenosis iliofemoral angiography demonstrated small amount of extravasation, stenosis at the level of the perclose site, and a non-flow limiting linear density in the external iliac concerning either for thrombus or dissection. thus, a 6.0 armada balloon was advanced to the area of stenosis and was inflated for 2 minutes. subsequent angiogram showed cessation of bleeding, along with improvement in the degree of stenosis. there was still the area of echolucency in the external iliac which was not ballooned due to concerns of distal embolization. thus, as it was not flow-limiting, 5,000 units of heparin was administered and heparin gtt was started. arterial doppler of the rle showed excellent dp and pt pulses, and thus no further intervention was performed.¿ the patient recovered with sequelae. the patient survived the event. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).
15729 | 1220648-2024-22024 | 20611325 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20611325&pc=OZD | 05/28/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 05/31/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1145434 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Nurse | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant, a nurse, called clinical support to report that the automated impella controller (aic) generated a battery failure alarm. the nurse reported that two days prior a case was going to be done and this console they were about to use showed battery failure. the nurse reported they used a different console for the case with no issues. reportedly after two days of charging, the system continued to receive the battery failure message. there was no patient harm reported, another console was used upon the failure of this device. | [5 empty] | The investigation into the battery failure issue has been completed. the impella automated controller was returned for investigation. the data analysis of the logs revealed around the time the complaint occurred date ((b)(6) 2022) was consistent with the reported symptom. upon boot-up of the console, the red alarm was issued and the console shut down. during the on-site power cycles, no ¿ac power disconnected¿ alarm was issued, which indicated the console was connected to ac power when the console was turned on. a review of the long-term logs showed descending battery voltages (from 5/28 to 5/31), which is contradictory to the statement of ¿2 days of charging¿ according to clinical staff. it was speculated that the console transport switch was not switched on during clinical use. without the battery switch engaged, the console would operate/ turn on with ac power, but unplugging ac would cause unexpected controller shutdown. furthermore, the console was not receiving charging even with ac cord plugged in for days. the returned pump was inspected, the battery switch was engaged during console check-in, and the console was powered on via batteries and ac power supplied. no alarms occurred. a batttest was performed, where no issues on batteries were found. the console had expected behavior both on ac power and battery power with no issues. it is likely that the console did not have the battery switch engaged in the field, and the failure was resolved by engaging the battery switch on the bottom of the console. the cause of the issue was most likely user technique related by the end user as a result of the battery switch not being engaged. this report is being filed as part of a retrospective review of historical records.
15730 | 1220648-2024-22028 | 20611355 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20611355&pc=OZD | 05/18/2022 | 12/11/2025 |  | 11/05/2024 |  |  | 05/18/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1414254 |  | 0042-0010-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [11 empty] | Other | [5 empty] | IMPELLA CP | [18 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported patient was placed on impella cp for hemodynamic support. while on support, the automated impella controller (aic) experienced battery charging/other issues. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the issue was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records. | Console logs from the reported day of event are consistent with complaint and show a single battery temperature high alarm (#49) that resolved in 10s, which was preceded by aic being plugged into ac power. ltlog data showed that during that time, battery temperature was normal, and never exceeded alarm threshold. it needs to be noted that ltlog data points are recorded every 1 minute, however battery parameters are monitored every 10 seconds. it's been observed in the past, in prior similar investigations, that a single erroneous sample of battery temperature would trigger alarm that would self-resolve once the next, correct, data sample is received. this is most likely due to a rare, momentary communication glitch caused by current/voltage spike when pbm is switching between battery power and external power (ac-dc power supply). there is no corrective action available, and no repair neede. console will remain at the site and is considered safe to use. corrected information provided in b1, b2, d2 common device name updated, d10 concomitant medical products and therapy dates updated, e4 should have been left blank on the initial report, g1 reporting contact email updated and h6.
15731 | 1220648-2024-22027 | 20611380 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20611380&pc=OZD | 06/27/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 06/27/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1159474 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller's (aic) purge clip was broken. no report of patient involvement, injury, or harm. a loaner device was sent to the customer. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, it was confirmed the purge disc retainer had broken/snapped off of the purge assembly. therefore, the root cause of the broken purge clip was due to a defective component. purge disc retainer and purge flag was replaced, the purge pressure sensor was validated, and the functional check was performed; before the console was returned to the customer. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15732 | 1220648-2024-22021 | 20611462 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20611462&pc=OZD | 07/05/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 07/05/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1399060 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller's (aic) purge disk insert was broken off. no patient involvement, harm, or injury reported. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, it was confirmed the purge disc retainer had broken/snapped off of the purge assembly. therefore, the root cause of the purge disc broken was due to a defective component. the purge disc retainer and purge flag were replaced, purge pressure sensor was validated, and a functional check was performed. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15733 | 1220648-2024-22025 | 20611471 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20611471&pc=OZD | 06/27/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 06/27/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1108092 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller's (aic) purge clip was broken. a loaner was sent to the customer. no report of patient involvement, injury, or harm. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, the reported broken purge clip was not confirmed, which contradicts the complaint description. the entire purge system was fully intact and functional. there were no issues priming a known good purge cassette. test were run to verify the purge system is within specification and perform a full functional check on the console and optical system; before returning the console to the customer. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15734 | 1220648-2024-22019 | 20611537 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20611537&pc=OZD | 06/17/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 06/17/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1105125 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The field service engineer reported that the automated impella controller (aic) had a broken touch button. a loaner was sent to the customer to initiate return of this device. there was no patient harm reported. | [5 empty] | E1: name of the initial reporter is unknown . the investigation into the keyboard rotary knob issue has been completed. the impella automated controller (aic) was returned for investigation. the data analysis of the logs from before the reported complaint date revealed that user had issues properly shutting down the console. the returned console was inspected on an engineering lab bench. the front enclosure was damaged where the rotary push button mounts. the hole for this component was countersunk too deep, which compromised the integrity of the plastic enclosure and allowed it break with less force. the root cause of the broken rotary encoder was due to normal use, but it was likely made weaker during servicing of the console to secure the encoder. this report is being filed as part of a retrospective review of historical records.
15735 | 1220648-2024-22029 | 20611539 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20611539&pc=OZD | 06/23/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 06/23/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1302150 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | Failure to Read Input Signal | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during a case set up with the automated impella controller (aic). a tear was noticed in the rubber near the purge disc and would not recognize the disc. a loaner device was sent to the customer. no report of patient involvement, injury, or harm. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, it was confirmed there was a broken blue retainer. the broken blue retainer was the root cause of reported failure mode, purge disc not detected. the purge disc retainer was replaced, and a functional check was performed; before the console was returned to the customer. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15736 | 1220648-2024-22039 | 20611555 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20611555&pc=OZD | 06/10/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 06/10/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1135098 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The clinical consultant reported that the automated impella controller (aic) experienced a controller error. it was reported that the error occurred on startup on two separate occasions and could not be resolved. there was no patient harm reported. | [5 empty] | E1: name of the initial reporter is unknown. the investigation into the controller error/loss of opm data issue has been completed. the impella automated controller was returned for investigation. the data analysis of the logs from the reported day of event are consistent with the complaint and show multiple controller error alarms due to a defective optical sensor lamp (1039) and ambient pressure out of range (131) on 2022-05-30 and 2022-06-10. on each day the console had been power-cycled, but the issue remained unresolved. the returned console was investigated the issue was reproduced. and on each boot-up the console's optical bench 5 signal parameters were out of specification. the invalid value of ambient pressure indicated failure of the optical bench electronics, rather than only failure of lamp assembly. when the optical bench was temporarily replaced with a known-good one, the issues was resolved. the cause of the aic issue was a defective optical bench. this report is being filed as part of a retrospective review of historical records.
15737 | 1220648-2024-22031 | 20611679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20611679&pc=OZD | 05/19/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 05/19/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1089665 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The clinical support coordinator reported that the automated impella controller (aic) generated a battery failure alarm during start up, when unplugged from the wall and turned on. the facility reportedly used another aic. there was no patient harm reported. | [5 empty] | The investigation into the battery failure issue has been completed. the impella automated controller (aic) was returned for investigation. the data analysis of the logs revealed a gap in time when the console was not powered on (~10 months). the batteries were depleted because the console was most likely not routinely charged as recommended by the end user. the investigation verified the cause of the aic issue was a depleted battery. this report is being filed as part of a retrospective review of historical records.
15738 | 1220648-2024-22036 | 20611723 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20611723&pc=OZD | 06/23/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 06/23/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1168795 |  | 0042-0000-US-L |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that they had a loaner aic in hospital when the purge disc was observed to be broken. another controller was used for support. there was no reported harm or injury to the patient. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, it was confirmed the purge disc retainer had broken/snapped off of the purge assembly. the purge disc retainer and purge flag was replaced, the purge pressure sensor was validated, and the functional check was performed; before the console was returned to the customer. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15739 | 1220648-2024-22051 | 20611731 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20611731&pc=OZD | 05/12/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 05/12/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1133727 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | Us complainant reported patient was placed on impella cp for hemodynamic support. while on support, the automated impella controller (aic) experienced purge issues. aic console changed and issue was resolved. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken (missing at blue disc retainer). before returning this console back to the customer, the purge flag was replaced and a full functional check on the console and optical system was performed. this report is being filed as part of a retrospective review of historical records.
15740 | 1220648-2024-22056 | 20611746 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20611746&pc=OZD | 04/13/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 04/13/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1084300 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The clinical coordinator reported that the automated impella controller (aic) generated a battery failure alarm. it was reported the battery switch was checked and toggled on and off multiple times with no resolution. there was no patient harm reported. | [5 empty] | The investigation into the battery failure issue has been completed. the impella automated controller (aic) was returned for investigation. the data analysis of the logs noted the results were consistent with the complaint and there were subsequent boots following the complaint date triggering these alarms. the long-term log analysis of the battery data revealed that beginning before the complaint date of occurrence, cell 3 voltage of battery 1 had been declining for an extended period of time. bench testing of the returned device indicated when the console was booted for the first time, console alarms were consistent with what was seen in the field. the battery failure alarm was due to a defective battery 1 (decline of individual cell voltage). the root cause of the battery cell failure was undetermined. the cause of the battery failure alarm is most likely due to single cell decline in a defective battery. this report is being filed as part of a retrospective review of historical records.
15741 | 1220648-2024-22054 | 20611794 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20611794&pc=OZD | 05/11/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 05/11/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1220358 |  | 0042-0000-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that automated impella controller (aic) console failed to recognize two different impella pumps during a preventive maintenance. there was no patient involved or impacted. | [5 empty] | E1: name of initial reporter is unknown. d9 - date device returned to manufacturer unknown. the investigation into the aic -pump not detected issue has been completed. the automated impella controller (aic) was not returned for investigation. imc logs are consistent with the complaint. through the case start wizard, pump data failed to be read once pump was connected. series of crc errors occurred, and an impella defective (error reading eeprom) alarm was issued every time pump was plugged in. the attempt of power-cycling failed to resolve the issue. console was fully investigated and tested on an engineer bench. the failure mode that pump was unable to be read and the yellow alarm, impella defective, was reproduced. impellatronic board was temporarily replaced with a known-good one and the issue was resolved after power cycling. the root cause of the yellow alarm was due to a defective impellatronic board. this report is being filed as part of a retrospective review of historical records.
15742 | 1220648-2024-22026 | 20611830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20611830&pc=OZD | 05/19/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 05/19/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1236850 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The field service logistics coordinator reported that the automated impella controller (aic) generated a battery failure alarm upon boot up. it was reported the console had been plugged in and the battery was turned on via the bottom of the aic. a loaner was sent to the customer to initiate return of this device. there was no patient harm reported. | [5 empty] | The investigation into the battery failure issue has been completed. the impella automated controller (aic) was returned for investigation. the data analysis of the logs revealed consistency with the complaint intake showing subsequent boots following the complaint date and triggering of these alarms. the long-term log analysis of the battery data did not contain enough data points, however, a telnet batttest revealed that cell 1 of battery 2 had drifted ~78mv from cell 3, tripping the cell imbalance mechanism within the battery. when the returned console was booted for the first time during bench testing the console alarms were consistent with what was seen in the field. when visually inspecting the batteries, it was observed that one of the battery status light-emitting diodes were completely blank. the battery failure alarm was due to a defective battery 2 (decline of individual cell voltage). the root cause of the battery cell failure was undetermined. the cause of the aic battery failure was due to the decline of single cell within battery 2. this report is being filed as part of a retrospective review of historical records.
15743 | 1220648-2024-22047 | 20611846 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20611846&pc=OZD | 04/25/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 04/25/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1573119 |  | 0042-0010-EU |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant had a battery failure. the battery is backup to the a/c power. there was no harm or injury and the team followed instructions for use and used a backup controller. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the aic issue was a defective battery. this report is being filed as part of a retrospective review of historical records.
15744 | 1220648-2024-22032 | 20611857 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20611857&pc=OZD | 10/11/2024 | 12/13/2024 |  | 11/05/2024 |  |  | 10/11/2024 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025561492 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that impella cp placement was aborted after impella did not cross at bifurcation of aorta. short and long sheath both attempted for placement but unsuccessful. torqued and buddy wire was also attempted to straighten out vessel. the impella was removed and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation into delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related since there was bifurcation of aorta. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-22032.
15745 | 1220648-2024-22068 | 20611877 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20611877&pc=OZD | 04/29/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 04/29/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1117231 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. while on support, the aic experienced a controller error yellow alarm that self-resolved. there was no report of patient harm or injury. | [5 empty] | The investigation into the aic -controller error issue has been completed. the automated impella controller (aic) was returned for investigation. data analysis: imc logs revealed the reported ¿impella position in aorta¿ alarms for multiple times but it resolved instantaneously. according to the description that pump was actually too deep, these alarms were not considered as issues. imc logs confirmed the reported failure mode given there was a controller error alarms (opm comm crc error ¿ event set # 1045) triggered during the clinical procedure. real time (rt) logs confirmed that this was a pattern failure mode in which all signals received from optical bench (placement, snr, contrast, lamp, gain) are interpreted as -16384 values. device analysis: console (ic1950) was sent back field service (fs) for further investigation. this known pattern failure, which is characterized by a temporary loss of optical data (placement, snr, contrast, lamp, gain) and triggering of a controller error alarm, has a low probability of reoccurrence. if needed, patient hemodynamics and impella position can be monitored with imaging. before being returned to the customer, fs performed a full functional check on the console and optical system. this report is being filed as part of a retrospective review of historical records.
15746 | 1220648-2024-22034 | 20611890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20611890&pc=OZD | 05/12/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 05/12/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1216754 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. near the conclusion of the case, the automated impella controller (aic) produced a controller failure alarm and stopped suddenly. attempts to restart the pump were unsuccessful, and the console was replaced. there was no reported patient harm. | [5 empty] | The device was returned and evaluated, and the investigation for the reported pump stop has been completed. console logs from the reported day of event are consistent with complaint and show that the pump was successfully started and ran without any issues for about 2 hours, a sudden pump stop occurred, as evidenced by alarm #6 (pump stop due to motor driver guard error or its supply voltage out of spec), preceded by error messages. two consecutive auto-restarts were unsuccessful and the alarm condition persisted. rtlog data shows that there were no motor current irregularities before the first pump stop. after the console was brought to the lab for testing, shortly after the console was powered on, there was a beeping sound from the impellatronic, indicating possible internal communication issue (between pmd and cpld). pbm was verified to be providing sufficient power to the epm and pmd sub-circuitries (5v and 20v respectively), but the beeping was persistent and did not resolve after aic was power-cycled. after impellatronic assembly was temporarily replaced with a know-good one, the issue was resolved. the cause of the aic issue was a defective impellatronics board. this report is being filed as part of a retrospective review of historical records.
15747 | 1220648-2024-22045 | 20611894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20611894&pc=OZD | 05/02/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 05/02/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1156151 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The field service logistics coordinator reported that the automated impella controller (aic) generated a battery failure alarm. there was no patient harm reported. | [5 empty] | The investigation into the battery failure issue has been completed. the impella automated controller (aic) was returned for investigation. the data analysis of the logs showed the console had not been powered on for roughly 11 months. the return product was inspected and revealed, upon arrival, the batteries were confirmed to be depleted. the batteries were likely depleted because the console was not routinely charged as recommended. the cause of the aic issue was a depleted battery set due to the end user not routinely charging the device. this report is being filed as part of a retrospective review of historical records.
15748 | 1220648-2024-22035 | 20611899 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20611899&pc=OZD | 10/14/2024 | 12/13/2024 |  | 11/05/2024 |  |  | 10/14/2024 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025564562 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, bleeding was noted at the access site. a stent was placed. ischemia was noted at the left leg. external femorofemoral bypass with 5 french ante grade sheath was used to manage the ischemia. the patient survived the explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into access site bleeding and limb ischemia ¿ reversible has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was not determined since the product was not returned and insufficient clinical details provided. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. g.1 revised reporting contact email since initial manufacturer device report 1220648-2024-22035 was submitted.
15749 | 1220648-2024-22030 | 20611925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20611925&pc=OZD | 06/28/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 06/28/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1125045 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the aic (automated impella controller) had purge disc damage observed at a point of non-patient use. the ifu was followed and a backup controller used. there was no reported harm or injury to the patient. | [5 empty] | The automated impella controller (aic) was returned for evaluation. upon inspection of the console, it was confirmed a ripped/ missing blue retainer. the purge transducer was replaced and a functional check was conducted; before the console was returned to the customer. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15750 | 1220648-2024-22050 | 20611935 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20611935&pc=OZD | 05/13/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 05/13/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 180730 |  | 0042-0000-JP |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had an aic fail and alarm for replacement. there was no harm or injury and instructions for use (ifu) followed and a backup controller was used. no patient was involved. | [5 empty] | E1: name of initial reporter is unknown. d8/d9: date of device return is unknown. the investigation into the aic -system self check error issue has been completed. the automated impella controller (aic) was returned for investigation. data analysis: the aic logs showed that the console would not boot up properly due to communication issues with pbm1. device analysis: the failure mode was reproduced by field service (fs) japan and resolved after replacing the pbm. the defective pbm had corrosion on the top and bottom. before sending this console back to the customer, field service was instructed to replace the pbm ((b)(6)) and perform a full functional check on the console and optical system ((b)(6)). this report is being filed as part of a retrospective review of historical records.
15751 | 1220648-2024-22052 | 20612004 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20612004&pc=OZD | 04/11/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 04/11/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1194751 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant had an automated impella controller (aic) with a battery failure. the aic was removed and replaced by a backup per instructions for use recommendations. there was no lasting harm or injury and no patient use at the time. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the aic issue was a defective battery. this report is being filed as part of a retrospective review of historical records.
15752 | 1220648-2024-22033 | 20612051 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20612051&pc=OZD | 06/24/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 06/24/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | XXXXXX |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on an automated impella controller (aic) for mechanical circulatory support. complainant reported that one of their controllers failed, switched to a new one with no issues. staff noted that the placement signal and left ventricle signal disappeared and only showed dashes. the impella flow and purge numbers were normal. there was no reported harm or injury to the patient. | [5 empty] | The investigation into the aic - optical signal issue has been completed. the automated impella controller (aic) was returned for investigation. data analysis: imc logs confirmed the reported failure mode given there was a controller error alarms (opm comm loss ¿ event set # 1043) triggered during the clinical procedure. with the lack of optical data, the pump disconnection detection mechanism was impacted, and the user had to perform an emergency shutdown. real time (rt) logs confirmed that this was a pattern failure mode in which all signals received from optical bench (placement, snr, contrast, lamp, gain) are interpreted as -16384 values. device analysis: console (b)(6) was sent back field service (fs) for further investigation. this known pattern failure, which is characterized by a temporary loss (resolved after power cycle) of optical data (placement, snr, contrast, lamp, gain) and triggering of a controller error alarm, has a low probability of reoccurrence. if needed, patient hemodynamics and impella position can be monitored with imaging before being returned to the customer, fs performed a full functional check on the console and optical system (0042-7316) and a known-good optical pump and purge. this report is being filed as part of a retrospective review of historical records.
15753 | 1220648-2024-22089 | 20612075 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20612075&pc=PYX | 04/13/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 04/13/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1213653 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in pulmonary embolism was implanted with an impella rp device for mechanical circulatory support. during support, device had an automated impella controller (aic) purge disc failure. no patient harm was reported. | [5 empty] | The console (aic) was returned for evaluation. upon visual inspection of the purge assembly area confirmed purge flag was missing/ broken. before returning this console back to the customer, the purge flag was replaced , a full functional check on the console and optical system was performed, and preventive maintenance was performed. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15754 | 1220648-2024-22108 | 20612096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20612096&pc=OZD | 08/10/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 08/10/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1558681 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a critical low battery alarm after the console was powered on, the staff tried charging the battery for three hours and still received an alarm. the customer was sent a loner device. no patient involvement. | [5 empty] | The investigation for the reported aic - battery failure (red alarm) issues has been completed. the controller device has been returned for evaluation. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15755 | 1220648-2024-22059 | 20612117 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20612117&pc=OZD | 04/01/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 04/01/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | ABIOMED, INC. |  |  | ABIOMED, INC. |  |  | US |  | 1566931 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant notified the clinical consultant to report that the automated impella controller (aic) generated a battery failure alarm. it was reported that the battery was turned on and the device was power cycled on/off several times, however, the device would not charge more than 50 percent. a loaner was sent to the customer to initiate return of this device. there was no patient harm reported. | [5 empty] | The investigation into the battery failure has been completed. the impella automated controller was returned for investigation. the data analysis of the logs revealed consistency with the complaint intake of battery failure, subsequent boots following the complaint date triggered these alarms. the long-term log analysis of the battery data revealed that beginning before the date of complaint occurrence, cell 1 voltage of battery 2 had been declining for an extended period of time. it was confirmed that the persistent battery failure (transport connection blown/missing) was not due to a disengaged battery switch. during bench testing when the returned console was booted for the first time, console alarms were consistent with what was seen in the field. when visually inspecting the batteries and power battery manager circuit board, it was observed that one of the batteries status light-emitting diodes was completely blank. the battery failure alarm was due to a defective battery 1 (decline of individual cell voltage). the root cause of the battery cell failure was undetermined. the most likely cause of the aic issue was a defective battery. this report is being filed as part of a retrospective review of historical records.
15756 | 1220648-2024-22041 | 20612169 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20612169&pc=OZD | 05/28/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 05/28/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1117653 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  |  | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a controller error while the pump was in use. there was no patient harm reported. | [5 empty] | The investigation into the controller error and loss of opm data issues have been completed. the impella automated controller was returned for investigation. the data analysis of the logs from the reported day of event are consistent with complaint and show that approximately 2h12m after pump was started, the placement signal became invalid, and console presented with controller error alarms due to placement signal not reliable (1036) and opm comm crc invalid (1045). after approx. 2 minutes, the placement signal returned and both alarms self-resolved. however, several minutes later the pump was stopped and disconnected, and the console was shut down. the real time log data revealed that events associated with alarm # 1036 were due to ¿transient loss of optical data¿ pattern failure, as all of optical bench signals had values of -16384 for 2 minutes. there's no need for repair as the issue self-corrects. the inspection of the returned device revealed a missing tamper evident sticker from the back of the console. inside the console the visual inspection observed that the ribbon cable between the optical bench and the 4-in-1 was misaligned, which might explain the communication issues evidenced by alarm #1045. also, the impellatronic ribbon cable was found to be kinked and should be replaced as a precaution, however it is unrelated to the observed malfunction and failure mode. the cause of the aic issue controller error was misaligned ribbon cable between optical bench and 4-in-1. the cause of the loss of opm data issue was not determined since the issue could not be reproduced this report is being filed as part of a retrospective review of historical records.
15757 | 1220648-2024-22040 | 20612213 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20612213&pc=OZD | 05/06/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 05/06/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1148812 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The production manager reported that the automated impella controller (aic) generated a persistent battery failure alarm when the aic was turned on. reportedly a yellow communication failure alarm also was noted even though the controller was plugged into ac power. there was no patient harm reported. | [5 empty] | The investigation into the battery failure issue has been completed. the impella automated controller (aic) was returned for investigation. the data analysis of the logs confirmed the reported failure. there was a battery failure alarm (transport connection blown/ missing ¿ event set # 360) and a battery 1 communication failure alarm (smbus communication loss or sda short ¿ event set # 352) due to low pack voltage. the failure mode was reproduced on an engineering bench. the battery 1 liquid crystal display indicator was blank (fully discharged). batt test was recorded for the single cell decline. while testing swapping battery 1 and battery 2 connectors, we were able to determine that the power battery manager circuit board (pbm) operated correctly, and battery 2 was able to be charged/ discharged through both channels on pbm. after battery 1 was momentarily replaced by a known-good battery, the battery failure alarm was resolved, battery test was recorded at each step. the battery failure was due to a defective battery 1. the root cause of the defective battery 1 was due to single cell failure. the cause of the aic issue was a defective battery. this report is being filed as part of a retrospective review of historical records.
15758 | 1220648-2024-22055 | 20612228 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20612228&pc=OZD | 05/10/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 05/10/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1226681 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Use of Device Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The clinical consultant reported that the automated impella controller (aic) experienced ¿fluid purge not detected¿ alarm. the purge bag and cassette were changed multiple times and after repeated alarms during purge cassette priming a new aic was utilized. there were no issues with new aic when using the same purge cassettes and purge bags. there was no patient harm reported. | [5 empty] | E1: name of the initial reporter is unknown. the investigation into the purge issue has been completed. the impella automated controller was returned for investigation. the data analysis of the logs indicated the complaint intake is consistent with the complaint date. the logs indicated, -piston blocked (w/ 38% range remaining), wizard for purge fluid not detected, and that several attempts at wizard for purge fluid not detected due to no resolution. additionally, several attempts to remove and reinsert the purge cassette with no resolution to purge fluid not detected. the real time logs from the event were reviewed, showing valid purge flow ticks, with no pressure build up. the returned console purge assembly was visually and physically inspected with no issues found. the console was run through a case start wizard with no issues found. the console was then tested on an engineering bench with a pump and purge setup, and ran for an extended period of time, with no issues occurring. in an attempt to reproduce the scenario that occurred in the field, tests were done after the successful period of time running the pump and purge: test 1: kink/obstruct the inlet purge line to the purge cassette, in order to attempt to reproduce the "purge fluid not detected" wizard to occur during case start: no issues found; console behaved as expected progressing through case start without issue. test 2: kink/ obstruct the outlet of the purge line from the purge cassette to the purge disc, in order to simulate a cassette piston block: no issues found, likely due to not enough clamping force to completely stop purge pressure fluid. unlikely to present itself in the field. test 3: opening luers during priming operation: produces purge fluid not detected however staff reported checking the purge bag, line and cassette multiple times during case start. the cause of the purge fluid not detected alarm is undetermined due to the inability to reproduce. this report is being filed as part of a retrospective review of historical records.
15759 | 1220648-2024-22077 | 20612298 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20612298&pc=OZD | 03/18/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 03/18/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1570900 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant notified the sales manager that the automated impella controller (aic) generated a battery failure alarm. it was reported that the controller indicated red and yellow battery failures. there was no patient harm reported. | [5 empty] | The investigation into the battery failure has been completed. the impella automated controller was returned for investigation. the data analysis of the logs confirmed the reported failure, subsequent boots following the complaint date triggered these alarms. the long-term log showed several of the cells within battery 1 had declined. the return product confirmed that the persistent battery failure (transport connection blown/missing) was not due to a disengaged battery switch. upon bench testing, when the console was booted for the first time, the console alarms were consistent with what was seen in the field. the visual inspection of the batteries and power battery manager circuit board revealed that one of the battery status light-emitting diodes was completely blank. the battery failure alarm was most likely due to a defective battery 1 (decline of individual cell voltage). the root cause of the battery cell failure was undetermined. the cause of the battery failure alarm was due to a defective battery. this report is being filed as part of a retrospective review of historical records.
15760 | 1220648-2024-22066 | 20612341 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20612341&pc=OZD | 04/29/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 04/29/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1215711 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The us complainant reported an automated impella controller on connect. the automated impella controller was observed on connect to have two malfunctions. the battery depleted and the display was distorted/rolling. there was no harm or injury noted from these issues. the instructions for use was followed, and the team used a backup controller for patient use. | [5 empty] | E1: name of initial reporter is unknown. the investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the issue was not determined since issue could not be reproduced. this report is being filed as part of a retrospective review of historical records.
15761 | 1220648-2024-22109 | 20612419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20612419&pc=OZD | 08/09/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 08/09/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1146849 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a battery failure alarm that could not be resolved with normal troubleshooting. there was no patient involved/impacted relating to this issue. | [5 empty] | The investigation for the reported aic - battery failure (red alarm) issues has been completed. the controller device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15762 | 1220648-2024-22053 | 20612463 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20612463&pc=OZD | 10/15/2024 | 12/12/2024 |  | 11/05/2024 |  |  | 10/15/2024 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025547099 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during removal of the impella 5.5 graft cut down to reveal clot at base of anastomosis. clot removed, graft oversewed, and pocket was closed. the patient is stable. the patient survived the event. | [5 empty] | The investigation into thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly on manufacturer device report 1220648-2024-22053. g1 reporting contact email. h6 code 4614 was reported incorrectly. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
15763 | 1220648-2024-22057 | 20612503 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20612503&pc=OZD | 04/08/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 04/08/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1588794 |  | 0042-0040 |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was on an automated impella controller (aic) for mechanical circulatory support. it was reported that customer support received a call for support due to a battery failure red alarm. reportedly support was unable to resolve the alarm. there was no patient harm reported. | [5 empty] | The investigation of the reported battery failure has been completed. the product was returned, and the battery failure was confirmed the data logs were not available due to a software upgrade performed post complaint date. it was confirmed that the persistent battery failure (transport connection blown/missing) was not due to a disengaged battery switch. when the console was booted for the first time during bench testing a battest showed that cell 3 of battery 1 was reading 0mv. the console interior connections were checked and verified. when visually inspecting the batteries and power battery manager circuit board (pbm), it was observed that one of the batteries status light-emitting diodes were completely blank. the battery failure alarm was due to a defective battery 1 (decline of individual cell voltage). the root cause of the battery cell failure was undetermined. the cause of the battery failure alarm is due to a single cell failure in a defective battery. this report is being filed as part of a retrospective review of historical records.
15764 | 1220648-2024-22086 | 20612516 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20612516&pc=OZD | 04/19/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 04/19/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1583419( |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported a controller failure alarm occurring prior to starting a case. there was no report of patient harm or injury. | [5 empty] | The investigation into the aic -controller error issue has been completed. the automated impella controller (aic) was returned for investigation. data analysis: complaint intake is accurate of case start date with the case starting on (b)(6) 2022 and pump explant. however, after reviewing imc logs, no log entries for "controller error" or "controller failure" were found. imc log analysis of the entire case showed no issues. device analysis: console arrived in danvers. ic8950 was powered on, where no alarms had occurred. alarm history was reviewed, where no controller error or controller failure alarms were present. console was ran on pump and purge with no issues occurring. this report is being filed as part of a retrospective review of historical records.
15765 | 1220648-2024-22087 | 20612625 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20612625&pc=OZD | 04/19/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 04/19/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1164186 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The us complainant had a successful impella support but at the end of the case it was noted that the purge pressure retainer dip was broken off. there was no harm or injury noted. the instructions for use (ifu) were followed and backup controller was used for patient use. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, the purge flag was confirmed to be broken. before returning this console back to the customer, the purge flag was replaced and a full functional check on the console and optical system was performed. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15766 | 1220648-2024-22061 | 20612635 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20612635&pc=OZD | 10/15/2024 | 10/13/2025 |  | 11/05/2024 |  |  | 10/15/2024 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1313442 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) had a controller error alarm. the aic was exchanged. there was no harm to the patient. the patient survived the event. | [5 empty] | Data analysis: the console displayed the controller error (power battery manager (pbm) voltage out-of-spec - event set #501), the console log also confirmed controller error (loss of comm (pbm2) - event set #1025), two controller error (loss of comm (epm) - event set #1021), loss of comm to opm - event set #1043, and multiple communication errors with other serial device. most importantly, the console log displayed the impella stopped (mechanical/electrical issue - event set #115) as the pump stop working due to the consequence of communication failures, but this issue was resolved by itself after 3 seconds. device analysis: the failure was not reproducible upon reboot in the lab. inspection of the internal circuitry of the console didn't reveal any anomalies. a known good pump and purge line were run with console for a significant amount of time with no anomalies. conclusion: the root cause for the controller error (pbm voltage out-of-spec) could not be determined since the failure could not be reproduced. pbm and impellatronics board precautionarily replaced. d9 device returned to manufacturer date added d2 common device name revised on manufacturer device report 1220648-2024-22061 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-22061 in accordance with updated procedures. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
15767 | 1220648-2024-22075 | 20612655 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20612655&pc=OZD | 04/14/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 04/14/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1219576 |  | 0042-0000-EU |  | Yes |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. during the preparation of the device and prior to the procedure, it was reported that the purge disc was not recognized by the system. the device was replaced with another aic that was at the customer site. no report of patient harm or injury. | [5 empty] | D10: the date of the devices return for evaluation is unknown. the console (aic) was returned for evaluation. upon inspection of the console, the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken (missing at blue disc retainer). before returning this console back to the customer, the purge flag was replaced, and preventive maintenance was performed. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15768 | 1220648-2024-22104 | 20612669 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20612669&pc=OZD | 08/16/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 08/16/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1266280 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission |  |  |  |  | 08/16/2022 |  |  |  |  | During a software update, it was discovered the battery of automated impella controller (aic) was at zero and indicated battery failure. no patient involvement. | [5 empty] | The investigation has been completed. the console (ic9269) was sent back for further investigation. the battery switch was engaged during console check-in, and the console was powered on via batteries and ac power supplied. no alarms occurred. ic4755 had expected behavior both on ac power and battery power with no issues. it is likely that the console did not have the battery switch engaged in field. and the failure was resolved by engaging the battery switch on bottom of the console. the cause of the issue was use related. the operator did not correctly engage the battery switch during use. this report is being filed as part of a retrospective review of historical records.
15769 | 1220648-2024-22037 | 20612717 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20612717&pc=OZD | 06/22/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 06/22/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1389732 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported there was an aic (automated impella controller) with a broken purge clip observed by the hospital biomed department. there was no patient use and backup aic was in place for support per ifu recommendations. there was no reported harm or injury to the patient. | [5 empty] | The automated impella controller (aic) was returned for evaluation. upon inspection of the console, it was confirmed the purge disc retainer had broken/snapped off of the purge assembly. the purge disc retainer and purge flag was replaced, the purge pressure sensor was validated, and a functional check was performed; before the console, was returned to the customer. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15770 | 1220648-2024-22063 | 20612761 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20612761&pc=OZD | 05/10/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 05/10/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1276778 |  | 0042-0000-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during a training at the customer site, the automated impella controller's (aic) purge disc was not detected. the staff replaced the cassette several times and this did not resolve the issue. no patient involvement. the customer had another aic onsite to utilize. | [5 empty] | D9: date of the device return for evaluation is unknown. e1, e2, e3: name and occupation of the initial reporter is unknown. the console (aic) was returned for evaluation. upon inspection of the console, the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken (missing at blue disc retainer). before returning this console back to the customer, the purge flag was replaced and a full functional check on the console and optical system was performed. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15771 | 1220648-2024-22118 | 20612766 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20612766&pc=OZD | 02/11/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 02/11/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1454057 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a battery failure red alarm that could not be resolved. there was no reported patient harm. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the battery failure alarm was due to the battery switch being unintentionally disengaged. this report is being filed as part of a retrospective review of historical records.
15772 | 1220648-2024-22044 | 20612809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20612809&pc=OZD | 10/15/2024 | 12/30/2024 |  | 11/05/2024 |  |  | 10/15/2024 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025521626 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that there was massive retroperitoneal bleeding. the artery was ruptured during the puncture for impella cp implantation. vessel was fixed by surgery. four units of red blood cells were administered. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation of vascular damage - peripheral vessel has been completed. no product was returned for evaluation. clinical details noted a massive retroperitoneal bleed due to ruptured artery from puncture during impella implantation. vessel was repaired in surgery and blood products given. the root cause of the vascular damage could not be definitively determined. g.1 revised reporting contact email since the initial manufacturer device report 1220648-2024-22044 was submitted in accordance with updated procedures.
15773 | 1220648-2024-22038 | 20612840 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20612840&pc=OZD | 05/09/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 05/09/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1186350 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The clinical trainer reported that the automated impella controller (aic) generated a battery failure alarm that did not resolve with charging or cycling the console battery on and off. reportedly there was no patient or pump involvement. | [5 empty] | The investigation into the battery failure issue has been completed. the impella automated controller (aic) was returned for investigation. the data analysis of the logs revealed consistency with the reported issue, subsequent boots following the complaint date trigger these alarms. the long-term logs were not available for review, however a batt test was performed via telnet, revealing that cell 1 of battery 1 had begun to decline, passing the operating parameters of the battery. during bench testing of the returned console, it was confirmed that the persistent battery failure (transport connection blown/missing) was not due to a disengaged battery switch. when the returned console was booted for the first time, console alarms were consistent with what was seen in the field. when visually inspecting the batteries and the power battery manager circuit board, it was observed that one of the batteries¿ status light-emitting diodes were completely blank. the battery failure alarm was due to a defective battery 2 (decline of individual cell voltage). the root cause of the battery cell failure was undetermined. the cause of the aic issue was a defective battery, due to the decline of single cell within battery two. this report is being filed as part of a retrospective review of historical records.
15774 | 1220648-2024-22043 | 20612909 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20612909&pc=OZD | 05/06/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 05/06/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1557111 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The clinical consultant reported that the automated impella controller (aic) generated a battery failure alarm the cord was checked and plugged into a red socket, however, the aic was still not charging. this aic was switched out due to turning off and then back on. there was no patient harm reported. | [5 empty] | The investigation into the battery failure issue has been completed. the impella automated controller (aic) was returned for investigation. the data analysis of the logs determined the complaint intake was consistent. an unexpected controller shutdown occurred, most likely due to staff checking if the battery switch was engaged/ disengaged while ac power was unstable. the console was powered on multiple times in the following days to likely check battery level while the console was left to charge, however each instance triggered a "battery level low" alarm despite remaining plugged into ac power. the inspection of the returned device indicated the battery switch was engaged, and ac power was connected, however the console still booted with "battery level low" and the batteries would not recharge. the console was opened, and interior connection integrity was verified. the power supply output voltage was checked via the j15 connector of the power battery manager circuit board (pbm), it was observed that the output read ~0v. the power supply unit input voltage was checked and observed at ~110v, a good power supply unit was placed into the console and the batteries began to properly charge. the cause of the lack of charge to the battery set was due to a defective power supply. this report is being filed as part of a retrospective review of historical records.
15775 | 1220648-2024-22083 | 20612992 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20612992&pc=OZD | 04/20/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 04/20/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1233946 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, automated impella controller (aic) failed to load the purge cassette disc. a backup console was used. there was no injury to the patient. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, the purge assembly of the console, it was confirmed that the blue receptacle for the purge disc had broken off. before returning this console back to the customer, the purge disc retainer was replaced and a functional check was performed. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15776 | 1220648-2024-22048 | 20613006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20613006&pc=OZD | 05/13/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 05/13/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER EU |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER EU |  |  | ABIOMED, INC. |  |  | US |  | 1231129 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 65 YR |  | Female |  |  | White | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) selector knob was no longer working properly. it was reported the device would not jump through the menus. this aic was replaced with another aic. there was no patient harm reported. | [5 empty] | E1: name of the initial reporter is unknown. the investigation into the keyboard rotary knob issue has been completed. the impella automated controller was not received from the customer. the data analysis of the logs indicated no directly linked failure modes. however, on the complaint date, it was observed and speculated that the operator kept accessing to ¿display¿ menu multiple times but failed to navigate and select the expected entry via the (defective) rotary knob. additionally, the operator could perform a normal shut-down on console ic4251 by pushing the rpb, which suggested the malfunction was limited to the rotary/ selection function of the rpb. the consoled was tested by the field service engineer and the visual inspection did not reveal any abnormality for internal cable connections (especially between 4 in1 and impellatronic. the engineer temporarily replaced the impellatronic board, and the issue was resolved. the root cause of the malfunction was due to the defective impellatronic. this report is being filed as part of a retrospective review of historical records.
15777 | 1220648-2024-22095 | 20613053 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20613053&pc=OZD | 04/08/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 04/08/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1391515 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported an automated impella controller (aic) console failed to recognize a purge cassette. no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken (missing at blue disc retainer). before returning this console back to the customer, the purge flag was replaced and a full functional check on the console and optical system was performed. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15778 | 1220648-2024-22071 | 20613060 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20613060&pc=OZD | 03/25/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 03/25/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1095638 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant notified the clinical consultant that the automated impella controller (aic) generated a battery failure alarm. it was reported that the controller indicated a battery failure message upon start up and attempts to rest the battery were unsuccessful. there was no patient harm reported. | [5 empty] | The investigation into the battery failure has been completed. the impella automated controller was returned for investigation. the data analysis of the logs confirmed the reported failure, there was a battery failure alarm (transport connection blown/ missing ¿ event set #360) due to low pack voltage. the failure mode was reproduced on an engineering bench. the battery 2 pack voltage was measured to be 3.5 v rather than the expected 16 v, and the battery liquid crystal display indicator was blank (fully discharged). batttest showed 400 mv cell voltage difference with battery 2. battery 2 was replaced to resolve the failure alarm. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
15779 | 1220648-2024-22058 | 20613116 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20613116&pc=OZD | 04/01/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 04/03/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1263253 |  | 0042-0000 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. the clinical consultant reported that the controller generated battery alarms when unplugged from the ac connector. the instructions for use was followed and the backup controller was used. there was no patient harm reported. | [5 empty] | The impella automated controller device was not received from the customer; therefore, investigation of the device was not possible. the logs did not require review given the failure mode was resolved by engaging the battery switch. the console was not returned for the investigation and remained in the field. the battery failure was due to the battery transport connection switch was not properly engaged once received due to improper technique usage by the end user. this report is being filed as part of a retrospective review of historical records.
15780 | 1220648-2024-22064 | 20613150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20613150&pc=OZD | 05/02/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 05/02/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1225726 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported an automated impella console (aic) had failed to recognize a purge cassette during use. a new console was used without any adverse event noted. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken (missing at blue disc retainer). before returning this console back to the customer, the purge flag was replaced and a full functional check on the console and optical system was performed. this report is being filed as part of a retrospective review of historical records.
15781 | 1220648-2024-22080 | 20613191 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20613191&pc=OZD | 04/13/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 04/13/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1126570 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported an automated impella controller (aic) enduring a purge cassette error. the console was returned, and a loaner was given. no patient was involved. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, the purge assembly was inspected, no issues were observed. the console was tested on pump and purge with no issues occurring. case start was tested with pump and purge several times, no issues occurred. therefore, the root cause is most likely that the issues with purging the pump are use related. a full functional check on the console and optical system was performed, before returning this console back to the customer, the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15782 | 1220648-2024-22067 | 20613213 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20613213&pc=OZD | 03/28/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 03/28/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1087127 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant had a battery failure on an automated impella controller (aic). biomed to report battery failure on aic ic1240. the instructions for use was followed, and a backup controller was used instead. no lasting harm or injury. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the root cause of the battery communication failure was due to a defective pbm pca. this report is being filed as part of a retrospective review of historical records.
15783 | 1220648-2024-22049 | 20613214 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20613214&pc=OZD | 04/23/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 04/23/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1468353 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a high temperature during use of the console. the pump was switched to a backup console with no harm to the patient. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the root cause of the battery temperature high alarm is most likely due to a misinterpretation of data within a single sample of communication from the batteries to the pbm. however, the overall root cause is undetermined. this report is being filed as part of a retrospective review of historical records.
15784 | 1220648-2024-22090 | 20613252 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20613252&pc=OZD | 04/13/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 04/13/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1322579 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report |  | Other | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had a battery failure on the automated impella controller (aic) for a coronary artery bypass graft (cabg)/5.5 patient. the battery is backup for the a/c power. the instructions for use states to use a backup console for failures. there was no lasting harm or injury and no patient use at the time. | [5 empty] | E1: name of initial reporter is unknown. e2/e3: unknown. the investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the aic issue was a defective battery. this report is being filed as part of a retrospective review of historical records.
15785 | 1220648-2024-22060 | 20613295 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20613295&pc=OZD | 05/09/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 05/09/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1289309 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the patient was placed onto impella cp support for high-risk percutaneous coronary intervention. while on support, the aic experienced purge disc/flag issues with no resolution specified. | [5 empty] | The investigation into the aic -purge disc/ flag issue has been completed. the automated impella controller (aic) was returned for investigation. data analysis: imc logs are not relevant for this failure mode. device analysis: console arrived in danvers. visual inspection of the purge assembly area confirmed a broken blue retainer. before returning to customer, fs will be instructed to: replace purge disc retainer (5500-0222), this report is being filed as part of a retrospective review of historical records.
15786 | 1220648-2024-22102 | 20613359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20613359&pc=OZD | 08/12/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 08/12/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1388031 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an aic (automated impella controller) experienced a battery failure. the aic was removed and replaced by a loaner device. there was no patient harm. | [5 empty] | The investigation for the reported aic - battery failure (red alarm) issues has been completed. the controller device has been returned for evaluation. clinical details were sufficient to determine no issues with console battery operation or the battery switch mechanism were found. this report is being filed as part of a retrospective review of historical records.
15787 | 1220648-2024-22116 | 20613362 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20613362&pc=OZD | 02/15/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 02/15/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1556635 |  | 0042-0040-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The automated impella controller (aic) controller failed to hold a charge after unplugging the device. this occurred without any patient involvement. the power supply was replaced, and the issue was resolved. there was no patient involvement. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the a/c power issues was a defective power supply. this report is being filed as part of a retrospective review of historical records.
15788 | 1220648-2024-22046 | 20613484 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20613484&pc=OZD | 05/11/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 05/12/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1254075 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Decrease in Pressure | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported that the automated impella controller (aic) experienced purge cassette/disc not detected. this aic was replaced with another aic. there was no patient harm reported. | [5 empty] | The investigation into the purge issue has been completed. the impella automated controller was returned for investigation. the data analysis of the logs from the complaint date revealed that the console issued the purge disc not detected alarm (event set # 402) several times throughout a long time. the change of new pc/ purge disc was observed, yet the console still had a problem in detecting purge disc. the returned console was tested the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken (or missing at blue disc retainer). the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15789 | 1220648-2024-22106 | 20613531 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20613531&pc=OZD | 04/04/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 04/04/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1417603 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. during a routine cassette and bag change, the automated impella controller (aic) would not move beyond the "insert purge cassette" screen although the cassette was clicked in. a new cassette was opened but the same issue persisted. the instructions for use (ifu) were followed and the backup controller was used. no patient harm was reported. | [5 empty] | The automated impella controller (aic) was returned for evaluation. upon inspection of the console, the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. before returning this console back to the customer, the purge flag was replaced and a full functional check on the console and optical system was performed. this report is being filed as part of a retrospective review of historical records.
15790 | 1220648-2024-22074 | 20613563 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20613563&pc=OZD | 04/24/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 04/24/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1182787 |  | 0042-0000-US-L |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The us complainant reported that an automated impella controller (aic) had a purge clip broken off. the damaged console was replaced and there was no patient involved. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, t the purge retainer was confirmed to be broken (completely missing from the console ¿ no image taken). before returning this console back to the customer, the retainer and purge flag were replaced and a full functional check on the console and optical system was performed. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15791 | 1220648-2024-22073 | 20613568 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20613568&pc=OZD | 04/25/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 04/25/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1561428 |  | 0042-4201 |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during an in-service at the customer site, the automated impella controller's (aic) purge disc would not engage. once purge disc was removed the purge disc holder was loose and a spring and small piece fell from behind the purge disc holder. no report of patient involvement. | [5 empty] | D9: the date of the devices return for evaluation is unknown. the console (aic) was returned for evaluation. upon inspection of the console, the purge retainer (flag also non-functional) was confirmed to be broken. before returning this console back to the customer, the purge retainer/ flag was replaced and a full functional check on the console and optical system was performed. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15792 | 1220648-2024-22069 | 20613598 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20613598&pc=OZD | 04/18/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 04/18/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1285254 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The us complainant reported that an automated impella controller (aic) purge disc failed. the team observed the issue and followed the instructions for use (ifu) and a backup aic was utilized. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, which was tested thoroughly on an engineering workbench. the exterior, there was no damage to the purge system, and it function as expected. the interior of the purge system was inspected. glucose build up was found around the purge pressure sensor. fluid ingress of any sort can impact the purge pressure sensors optical sensor which would affect the ability to detect the disc being inserted. therefore, the root cause of the purge disc detection issues was most likely contamination on the purge pressure sensor.). before returning this console back to the customer, the purge system components were cleaned specifically the around the purge pressure senso, the purge system was tested through internal procedures, and a full functional check on the console and optical system was performed. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15793 | 1220648-2024-22084 | 20613622 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20613622&pc=OZD | 04/19/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 04/19/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1092252 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the complainant had an aic that failed with a red alarm. the instructions for use (ifu) were followed and a backup console was used for patient support. there was no report of patient harm or injury. | [5 empty] | The investigation into the aic -controller error issue has been completed. the automated impella controller (aic) was returned for investigation. data analysis: imc logs are consistent with the controller failure alarms on the complaint occurred date. after console booted, the purge sensor defective error message appeared. according to the logs, attempts of power cycling the console failed in solving the controller failure alarm. device analysis: console was investigated on an engineering bench, and the controller failure alarm was reproduced with pump simulator. console was opened, no abnormality or misalignment was found. replacement of pud failed to resolve the controller failure red alarm. eventually, the attempt of replacing purge pressure sensor cleared the alarm. the controller failure alarm was due to defective purge pressure sensor, but the exact root cause of pps malfunction was unable to determine. during investigation, another failure mode of purge disk not recognized was discovered. and the root cause was due to the broken blue retainer. before sending this console back to the customer. field service replaced the purge pressure sensor and blue retainer due to damage. also, fs was instructed to perform a full functional check on the console. this report is being filed as part of a retrospective review of historical records.
15794 | 1220648-2024-22107 | 20613643 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20613643&pc=OZD | 08/11/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 08/11/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1213648 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had contacted field service (fs) for a battery failure. the device was replaced with a backup console will be used. it was reported that there was no patient involvement. | [5 empty] | The investigation for the reported aic - battery failure (red alarm) issues has been completed. the controller device has been returned for evaluation. the cause of the aic issue was a defective battery. this report is being filed as part of a retrospective review of historical records.
15795 | 1220648-2024-22062 | 20613673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20613673&pc=OZD | 04/01/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 04/01/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1255124 |  | 0042-0000 |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant notified the senior therapy and business development consultant that the automated impella controller (aic) generated a battery failure alarm. it was reported that the controller indicated a battery failure message upon start up. there was no patient harm reported. | [5 empty] | The investigation into the battery failure has been completed. the impella automated controller was returned for investigation. the data analysis of the logs revealed consistency with the complaint intake, subsequent boots following the complaint date triggered the alarms. the long-term log analysis of the battery data revealed that beginning before the complaint date of occurrence, cell 2 voltage of battery 2 had been declining for an extended period of time. it was confirmed that the persistent battery failure (transport connection blown/missing) was not due to a disengaged battery switch. when the console was booted for the first time, console alarms were consistent with what was seen in the field. the battery failure alarm was due to a defective battery 2 (decline of individual cell voltage). the root cause of the battery cell failure was undetermined. the cause of the single cell failure is most likely due to a defective battery. this report is being filed as part of a retrospective review of historical records.
15796 | 1220648-2024-22076 | 20613743 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20613743&pc=OZD | 03/21/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 03/21/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1186531 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported an automated impella controller (aic) was not detecting purge fluid during priming process. the site had a backup unit that was used. no patient was involved. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, with the device powered off, the stepper motor spindle should move freely. however, it could not spin at all. glucose build-up was found between the purge insert and the stepper motor gasket. once cleaned, the system operated as expected. before returning this console back to the customer, the entire purge system was cleaned for glucose ingress, the purge stepper motor was disconnected to check for buildup underneath the rubber gasket and realigned, a check of the stepper motor torque, a good pump and purge was tested, and the final review of the logs were sent. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15797 | 1220648-2024-22103 | 20613843 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20613843&pc=OZD | 04/05/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 04/05/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1323598 |  | 0042-0010-EU |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, automated impella controller (aic) alarmed for an error. a back up console was used to complete the procedure. no patient harm was reported. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, the failure mode was not reproduced. inspection of optical pump receptacle and measurement of 5s parameters did not show any abnormalities. however, during short-term testing with simulator pump and pc, the console functioned well without any failure. the optical relating alarms were not reproduceable. before returning this console back to the customer, a full functional check on the console and optical system was performed. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15798 | 1220648-2024-22070 | 20613898 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20613898&pc=OZD | 04/27/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 04/27/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1087127 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  |  | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was on an automated impella controller (aic) for mechanical circulatory support. the us complainant had an automated impella controller (aic) alarm that prompted the team to replace the aic. the instructions for use (ifu) was followed and backup console. there was no report of patient harm or injury. | [5 empty] | The investigation into the aic -controller error issue has been completed. the automated impella controller (aic) was returned for investigation. data analysis: aic logs showed potential issues with detecting the second pump connected after the first pump (impella cp sn (b)(6) experienced issues with reoccurring suction alarms. device analysis: the console was thoroughly tested on an engineering lab bench. there were no issues reproduced which impacted the ability to properly detecting an impella cp. the pump detection issue was most likely due to the rarely occurring epm crystal issue. the impellatronic pca does not have permanent damage, but this intermittent pump detection failure can occur infrequently and already has a sw mitigation in place. the detection issues likely occurred after the epm was power cycled during connection of the second pump and powered back up in a failed state. aic software release v8.3 brought mitigation for this issue to reset the epm sau but did not recover the system. in this case, it is likely that the epm reset induced the failure state. persistent failed handshakes were identified in the aic logs. inspection of the rlm revealed that l1 was knocked off. see image attached. the micro usb cable was also strained and damaged. before sending this console back to the customer, field service was instructed to replace the impella connect module, the ¿backstrap¿ cable, and transfer the customer configuration. perform a full functional check on the console and optical system. this report is being filed as part of a retrospective review of historical records.
15799 | 1220648-2024-22065 | 20613920 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20613920&pc=OZD | 05/05/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 05/05/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1140919 |  | 0042-0000-US |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on an automated impella controller (aic) for mechanical circulatory support. aic was pulled from service by the biomed. there was documented sensor and purge disc issues. the instructions for use (ifu) was followed and a backup console was utilized. | [5 empty] | E1: name of initial reporter is unknown. date device returned to manufacturer unknown. the investigation into the aic -purge disc/ flag issue has been completed. the automated impella controller (aic) was returned for investigation. data analysis: biomed reported an additional issue with a dirty optical ferrule sensor on the front panel optical connector, however no issues were found pertaining to this. imc logs from the complaint date revealed that the console issued the purge disc not detected alarm (event set #402) several times. device analysis: the console was tested after arriving in field service (fs). the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken (missing at blue disc retainer - see attached image). front panel optical connector ferrule was cleaned and inspected with no issues found. additionally the bulk head connection point between the optical bench and front panel optical connector was cleaned and inspected with no issues found. repair task: before returning to customer, fs will be instructed to: replace purge disc flag (3100-0256) verify purge pressure sensor via section 12 of preventive maintenance. this report is being filed as part of a retrospective review of historical records.
15800 | 1220648-2024-22085 | 20613935 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20613935&pc=OZD | 03/03/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 03/03/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1281597 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant notified the clinical trainer that the automated impella controller (aic) generated battery failure and battery level low alarms. there was no patient harm reported. | [5 empty] | The investigation of the reported battery failure has been completed. the impella automated controller was returned for investigation. the data log review confirmed the reported failure. there was a battery failure alarm (transport connection blown/ missing ¿ event set # 360) and a battery 2 communication failure alarm (smbus communication loss or sda short ¿ event set # 353) due to low pack voltage. the failure mode was reproduced on an engineering bench. the battery 2 liquid crystal display indicator was blank (fully discharged). battest was recorded for the single cell decline. by swapping battery 1 and battery 2 connectors, we were able to determine that the power battery manager circuit board (pbm) operated correctly, and battery 1 was able to be charged/ discharged through both channels on pbm. after battery 2 was momentarily replaced by a known-good battery, the battery failure alarm was resolved, battest was recorded at each step. the battery failure was due to a defective battery 2. the root cause of the defective battery 2 was due to single cell failure. the exact cause of the defective battery was unable to be determined. the cause of the battery failure was due to a defective battery (rapid decline in cell voltage). this report is being filed as part of a retrospective review of historical records.
15801 | 1220648-2024-22097 | 20614006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20614006&pc=OZD | 03/01/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 03/01/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1429644 |  | 0042-0010 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant in the had an automated impella controller (aic) with battery issues. the aic was removed and replace per the instructions for use recommendations. a backup battery was utilized for a/c power. there was no harm or injury reported. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the console was tested while in field service. the battery issues were not reproduced. the root cause of the battery communication failure was not determined due to the inability to reproduce. this report is being filed as part of a retrospective review of historical records.
15802 | 1220648-2024-22111 | 20614045 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20614045&pc=OZD | 12/09/2021 | 11/05/2024 |  | 11/05/2024 |  |  | 12/09/2021 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1219726 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automatic impella controller (aic) had a battery failure, of a red alarm. a back up console was used. no patient involved. | [5 empty] | The investigation for the reported aic - battery failure (red alarm) issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the battery failure was due to a defective battery 1. the root cause of the defective battery 1 was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
15803 | 1220648-2024-22079 | 20614072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20614072&pc=OZD | 03/11/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 03/11/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2022110806 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant notified the field service logistics coordinator that the automated impella controller (aic) generated a battery failure alarm upon startup. reportedly the alarm could not be cleared, and this device was replaced with a backup device. there was no patient harm related to the battery failure as a backup device was used. | [5 empty] | The investigation of the reported battery failure has been completed. the suspected console ( (b)(6) ) was sent to abiomed japan field service by the customer for further investigation. the japan investigation of the returned product identified the issue as a defective battery. the data analysis of the logs were reviewed by the manufacturer and were consistent with the reported complaint and showed battery-related alarms upon each console boot-up: battery failure (603 due to blown fuse, 360 due to low voltage), battery 1 communication failure (352), and battery level low (#23). batttest diagnostic information from battery 1 (sn (b)(6) ) showed that 3 out of 4 cells had voltage of 0v, and the battery pack was internally shut down to the point that it would not communicate with the power battery manager circuit board or diagnostic equipment (ata interface cable). this is indicative of a cell imbalance pattern failure. after the batteries were replaced, the issue was resolved. the cause of the aic issue was a defective battery. this report is being filed as part of a retrospective review of historical records.
15804 | 1220648-2024-22115 | 20614078 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20614078&pc=OZD | 07/10/2021 | 11/05/2024 |  | 11/05/2024 |  |  | 07/10/2021 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1570900 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 20 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support, the automated impella controller (aic) experienced an unexpected shutdown. the console was replaced to address the issue. no patient harm reported. | [5 empty] | Console logs from the reported day of event are consistent with complaint and show that alarm#: 603 (unexpected controller shutdown) was shown, 13 days after support started (on (b)(6) 2021). this alarm was a consequence of console self-reset (and self-restart) following crash of io-graphics process (part of qnx), which was detected by aic software: on (b)(6) 2021 08:29:56: systemidle: process io-graphics 339995 died abnormally this self-reset is a mitigation mechanism for when io-graphics crashes, that has been implemented in aic_v8.2 in response to a series of complaints regarding that failure mode (jama: gid-dft-863138). during and after the reset, hemodynamic support continued. console operated within specification, and the issue was resolved. this report is being filed as part of a retrospective review of historical records.
15805 | 1220648-2024-22088 | 20614135 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20614135&pc=OZD | 04/19/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 04/19/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1117659 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the automated impella controller (aic) exhibited a battery communication failure alarm. the console was plugged into the ac and the device showed that the console was plugged in, however the battery level was at 50%. the device was replaced. no report of patient harm or injury. | [5 empty] | The investigation for the reported aic - battery comm. failure (yellow alarm) issues has been completed. the impella device has been returned for evaluation. the root cause of the battery communication failure alarm was most likely due to an intermittent malfunction of the pbm board. this report is being filed as part of a retrospective review of historical records.
15806 | 1220648-2024-22148 | 20614164 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20614164&pc=OZD | 03/29/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 03/29/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1543580 |  | 0042-0010-EU |  | Yes |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that after turning on the automated impella controller (aic) a blue screen appeared with the words "system check failed". the resolution for this issue is unknown. no report of patient involvement, injury, or harm. | [5 empty] | D9: the date of the device return for evaluation is unknown e1, e2, e3: the name and occupation of the initial reporter is unknown. the console (aic) was returned for evaluation. upon inspection of the console, the blue screen and self-check failure was not reproduced. the impellatronic board, pbm pca (especially channel 1), and the cable connection was inspected and tested. no abnormalities were noticed and the pbm was replaced as a precaution. therefore, the root cause of system check failed issue could not be determined, due to the failure mode was not reproduced. a functional check and preventive maintenance was performed; before the console was returned to the customer. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15807 | 1220648-2024-22081 | 20614166 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20614166&pc=OZD | 04/21/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 04/21/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1216751 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The us complainant had an automated impella controller (aic) with a purge disc breakage/failure. a backup console was used. no patient was involved. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, the purge assembly was inspected, it was observed that the purge disc retainer had broken off of the purge assembly. the purge disc retainer was replaced, and a functional check was performed; before returning this console back to the customer, the purge flag was replaced and a full functional check on the console and optical system was performed. this report is being filed as part of a retrospective review of historical records.
15808 | 1220648-2024-22082 | 20614175 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20614175&pc=OZD | 03/07/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 03/07/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1145540 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant notified the field service logistics coordinator that the automated impella controller (aic) generated a controller fault alarm. a loaner was sent to the customer to initiate return of this device. there was no patient harm reported. | [5 empty] | The investigation of the reported controller failure has been completed. the impella automated controller was returned for investigation. the data log review were consistent with the controller error alarms on the complaint occurred date. after a few days of support, the purge sensor defective error message appeared. the returned console was investigated on an engineering bench, but the controller failure alarm was not reproduced during power cycling and overnight running with the pump simulator. the console was opened, and the ribbon cable from the purge pressure sensor was found slightly bent. it was likely due to the bent pps ribbon cable, leading to almost insufficient voltage supply and triggering controller error alarms. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15809 | 1220648-2024-22096 | 20614202 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20614202&pc=OZD | 03/31/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 03/31/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1147479 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported an automated impella controller (aic) emitting an error alarm. the aic was power cycled 3 times and still gave the same error message. the instruction for use (ifu) was followed and a backup controller was used instead. there was no patient involved or impacted. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, the device was booted with controller error, which was confirmed via logs to be the same controller error seen in field. optical bench connection points were checked with no issues observed, a optical bench power was unplugged and pins were measured which read 5v and optical voltage was checked. a good optical bench was swapped in and the controller error cleared. therefore, the root cause of the controller failure was due to the optical bench is drawing too much current, and pulling down the voltage supply from the pbm. this report is being filed as part of a retrospective review of historical records.
15810 | 1220648-2024-22149 | 20614333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20614333&pc=OZD | 04/12/2021 | 11/05/2024 |  | 11/05/2024 |  |  | 04/12/2021 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1219722 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Display or Visual Feedback Problem | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. while on support, the automated impella controller experienced a display issue and controller error yellow alarm that was resolved by switching to a backup controller. there was no reported patient harm. | [5 empty] | The investigation for the reported controller error (yellow alarm) has been completed. clinical staff reported that the placement signal was not visible on the console. switching to the backup controller resolved the issue. this complaint was related to a clinical procedure, but there were no adverse events related to this product malfunction. the logs revealed there was a loss of communication to the optical bench given a controller error alarm was triggered. real time logs showed there was only a few samples of optical-related data while the impella cp was running. after multiple other errors were visible in the automated impella controller¿s user interface, the user attempted to disconnect the pump but it was not recognized. an emergency shutdown had to be performed to turn the console off. the console was tested thoroughly on an engineering bench, but the failure mode could not be reproduced. there were no optical bench-related alarms triggered during more than 24 hours of testing with an optical pump simulator and purge. the root cause of the controller error (opm comm. loss) could not be determined because the failure mode was not reproduced during testing. this report is being filed as part of a retrospective review of historical records.
15811 | 1220648-2024-22132 | 20614352 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20614352&pc=OZD | 05/30/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 05/30/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1145132 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Use of Device Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced a purge issue. no clinical details regarding resolution or patient involvement / condition were provided. | [5 empty] | The investigation has been completed. the console (ic2751) was sent back for further investigation. the root cause of the aic issue was a loose connector. this report is being filed as part of a retrospective review of historical records.
15812 | 1220648-2024-22125 | 20614369 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20614369&pc=OZD | 04/19/2021 | 11/05/2024 |  | 11/05/2024 |  |  | 04/19/2021 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1194756 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Nurse | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's registered nurse reported the automated impella controller experienced a controller error yellow alarm in which the screen froze while trying to access the purge menu. the aic was then replaced with a backup console. it is unknown if there was patient involvement but there was no reported patient harm. | [5 empty] | The investigation for the reported controller error (yellow alarm) has been completed. logs showed that the purge bag change wizard was started and completed. however, it was canceled when completed. several hours later, a second purge bag change occurred and at the point when the bag was supposed to be spiked, there were several entries for "keyboard: high guard level (esd)." the console interior and exterior were examined for fluid ingress and none was observed. the console was ran on a different pump for several hours without issue. the console screen, softkey and rpb were tampered with during pump operation. no issues were observed. the issue was unable to be reproduced. it is likely that during the second purge bag change that, when the staff attempted to spike the bag, the bag was spilt onto the console. the keyboard: high guard level entries were in response to fluid falling onto the softkeys in conjunction with the staff attempting to press the softkeys. the cause of the failure mode was use related. this report is being filed as part of a retrospective review of historical records.
15813 | 1220648-2024-22092 | 20614381 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20614381&pc=OZD | 04/11/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 04/11/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1496398 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | Us complainant reported that the automated impella controller (aic) experienced purge disc/flag issues. no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken (missing at blue disc retainer). before returning this console back to the customer, the purge flag was replaced and a full functional check on the console and optical system was performed. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15814 | 1220648-2024-22155 | 20614382 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20614382&pc=OZD | 03/16/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 03/16/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1172385 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) that there was a display issue, and that a backup console was used. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported display issue has been completed. the product was returned, and the issue was confirmed. the console logs from the reported day of event were not available, as the console was unable to be started. tested in danvers, upon pressing the power button, fans started spinning, indicating pbm powering up, but no other activity was observed and the console's screen remained blank. measurements of the pc_on signal on j13 port on pbm and v_2_pc on j16 port of pbm indicated them to be within specification. temporarily replacing 4-in-1 assembly resolved the issue and console was able to be powered up and its display (screen) was confirmed to be in working condition. per the results of the clinical review and investigation: the root cause was determined to be a defective 4-in-1 board. this report is being filed as part of a retrospective review of historical records.
15815 | 1220648-2024-22099 | 20614397 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20614397&pc=OZD | 02/28/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 02/28/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1209412 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant had a battery failure on an automated impella controller (aic). the battery did not charge. the aic was not involved during patient use. no harm or injury. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the root cause of the failure to charge the console battery set is most likely due to a defective power supply. this report is being filed as part of a retrospective review of historical records.
15816 | 1220648-2024-22119 | 20614454 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20614454&pc=OZD | 04/15/2021 | 11/05/2024 |  | 11/05/2024 |  |  | 04/21/2021 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1263274 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The field service engineer reported that the automated impella controller generated a "purge line click-on not detected" error when the purge cassette was connected. the reported indicated the lever for the transmitter slot was damaged. the purge sensor unit was replaced and the error reportedly was resolved. there was no patient harm reported. | [5 empty] | E.1 name prefix/title, first and last name of the initial reporter are unknown. the investigation into purge disc/flag issues has been completed. the impella automated controller was returned for investigation. the data analysis of the logs for this failure mode was not necessary. a visual inspection of the returned device indicated that the purge flag in the purge disc retainer was broken. the field service engineer replaced the purge flag. the cause of the "purge disc not detected" alarm was due to a broken purge flag. the cause of the broken purge flag was likely due to excessive force inserting and removing previous purge flags. this excessive force can potentially cause harm to the plastic componentry. this report is being filed as part of a retrospective review of historical records.
15817 | 1220648-2024-22122 | 20614659 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20614659&pc=OZD | 02/21/2021 | 11/05/2024 |  | 11/05/2024 |  |  | 04/20/2021 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1095483 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced a controller error yellow alarm. there was no reported patient involvement. | [5 empty] | The device was returned and evaluated, and the investigation for the reported controller error (yellow alarm) has been completed. field service observed there was a controller error present in the logs during service. this alarm was triggered during the setup of a case using the impella cp. the console was investigated. the logs revealed there was a loss of communication to the optical bench given a controller error alarm was triggered. this presented issues getting through the case start procedure. the user attempted to disconnect the pump but it was not recognized. an emergency shutdown had to be performed to turn the console off. the console was tested thoroughly on an engineering bench, but the failure mode could not be reproduced. there were no optical bench-related alarms triggered during more than 24 hours of testing with an optical pump simulator and purge. the root cause of the controller error could not be determined because the failure mode was not reproduced during testing. this report is being filed as part of a retrospective review of historical records.
15818 | 1220648-2024-22112 | 20614739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20614739&pc=OZD | 04/27/2021 | 11/05/2024 |  | 11/05/2024 |  |  | 04/27/2021 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1220360 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced a controller error yellow alarm. there was no reported patient involvement. | [5 empty] | The investigation for the reported controller error (yellow alarm) has been completed. the console logs from the reported day of event are consistent with the complaint and showed an alarm (impella defective) during boot-up, presumably while the pump was not connected to the console. several power-cycles were performed, but the issue was not resolved. analysis of console logs prior to the reported event revealed that on the same day the console failed to detect the pump being disconnected (which was verified in real time logs) and as a result was unable to be shut down (as the console would not allow to be shut down with the pump attached). there were also some communication issues (framing errors) days prior to the event. after console arrived for investigation, the issue was reproduced. because the evidence contained in the console logs pointed to impellatronic malfunction, the board was temporarily replaced with a known-good one, and the issue was resolved. after impellatronic was replaced, the console successfully passed functional check. the cause of the controller alarm was a defective impellatronic board. the cause of the defective impellatronic board is unknown. this report is being filed as part of a retrospective review of historical records.
15819 | 1220648-2024-22131 | 20614780 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20614780&pc=OZD | 03/09/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 03/09/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1127432 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant notified the field service logistics coordinator that the automated impella controller (aic) generated a battery failure alarm. a loaner was sent to the customer to initiate return of this device. no reported patient involvement. | [5 empty] | The investigation of the reported battery failure has been completed. the impella automated controller was returned for investigation. the data analysis of the logs confirmed the reported failure. the root cause of the controller failure was due to a defective battery 1. this report is being filed as part of a retrospective review of historical records.
15820 | 1220648-2024-22134 | 20614931 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20614931&pc=OZD | 03/06/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 03/06/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1245452 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 65 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. while on support, the aic alarm went off for controller failure. the console was switched out without any adverse impact to the patient. | [5 empty] | The investigation for the reported controller failure (red alarm) issue has been completed. the product was returned, and the issue was confirmed. data analysis: complaint date is consistent with complaint intake. imc logs from the complaint show the following: - several power_mod_1 comm errors - controller error (loss of comm (pbm1)) device analysis: the console arrived in danvers. the controller alarm was caused by pbm communication failure resulting from corrosion due to electrolyte leak from supercaps c69, c70. per the results of the clinical review and investigation, the root cause was determined to be due to corrosion from components c69 and c70. this report is being filed as part of a retrospective review of historical records.
15821 | 1220648-2024-22129 | 20614993 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20614993&pc=OZD | 06/17/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 06/17/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1363939 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient with high-risk percutaneous coronary intervention was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) experienced optical signal issues with no resolution specified. no patient harm reported. | [5 empty] | The investigation has been completed. the console (ic6106) was sent back for further investigation. the root cause of the aic issue was a defective optical bench. this report is being filed as part of a retrospective review of historical records.
15822 | 1220648-2024-22147 | 20615162 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20615162&pc=OZD | 03/29/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 03/29/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1230748 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported an automated impella controller (aic) console not recognizing a purge cassette. upon visual inspection the purge disc and purge sensor did not appear to be broken/missing. no patient was involved. | [5 empty] | The automated impella controller (aic) was returned for evaluation. upon inspection of the console, the purge disk detect flag and retainer were not compromised and the purge motor shaft would rotate freely when unpowered. when tested with a purge cassette the issue was reproduced and the motor would not advance. however, when the pud pca was temporarily replaced, the issue was resolved and the purge pope rated within specification. the purge pud board was replaced and a functional check of the console was performed. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15823 | 1220648-2024-22139 | 20615189 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20615189&pc=OZD | 01/17/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 01/17/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1565140 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported automated impella controller (aic) battery and battery communication failures occurred. the aic was replaced, and the case continued. there was no reported patient harm. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the battery failure was due to a defective battery (rapid decline in cell voltage). the exact cause of the defective battery was utd. this report is being filed as part of a retrospective review of historical records.
15824 | 1220648-2024-22181 | 20615208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20615208&pc=OZD | 01/26/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 01/26/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1287975 |  | 0042-0010-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. during preparation of the device, prior to backloading on the 00.18inch guidewire, a yellow alarm on the aic informed that the console couldn't detect the optical sensor. the device was replaced with another aic at the customer's site. no report of patient harm or injury. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, there were no error/failure was reproduced during long term simulation with test pump and purge. the optical bench was also tested without problem in parameters. given the testing and data analysis for the console, it was concluded that the system failure optical sensor error was highly likely due to pump not fully connected to blue receptacle to make enough optical connection when the event occurred. however, the root cause was unable to be determined. before returning this console back to the customer, the optical bench was replaced and a functional check on the console was performed. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15825 | 1220648-2024-22138 | 20615248 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20615248&pc=OZD | 04/15/2021 | 11/05/2024 |  | 11/05/2024 |  |  | 04/15/2021 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1153951 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced a controller error yellow alarm which occurred upon boot up and did not resolve after multiple reboots. there was no reported patient involvement. | [5 empty] | The investigation for the reported controller error (yellow alarm) has been completed. the reported day of event was 2021-04-15, but the actual day of event represented in logs was 2021-04-13, and console logs from that day are consistent with complaint and show an alarm (purge pressure sensor defective) immediately after console boot-up. the logs also showed that the purge pressure sensor issue occurred for the first time at the end of prior case on (b)(6) 2020, when purge pressure disk was being removed. during testing of the console in engineering, the issue was reproduced and purge diagnostics showed sensor defect 2, despite purge pressure disk not being inserted. visual inspection of the purge assembly revealed misaligned ribbon cable between pressure transducer and the pud pca. after realigning the cable, the issue was not resolved, so the transducer assembly was temporarily replaced with a known-good one, and then the issue was resolved. the cause of the controller alarm on boot-up was a defective purge pressure transducer assembly (pressure sensor board). the reason for purge sensor board defect was not determined. this report is being filed as part of a retrospective review of historical records.
15826 | 1220648-2024-22128 | 20615260 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20615260&pc=OZD | 03/10/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 03/11/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, JP |  |  | ABIOMED, INC. |  |  | US |  | 1459449 |  | 0042-0000-JP |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. after patient support, the automated impella controller (aic) did not recognize when the impella pump was disconnected. there was no patient involvement, as the device was no longer in patient use. | [5 empty] | D8/9: device returned to field service, date unknown. the device was returned and evaluated, and the investigation for the reported shutdown/restart issue has been completed. console logs from the reported day of event are consistent with complaint and show that after 9 days of support with pump 352692, after p-level was reduced to 0 and pump stopped, console presented with error message "impella connected. please disconnect impella before shutdown." in an attempt to troubleshoot on-site, pump was momentarily re-connected then disconnected 3 times (as evidenced by placement signal values recorded in rtlog), but the issue persisted and console was unable to register the fact that pump was being disconnected. multiple additional alarms were recorded in logs, all associated with the console erronuously detecting the pump, and eventually alarm #50 "emergency shutdown imminent" was presented right before log entries ended abruptly, indicating a shutdown. after console restarted, it presented with unexpected controller shutdown alarm (# 603, which naturally follows emergency shutdown procedure), but console was able to be properly shut down after that. after receiving the console at abiomed japan field service, the issue was unable to be reproduced during testing, however these symptoms have been observed in the past and are known to be caused by malfunction of one of pump detection mechanisms. the cause of the aic issue was a defective impellatronics board. this report is being filed as part of a retrospective review of historical records.
15827 | 1220648-2024-22091 | 20615273 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20615273&pc=OZD | 04/06/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 04/06/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1213827 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The us complainant had an automated impella controller (aic) fail at case startup. the instructions for use (ifu) were followed, and the aic was replaced. no patient harm was reported. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, the failure mode was persistently reproduced on an engineering workbench. the pmd circuitry was supplied the proper voltage (supply led was illuminated) but the communication issues persisted. momentarily replacing the impellatronic board with a known-good resolved the issue. the root cause of the properly booting up issue was due to a defective impellatronic board. the impellatronic pca was replaced and a functional check on the console and optical system was performed. this report is being filed as part of a retrospective review of historical records.
15828 | 1220648-2024-22140 | 20615300 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20615300&pc=OZD | 03/01/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 03/01/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1075036 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) produced an unexpected controller shutdown alarm after startup. a backup unit was employed prior to connecting a pump, and there was no reported patient involvement. | [5 empty] | The device was returned and evaluated, and the investigation for the reported shutdown or restart issue has been completed. data analysis: all boots recorded through imc logs on the complaint date contained no adverse log events that would lead or exhibit a condition that would trigger an "unexpected controller shutdown" alarm. it was observed that the boot containing that alarm was "skipped." device analysis: console arrived in danvers. console booted with no issues observed. console was power-cycled over 100 times, where it was observed that the console again skipped boots. additionally, while the console was powered on and sitting idle, the imclogger and rtlogger processes were intentionally stopped, however the console behaved as expected. the cause of the unexpected controller shutdown could not be determined due to the inability to reproduce. this report is being filed as part of a retrospective review of historical records.
15829 | 1220648-2024-22168 | 20615327 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20615327&pc=OZD | 02/24/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 02/24/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1167092 |  | 0042-0000-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) displayed a "system check failed, replace console immediately" alarm during startup. there was no reported patient involvement. | [5 empty] | The investigation for the reported system self-check error issue has been completed. the product was returned (return date unknown), and the issue was confirmed. data analysis: imc logs during the complaint date and time are not available. device analysis: console arrived at (b)(6), the failure mode was reproduced. further investigation revealed that there was a problem with the purge unit, fwcheck revealed ¿n.a.¿ for hw revision and version. power measurement was taken for pud voltage supply. 6v and 20v respectively were present. fs replaced the old pud pcb with a new one and sw was upgraded. the console returned to normal working status. per the results of the clinical review and investigation, the root cause was determined to be a defective component. this report is being filed as part of a retrospective review of historical records.
15830 | 1220648-2024-22136 | 20615334 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20615334&pc=OZD | 10/10/2021 | 11/05/2024 |  | 11/05/2024 |  |  | 10/10/2021 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1119116 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) had a failure and pump stop. it was reported the pump was restarted without issues and the controller was changed out. this aic was replaced with another aic resulting in no issues with the replacement aic. there was no patient harm reported. | [5 empty] | The investigation into the optical signal has been completed. the impella automated controller was returned for investigation. the data analysis of the logs confirmed the controller error on the complaint date. the real time lot showed that for a few minutes all signals from optical bench became invalid, with reported values of -16384. however, the pump stop issue reported was not present in the logs around the complaint date. it was determined that the reported issue is a known pattern failure caused by a temporary loss of optical placement signal. abnormal impella stop was not observed through the case. however, it was observed that there was an error of invalid optical placement signal before the pump was recorded for unplugging from console in the logs. it was observed that the patient cable was not fully removed from (barely attached to) the optical receptacle in aic at that moment, which highly likely caused the pump stop reported by staff who thought the pump was stopped while connected to the aic. the pump stop was not due to malfunction, but the disconnection to the aic. the console (b)(6) was tested thoroughly by the post market engineering team, and the failure was not reproduced. this known pattern failure, which is characterized by a temporary loss of optical data and triggering of a controller error alarm, has a low probability of recurrence. the cause of the transient loss of opm-related signals was unable to be determined because the issue was unable to be reproduced. this report is being filed as part of a retrospective review of historical records.
15831 | 1220648-2024-22101 | 20615345 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20615345&pc=OZD | 04/27/2021 | 11/05/2024 |  | 11/05/2024 |  |  | 04/27/2021 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1337492 |  | 0042-0010-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) did not recognize the purge cassette resulting in the display not showing the next step. the complainant tried reinserting the cassette but was unsuccessful. the aic was replaced with another aic without any issues. there was no patient harm reported. | [5 empty] | The investigation into purge disc/flag issues has been completed. the impella automated controller was not received from the customer. the data logs from the date relevant to the complaint date were unable to be provided due to the data not being downloaded. the field service engineer noted that the purge assembly, mainly the purge membrane underneath the purge disc retainer had dextrose build up on it. the engineer cleaned the purge membrane and purge assembly of all built up dextrose and reported that the console successfully passed preventive maintenance. in conclusion, when the staff reported the "purge cassette" was not recognized, the term may have been swapped for "purge disc" as the root cause of the consoles inability to progress through the case start wizard was due to dextrose build up on the purge membrane. the cause of the complaint was most likely due to dextrose build up on the purge membrane stopping the purge disc from making a correct connection with the purge disc retainer. this report is being filed as part of a retrospective review of historical records.
15832 | 1220648-2024-22113 | 20615377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20615377&pc=OZD | 08/03/2021 | 05/21/2025 |  | 11/05/2024 |  |  | 08/03/2021 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1563183 |  | 0042-0010-EU |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) experienced a e power supply issue. the console was replaced to address the issue. no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | Data analysis: the imc logs confirmed the reported failure. there was a battery failure alarm (transport connection blown/ missing ¿ event set #360). a day after the reported complaint date, there was also a battery 1 communication failure alarm (smbus communication loss or sda short ¿ event set #352) due to low pack voltage which could be a result of the battery failing to request charging. the ltpowerdata from the complaint date showed an unknown error for battery 1 which occurred as the battery failure alarm triggered. device analysis: the failure mode was reproduced in fs aachen. the battery 1 pack voltage was measured to be 3.5 v rather than the expected 16 v. the battery set was replaced to resolve the failure alarm, and batttest was recorded to validate charging. the battery charging issues were due to a defective battery 1. the exact root cause of the defective battery 1 was not determined.
15833 | 1220648-2024-22137 | 20615381 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20615381&pc=OZD | 03/03/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 03/03/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1210880 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. after insertion, the automated impella controller (aic) produced a placement signal failure. the aic was replaced, which resolved the issue. there was no reported patient harm. | [5 empty] | The device was returned and evaluated, and the investigation for the reported optical signal issue has been completed. data analysis: aic logs revealed a placement signal not reliable alarm (opm signal invalid) was triggered, followed by log entries indicating the snr was below minimum. rt logs associated with the pump showed optical placement signal spiking which eventually raised to the max value. device analysis: the console was tested thoroughly on an engineering lab bench and did not pass the 5s parameter check in the condition as it was received. bypassing the blue receptacle did not resolve the out-of-spec contrast reading. the low contrast was isolated to the optical bench. replacing the optical bench resolved the low contrast. the root cause of the optical placement signal issues was due to a malfunctioning optical bench. the exact root cause of the malfunctioning optical bench was not determined. this report is being filed as part of a retrospective review of historical records.
15834 | 1220648-2024-22141 | 20615388 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20615388&pc=OZD | 02/24/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 02/24/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1273696 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) was discovered to have a cracked housing and produced alarms upon startup. there was no reported patient involvement. | [5 empty] | The device was returned and evaluated, and the investigation for the reported damage has been completed. device analysis: console arrived in lab. console housing damage was confirmed. when attempting to power on the console, it does not boot correctly. console was opened and interior connections were checked with no issues found. however, damage to the front housing and its major components was observed. this includes: front housing, keyboard, lcd bracket, 4-in-1, power switch. these components were swapped out and the console booted and behaved as expected. the root cause of the power on issues is most likely related to console damage due to misuse of the console. this report is being filed as part of a retrospective review of historical records.
15835 | 1220648-2024-22160 | 20615534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20615534&pc=OZD | 03/30/2021 | 11/05/2024 |  | 11/05/2024 |  |  | 03/30/2021 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1321785 |  | 0042-0000-JP |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller experienced a purge system alarm. the console was replaced to address the issue. no patient harm reported. | [5 empty] | The date of the device return is unknown. the investigation was completed. the device was returned for evaluation. console logs from the reported day of event were consistent with the complaint and confirm alarm (purge system failure - piston blocked). during troubleshooting, the console successfully passed purge system testing and no anomalies were observed. according to local clinical consultant, the root cause of the irregular purge operation was most likely use related, as the personnel operating the console was not familiar with the system. the cause of the failure mode was most likely use related. this report is being filed as part of a retrospective review of historical records.
15836 | 1220648-2024-22150 | 20615578 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20615578&pc=OZD | 03/28/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 03/28/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1121686 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. while on support, the aic failed at start. when impella console was turned on for use during a case, the controller had an error message that said, ¿controller error, switch to back-up controller,¿ and none of the soft keys or rotary knob worked. a backup controller was used and there was no report of patient harm or injury. | [5 empty] | The investigation for the reported controller error (yellow alarm) issue has been completed. the product was returned, and the issue was confirmed. console logs from the reported day of event are consistent with the complaint and show kbd communication errors and associated alarm (# 24) on the day believed to be (b)(6) 2022. the entries in console logs had incorrect timestamps ( (b)(6) 2011 ) due to, what was later determined, rtc coin cell battery issue. on the day of reported issues and kbd comm alarms, there is no entries in the log that would indicate soft-key presses, however aic shutdown was invoked by power-button press and confirmed by rotary push-button press which is inconsistent with complaint claim that rpm did not work at all. on the day prior to (b)(6) 2022, log data shows that soft-key presses were registered, and also timestamps were correct ("(b)(6) 2022"). console was tested in danvers, but the kbd communication issue was not reproduced during repeated power-cycling. prior to testing, the rtc battery was replaced to assure proper timestamps, but also it's been discovered that one of the terminals of the coin cell battery holder on the 4-in-1 carrier was not making proper mechanical and electrical contact with the battery. the rtc issue is unrelated to any kbd communication issues. the communication error and alarm could have been due to issue with either of the components involved in said communication. per the results of the clinical review and investigation: the root cause was determined to be most likely a defective 4-in-1 board. this report is being filed as part of a retrospective review of historical records.
15837 | 1220648-2024-22152 | 20615660 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20615660&pc=OZD | 01/07/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 01/07/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1116077 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) exhibited controller error alarm. no further information available. no report of patient involvement. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, the issue was not reproduced. monitoring of kbd communication during several aic start-up cycles did not reveal any irregularities. nevertheless, due to the age of components that could be responsible for the communication failure (4-in-1 assembly, kbd assembly, ribbon cable), it had been recommended that all the above components be replaced to minimize chance of aic alarm re-occurring. therefore, the root cause of the controller error was due to a defective component(s). the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15838 | 1220648-2024-22105 | 20615701 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20615701&pc=OZD | 08/15/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 08/15/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1147689 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a red alarm upon power up of this console for a patient case. a new console was used without any harm to the patient. no additional information available. | [5 empty] | The investigation for the reported aic - battery failure (red alarm) issues has been completed. the controller device has been returned for evaluation. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15839 | 1220648-2024-22130 | 20615722 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20615722&pc=OZD | 04/19/2021 | 11/05/2024 |  | 11/05/2024 |  |  | 04/19/2021 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1556637 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller had a system self check failure during a functional check. the software was upgraded, but the issue persisted. there was no reported patient involvement. | [5 empty] | The investigation for the reported system self-check error has been completed. clinical review during preventative maintenance in field service the console failed to properly boot up after being upgraded. the logs revealed that the console failed to properly boot up (resulting in system self-check failure) because of communication issues with the power battery manager (pbm) circuit board. the failure mode was reproduced during power cycling the console. check showed that it did not initialize properly. all cable connections were confirmed to be properly aligned in their respective connectors. replacing the pbm resolved the bootup issue. the cause of boot up issue was a defective pbm. the cause of the defective pbm is unknown. this report is being filed as part of a retrospective review of historical records.
15840 | 1220648-2024-22117 | 20615761 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20615761&pc=OZD | 06/28/2021 | 11/05/2024 |  | 11/05/2024 |  |  | 06/28/2021 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1539910 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced an unexpected shutdown during training. no clinical details regarding resolution or patient condition were provided. no patient was involved. | [5 empty] | The investigation into the reported issue has been completed. logs revealed an abrupt stop of logging on the complaint date. the console was powered on a couple more times and continued to unexpectedly shutdown, indicating the issue was persistent. see the attachments section for the logs from the complaint date. the failure mode could be easily reproduced. after isolating the suspect 4-in-1 and compact flash cards and installing them with other known-good subcomponents, the failure could be again reproduced on an aic test fixture. the cause of the unexpected controller shutdown with no successful reboot was due to the 4-in-1. this report is being filed as part of a retrospective review of historical records.
15841 | 1220648-2024-22143 | 20615772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20615772&pc=OZD | 03/15/2021 | 11/05/2024 |  | 11/05/2024 |  |  | 03/30/2021 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1350280 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The clinical consultant reported that the automated impella controller purge disc was not able to be recognized despite multiple attempts. there was no patient harm reported. | [5 empty] | E.1 first and last name of the initial reporter are unknown. the investigation into purge disc/flag issues has been completed. clinical staff reported that the console would not detect the purge disc after multiple attempts. the complaint intake form reported that this complaint was related to a clinical procedure, but there were no adverse events related to this product malfunction. the logs from the complaint date revealed that the console issued the purge disc not detected alarm several times. the console was tested. the purge flag was confirmed to be broken (missing at blue disc retainer). the cause of the purge disc not detected was confirmed to be a broken flag. this report is being filed as part of a retrospective review of historical records.
15842 | 1220648-2024-22144 | 20615789 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20615789&pc=OZD | 11/16/2021 | 11/05/2024 |  | 11/05/2024 |  |  | 11/16/2021 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1091704 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported an automated impella controller (aic) battery failure occurred. the aic was replaced, and the case continued. there was no reported patient harm. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the battery failure alarm was due to a defective battery (rapid decline in cell voltage). the exact cause of the defective battery was utd. this report is being filed as part of a retrospective review of historical records.
15843 | 1220648-2024-22078 | 20615821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20615821&pc=OZD | 03/17/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 03/17/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1094060 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant notified the surgical account manager that the automated impella controller (aic) generated a battery failure alarm. it was reported that the backup console indicated battery failure, would not charge, and the button on the bottom was pushed to confirm the battery was on. there was no patient harm reported. | [5 empty] | The investigation into the battery failure has been completed. the impella automated controller was returned for investigation. the data analysis of the logs confirmed the reported failure, and showed battery failure alarm (360), battery level low alarm (23), battery 2 communication failure alarm (353) and battery failure due to low voltage alarm (350). the issue was not resolved after power-cycling or re-connecting the ac cord. the returned device was tested, and the issue was reproduced. diagnostic information from battest revealed cell 4 of battery 2 having only 400mv (with all other cells over 4130mv), which is indicating of cell imbalance failure mode. the long-term log power data, analyzed later, showed that the cell imbalance was present already over a week earlier, but did not result in any alarms that would be visible to the operator. the cause of the aic issue was a defective battery. this report is being filed as part of a retrospective review of historical records.
15844 | 1220648-2024-22153 | 20615917 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20615917&pc=OZD | 03/22/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 03/22/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1102527 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) had a software update failure and system self check failure, along with constant beeping when powered on. there was no reported patient involvement. | [5 empty] | The investigation for the reported system self-check error issue has been completed. the product was returned, and the issue was confirmed. data analysis: aic logs revealed that there were pmd communication issues after the failed software installation, which resulted in failure of system self check. the logs before the sw installation were not available for review. device analysis: the console troubleshooting was conducted on an engineering lab bench. the impellatronic was momentarily replaced with a known-good to resolve the console start up issue (failing system self check). the console¿s sysconfig file was corrupted because of the failed sw installation and was corrected. per the results of the clinical review and investigation, the root cause of the software upgrade issues was due to a malfunction (likely communication) during the software installation process. this report is being filed as part of a retrospective review of historical records.
15845 | 1220648-2024-22146 | 20616077 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20616077&pc=OZD | 04/02/2022 | 11/05/2024 |  | 11/05/2024 |  |  | 04/02/2022 | 11/05/2024 |  |  |  |  | 11/05/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1445594 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. while on support, a yellow alarm emitted from the automated impella controller (aic) console. the medical staff was unable to change the cassette due to yellow screen ¿unable to detect cassette." nurse tried a second cassette and was unable to load in the console. a new console was turned on and cassette loaded without issue. no patient harm reported. | [5 empty] | The automated impella controller (aic) was returned for evaluation. upon inspection of the console, the purge system contained evidence of purge fluid spill, compromising the hinge of purge lid. during testing, it's been determined that occasionally a good purge cassette is not being detected by the rfid, therefore not activating purge motor. if the gear of the purge cassette happened to be already aligned with the gear of purge motor before and after the insertion, the purge motor was activated. however, if the gears were not initially aligned, the motor would not activate, even when the purge cassette was placed near the purge assembly. this appeared to be the cause of the unusually short range of the rfid circuitry. once temporarily replacing the rfid pca, the issue was resolved and the purge cassette was detected each time was able to be installed. therefore, the root cause of the purge issue was most likely, due to user error in relation to the purge bag and cassette change; as well as the rfid malfunctioning which prevented the purge cassette from being detected and installed. the purge insert and lid were cleaned, the purge pressure sensor was tested, parameters were verified to be within specification. and a functional check of the console was performed; before the console was returned to the customer. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records. issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken (missing at blue disc retainer). before returning this console back to the customer, the purge flag was replaced and a full functional check on the console and optical system was performed. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15846 | 1220648-2024-21794 | 20597106 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20597106&pc=OZD | 10/13/2022 | 11/03/2024 |  | 11/04/2024 |  |  | 10/13/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1381389 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a controller error alarm. the clinician assisted the nurse with replacing this aic with another aic. there was no patient harm reported. | [5 empty] | The investigation into the controller error/loss of opm data has been completed. the impella automated controller was returned for investigation. the data logs confirmed the reported failure mode given there was a controller error alarms (opm comm crc error ¿ event set # 1045) triggered during the clinical procedure. real time (rt) logs confirmed that this was a pattern failure mode in which all signals received from optical bench (placement, snr, contrast, lamp, gain) are interpreted as -16384 values. the analysis of the returned console indicated that this is a known pattern failure, which is characterized by a temporary loss of optical data (placement, snr, contrast, lamp, gain) and triggering of a controller error alarm, has a low probability of reoccurrence. if needed, patient hemodynamics and impella position can be monitored with imaging. the cause of the issue was not determined since the failure was not reproduced in the laboratory. this report is being filed as part of a retrospective review of historical records.
15847 | 1220648-2024-21801 | 20597107 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20597107&pc=OZD | 09/18/2022 | 11/03/2024 |  | 11/04/2024 |  |  | 09/23/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1170092 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that the staff was prepared for a 5.5 implant. as they put the plug into the automated impella controller (aic), failure mode "optical sensor not detected" appeared. they repeat the plug in, but same failure occurred. the aci was replaced onsite. no report of patient harm or injury. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the console was tested, and failure mode was reproduced. the pmd hw revision number was set to 0, which caused issued with supporting optical pumps (compatibility). after issuing a telnet command to correct the number to 257, the console was able to support optical pumps once again. this issue occurred due to a technician error during the preventative maintenance process. this report is being filed as part of a retrospective review of historical records.
15848 | 1220648-2024-21795 | 20597113 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20597113&pc=OZD | 10/07/2022 | 11/03/2024 |  | 11/04/2024 |  |  | 10/07/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1221446 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) failed the self check during maintenance. rebooting the console did not resolve the issue. there was no reported patient involvement. | [5 empty] | The investigation for the reported system self-check error issue has been completed. the product was returned, and the issue was confirmed. data analysis: aic logs showed sau_powermod1 communication errors. see ic4136 aic logs complaint date.pdf. device analysis: pbm corrosion was identified on the top side of the pbm. fcn 71 had been implemented to address other pbm boards manufactured between march and november 2013 that do not yet show signs of damage due to leaking capacitors. pbm super cap is a pattern failure mode and usually occurs with pbms that have a sn between (b)(6) - (b)(6) (inclusive). this pbms sn was (b)(6) which was outside of the date range. per the results of the clinical review and investigation, the root cause was determined to be due to leaking electrolyte on the pbm. this report is being filed as part of a retrospective review of historical records.
15849 | 1220648-2024-21805 | 20597114 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20597114&pc=OZD | 09/06/2022 | 11/03/2024 |  | 11/04/2024 |  |  | 09/06/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1078424 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported that the automated impella controller (aic) experienced a self-check failure. there was no reported patient involvement. | [5 empty] | The investigation for the reported system self-check error issue has been completed. the product was returned, and the issue was confirmed. data analysis: imc logs on the complaint date were consistent. logs show the same repeated entries because console was left on in failed state, they show the following: (b)(6) 2022 12:05:18 send_heartbeat_messages: name_open ncc_sau_simul did not succeed err=no such file or directory. (b)(6) 2022 12:05:19 shmemtools:scumain:: timeout wait for communication channel /dev/name/local/sau_sens_eep_1 10sec. these log lines are representative of communication errors occurring with the epm of the impellatronic board. device analysis: console arrived in danvers. failure mode was reproduced on first boot, console booted into a system self-check error. a known good impellatronic board was swapped into console. console booted without issue or alarms. during replacement, the ribbon cable from impellatronic was found decolorated. the ribbon cable was degenerated but has no contribution to the failure mode. per the results of the clinical review and investigation, the root cause was determined to be a defective component. this report is being filed as part of a retrospective review of historical records.
15850 | 1220648-2024-21797 | 20597116 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20597116&pc=OZD | 10/05/2022 | 11/03/2024 |  | 11/04/2024 |  |  | 10/05/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1257550 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The biomedical engineer reported that the automated impella controller (aic) experienced a controller error alarm when performing a routine check of the console. reportedly unsuccessful attempts were made to restart the console and by pushing the button for the battery kill switch. there was no patient involvement reported. | [5 empty] | E1: name of the initial reporter is unknown. the investigation into the controller error has been completed. the impella automated controller was returned for investigation. the data log reports from the reported day of event are consistent with complaint and show persistent alarm ¿controller error due to opm communication stopped¿ (#1035) upon each boot-up of the console. further analysis of logs prior to that day revealed the same alarm intermittently triggering on (b)(6) 2022 while console was in danvers for its preventive maintenance (pm) procedure. the returned console was tested, and the issue was reproduced. visual inspection of the internal components of the console revealed a misaligned communication ribbon cable between optical bench and 4-in-1 carrier. it is most likely that the cable was accidentally misaligned during prior pm procedure, during which time the aic housing had to be opened in order to replace the console batteries. after the cable was re-seated, communication errors subsided. the cause of the aic issue was a loose connector. this report is being filed as part of a retrospective review of historical records.
15851 | 1220648-2024-21800 | 20597117 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20597117&pc=OZD | 09/26/2022 | 11/03/2024 |  | 11/04/2024 |  |  | 09/26/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1383039 |  | 0042-0000-JP |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  | Asian | Death | No | Yes | Application Program Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. the complainant reported a controller error alarm during use of the device. although no flow rate or purge issues were impacted, the pump was switched proactively to a new console. the procedural outcome noted that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: the date of patient death is unknown. e1: the name of the initial reporter is unknown. the investigation into the controller failure (red alarm) issue has been completed. the automated impella controller (aic) was not returned for investigation. the device logs confirmed the reported failure mode given there was a controller error alarms triggered during the clinical procedure. this known pattern failure, which is characterized by a temporary loss of optical data (placement, snr, contrast, lamp, gain) and triggering of a controller error alarm, has a low probability of reoccurrence. if needed, patient hemodynamics and impella position can be monitored with imaging. the root cause is unable to be determined. this report is being filed as part of a retrospective review of historical records.
15852 | 1220648-2024-21804 | 20597122 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20597122&pc=OZD | 09/15/2022 | 11/03/2024 |  | 11/04/2024 |  |  | 09/15/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1239556 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) would not move past the purge cassette setup step. troubleshooting with field service though impella connect attempts were unsuccessful. a loaner was sent to the customer to initiate return of this device. there was no patient harm reported. | [5 empty] | E1: name of the initial reporter is unknown the investigation into the purge issue has been completed. the impella automated controller was returned for investigation. the data analysis of the logs from the complaint date revealed that the console issued the purge disc not detected alarm several times. the returned console was tested and the issue of properly detecting the purge pressure disc was reproduced, and the purge flag was confirmed to be broken. the cause of the issue was a defective component, a broken purge flag. this report is being filed as part of a retrospective review of historical records.
15853 | 1220648-2024-21796 | 20597123 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20597123&pc=OZD | 09/22/2022 | 11/03/2024 |  | 11/04/2024 |  |  | 09/22/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1651060 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had an automated impella controller plugged in and found to have a made an "explosion" occurred. a loaner device was already onsite to utilize. there was no noted harm or injury to the operator or team. | [5 empty] | E1: the name of the initial reporter is unknown. the investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. console logs from the reported day of event did not contain any information relevant to the reported failure. after console arrived in danvers, it¿s been observed that its ac power cord was missing. no damage to the console was observed during testing, inline fuses were intact, and internal power supply operated within specification. aic operated on ac power as well as on batteries. the issue described in complaint must have originated either at the ac wall outlet or the power cord, or could have been caused by misuse, however because the cord set had not been returned for analysis, the root cause could not be determined. this report is being filed as part of a retrospective review of historical records.
15854 | 1220648-2024-21802 | 20597124 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20597124&pc=OZD | 09/22/2022 | 11/03/2024 |  | 11/04/2024 |  |  | 09/22/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1313641 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | No | Yes | Application Program Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. the complainant reported automated impella controller (aic) failure during support. the patient was successfully switched to a new aic device. the patient was on impella cp support for 2 hours prior to expiring on support. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation into the controller failure (red alarm) issue has been completed. the automated impella controller (aic) was returned for investigation. data analysis: the device logs show that alarm #503 was set while the pump was still connected. this alarm was reproduced on a lab bench when logs were being downloaded from the aic, during this time the pump was not connected. alarm #503 refers to voltage-out-of-spec for pump motor controller board. device analysis: the device was tested in the lab and the controller failure alarm was reproduced. according to the analysis of the aic logs and the schematic for impellatronic, it was speculated that the reported failure was due to malfunction of impellatronic board. most likely, a component between the 20v supply coming from the pbm board to dsp on impellatronic malfunctioned. the aic was further tested by replacing the impellatronic pcb with a known good board. a known good pump was ran for 1 hour. the failure did not reproduce with the known good impellatronic pcb. thus, the root cause for this failure was malfunction of a subcomponent (most likely varistor r4162) on the impellatronic board. this report is being filed as part of a retrospective review of historical records.
15855 | 1220648-2024-21799 | 20597125 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20597125&pc=OZD | 09/27/2022 | 11/03/2024 |  | 11/04/2024 |  |  | 09/27/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1341199 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) exhibited a controller failure (red alarm). a loaner was sent to the customer. no report of patient involvement. | [5 empty] | The name of the initial reporter is unknown. the investigation into the controller failure (red alarm) issue has been completed. the automated impella controller (aic) was returned for investigation. the console was booted, and the controller error was reproduced. a new lamp assembly was replaced, and the controller error issue resolved. this report is being filed as part of a retrospective review of historical records.
15856 | 1220648-2024-21803 | 20597126 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20597126&pc=OZD | 09/18/2022 | 11/03/2024 |  | 11/04/2024 |  |  | 09/18/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1353540 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  |  | Malfunction | No | Yes | Battery Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) alarmed "controller failure". placement signal and left ventricle waveform disappeared. flows dropped to the twos. alarms were able to be resolved and all numbers came back to normal within a couple minutes. bedside nurse quickly transferred patient to new aic. no patient harm was reported. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the issue was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records.
15857 | 1220648-2024-21798 | 20597127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20597127&pc=OZD | 10/04/2022 | 11/03/2024 |  | 11/04/2024 |  |  | 10/04/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1193567 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that an aic alarmed with a controller failure alarm. the alarm initially resolved itself; however, the failure reoccurred, and the decision was made to swap the aic with another unit. there was no reported harm to the patient. | [5 empty] | The investigation into the controller failure (red alarm) issue has been completed. the automated impella controller (aic) was returned for investigation. console was tested with demo pump and purge cassette for long time, but no issue was reproduced. purge pressure was able to be maintained and increased as piston forward strokes. purge pressure sensor area was visually inspected without issue. the root cause of the controller failure was not determined because the symptom was unable to reproduce. this report is being filed as part of a retrospective review of historical records.
15858 | 1220648-2024-21819 | 20597747 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20597747&pc=OZD | 10/29/2024 | 12/12/2024 |  | 11/04/2024 |  |  | 10/09/2024 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025555675 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 insertion was aborted due to the 43-year-old male patient small vessel size and tortuosity. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. | The investigation into delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determine to be patient condition because of tortuosity of vessels and the pump was unable to pass through vasculature. the following have been reported incorrectly on manufacturer device report 1220648-2024-21819. b3 date of event. g1 reporting contact email.
15859 | 1220648-2024-21820 | 20597812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20597812&pc=OZD | 10/10/2024 | 12/09/2024 |  | 11/04/2024 |  |  | 10/10/2024 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025551771 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Death | Required Intervention |  |  |  |  | PROTEK DUO WITH OXYGENATOR. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, left ventricle (lv) perforation was visible during repositioning. lv perforation was repaired in the operating room. patient expired while on support in the operating room. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿. ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿. section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿. | The investigation for the ventricle perforation has been completed. neither data logs nor product were returned for investigation. clinical details report that echo showed the impella had shifted toward the mitral valve at the time of the complication, potentially due to repositioning attempts. additionally, there is potential the lv wall was diseased/damaged prior to impella support. due to incomplete and unclear clinical detail and product not being returned, the root cause of the ventricle perforation could not be determined. correction to the initial mfr report: g1 mdr reporting contact email.
15860 | 1220648-2024-21821 | 20597862 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20597862&pc=OZD | 12/29/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 12/29/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1373062 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a broken purge clip. there was no patient involvement. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the controller device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the issue was broken/ missing purge retainer. this report is being filed as part of a retrospective review of historical records.
15861 | 1220648-2024-21822 | 20597969 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20597969&pc=OZD | 10/09/2024 | 01/02/2025 |  | 11/04/2024 |  |  | 10/09/2024 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025527973 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a female patient in an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient developed compartment syndrome on their right side following impella cp explant. the patient survived the explant. no information was provided in regards to treatment. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email revised on manufacturer device report 1220648-2024-21822 in accordance with updated procedures.
15862 | 1220648-2024-21823 | 20597990 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20597990&pc=OZD | 10/10/2024 | 12/06/2024 |  | 11/04/2024 |  |  | 10/10/2024 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025552629 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the physician could not get the impella cp across the aortic valve when repositioning. the impella had been in and out of the body twice and decision was made to exchange the impella cp for a new cp. there was no harm to the patient. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. | The investigation for the positioning issue has been completed. the pump was not returned for investigation. data log review was not required for this case run. according to the clinical details, the impella was unable to cross the aortic valve during repositioning of the device. the cause of the positioning issue was most likely use related, based on given clinical details. corrections to the initial mfr report: g1 mdr reporting contact email
15863 | 1220648-2024-21824 | 20598057 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20598057&pc=OZD | 10/09/2024 | 11/11/2024 |  | 11/04/2024 |  |  | 10/09/2024 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1132457 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller was replaced due to purge issues. | [5 empty] | D.2 revised common device name as it was entered incorrectly on the initial manufacturer device report number 1220648-2024-21824. | The automated impella controller was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
15864 | 1220648-2024-21825 | 20598074 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20598074&pc=OZD | 01/03/2023 | 11/04/2024 |  | 11/04/2024 |  |  | 01/03/2023 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1465870 |  | 0042-0040 |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the aic (automate impella controller) experienced purge disc/flag issues that were resolved by swapping out the console. no patient harm was reported. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the controller device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the issue was broken/ missing purge retainer. this report is being filed as part of a retrospective review of historical records.
15865 | 1220648-2024-21827 | 20598120 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20598120&pc=OZD | 10/10/2024 | 11/26/2024 |  | 11/04/2024 |  |  | 10/10/2024 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025502108 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Difficult to Remove | [6 empty] | Insufficient Information | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 pump placed via the axillary graft to support a patient admitted in for a high-risk surgery/coronary artery bypass graft. the pump support was for 24 hours. the pump was explanted, but the explant could not be done via the graft due to resistance felt. the team had to surgically explant. the patient survived the event. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the reported removal issues has been completed. the product was not returned for analysis. the clinical mentions that there was resistance when trying to remove the impella from the graft. there was no mention of kink or patient related issues. the root cause of the removal issue could not be determined.
15866 | 1220648-2024-21826 | 20598122 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20598122&pc=OZD | 12/29/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 12/29/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1464036 |  | 0042-0010 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the aic (automated impella controller) experienced purge disc/flag issues. no patient harm was reported. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the controller device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15867 | 1220648-2024-21829 | 20598148 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20598148&pc=OZD | 11/17/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 11/21/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1127921 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) produced a battery failure (red alarm). there was no reported patient involvement. | [5 empty] | The device was returned and evaluated, and the investigation for the reported battery failure (red alarm) has been completed. lt log analysis of the battery data reveals that beginning before the complaint date of occurrence, cell 3 voltage of battery 1 had been declining for an extended period of time. it was confirmed that persistent battery failure (transport connection blown/missing) was not due to a disengaged battery switch. when console was booted for the first time in the lap, console alarms are consistent with what was seen in the field. console interior connections were checked and verified. when visually inspecting the batteries and pbm, it was observed that one of the batteries status leds were completely blank. the battery failure alarm was due to a defective battery 1 (decline of individual cell voltage). the root cause of the battery cell failure was undetermined. this report is being filed as part of a retrospective review of historical records.
15868 | 1220648-2024-21830 | 20598253 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20598253&pc=OZD | 11/21/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 11/21/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1432818 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. the complainant reported an unexpected automated impella controller (aic) shutdown during use. the patient was stable throughout and the pump was restarted on a new console. no patient harm was reported. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the issue was not determined since failure mode was unable to be reproduced. this report is being filed as part of a retrospective review of historical records.
15869 | 1220648-2024-21833 | 20598347 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20598347&pc=OZD | 10/10/2024 | 11/27/2024 |  | 11/04/2024 |  |  | 10/10/2024 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025524837 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, the patient was bleeding from chest tubes. additionally, pulsatility was lost. unable to doppler pedal signals in lower left extremity (lle), limb is cooler and paler despite reduced vasopressor burden. fasciotomy to lle. additionally, the patient received blood products and axillary site received washout/exploration, due to copious jackson-pratt drain output. patient went into ventricular tachycardia (vt) and was converted with 200 j, amio bolus, glucose tolerance test up to 1, lidocaine bolus. the patient survived the event. | [5 empty] | The investigation for the limb ischemia, vascular damage, access site bleed, and ventricular tachycardia has been completed. the product was returned and there were do abnormalities noted. due to limited clinical details provided and no damages with the returned pump, the root cause of the access site bleeding - major cannot be determined. the root cause of the vascular damage, limb ischemia, and the vt could not be determined without additional clinical details. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
15870 | 1220648-2024-21834 | 20598365 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20598365&pc=OZD | 11/18/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 11/18/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1651060 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) soft touch buttons were not functioning. the team was decannulating in the operating room and attempting to ramp up the impella support then none of the buttons would respond. this aic was replaced with another aic resulting in the console appropriately ramping for the impella 5.5 support. there was no patient harm reported. | [5 empty] | The investigation into the keyboard rotary knob issue has been completed. the impella automated controller (aic) was returned for investigation. the data analysis of the logs showed that upon pressing the mute button there were kbd communication errors and high guard and eventually a controller error (event set #24). these were uniquely paired with errors indicating both subsystems were affected simultaneously. the failure mode was not reproduced. the cable/ connector misalignment was not observed between the keyboard (kbd) and 4-in-1. errors related to kbd and radio frequency identification (rfid) can be triggered by eternal sources of electromagnetic interference (emi), but no such link could be established from the evidence collected. the root cause of the kbd-related alarm was not determined due to the inability to reproduce the issue. this report is being filed as part of a retrospective review of historical records.
15871 | 1220648-2024-21828 | 20598373 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20598373&pc=OZD | 01/03/2023 | 11/04/2024 |  | 11/04/2024 |  |  | 01/03/2023 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1684985 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the aic (automated impella controller) experienced purge disc/flag issues. no patient harm was reported. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the controller device has been returned for evaluation. clinical details were sufficient to determine the root cause. the root cause of the purge issues was due to a broken retainer. this report is being filed as part of a retrospective review of historical records.
15872 | 1220648-2024-21832 | 20598412 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20598412&pc=OZD | 11/17/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 11/17/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1211529 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | Us complainant reported that the patient was placed onto impella cp support due to cardiomyopathy. while on support, the automated impella controller (aic) experienced a purge issue with no resolution specified. no patient harm was reported. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the console alarms for purge liquid low/ critically low was due to bag sized entered wrongly. this report is being filed as part of a retrospective review of historical records.
15873 | 1220648-2024-21831 | 20598484 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20598484&pc=OZD | 10/14/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 10/14/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1116483 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  |  | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. while on support, the aic experienced controller failure (red alarm) issues that were resolved by replacing the aic. there was no report of patient harm or injury. | [5 empty] | The investigation into the controller failure (red alarm) issue has been completed. the automated impella controller (aic) was returned for investigation. console was tested on engineering lab bench, but the errors were unable to be reproduced. this known pattern failure, which is characterized by a temporary loss of optical data (placement, snr, contrast, lamp, gain) and triggering of a controller error alarm, has a low probability of reoccurrence. if needed, patient hemodynamics and impella position can be monitored with imaging. this report is being filed as part of a retrospective review of historical records.
15874 | 1220648-2024-21839 | 20598494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20598494&pc=OZD | 08/20/2024 | 11/11/2024 |  | 11/04/2024 |  |  | 10/10/2024 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025528109 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | Tachycardia | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured intraprocedural episodes of bradycardia with a possible relationship to the impella device used: during the procedure patient had episodes of bradycardia with rates in the lows 40 bpm and tachycardia with rates in the 110s-120s with a left bundle branch block morphology. the patient recovered with no sequelae. additionally, the complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured bleeding from right groin access site with a probable relationship to the impella device used: bleeding from right groin access site, pressure applied for 20 minutes and pressure dressing applied per interventional cardiologist (ic) fellow. ic fellow reassessed the site and appeared normal. fellow reassessed the next morning and access site appeared normal 8/21/24: i did examine her right femoral access site and noted that there was no significant ecchymosis or hematoma and she had palpable pulses. the patient recovered with no sequelae. furthermore, the complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured right groin hematoma with a probable relationship to the impella device used: no evidence of pseudoaneurysm of the right lower extremity. incidental finding: well-defined hypoechoic nonvascular mass located in the right groin measuring 0.36 x 0.77 cm. patient presents to catheterization lab for evaluation of right groin. patient is hemodynamically stable. groin noted to have small approximately 2cm long by 1 cm wide hard, nontender area of swelling. femoral pulse 2/2. area of bruising on groin noted but healing. patient had ultrasound this afternoon which showed small hematoma but was negative for pseudoaneurysm. the patient outcome is unknown. the patient survived the events. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | G.2: revised report source as study was inadvertently omitted on the initial manufacturer device report number 1220648-2024-21839.
15875 | 1220648-2024-21841 | 20598517 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20598517&pc=OZD | 12/27/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 12/27/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1207294 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the aic (automated impala controller) experienced purge disc/flag issues. no patient harm was reported. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the controller device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15876 | 1220648-2024-21840 | 20598531 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20598531&pc=OZD | 11/16/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 11/16/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1586601 |  | 0042-0010 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Unknown |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported patient was placed on impella cp with smart assist for hemodynamic support. it was no-ted after the connection of a current cable to automated impella controller (aic), the placement signal of a cpsa pump was flas 0/0 mmhg. issue resolved after switching to another aic. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the aic issue was a software defect that has been documented. this report is being filed as part of a retrospective review of historical records.
15877 | 1220648-2024-21837 | 20598534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20598534&pc=OZD | 10/06/2024 | 12/05/2024 |  | 11/04/2024 |  |  | 10/06/2024 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025522241 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Ischemia | [11 empty] | Other | Death | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported not able to find distal pulses on the patient with impella cp support. the patient was taken to the intensive care unit after placement of the impella where patient coded for 20 minutes and expired on support. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email.
15878 | 1220648-2024-21843 | 20598543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20598543&pc=OZD | 11/13/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 11/13/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1261093 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | Insufficient Information | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported an aic controller failure alarm during support. the pump continued to operate; however the aic device was replaced with new controller. there was no report of patient harm or injury. | [5 empty] | E1: name of initial reporter is unknown. the investigation into the controller failure (red alarm) issue has been completed. the automated impella controller (aic) was returned for investigation. the failure was reproduced during investigation causing an unexpected controller shutdown and additional epm errors. opening up the console, the impellatronic board should have two leds illuminated, but neither led was ever illuminated. the impellatronic board was swapped with the burn-in fixture, and while both consoles ran with no failures, the leds on impellatronic boards never lit up. the cause of the aic issue was a defective impellatronics board. this report is being filed as part of a retrospective review of historical records.
15879 | 1220648-2024-21838 | 20598562 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20598562&pc=OZD | 12/29/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 12/29/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1117229 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the aic (automated impella controller) experienced purge disc/flag issues. no patient harm was reported. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the controller device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15880 | 1220648-2024-21842 | 20598620 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20598620&pc=OZD | 11/15/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 11/15/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1166648 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Decrease in Pressure | Defective Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support. the complainant had a console fall and damage occurred during 5.5 support that was escalated from the cp. a backup console used to complete the procedure. multiple purge disc ruptures resulted in leaks. no intervention was specified. no patient harm was reported. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. damage was caused by improper use. this report is being filed as part of a retrospective review of historical records.
15881 | 1220648-2024-21844 | 20598621 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20598621&pc=OZD | 10/11/2024 | 02/11/2025 |  | 11/04/2024 |  |  | 10/11/2024 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1132457 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 31 YR |  | Female |  |  |  | Malfunction | No | Yes | Defective Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an automated impella controller displayed a suction alarm, any time the support level was set above p2. the failure occurred during preparation with no resulting patient harm. | [5 empty] | The investigation of automated impella controller (aic) - false buzzer/ alarm has been completed. the logs from the console confirmed there were numerous suction level alarms issued by the console. during the device analysis, the console's internal circuitry was inspected for any anomalies and no anomalies were found. the console ran from a known good pump and purge for a sufficient amount of time with no anomalies. there was no problem found with the console. it operated to specification. a.1, a.2 and a.3 sections were revised as information was inadvertently omitted from the initial manufacturer device report number 1220648-2024-21844. d.2 revised common device name since the initial manufacturer device report number 1220648-2024-21844 was submitted in accordance with updated procedures. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2024-21844 was submitted in accordance with updated procedures. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
15882 | 1220648-2024-21835 | 20598636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20598636&pc=OZD | 10/10/2024 | 11/26/2024 |  | 11/04/2024 |  |  | 10/10/2024 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025522219 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Vascular Problem | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, bleeding was noted but the source of the bleeding was not found. cangelor and heparin off. the patient received two units of packed red blood cells. the patient survived the event. | [5 empty] | The investigation for the reported vascular damage has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vascular damage could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
15883 | 1220648-2024-21836 | 20598643 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20598643&pc=OZD | 11/18/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 11/18/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1114576 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) had a battery failure while powering on. no patient involvement was reported. | [5 empty] | The device was returned and evaluated, and the investigation for the reported battery failure (red alarm) has been completed. following imc logs, it can be seen that each subsequent power cycle, including the initial power cycle, does not have the "ac power disconnected" alarm. ac power is active due to the absence of this alarm. console battery switch was engaged and the console powered on via batteries without issue. the console was also tested with ac power connected and the battery switch was toggled multiple times to ensure functionality. customer likely did not toggle the battery switch to properly engage the batteries after transport. this report is being filed as part of a retrospective review of historical records.
15884 | 1220648-2024-21845 | 20598732 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20598732&pc=OZD | 11/14/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 11/14/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | XXXXXX |  |  |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a red alarm battery failure on the automated impella controller (aic) controller was observed. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the aic issue was a defective battery. this report is being filed as part of a retrospective review of historical records.
15885 | 1220648-2024-21848 | 20598811 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20598811&pc=OZD | 11/17/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 11/17/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1350280 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant had automated impella controller (aic) that failed to power down after pump removal. there was no harm or injury to the patient. a new second back up console was used. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the issue was not determined since issue could not be reproduced. this report is being filed as part of a retrospective review of historical records.
15886 | 1220648-2024-21847 | 20598864 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20598864&pc=OZD | 10/10/2024 | 12/10/2024 |  | 11/04/2024 |  |  | 10/10/2024 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025467963 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during the implant, the patient went into supraventricular tachycardia. the patient was cardioverted and support continued. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation into vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-21847. d6a and d6b entered incorrectly. g1 reporting contact email.
15887 | 1220648-2024-21851 | 20598900 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20598900&pc=OZD | 10/11/2024 | 12/04/2024 |  | 11/04/2024 |  |  | 10/11/2024 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1210718 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller (aic) was being serviced when there was a"system check failed error. the aic was rebooted several times before displaying system check failed twice. no patient harm has been reported. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the console (aic) self check issue has been completed. data logs were reviewed which showed that initial boot up, the console restarted automatically due to a pud communication issue. due to a persistent pud communication issue, the console rebooted several times and displayed ¿system self check failed.¿ upon manual shutdown and reboot, the console rebooted again due to the pud communication issue and displayed ¿system self check failed¿. this confirms the reported failure mode. this console was returned for evaluation and the issue was reproduced. during the initial boot-up, the console restarted automatically due to the pud communication issue, resulting in the ¿system self check failed¿ screen. the fwcheck via telnet confirmed the problem was due to the pud board, as the hw revision was n.a. for the sau_purge. it was identified that the firmware was corrupted upon comparing the extracted firmware with a known good (curroupted pud fw vs known good.png). the reported failure mode was resolved after re-flashing the production firmware. the root cause for the ¿system self check failure¿ was corrupted firmware on the pud board.
15888 | 1220648-2024-21849 | 20598908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20598908&pc=OZD | 11/14/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 11/14/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2023250798 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. a controller failure alarm was noted on the aic. the initial device was replaced with a new controller. there was no report of patient harm or injury. | [5 empty] | The investigation into the controller failure (red alarm) issue has been completed. the automated impella controller (aic) was returned for investigation. console logs from the reported day of event are consistent with complaint and show a controller failure alarm after 8 days of support with pump 395280. the alarm was preceded by several sau_purge communication error messages, and followed by alarms #501 (controller error due to pbm voltage out-of-spec), and failed resets of pud. the issue was reproduced. inspection of the electronics revealed that although pbm was providing nominal voltage to the pud (24v between pin5 and pin4 on j1.1 on pud pca), the green led on pud was not lit, indicating failure of pud. after the board was temporarily replaced with a known-good one, the issues was resolved, and the pud (sau_purge) communication was active. any pbm-related alarms (#501) observed in logs were related to the pud malfunction and were also resolved with a pud replacement. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15889 | 1220648-2024-21850 | 20598917 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20598917&pc=OZD | 11/13/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 11/13/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1417971 |  | 0042-0010 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | Us complainant reported patient was placed on impella cp with smart assist for hemodynamic support. while on support an automated impella controller (aic) placement signal erratic signal was noted. successfully switched aic. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the tight fit for the front panel optical connector may have caused an imperfect connection with the potential for an airgap which would have triggered the errors seen clinically as well as in the lab. this report is being filed as part of a retrospective review of historical records.
15890 | 1220648-2024-21852 | 20598930 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20598930&pc=OZD | 11/12/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 11/12/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1656412 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported patient was placed on impella cp with smart assist for hemodynamic support. while on support the placement signal unreliable alarm went off. the replace console alarm came on, backup automated impella controller (aic) was utilized. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the likely cause for placement signal issues was detritus causing an airgap at the optical connector. this report is being filed as part of a retrospective review of historical records.
15891 | 1220648-2024-21853 | 20598973 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20598973&pc=OZD | 10/10/2024 | 12/23/2024 |  | 11/04/2024 |  |  | 10/10/2024 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025513032 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Female |  | Non Hispanic |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 53-year-old female patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. the pump support was ongoing when there were concerns of high purge pressure. the team consulted for adding the tpa lytic to the purge fluid. additionally there was noted heparin-induced thrombocytopenia (hit). | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation for the high purge pressure has been completed. impella 5.5 returned for evaluation. upon visual inspection, dried blood and mold was present on the pump and biomaterial was present on the outside of the pump. no biomaterial was present in the purge gap. pump was soaked in warm water while purging for approximately 5 minutes. high purge pressure did not reproduce, pump was then run at p-9 and high purge pressure again did not reproduce. biomaterial from the pump was kicked off when running as well. no data logs were returned for evaluation. clinical details noted that tissue plasminogen activator (tpa) protocol was initiated after high purge pressure alarms and was able to resolve high purge pressure. the root cause of the high purge pressure was most likely due to biomaterial.
15892 | 1220648-2024-21854 | 20598989 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20598989&pc=OZD | 10/10/2024 | 12/10/2024 |  | 11/04/2024 |  |  | 10/10/2024 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025567114 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The us complainant had a cp removed after clotting off. the pump had been placed to support a patient admitted in for a high-risk percutaneous coronary intervention. the cp was replaced. the patient¿s outcome is unknown. | [5 empty] | The investigation into thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-21854: e4 should have been left blank. g1 reporting contact email. g1 reporting contact fax number missing. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
15893 | 1220648-2024-21856 | 20599015 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20599015&pc=OZD | 10/10/2024 | 12/06/2024 |  | 11/04/2024 |  |  | 10/10/2024 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025554000 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported that during prep the impella cp was noted to be damaged. it is unclear what type of damage was present. the decision was made by the doctor to use a different pump. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation into product damage (sterile sleeve damage) has been completed. the impella device was not returned for evaluation. as per the given clinical details, sterile sleeve was ripped during pump preparation. the cause of the sterile sleeve damage issue was use related, based on given clinical details. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-21856. e4 should have been left blank. g1 reporting contact email. g1 reporting contact fax number missing.
15894 | 1220648-2024-21859 | 20599016 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20599016&pc=OZD | 10/11/2024 | 12/11/2024 |  | 11/04/2024 |  |  | 10/11/2024 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025552447 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support had distal pulses absent. the patient was cool to touch with three second cap refill. the patient was assessed for escalation to impella 5.5 and impella cp support was for eight hours. the patient survived the event. | It was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. corrections to the initial mfr report: b2-should have selected death and added date of death. b5-mso review. g1 mdr reporting contact email. h1-changed to death. h6 impact code 1802 added.
15895 | 1220648-2024-21860 | 20599021 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20599021&pc=OZD | 10/11/2024 | 12/19/2024 |  | 11/04/2024 |  |  | 10/11/2024 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025553445 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 47 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient had an attempted impella 5.5 implant and it was aborted due to delivery issues. the patient's right clavicle artery was slightly too small. the catheter was kinked and implant was aborted. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings and cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings and cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings and cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings and cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation into delivery issues has been completed. the root cause of the delivery issue was most likely patient condition related, the diameter of the right clavicle artery was slightly insufficient, so impella could not be inserted as per the clinical description provided. the following have been reported incorrectly or missing from manufacturer device report: e4 should have been left blank in the initial report. g1 reporting contact email revised in accordance with updated procedures.
15896 | 1220648-2024-21855 | 20599027 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20599027&pc=OZD | 10/10/2024 | 12/11/2024 |  | 11/04/2024 |  |  | 10/10/2024 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1734780 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The us complainant had an automated impella controller (aic) failure of the optical sensor. the team replaced the aic and the pump functioned and support began. no harm or injury from issue. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the optical signal issues has been completed. data logs were reviewed which confirm that the optical sensor system error message was displayed by console during case start after the console failed optical bench calibration with low snr. additionally, a placement signal not reliable alarm was issued twice on the date of the occurrence. the pump was unplugged and plugged into the console twice before the console was rebooted and the pump re-plugged with no resolution of the issue. the device was returned for evaluation. the front panel optical connector was inspected. the contamination cleared from the connector after cleaning. when a known good pump and purge line was run with console. the console operated as expected. the root cause for optical sensor system error was most likely contamination of the front optical connector creating an air gap upon insertion of pump into console.
15897 | 1220648-2024-21861 | 20599035 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20599035&pc=OZD | 10/12/2024 | 12/13/2024 |  | 11/04/2024 |  |  | 10/12/2024 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025564556 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic | Asian | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | It was noted that the patient developed hemolysis. | The complainant reported that the impella cp could not flow higher than p2. console was exchanged and the flows were higher. additionally, the impella was repositioned several times. impella was still in mitral position. the plan is to escalate to 5.5 the next day in the morning. the patient survived the event. |  |  |  |  | The investigation into low pump flow, positioning issues, and hemolysis has been completed. product was not returned for review. based on clinical description and data review, the root cause of low flow was most likely due to biomaterial ingestion. the root cause of hemolysis was most likely due to biomaterial ingestion. the root cause of positioning issue was unable to be determined as limited clinical details were provided and no product was returned for review. instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-21861. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
15898 | 1220648-2024-21862 | 20599041 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20599041&pc=OZD | 10/13/2024 | 12/16/2024 |  | 11/04/2024 |  |  | 10/13/2024 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025555005 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  |  | Asian | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Arrhythmia | Tachycardia | [11 empty] | Required Intervention | [5 empty] | EXTRA CORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, the patient was having supraventricular tachycardia (svt) and heart failure. physician repositioned. additionally, the sterile sleeve was found to be damaged. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation into product damage (sterile sleeve damage) and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-21862. e4 should have been left blank. g1 reporting contact email. g1 reporting contact fax number missing.
15899 | 1220648-2024-21865 | 20599050 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20599050&pc=OZD | 10/10/2024 | 12/09/2024 |  | 11/04/2024 |  |  | 10/10/2024 | 11/04/2024 |  |  | 01/31/2025 |  | 11/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024332110 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, the patient had ventricular tachycardia runs overnight. an echocardiogram was done at bedside and the team repositioned pigtail away from lateral wall and ectopy was better. support continued until weaning was successful, the device was explanted, and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation into vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly on manufacturer device report 1220648-2024-21865.
15900 | 1220648-2024-21863 | 20599058 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20599058&pc=OZD | 10/24/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 10/24/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | XXXXXX |  | 0042-0010 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a battery failure error during testing of the automated impella controller (aic). no further information was provided. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the battery failure was due to depleted batteries due to infrequent use and charge. this report is being filed as part of a retrospective review of historical records.
15901 | 1220648-2024-21846 | 20599067 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20599067&pc=OZD | 10/06/2024 | 11/27/2024 |  | 11/04/2024 |  |  | 10/06/2024 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025507251 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support had a gastrointestinal bleed. heparin was stopped, volume resuscitation, and two units of packed red blood cells were given. it was further reported that care was withdrawn and patient expired on impella support. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the vascular damage has been completed. data logs are not relevant to the investigation and product was not returned for investigation. clinical details provided no information regarding why/how the complication could have occurred. due to lacking clinical detail and product not being returned, the root cause of the vascular damage could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
15902 | 1220648-2024-21857 | 20599116 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20599116&pc=OZD | 11/10/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 11/10/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1108091 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a controller error, with notification to switch to a new controller. it was reported the aic was turned off and then re-started with no issues. however, the device was removed from use regardless of the console restarting without any issues. there was no patient harm reported. | [5 empty] | E1: name of the initial reporter is unknown. the investigation into the controller error has been completed. the impella automated controller was returned for investigation. the data logs revealed on (b)(6) 2022, on start up the optical bench recorded 5 signal measurements with ~0.5 v twice before triggering a controller error (defective optical sensor lamp) and console was user shut down. the console was started up again about 10 hours later with the same deficient light voltages and controller error. the console was power cycled at this point and the console started up fine and was started up again five days later with no errors. visual inspection of the returned optical bench showed no loose wires (attempted jiggling wires for connectivity). the console was started up and received normal 5 signal measurements. removing power to the lamp provides the same 5 signal results as received in the field. upon inspection of the front panel optical connector (0042-2750), the connector had vertical smudges that implied attempted cleaning at some point in the past. despite the introduction of dirtiness, the lamp voltage only dropped to the ~1.7v range and did not go down to 0.5v. the cause of the aic issue was a defective optical bench. this report is being filed as part of a retrospective review of historical records.
15903 | 1220648-2024-21858 | 20599119 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20599119&pc=OZD | 10/11/2024 | 12/11/2024 |  | 11/04/2024 |  |  | 10/11/2024 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025570565 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient had an attempted impella 5.5 implant and during the insertion, the pump kinked at the junction of the motor and the catheter. the pump and wire were removed and a different pump was implanted. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation for the delivery issue has been completed. the device was returned for evaluation. no issues were found in the returned product. however, clinical details states that the wire prolapsed into the descending aorta and kinking the catheter at junction of motor and catheter. the root cause of the delivery issue was most likely patient condition.
15904 | 1220648-2024-21864 | 20599127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20599127&pc=OZD | 10/11/2024 | 12/05/2024 |  | 11/04/2024 |  |  | 10/13/2024 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1585312 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant in israel had an impella cp-automated impella controller (aic) support ongoing for a patient admitted in for a high risk percutaneous coronary intervention. the aic was replaced due to optical signal issues. the patient survived the event. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into optical signal issue has been completed. after reviewing the console logs, the reported optical sensor defect issue was confirmed. there were numerous placement signal not reliable (snr) and controller error alarms observed in the logs from the reported complaint date. the console was tested and the reported optical sensor defect was unable to be reproduced. however, the snr was found to be well below specifications after checking the parameters of the optical bench. the root cause of placement signal not reliable issue was a defective optical bench (low snr). the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-21864. b3 date of event e4 should have been left blank. g1 reporting contact fax number missing. h6 code 4629 and 925 were reported incorrectly. new code was added to health effect - impact code.
15905 | 1220648-2024-21868 | 20599311 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20599311&pc=OZD | 11/04/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 11/09/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1422863 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a controller error. there was no patient harm reported. | [5 empty] | The investigation into the controller error/loss of opm data has been completed. the impella automated controller was returned for investigation. the data logs revealed the reported day of event are consistent with the complaint and indicate that the reported issue is a known pattern failure caused by a temporary loss of optical placement signal. there¿s no need for repairs if the condition self-correct. this is a low probability event and lasts only a few minutes. if needed, the patient¿s hemodynamics and impella position can be monitored with imaging. the cause of the issue was not determined since the issue could not be reproduced. this report is being filed as part of a retrospective review of historical records.
15906 | 1220648-2024-21866 | 20599332 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20599332&pc=OZD | 10/12/2024 | 12/12/2024 |  | 11/04/2024 |  |  | 10/12/2024 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025547738 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a patient admitted in acute myocardial infarction/cardiogenic shock and taken to the cardiac catheterization lab for impella cp implant. the patient was bilaterally accessed but the femorals did not allow for pump delivery. the pump was aborted. no harm or injury was noted from the attempt. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation into delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was determined to be patient condition because of patient anatomy. g.1 revised reporting contact email since initial manufacturer device report number 1220648-2024-21866 was submitted.
15907 | 1220648-2024-21870 | 20599464 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20599464&pc=OZD | 10/08/2024 | 12/02/2024 |  | 11/04/2024 |  |  | 10/08/2024 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025551760 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Ventricular Fibrillation | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 70-year-old male patient noted as high risk intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient went into atrial fibrillation with rapid ventricular response with hypotensive episode requiring increased impella support and increasing medication. the following day, the patient coded ventricular fibrillation arrest. heparin was stopped after first code. the patient had a poor prognosis. the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the reported ventricular fibrillation/ventricular tachycardia (vf/vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf/vt could not be determined.
15908 | 1220648-2024-21871 | 20599510 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20599510&pc=OZD | 10/12/2024 | 12/10/2024 |  | 11/04/2024 |  |  | 10/13/2024 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025541861 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that purge line block alarm occurred on the impella cp. tissue plasminogen activator was used but alarm continued. the patient was taken to the catheterization lab and impella was exchanged. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. | The investigation for the high purge pressure has been completed. neither data logs nor product were returned for investigation. per clinical details, tissue plasminogen activator (tpa) procedure did not resolve the complication, but the pump was explanted shortly after the complication to avoid a pump stop. based on limited clinical information and product not being returned, the root cause of the high purge pressure could not be determined. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high-risk cpi states the following: section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ added h6 impact code 4644 corrections to the initial mfr report: g1 mdr reporting contact email.
15909 | 1220648-2024-21867 | 20599541 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20599541&pc=OZD | 10/10/2024 | 12/06/2024 |  | 11/04/2024 |  |  | 10/10/2024 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025471072 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump placed to support a patient admitted in with heart failure and valvular disease, and awaiting final decisions and heart transplant. it was reported that a resident repositioned the device and pulled inflow across aortic valve into aorta, unable to recross aortic valve. outflow noted in vascular graft and a decision was made to bring the patient to the operating room for device replacement. the patient remained stable and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the positioning issue has been completed. the pump was not returned for investigation. data log review was not required for this case run. according to the clinical details, the pump was pulled out of the ventricle and was unable to cross again during repositioned the device. the pump was replaced with a new one. the cause of the positioning issue was most likely use related, based on given clinical details. corrections to the initial mfr report: g1 mdr reporting contact email h8 usage of device is initial use of device | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
15910 | 1220648-2024-21869 | 20599554 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20599554&pc=OZD | 11/03/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 11/03/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1191151 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) failed to turn on. the device was removed from service as a result of the noted failure. a loaner was subsequently provided to the customer site for support as needed. no report of patient involvement. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the console was tested in the lab, the cause was not determined due to inability to reproduce the reported issue. this report is being filed as part of a retrospective review of historical records.
15911 | 1220648-2024-21876 | 20599775 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20599775&pc=OZD | 11/24/2021 | 11/04/2024 |  | 11/04/2024 |  |  | 11/24/2021 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1554466 |  | 0042-0010-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) exhibited a battery failure; red alarms and yellow alarm, and there was no mention of patient involvement, harm, or medical intervention. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, the device logs were analyzed and confirmed the failure. the cause of the battery failure was due to a defective battery (rapid decline in cell voltage). the exact cause of the defective battery was utd. this report is being filed as part of a retrospective review of historical records.
15912 | 1220648-2024-21872 | 20599800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20599800&pc=OZD | 11/12/2021 | 11/04/2024 |  | 11/04/2024 |  |  | 11/12/2021 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | IMPELLA CP |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1431278 |  | 0042-0040-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automated impella controller (aic) experienced a battery failure (red alarm). no patient involvement. | [5 empty] | E1: name of initial reporter is unknown. e2/e3: unknown the investigation into the battery power issue has been completed. the automated impella controller (aic) was returned for investigation. the cause of the battery failure alarm was due to a defective battery (rapid decline in cell voltage). the exact cause of the defective battery was undetermined. this report is being filed as part of a retrospective review of historical records.
15913 | 1220648-2024-21873 | 20599812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20599812&pc=OZD | 11/21/2021 | 11/04/2024 |  | 11/04/2024 |  |  | 11/21/2021 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1323489 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported patient was placed on impella cp for hemodynamic support. while on support, the automated impella controller (aic) experienced battery communication failure (yellow alarm). a loaner was sent to the customer, no report of patient harm or injury. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. additionally, the device logs were reviewed and confirmed the failure. the cause of the battery failure was due to a defective battery (rapid decline in cell voltage). this report is being filed as part of a retrospective review of historical records.
15914 | 1220648-2024-21877 | 20599823 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20599823&pc=OZD | 12/01/2021 | 11/04/2024 |  | 11/04/2024 |  |  | 12/01/2021 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1113656 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an automated impella controller (aic) had battery failure. no patient was involved. should an automated impella controller (aic) fail in patient use the instructions for use is to be followed, and a backup console utilized for support. | [5 empty] | E1: name of initial reporter is unknown. e2/e3: unknown. the investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of intermittent battery communication error was a defective pbm. the cause of the defective pbm is unknown. this report is being filed as part of a retrospective review of historical records.
15915 | 1220648-2024-21875 | 20599884 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20599884&pc=OZD | 09/07/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 09/07/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1462131 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that during preparation, a failure occurred, and the automated impella controller (aic) did not recognize the purge cassette. a replacement controller was required, and the case resumed. no patient was involved. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. there was no issue found during the case that aligned with the reported failure mode. the device functioned/operated to specification. this report is being filed as part of a retrospective review of historical records.
15916 | 1220648-2024-21874 | 20599909 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20599909&pc=OZD | 11/24/2021 | 11/04/2024 |  | 11/04/2024 |  |  | 11/24/2021 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1554466 |  | 0042-0010-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant was testing automated impella controller (aic) in non-patient use and found a battery failure. no patient involvement. | [5 empty] | The investigation into the battery power issue has been completed. the automated impella controller (aic) was returned for investigation. the cause of the battery failure alarm was due to a defective battery (rapid decline in cell voltage). the exact cause of the defective battery was undetermined. this report is being filed as part of a retrospective review of historical records.
15917 | 1220648-2024-21878 | 20599975 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20599975&pc=OZD | 11/07/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 11/07/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1567887 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that during preparation, a purge system alarm failure was observed post changing the purge cassette/tubing. a replacement automated impella controller (aic) was used, and the case resumed. no patient harm was reported. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the purge system failure was not determined. this report is being filed as part of a retrospective review of historical records.
15918 | 1220648-2024-21881 | 20599995 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20599995&pc=OZD | 06/02/2023 | 11/04/2024 |  | 11/04/2024 |  |  | 06/02/2023 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1108088 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Emits Odor | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) had a purge flow malfunction and smoke smell coming out of the device during patient support. no patient harm reported. | [5 empty] | E1: name of initial reporter is unknown. the investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the issue was not determined since issue could not be reproduced. this report is being filed as part of a retrospective review of historical records.
15919 | 1220648-2024-21879 | 20600049 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20600049&pc=OZD | 10/06/2024 | 11/20/2024 |  | 11/04/2024 |  |  | 10/06/2024 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025507357 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella 5.5 support to assist in recovery of the kidneys and stabilize the patient with lymphatic malformation and right coronary artery coronary artery disease. it was reported that the patient had three left chest hematomas which all required evacuation, multiple blood products given, and stopping heparin. it was further reported that due to complication of coagulopathy and internal bleeding from gastrointestinal bleed that was unsuccessfully resolved, the family decided to withdraw care and utilized comfort measures. the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into access site bleeding and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding - major was most likely patient condition related. the root cause of the vascular damage - peripheral vessel was not determined as insufficient clinical information was provided. e4 should have been left blank on manufacturer device report 1220648-2024-21879.
15920 | 1220648-2024-21880 | 20600090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20600090&pc=OZD | 10/06/2024 | 12/10/2024 |  | 11/04/2024 |  |  | 10/06/2024 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025510585 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported that a patient was on impella cp support. while on support, the patient went in atrial fibrillation (af) with rapid ventricular response (rvr). the patient was given amio bolus. additionally, the patient had some bleeding at the insertion site. site was re-sutured and new dressing was placed. nurse held slight manual pressure. care was withdrawn and patient eventually expired. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed and atrial fibrillation (af) has been completed. the device was not returned for evaluation. however clinical details were sufficient to determine root cause. the root cause of the access site bleeding issue was most likely patient movement, some oozing and bleeding at insertion site due to patient movement upon waking up as per the clinical description provided. the root cause of the access site bleed and af could not be determined. corrections to the initial mfr report: g1-mdr reporting contact email.
15921 | 1220648-2024-21882 | 20600202 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20600202&pc=OZD | 06/15/2023 | 09/22/2025 |  | 11/04/2024 |  |  | 06/16/2023 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1651581 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Read Input Signal | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported that their automated impella controller (aic) was unable to detect an installed impella cp pump and showed no aortic placement signal. the aic was swapped to resolve the noted issue. the patient survived the procedure. | [5 empty] | Data analysis revealed there was no evidence of a pump not being detected by the console on the date of complaint. retuned device was analyzed and the pump-not-detected failure mode could not be confirmed because there was no indication that the pump was unrecognized by the console on the complaint date. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. data logs confirm the reported failure mode of optical signal issue. the console was inspected for contamination of optical connector ferrule and revealed to have a high build up of contaminant on the optical connector surface. after cleaning the optical connector, measurement of the consoles 5 s parameters were performed that showed satisfactory optical bench parameter. the console was confirmed for placement signal not reliable alarm resulting from low snr of the optical bench system. the root cause of the optical signal failure was an air gap from between the aic and the patient cable plug or within the middle of the red handle ferrule. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d2 updated with correct common device name. g1 updated with correct mdr reporting contact email from the initial manufacturer device report number.
15922 | 1220648-2024-21883 | 20600381 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20600381&pc=OZD | 09/20/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 09/20/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1454040 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The field service engineer reported that the automated impella controller (aic) would not recognize the purge cassette. a return box was sent to the customer to initiate return of this device. there was no patient involvement, harm, or injury reported. | [5 empty] | E1: name of the initial reporter is unknown the investigation into the purge issue has been completed. the impella automated controller was returned for investigation. the data logs from the complaint date revealed that the console issued the purge disc not detected alarm (event set # 402) occurred once as per the logs, which got resolved in <10 sec. as per the logs, the purge cassette was detected by the aic. per review of the real time logs it can be seen that the purge pressure was not holding during the first few "purge_system_priming_fluid_not_detected" errors. finally, the purge pressure becomes a flat line which might be indicative of a leak. the testing of the returned console indicates the issue with properly detecting the purge pressure disc was not reproduced, and purge assembly had no visible damage. the console was detecting purge disc and cassette when evaluated multiple times on the laboratory bench. the device functioned/operated to specification. this report is being filed as part of a retrospective review of historical records.
15923 | 1220648-2024-21886 | 20600792 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20600792&pc=OZD | 11/19/2021 | 11/04/2024 |  | 11/04/2024 |  |  | 11/19/2021 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1558681 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automated impella controller (aic) experienced a battery failure (red alarm). no patient involvement. | [5 empty] | E1: name of initial reporter is unknown. e2/e3: unknown. the investigation into the battery power issue has been completed. the automated impella controller (aic) was returned for investigation. the cause of the battery failure was due to a defective battery (rapid decline in cell voltage). the exact cause of the defective battery was undetermined. this report is being filed as part of a retrospective review of historical records.
15924 | 1220648-2024-21891 | 20600889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20600889&pc=OZD | 08/24/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 08/24/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1213834 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Use of Device Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) was turned on for clinical use and the device exhibited a red controller error message and prompted the user to use another aic device. the staff tried to reset the aic and reinstall the software, this did not fix the issue. the resolution is not specified; however, it was reported that there was not harm to the patient. | [5 empty] | The console (aic) was returned for evaluation. upon evaluation of the console, the reported issue was duplicated. temporarily replacing the purge sensor (transducer) assembly with a known-good one resolved the issue. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15925 | 1220648-2024-21890 | 20601460 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20601460&pc=OZD | 08/25/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 08/25/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER EU |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER EU |  |  | ABIOMED, INC. |  |  | US |  | 1273766 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that during annual preventive maintenance the automated impella controller (aic) experienced a controller error alarm and could not be restarted within the next five seconds. there was no patient involvement. | [5 empty] | E1: name of the initial reporter is unknown the investigation into the controller error has been completed. the impella automated controller was returned for investigation. the data log reports from the reported day of event were analyzed but it was very difficult to distinguish between alarms and errors that are expected during pump short circuit testing, from those that rarely happen (or those that are not supposed to occur). for that reason, a ¿live test session¿ was scheduled to repeat the testing and help eliminate possibly defective service pump. during testing of the returned device, the issue was not reproduced, however it¿s been recommended to replace the impellatronic assembly as a precaution. the aic behavior is considered acceptable, as long as the console¿s full operation is restored after power-cycle (which was the case). the console later passed the preventive maintenance procedure and electrical safety testing. there was no issue found during the investigation. the device functioned/operated to specification. this report is being filed as part of a retrospective review of historical records.
15926 | 1220648-2024-21889 | 20601536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20601536&pc=OZD | 08/24/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 08/24/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1583419 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on an automated impella controller (aic) for mechanical circulatory support. while on support, there was an aic with a controller error. the aic was replaced, and support proceeded with another aic. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported red alarm issue has been completed. the product was returned, and the red alarm issue was confirmed. data analysis: imc logs are incomplete around the complaint occurred date. the reported controller error alarm was unable to be confirmed. on the other hand, there was an unexpected shutdown observed on the boot after the complaint date. the reason of unexpected shutdown was unable to confirmed, but it was likely related to the observation of frozen screen which suggested operator to force shutdown the console. rtlogs were also incomplete, but the recorded rtlogs did not reveal any abnormalities. device analysis: after console was received in danvers it's been thoroughly tested, including manual power cycles and long-term testing with pump and purge, but no hardware or software issues were discovered. inspection on the cabling found that the ribbon cable between 4in1 and impellatronic was mildly kinked, which might result in communication issue and triggered controller error alarm. per the results of the clinical review and investigation: the root cause of the issue was likely due to a kinked ribbon cable. this report is being filed as part of a retrospective review of historical records.
15927 | 1220648-2024-21887 | 20601635 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20601635&pc=OZD | 11/16/2021 | 11/04/2024 |  | 11/04/2024 |  |  | 11/16/2021 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1321472 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported the automated impella controller (aic) experienced a battery failure (red alarm). this issue did not occur during clinical support, no patient involvement. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the battery failure alarm was due to a defective battery (rapid decline in cell voltage). this report is being filed as part of a retrospective review of historical records.
15928 | 1220648-2024-21884 | 20602314 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20602314&pc=OZD | 09/22/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 09/22/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1464426 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The biomedical engineer reported that the automated impella controller (aic) purge cassette stopped and was leaking. a loaner was sent to the customer to initiate return of this device. there was no patient harm reported. | [5 empty] | E1: name of the initial reporter is unknown. the investigation into the purge issue has been completed. the impella automated controller was returned for investigation. the data logs from the complaint date revealed that the console issued purge pressure related alarms. after the purge cassette was inserted, the purge system open alarm was triggered, indicating the average purge pressure was below 100mmhg for a period when the purger was running. later, purge pressure low alarm also triggered. the real time logs also revealed purge pressure was unable to build up even with piston forward strokes. the returned console was tested, the issue with properly detecting the purge pressure disc was reproduced, and the purge retainer stand was confirmed to be broken. the root cause of the reported purge cassette stopped and was leaking was unable to determined, as the purge cassette was not returned. the cause of the purge pressure issue was a broken retainer stand. this report is being filed as part of a retrospective review of historical records.
15929 | 1220648-2024-21888 | 20602897 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20602897&pc=OZD | 08/29/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 08/29/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1164693 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was on an automated impella controller (aic) for mechanical circulatory support. while on support the aic failed to recognize an installed catheter. device was removed and replaced with a backup console. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported pump not detected has been completed. the product was returned and the pump not detected issue was not confirmed. data analysis: console logs from the reported day of event are consistent with the complaint and show multiple resets of the system in response to ossiopsens process crashes. device analysis: after the console was received and subject to testing, the issue was not reproduced in long term testing on engineer bench. there were no observed eeprom reading issues, or spontaneously ossiopsens process (or any other processes) crashes. per the results of the clinical review and investigation: the root cause of the issue was not determined since product was not reproduced. this report is being filed as part of a retrospective review of historical records.
15930 | 1220648-2024-21893 | 20603386 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20603386&pc=OZD | 08/23/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 08/24/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1287215 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Audible Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported an automated impella controller (aic) was not emitting alarm sounds before mechanical circulatory support. no alarms are suppressed in the service menu. the following alarms were shown in the display (demo pump is connected): impella stopped, retrograde flow and purge system blocked and the console still remained silent. there were no reports of patient harm or injury. | [5 empty] | The investigation for the reported buzzer/ alarm inaudible issue has been completed. the product was returned, and the buzzer/ alarm inaudible issue was confirmed. data analysis: the imc logs is not related to this issue. the mentioned alarms were triggered because of the demo pump, but the staff did not hear an audible tone. device analysis: during inspection, fs found the speaker assembly cable was not connected to epc. after the cable was properly seated, the speaker¿s alarming function was recovered. per the results of the clinical review and investigation: the root cause was determined to be a loose connector. this report is being filed as part of a retrospective review of historical records.
15931 | 1220648-2024-21935 | 20603463 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20603463&pc=OZD | 02/24/2022 | 10/06/2025 |  | 11/04/2024 |  |  | 02/24/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1229948 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient of unknown race/ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. during preparation, the automated impella controller (aic) displayed a "console failure" message and could not be turned off the correct way, prompting an emergency shutdown. a backup console was used, and there was no reported patient harm. | [5 empty] | D8/9: device returned to field service, unknown date. the device was returned and evaluated, and the investigation for the reported controller failure - red alarm has been completed. data analysis: imc logs confirmed the reported failure mode given there was an optical sensor system error window and placement signal not reliable (opm invalid signal, optical pump connected] ¿ event set # (b)(4)) triggered during the case start procedure. real time (rt) logs revealed that the placement signal was invalid during the case start procedure which resulted in the error window. despite troubleshooting, the failure condition could not be resolved. on the last pump connection, the console alarmed controller error alarms (opm comm stopped ¿ event set # (b)(4)), and rt logs confirmed the optical bench parameters were invalid. given the complaint description states "pump could not be turned off", it was interpreted as the console could not be properly shutdown, which was related to the alarm #(b)(4). device analysis: console (ic4242) was sent back fs for further investigation. the failure mode was not reproduced when running a known-good optical pump and purge. the 5s parameters and optical power were within specification according to the functional check procedure (0042-7316). inspection of the pump connecter revealed contamination on the ferrule surface that could not be removed with conventional cleaning methods. the root cause of the invalid placement signal was most likely due to an intermittent air gap between the two ferrules that interface at the pump connector. the root cause of the transient loss of optical data was unable to be determined due to the inability to reproduce. the root cause of the loss of optical data was unable to be determined due to the inability to reproduce. this report is being filed as part of a retrospective review of historical records. | The investigation for aic - controller failure (red alarm) has been completed. imc logs confirmed the reported failure mode given there was an optical sensor system error window and placement signal not reliable (opm invalid signal, optical pump connected] ¿ event set # 1036) triggered during the case start procedure. real time (rt) logs revealed that the placement signal was invalid during the case start procedure which resulted in the error window. despite troubleshooting (as recommended by aic ui), the failure condition could not be resolve. on the last pump connection, the console alarmed controller error alarms (opm comm stopped ¿ event set # 1043), and rt logs confirmed the optical bench parameters were invalid (16000).given the complaint description states "pump could not be turned off", it was interpreted as the console could not be properly shutdown, which was related to the alarm #1043 (most likely the result of optical fiber being unplugged while console did not recognize pump being unplugged). inspection of the pump connecter revealed contamination on the ferrule surface that could not be removed with conventional cleaning methods. the root cause of the invalid placement signal was most likely due to an intermittent air gap between the two ferrules that interface at the pump connector. the root cause of the transient loss of optical data was unable to be determined due to the inability to reproduce. corrections to the initial mfr report: g1 mdr reporting contact email. d2 device name updated . d7 and 8 updated.
15932 | 1220648-2024-21912 | 20603607 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20603607&pc=OZD | 08/06/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 08/06/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1136694 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | It was reported that the automated impella controller (aic) had battery issues. console not involved in patient case. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the issue was use related. the battery transport switch was probably not engaged by the customer when console returned from pm in december. further, the console was not charged from a long time. this report is being filed as part of a retrospective review of historical records.
15933 | 1220648-2024-21885 | 20603627 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20603627&pc=OZD | 09/16/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 09/16/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER EU |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER EU |  |  | ABIOMED, INC. |  |  | US |  | 1378679 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) exhibited the red alarm, bleeding system blocked error. no additional information was provided, no report of harm or injury to the patient. | [5 empty] | E1: name of the initial reporter is unknown. the investigation into the purge issue/controller failure red alarm has been completed. the impella automated controller was returned for investigation. the data logs from the complaint date revealed on (b)(6) 2022 when the console was booted up, controller error #418 (purge pressure sensor defective) triggered. the reported ¿red alarm, bleeding system blocked¿ might be mistranslated because the original complaint was not in english. the reported red alarm was consistent with the controller failure alarm. the returned device was tested, when booting up the returned console, the controller failure was not reproduced. the purge assembly was inspected, specifically the connection between the purge unit driver (pud) and purge pressure sensor. visual inspection indicated the ribbon cable was connected and well-aligned without any discoloration or abnormalities on the cable. the cause of the issue was not determined since the failure mode was not reproduced after testing and simulation. this report is being filed as part of a retrospective review of historical records.
15934 | 1220648-2024-21911 | 20603639 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20603639&pc=OZD | 12/10/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 12/10/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2023252150 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with postcardiotomy cardiogenic shock (pccs)/low cardiac output syndrome (lcos) was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the automated impella controller (aic) experienced purge disc/flag issues that were resolved by replacing the aic. no patient harm was reported. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15935 | 1220648-2024-21946 | 20603679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20603679&pc=OZD | 02/17/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 02/17/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1548488 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported an automated impella controller (aic) battery failure and battery communication failure. the initial aic was replaced, and the case continued. there was no reported patient harm. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the aic issue was a defective battery. this report is being filed as part of a retrospective review of historical records.
15936 | 1220648-2024-21950 | 20603738 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20603738&pc=OZD | 08/08/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 08/08/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1276707 |  |  |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller's (aic) purge clip was broken. a loaner device was shipped to the customer. no indication of patient involvement, harm, or injury. | [5 empty] | The console (aic) was returned for evaluation. once the logs were reviewed, there was consistency with the failure mode at the complaint date. during the case start wizard, purge disk not detected alarm was issued several times despite purge cassette was re-inserted and another pc and disk was utilized. upon visual inspection of the console, the purge assembly area confirmed a ripped/ missing blue retainer. the root cause of the purge disc/ flag issues was due to a defective component. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15937 | 1220648-2024-21949 | 20603755 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20603755&pc=OZD | 08/11/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 08/11/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1087127 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller's (aic) purge disc retainer was broken. this was identified during preventative maintenance. no patient involvement. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, the purge disc retainer was confirmed to be broken. therefore, the root cause of the purge disc/flag issues was due to a broken retainer. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15938 | 1220648-2024-21929 | 20603766 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20603766&pc=OZD | 02/26/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 02/26/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1102523 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a console emitting a "battery temperature high" alarm during use. a new console was used without any adverse impact to the patient. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the false alarm for high battery temperature is due to the pbm misreading battery data during a single sample. this report is being filed as part of a retrospective review of historical records.
15939 | 1220648-2024-21913 | 20603825 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20603825&pc=OZD | 12/09/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 12/09/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1108702 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automated impella controller (aic) experienced purge disc/flag issues during clinical use. no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15940 | 1220648-2024-21907 | 20604006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20604006&pc=OZD | 12/22/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 12/22/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1450938 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller experienced purge disc/flag issues. no patient harm was reported. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the controller device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15941 | 1220648-2024-21938 | 20604056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20604056&pc=OZD | 06/20/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 06/20/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1565140 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant had and automated impella controller (aic) battery alarm. the team made decision to follow instructions for use and use a 2nd console. there was no reported harm or injury to the patient. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. when visually inspecting the batteries and pbm, it was observed that one of the batteries statuses leds were completely blank. the battery failure alarm was due to a defective battery 1 (decline of individual cell voltage). the root cause of the battery cell failure was undetermined. this report is being filed as part of a retrospective review of historical records.
15942 | 1220648-2024-21940 | 20604067 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20604067&pc=OZD | 08/11/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 08/11/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1368535 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that while the device was being prepped, the automated impella controller (aic) would not detect two brand new purge cassettes. a loaner was sent to the customer and there was no report of patient harm or injury. | [5 empty] | The console (aic) was returned for evaluation. after review of the logs, data showed no ¿purge cassette ticks¿ were recorded, confirming that purge motor was not advancing the piston. upon inspection of the console and despite installing a known-good purge cassette and purge disk, the priming would not start. also, diagnostic data would show very low values of purge flow, the purge pressure was not constant (not increasing) the number of motor ticks. temporarily replacing only the purge motor did not resolve the issue. temporarily replacing only the pud pca did not resolve the issue either. once both components were simultaneously replaced, the issue was resolved, and the purge system was working within specification. therefore, the root cause of the purge disc/flag issues is due to a defective purge assembly (stepper motor and pud pca). the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15943 | 1220648-2024-21924 | 20604108 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20604108&pc=OZD | 08/16/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 08/16/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1145539 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller's (aic) purge disk plastic insert holder was missing from the console. no report of patient involvement, harm, or injury. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, the purge assembly area was confirmed to be ripped/ missing blue retainer. therefore, the root cause of the purge disc/flag issues was due to a broken retainer. the purge disc retainer was replaced, and the purge assembly was cleaned before being returned to the customer. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15944 | 1220648-2024-21892 | 20604153 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20604153&pc=OZD | 01/29/2023 | 11/04/2024 |  | 11/04/2024 |  |  | 01/29/2023 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1315811 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | It was reported that after cp pump stops during case believed to be caused by defective cp, local team went back next day to confirm that ic3254 and ic5608 were working correctly. the field team operations tested with demo pump and discovered both automated impella controller (aic)s were causing pump stops. the same alarms occurred during the case before both the cp and a third aic was swapped in. field team successfully tested demo pump on two other aics and confirmed no issue with demo pump. field team contacted field service and is waiting on two loaner aics to be flown out and aics will be sent in for troubleshooting. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the issue was a defective demo pump. this report is being filed as part of a retrospective review of historical records.
15945 | 1220648-2024-21945 | 20604274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20604274&pc=OZD | 02/23/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 02/23/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1112624 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a battery failure alarm occurring, despite being plugged in and charging. there was no impact to any patient reported because of this. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the battery failure alarm is due to the decline of voltage of a single cell. this report is being filed as part of a retrospective review of historical records.
15946 | 1220648-2024-21904 | 20604284 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20604284&pc=OZD | 03/01/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 03/01/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1154172 |  | 0042-0000-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  | White | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. while on support, the automated impella controller (aic) experienced a purge issue that was resolved by switching to a backup console. there was no patient harm reported. | [5 empty] | D8/9: device returned to field service, unknown date. the device was returned and evaluated, and the investigation for the reported purge issue has been completed. data analysis: purge system failure (piston blocked) was observed in the imc logs. device analysis: failure mode was not reproduced. a known-good pump and purge was run overnight without errors. inspection of purge assembly found dry glucose between motor gasket and housing. the stepper motor gasket was cleaned thoroughly. the cause of the aic issue was contamination. this report is being filed as part of a retrospective review of historical records.
15947 | 1220648-2024-21914 | 20604286 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20604286&pc=OZD | 12/07/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 12/07/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1496398 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automated impella controller (aic) experienced purge disc/flag issues prior to case start. automated impella controller (aic) was unable to recognize inserted purge disk and set up process was unable to be completed. console was switched out and catheter set up with new case start without issue on the new console. no patient harm was reported. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken (missing at blue disc retainer). the cause of the issue was due to broken (missing) purge flag. this report is being filed as part of a retrospective review of historical records.
15948 | 1220648-2024-21902 | 20604297 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20604297&pc=OZD | 03/03/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 03/03/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1316951 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  |  | Asian | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella for mechanical circulatory support. during priming, the automated impella controller (aic) produced a purge disc/flag issue. replacing the purge cassette did not resolve the issue, so the console was replaced. there was no reported patient harm. | [5 empty] | The device was returned and evaluated, and the investigation for the reported purge disc/flag issue has been completed. console logs from the reported day of event are consistent with complaint. the issue was reproduced during troubleshooting and aic was unable to detect purge disk. visual inspection of the purge retainer area revealed that purge disk detect flag was broken. after the flag was replaced, the console was able to detect purge disk, and console operated within specification. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15949 | 1220648-2024-21910 | 20604306 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20604306&pc=OZD | 12/14/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 12/14/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1464780 |  | 0042-0010 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the aic (automated impella controller) experienced purge disc/flag issues. no patient harm was reported. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the controller device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the issue was a completely broken component. this report is being filed as part of a retrospective review of historical records.
15950 | 1220648-2024-21898 | 20604322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20604322&pc=OZD | 10/06/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 10/06/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1320563 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The biomedical engineer reported that the automated impella controller (aic) experienced a controller error alarm. a loaner was sent to the customer to initiate return of this device. there was no patient harm reported. | [5 empty] | E1: name of the initial reporter is unknown. the investigation into the controller error has been completed. the impella automated controller was returned for investigation. the data log reports from the reported day of event were not available for analysis due to log file corruption as a result of this malfunction. events leading to the malfunction were, however, evidenced by prior log entries and are consistent with the complaint, showing controller error alarms due to loss of communication with power battery manager (pbm). the returned console was visually inspected, it was observed that the pbm was corroded near the supercaps c69, c70 - this is a known pattern failure, resulting in sau_powermod_1 (pbm1) communication loss. the cause of the aic issue was contamination. (pbm corrosion). this report is being filed as part of a retrospective review of historical records.
15951 | 1220648-2024-21928 | 20604363 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20604363&pc=OZD | 11/30/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 11/30/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1320103 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported patient was placed on impella 5.0 eu for hemodynamic support. while on support, the aic (automated impella controller) experienced purge disc/flag issues that were resolved by replacing the aic. no patient harm was reported. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15952 | 1220648-2024-21897 | 20604421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20604421&pc=OZD | 08/15/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 08/15/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1427009 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) exhibited purge cassette leaking error. troubleshooting was performed without success. no report of patient involvement, harm, or injury. | [5 empty] | E1: the name of the initial reporter is unknown. the investigation into the purge disc/flag issues issue has been completed. the automated impella controller (aic) was returned for investigation. the device logs were consistent with the complaint. once the purge cassette was inserted, alarms associated with low purge pressure would be triggered. visual inspection revealed that the purge flag was broken, and there was a hair line crack on the blue purge retainer. it was speculated that the broken purge retainer was unable to hold the purge pressure disk firmly to the purge pressure sensor, therefore the measured purge pressure sensor was lower than speculation. also, before disassembling the purge unit, the purge flag was almost broken but still able to function as detecting purge disk insertion. this report is being filed as part of a retrospective review of historical records.
15953 | 1220648-2024-21932 | 20604503 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20604503&pc=OZD | 11/21/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 11/21/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1139979 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the aic (automated impella controller) experienced purge disc/flag issues. no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has not been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15954 | 1220648-2024-21895 | 20604506 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20604506&pc=OZD | 10/30/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 10/31/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1383039 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 34 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a controller error. there was no patient harm reported. | [5 empty] | The investigation into the controller error/loss of opm data has been completed. the impella automated controller was returned for investigation. the data logs confirmed the reported failure mode given there were multiple controller error alarms (opm comm crc error ¿ event set # 1045) triggered and resolved during the clinical procedure. real time (rt) logs confirmed that this was a pattern failure mode in which all signals received from optical bench placement, snr, contrast, lamp) are interpreted as -16384 values. this known pattern failure, which is characterized by a temporary loss of optical data (placement, snr, contrast, lamp, gain) and triggering of a controller error alarm, has a low probability of reoccurrence. if needed, patient hemodynamics and impella position can be monitored with imaging. the cause of the controller error and transient loss of optical data was not determined due to the inability to reproduce the failure. this report is being filed as part of a retrospective review of historical records.
15955 | 1220648-2024-21936 | 20604625 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20604625&pc=OZD | 11/15/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 11/15/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | XXXXXXX |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a broken purge disc clip was observed. there was no harm to any patient. a replacement aic was forwarded to the complainant. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15956 | 1220648-2024-21900 | 20604640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20604640&pc=OZD | 09/19/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 09/19/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1471020 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a controller error alarm. this aic was replaced with a new aic resolving the issue, and the procedure was successfully completed. there was no patient harm reported. | [5 empty] | The investigation into the controller error/purge issue data have been completed. the impella automated controller was returned for investigation. the data log reports confirmed the reported failure mode as there were multiple instances of the controller failure alarm (purge pressure sensor defective ¿ event #414). before this alarm would trigger, the console would alarm purge system open. real time logs for the complaint date revealed that purge pressure sensor value (purge pressure) was static during a forward stroke. however, the purge pressure was normal during the time the console was not in purge change wizard (before and after purge change). the returned console underwent long term simulation with the pump and purge without any issues. the failure mode was not reproduced during the simulation. it was observed that the membrane was contaminated/ damaged. this observation may have not been related to the inability to build/ maintain pressure and the triggering of the controller failure alarm, which occurred in the clinical setting. the root cause of the reported controller failure was most likely due to an open line during purge change as a result of improper technique by the end user. this report is being filed as part of a retrospective review of historical records.
15957 | 1220648-2024-21939 | 20604696 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20604696&pc=OZD | 08/12/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 08/12/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1431282 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller's (aic) purge clip was broken. the customer stated that they had another device onsite to utilize if necessary. no report of patient involvement, harm, or injury. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, it was confirmed the purge disc retainer had broken/snapped off of the purge assembly. therefore, the root cause of the purge disc/flag issues was due to a defective component. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15958 | 1220648-2024-21899 | 20604709 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20604709&pc=OZD | 09/28/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 09/28/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1135599 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The biomedical engineer reported that the automated impella controller (aic) experienced a controller error alarm. a loaner was sent to the customer to initiate the return of this device. there was no patient harm reported. | [5 empty] | E1: name of the initial reporter is unknown. the investigation into the controller error/loss of opm data has been completed. the impella automated controller was returned for investigation. the data log reports confirmed the reported failure mode given there was a controller error alarms (opm comm crc error ¿ event set # 1045) triggered during the clinical procedure. real time (rt) logs confirmed that this was a pattern failure mode in which all signals received from optical bench (placement, snr, contrast, lamp) are interpreted as -16384 values. this known pattern failure, which is characterized by a temporary loss of optical data (placement, snr, contrast, lamp, gain) and triggering of a controller error alarm, has a low probability of reoccurrence. if needed, patient hemodynamics and impella position can be monitored with imaging. the cause of the transient loss of optical data was not determined due to the inability to reproduce. this report is being filed as part of a retrospective review of historical records.
15959 | 1220648-2024-21925 | 20604741 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20604741&pc=OZD | 07/19/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 07/19/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1116352 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a controller error. the specific failure was reported as ¿"system self check failed, change console immediately" when it was turned on during a software update. there was no patient harm reported. | [5 empty] | E1: name of the initial reporter is unknown. the investigation into the controller error issue has been completed. the impella automated controller was returned for investigation. the data analysis of the logs revealed from the complaint date that the console issued the purge disc not detected alarm (event set #402) several times. the complainant incorrectly reported the issue as a "controller error". no "controller errors" were present on the complaint date, or within the logs. occurring on the complaint date of 7/1/22, imc logs reported several triggers of "purge disc not detected". the returned console was tested and the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken (missing at blue disc retainer).the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15960 | 1220648-2024-21948 | 20604746 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20604746&pc=OZD | 08/11/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 08/11/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1215718 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Peeled/Delaminated | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller's (aic) unit where the purge disc is inserted in the console is cracked and coming off. no further information is available. this issue was discovered after the removal of the device. no reported harm or injury to the patient. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, the purge disc retainer was confirmed to be broken. the root cause is defective component. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15961 | 1220648-2024-21916 | 20604772 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20604772&pc=OZD | 07/19/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 07/19/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1145129 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. it was reported that the aic experienced a/c power issues. no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | The investigation into the a/c power issue has been completed. the automated impella controller (aic) was returned for investigation. after console was tested, the issue was not immediately reproduced. we confirmed that the pbm and internal power supply operated within specification. however, while manipulating the ac power cord, the operation of power supply was intermittent. inspection of the power cord revealed a wire defect resulting in intermittent electrical contact. replaced ac power cord set, and performed a functional check. this report is being filed as part of a retrospective review of historical records.
15962 | 1220648-2024-21931 | 20604794 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20604794&pc=OZD | 08/14/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 08/14/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1619563 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on an automated impella controller (aic) for mechanical circulatory support. while on support, the aic showed ¿controller failure¿ alarm twice. they successfully switched to a new console. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported controller failure (red alarm) issue has been completed. the product was returned, and the controller failure (red alarm) issue was not confirmed. console logs from the reported day of event show multiple controller failure (red) alarms, of varying causes, which occurred between (b)(6) 2022 (case start) and (b)(6) 2022 (case end): ¿adjust lv signal¿ (#437), ¿impella position in aorta¿ (#128), ¿enter cardiac output¿ (#438), ¿placement monitoring disabled¿ (#140), ¿purge pressure high¿ (#408), ¿purge system blocked¿ (# 409), ¿purge disk not detected¿ (#402), and ¿impella in ventricle¿ (#139). there were also several controller error (yellow) alarms, most notably ¿placement signal not reliable / opm signal invalid¿ (#1036)¿ and ¿optical bench communication crc invalid¿ (#1045). most likely these were the alarms the staff complained about, due to the prevalence of optical signal issues (multiple events 1036 without alarms) throughout the case (as evidenced by imc and rtlog data). these optical signal issues associated with alarms # 1036 and # 1045 were identified as "transient loss of optical data". this known pattern failure is caused by a temporary loss of optical placement signal. no deficiencies of optical system were discovered during testing, and optical bench 5s parameters were within specification. per the results of the clinical review and investigation: the root cause was not determined because issue could not be reproduced. this report is being filed as part of a retrospective review of historical records.
15963 | 1220648-2024-21908 | 20604806 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20604806&pc=OZD | 12/14/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 12/14/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1254276 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the aic experienced purge disc/flag issues. no patient harm was reported. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the controller device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the issue was broken blue purge retainer. this report is being filed as part of a retrospective review of historical records.
15964 | 1220648-2024-21909 | 20604850 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20604850&pc=OZD | 12/20/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 12/20/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1145436 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the aic (automated impella controller) experienced purge disc/flag issues. no patient harm was reported. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the controller device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15965 | 1220648-2024-21930 | 20604859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20604859&pc=OZD | 11/24/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 11/24/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1233232 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported patient was placed on impella cp for hemodynamic support. while on support, the aic (automated impella controller) experienced purge disc/flag issues that were resolved by replacing the aic. no patient harm was reported. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15966 | 1220648-2024-21915 | 20604878 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20604878&pc=OZD | 08/11/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 08/11/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1108316 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on an automated impella controller (aic) for mechanical circulatory support. while on support, the aic experienced a controller error (yellow alarm). the patient never experienced a drop in flows or hemodynamic compromise and there was no report of patient harm or injury. | [5 empty] | The investigation for the reported controller error (yellow alarm) and optical signal issue has been completed. the product was returned, and the controller error (yellow alarm) and optical signal issue was not confirmed. console logs from the reported day of event are consistent with complaint and show optical placement signal issues approx. 6 hours after pump was started, evidenced by optical bench errors #1040 (ambient pressure out of range) and optical bench alarms #1045 (opm communication crc errors). the alarms kept reoccurring, but also self-resolving within minutes. further analysis of rtlog data confirmed ¿transient loss of optical data¿ pattern failure. per the results of the clinical review and investigation: the root cause of the issue was not determined since the issue could not be reproduced. this report is being filed as part of a retrospective review of historical records.
15967 | 1220648-2024-21922 | 20604916 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20604916&pc=OZD | 07/07/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 07/07/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1104309 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Decrease in Pressure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced optical signal issues that were resolved by unplugging the device, followed by unspecified troubleshooting. additionally, the aic experienced purge issues that were resolved by switching to an alternative console. no clinical details regarding patient condition were provided. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the issue was use related. during implementing the pump, clinical staff did not push the pump all the way in and triggered the optical sensor failure and led to potential confusion of the priming wizard. there was no issue found during the case. the device functioned/operated to specification. there was no issue found during the case that could cause the purge issue. the device functioned/operated to specification. this report is being filed as part of a retrospective review of historical records.
15968 | 1220648-2024-21926 | 20604921 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20604921&pc=OZD | 08/12/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 08/12/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1153713 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller's (aic) purge cassette was not recognized. no report of patient involvement, harm, or injury. | [5 empty] | The console (aic) was returned for evaluation. after review of the logs, the complaint was incorrectly reported as a detection issue with the purge cassette, however the issue was with purge disc detection issue. upon inspection of the console, the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. therefore, the root cause was due to a defective component. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15969 | 1220648-2024-21947 | 20605040 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20605040&pc=OZD | 08/11/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 08/11/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1266539 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Other | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | Failure to Read Input Signal | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) does not detect the purge disc. several cassettes were replaced; however, this issue was not resolved. additionally, it was observed that a small black plastic piece broke-off from the inside the purge disc fixation. the customer replaced this aic with another device, and it was reported that the patient was stable. no report of patient harm or injury. | [5 empty] | E1: the name of the initial reporter is unknown. e2 and e3: unknown. the console (aic) was returned for evaluation. the device logs from the complaint date revealed that the console issued the purge disc not detected alarm several times throughout the case. re-insertion of the pc/ purge disc was observed, yet the console still had persistent problem in detecting purge disc. upon inspection of the console, the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. the root cause of purge disc not detected was broken purge flag. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15970 | 1220648-2024-21919 | 20605052 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20605052&pc=OZD | 08/11/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 08/11/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1084303 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on an automated impella controller (aic) for mechanical circulatory support. while on support, an aic alarmed with a controller failure alarm. the alarm resolved itself after six seconds and the pump maintained proper flows throughout the duration. the unit was removed and replaced by a new aic. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported controller failure (red alarm) issue has been completed. the product was returned, and the controller failure (red alarm) issue was confirmed. console logs from the reported day of event are consistent with complaint and show purge-related alarms during time frame specified in complaint description. the particular alarms mentioned, were identified as ¿controller failure due to purge pressure sensor defective¿ (b)(6), that occurred during purge bag change. during testing, the purge system was operating within specification - there were no issues with purge motor and purge pressure sensor calibration was confirmed. although some irregularities of the purge pressure sensor area were discovered (purge disk retainer crack, flag contamination with dextrose, and foil crease/tear), these issues did not contribute to the reported malfunction. per the results of the clinical review and investigation: the root cause was determined to be most likely use related. this report is being filed as part of a retrospective review of historical records.
15971 | 1220648-2024-21920 | 20605089 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20605089&pc=OZD | 02/28/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 02/28/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1494427 |  | 0042-0040-JP |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that a complainant had an observed battery failure in a non-patient use. the console was in the hospital offices. the ifu states should the battery or automated impella controller (aic) fail in patient use to use a backup console. | [5 empty] | D9: unknow device return date. the investigation into the a/c power issue has been completed. the automated impella controller (aic) was returned for investigation. operational status was found to be ¿discharged disabled due to a current issue¿ and ¿discharge fault¿. the battery failure alarm was due to a defective battery. this report is being filed as part of a retrospective review of historical records.
15972 | 1220648-2024-21906 | 20605102 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20605102&pc=OZD | 12/26/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 12/26/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1375416 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 38 YR |  | Male |  |  | Asian | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was placed onto impella 5.5 support due to cardiomyopathy. while on support, the aic (automated impella controller) experienced purge disc/flag issues that were resolved by replacing the aic. no patient harm or injury noted. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the controller device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the issue was broken/ missing purge flag. this report is being filed as part of a retrospective review of historical records.
15973 | 1220648-2024-21905 | 20605155 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20605155&pc=OZD | 08/20/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 08/20/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1162032 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a controller error alarm for three minutes. it was reported the alarm resolved, the console did not alarm again, and the procedure was successfully completed with this aic. there was no patient harm reported. | [5 empty] | The investigation into the controller error/loss of opm data has been completed. the impella automated controller was returned for investigation. the data log reports confirmed the reported failure mode given there was a controller error alarms (opm comm crc error ¿ event set # 1045) triggered during the clinical procedure. there were also multiple occurrences of opm data crc errors, two of which did not last long enough to result in a controller error. there were no instances of impella defective alarms on the complaint occurrence date. real time (rt) logs confirmed that this was a pattern failure mode in which all signals received from the optical bench (placement, snr, contrast, lamp, gain) are interpreted as -16384 values. the returned console was further investigated. the failure mode was not reproduced during normal operation. this is a known pattern failure, which is characterized by a temporary loss of optical data (placement, snr, contrast, lamp, gain) and triggering of a controller error alarm, has a low probability of reoccurrence. if needed, patient hemodynamics and impella position can be monitored with imaging. the cause of the transient loss of optical was not determined due to the inability to reproduce. this report is being filed as part of a retrospective review of historical records.
15974 | 1220648-2024-21952 | 20605248 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20605248&pc=OZD | 08/04/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 08/04/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1300700 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the purge disc connector was damaged. unable to move past start case screen with soft keys. attempted to turn off and turn console back on without result. there was no reported harm or injury to the patient. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, the purge assembly area confirmed a ripped/ missing blue retainer. the missing/ broken retainer led to console inability of detecting purge fluid. therefore, the root cause of the purge disc/flag issues was due to a defective component. before returning the console, the purge disc retainer was replaced and a functional check was performed. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15975 | 1220648-2024-21941 | 20605251 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20605251&pc=OZD | 06/14/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 06/14/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1382144 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a controller error. there was no patient harm reported. | [5 empty] | E1: name of the initial reporter is unknown the investigation into the controller error issue has been completed. the impella automated controller was returned for investigation. the data analysis of the logs indicated that multiple purge related alarms were issued on event occurred date (b)(6) 2022). according to the logs, the pump was connected before the purge cassette was connected, which triggered purge system open (event set - 407) alarm as designed. many entries of errors occurred during insertion of the purge cassette (pc), indicating the pc was likely not fully inserted and caused multiple read errors of the purge cassette radio-frequency identification. later, the purge pressure low alarm was triggered and cleared due to reinsertion of the pc. based on real time logs, the purge pressure value was restored into normal range, however the purge system open alarm stuck, and user was acknowledged by pressing the mute button frequently. the stuck purge system open alarm was the ¿controller error¿ alarm reported by the complainant. the returned console was fully investigated on an engineer bench. inspection of the console, especially the purge system, did not reveal any abnormality. in conclusion, the purge system open alarm should clear within 20 seconds after the purge pressure increases above the alarm detection threshold. however, the persistent purge system open (false) alarm was not cleared. it was speculated that the repetitive ahp warnings due to pc not being fully inserted caused the clearing of the ¿outlett_open¿ flag in console software and eventually lead to the alarm stuck and attention of user. the purge system open alarm that would not clear was highly likely due to an abnormal software reaction to a pc that was not fully inserted. the cause of the aic issue was a software issue caused by the end user not properly inserting the pc, the product did not fail to meet specifications during testing. this report is being filed as part of a retrospective review of historical records.
15976 | 1220648-2024-21943 | 20605280 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20605280&pc=OZD | 08/12/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 08/12/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1401649 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) exhibited that the pressure transmitter was not recognized during a cassette change. the device was replaced with another aic the customer had onsite. it was reported that this occurred during a procedure, the procedure was successful with no harm to the patient. | [5 empty] | The console (aic) was returned for evaluation. once the logs were reviewed, the complaint date revealed that the console issued the purge disc not detected alarm several times. re-insertion of the pc/ purge disc was observed, yet the console still had problem in detecting purge disc. upon inspection of the console, the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. the root cause is defective component. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15977 | 1220648-2024-21927 | 20605417 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20605417&pc=OZD | 08/15/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 08/15/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1124582 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. during support, the device exhibited a purge disc undetected alarm. the staff replaced the purge cassette, the aic stopped alarming and then started to alarm with the same message. the device was replaced with another aic. no report pf patient harm or injury. | [5 empty] | The console (aic) was returned for evaluation. a review of the logs from the complaint date revealed that the console issued the purge disc not detected alarm several times. upon inspection of the console, the issue with properly detecting the purge pressure disc was reproduced, and blue purge retainer was confirmed to be broken. before returning this console back to the customer, the purge retainer was replaced, and a full functional check was performed on the console and optical system. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15978 | 1220648-2024-21921 | 20605472 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20605472&pc=OZD | 12/02/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 12/02/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1241682 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the aic (automated impella controller) experienced purge disc/flag issues. no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15979 | 1220648-2024-21955 | 20605505 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20605505&pc=OZD | 08/05/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 08/05/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | AUTOMATED IMPELLA CONTROLLER |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1407963 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) experienced "system error optical sensor" occurred, which was resolved through new console. no patient harm reported. | [5 empty] | The investigation has been completed. the console (ic7048) was sent back for further investigation. after cleaning the receptacle, the console was tested without reproducing the optical signal issue. the root cause of the aic issue was contamination on the subcomponent. this report is being filed as part of a retrospective review of historical records.
15980 | 1220648-2024-21918 | 20605813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20605813&pc=OZD | 12/03/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 12/03/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1690147 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the aic (automated impella controlller) experienced purge disc/flag issues. no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause of the issue could not be determined due to the insufficient data and the inability to reproduce these issues. this report is being filed as part of a retrospective review of historical records.
15981 | 1220648-2024-21934 | 20605820 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20605820&pc=OZD | 06/28/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 06/28/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1456795 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported an automated impella controller (aic) failing to hold a charge. no patient harm reported. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. console purge assembly was inspected. upon opening the purge assembly door, it was confirmed the purge disc retainer had broken/snapped off of the purge assembly. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15982 | 1220648-2024-21954 | 20605831 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20605831&pc=OZD | 08/01/2022 | 11/04/2024 |  | 11/04/2024 |  |  | 08/01/2022 | 11/04/2024 |  |  |  |  | 11/04/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1148752 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a controller error. a loaner was sent to the customer to initiate return of this device. there was no patient harm reported. | [5 empty] | E1: name of the initial reporter is unknown. the investigation into the controller error issue/loss of opm data has been completed. the impella automated controller was returned for investigation. the data analysis of the logs from the reported day of event are consistent with complaint and show ¿opm communication crc invalid¿ alarm (#1045) that self-resolved after less than 7 minutes, while all values of data from optical bench were -16384. this is a known pattern failure caused by a temporary loss of optical placement signal. there¿s no need for repairs if the condition self-corrects. this is a low probability event and lasts only a few minutes. if needed, the patient¿s hemodynamics and impella position can be monitored with imaging. the cause of the issue was not determined because the issue was unable to be reproduced. this report is being filed as part of a retrospective review of historical records.
15983 | 1220648-2024-21806 | 20595240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20595240&pc=OZD | 09/09/2022 | 11/03/2024 |  | 11/03/2024 |  |  | 09/09/2022 | 11/03/2024 |  |  |  |  | 11/03/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1153947 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Damaged Prior to Use | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that the automated impella controller (aic) was unboxed and found to be damaged. the automated impella controller will be returned for service and evaluation of the damage. no patient involvement. | [5 empty] | E1: name of initial reporter is unknown. the investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the issue was mishandling during transportation. this report is being filed as part of a retrospective review of historical records.
15984 | 1220648-2024-21807 | 20595253 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20595253&pc=OZD | 09/08/2022 | 11/03/2024 |  | 11/03/2024 |  |  | 09/08/2022 | 11/03/2024 |  |  |  |  | 11/03/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1251267 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant was replacing a pc and found automated impella controller (console) damaged. the team made the choice to replace the automated impella controller (aic) device and following the instructions for use recommendations by utilization of a backup controller. no patient was involved. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
15985 | 1220648-2024-21808 | 20595258 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20595258&pc=OZD | 09/08/2022 | 11/03/2024 |  | 11/03/2024 |  |  | 09/08/2022 | 11/03/2024 |  |  |  |  | 11/03/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1222703 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | Insufficient Information | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. it was reported that the aic console no longer accepts mains power and runs only battery only. no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | The investigation into the a/c power issue has been completed. the automated impella controller (aic) was returned for investigation. the failure mode was reproduced by plugging in the ac cord to mains supply. console would not accept charging, and the amber led stayed off. further investigation found the fuses blown. please see the attached picture in attachment section. field service replaced the fuses, but the new fuses would blow again. the power supply was temporarily replaced. with new fuses intact, console would recognize the ac power. the investigation led to the conclusion that the defective power supply as the root cause. this report is being filed as part of a retrospective review of historical records.
15986 | 1220648-2024-21809 | 20595366 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20595366&pc=OZD | 09/04/2022 | 11/03/2024 |  | 11/03/2024 |  |  | 09/04/2022 | 11/03/2024 |  |  |  |  | 11/03/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1147986 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was on an automated impella controller (aic) for mechanical circulatory support. the complainant had a 5.5 support ongoing when the aic failed after a pc replacement was done. the team followed the ifu and utilized a backup aic. no harm or injury from the aic replacement and interruption in the support. | [5 empty] | Impella 5.5 sn: (b)(6) began on console (b)(6) on (b)(6)2022. although not technically persistent, several purge issues occurred with the console across multiple purge cassettes over the course of the case duration on (b)(6) from (b)(6) 2022 to (b)(6) 2022. on (b)(6) 2022 staff swapped impella 5.5 sn: (b)(6) and purge cassette sn: (b)(6) to (b)(6). purge issues, again, are not persistent but do occur on the replacement console. data analysis: occurring before the reported complaint date, on (b)(6)2022, several unreported purge related alarms occurred across multiple purge cassettes (sn's (b)(6)): (b)(6)2022 19:52:59 imcgui: event set: #409 from 13 (0) - purge system blocked (b)(6)2022 19:53:59 imcgui: event set: #431 from 13 (0) - purge pressure low (b)(6)2022 19:15:01 imcgui: event set: #408 from 13 (0) - purge pressure high (b)(6)2022 19:15:31 imcgui: event set: #409 from 13 (0) - purge system blocked (b)(6) 2022 19:17:28 imcgui: event set: #431 from 13 (0) - purge pressure low (b)(6) 2022 16:24:11 imcgui: event set: #409 from 13 (0) - purge system blocked on the reported complaint date of (b)(6) 2022, staff began a purge bag and cassette change with purge cassette sn: (b)(6) when they experienced a controller error. 19:12:49 purgeadmin: purge sensor_defect2 set - purge: sensor defective (voltage out of range) 19:12:49 purgeadmin: purge_stop - purge: stop 19:12:49 imcgui: event set #418 from 13 (0) - controller error (purge pressure sensor defective) these logs error messages repeated several times through the purge wizard, until the pump was disconnected from (b)(6), as the controller failure would not resolve. rt logs were reviewed at the time of these events occurring. purge related events occurring from (b)(6) caused fluctuation in purge pressure however not an overall loss of pressure. the controller failure related events caused large spikes in pressure, however resulted in an overall loss of purge pressure during the purge cassette and bag change. after experiencing several unresolvable controller failures, staff swapped the pump and last used purge cassette (sn (b)(6)) to (b)(6). imc logs from (b)(6) were reviewed. although for several days there were no issues with purge related function, the same alarms previously seen on (b)(6) were appearing across several different cassettes again. device analysis: (b)(6) arrived in (b)(6). console purge assembly was inspected, with special attention to the purge pressure sensor pca and ribbon cable. no observable issues were found. console was tested running a pump and purge system both inside and outside of the burn-in room. no issues occurred in either test scenario. during testing outside of the burn-in room, on an engineering lab bench, a purge cassette and bag change was attempted. console successfully went through the associated purge wizard with no issues. purge cmd bolus command was also used via telnet, and purge cassette and disc were manipulated in and out of their respective spots purposely through the wizard but no issues occurred. controller failure could not be reproduced on (b)(6). this report is being filed as part of a retrospective review of historical records.
15987 | 1220648-2024-21810 | 20595373 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20595373&pc=OZD | 09/06/2022 | 11/03/2024 |  | 11/03/2024 |  |  | 09/06/2022 | 11/03/2024 |  |  |  |  | 11/03/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1401386 |  | 0042-0010-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  |  | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a console shutting down during use. the pump was moved to a new console with no harm to the patient. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the issue was not determined since product was not returned, and the available data was insufficient to determine root cause. this report is being filed as part of a retrospective review of historical records.
15988 | 1220648-2024-21811 | 20595428 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20595428&pc=OZD | 09/03/2022 | 11/03/2024 |  | 11/03/2024 |  |  | 09/03/2022 | 11/03/2024 |  |  |  |  | 11/03/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1117656 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was placed onto impella cp support due to acute myocardial infarction / cardiogenic shock. while on support, the automated impella controller (aic) experienced a purge issue that was resolved by exchanging for another console. no patient harm was reported. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the root cause of the purge bag change that could not be completed was most likely due to cfc file system corruption. however, this could not be verified due to the absence of logs from the complaint date. this report is being filed as part of a retrospective review of historical records.
15989 | 1220648-2024-21812 | 20595461 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20595461&pc=OZD | 09/02/2022 | 11/03/2024 |  | 11/03/2024 |  |  | 09/02/2022 | 11/03/2024 |  |  |  |  | 11/03/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1251270 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) did not recognize the purge cassette disc during preparation. it was reported that multiple cassettes were unsuccessfully tried. there was no patient harm reported. | [5 empty] | E1: name of the initial reporter is unknown. the investigation into the purge issue has been completed. the impella automated controller was returned for investigation. the data analysis of the logs indicated they were found to be not relevant with the failure mode. the investigation of the returned device determined the issue of properly detecting the purge pressure disc was reproduced, and the blue purge retainer was confirmed to be broken. the cause of the issue was a defective component, broken retainer. this report is being filed as part of a retrospective review of historical records.
15990 | 1220648-2024-21813 | 20595551 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20595551&pc=OZD | 09/02/2022 | 11/03/2024 |  | 11/03/2024 |  |  | 09/02/2022 | 11/03/2024 |  |  |  |  | 11/03/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1164956 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an automated impella controller (aic) experienced power up issues during preventative maintenance. it was documented that when powered on, the aic would display a blue screen and then reboot. no patient involvement. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the issue was not determined since failure mode was not reproduced. this report is being filed as part of a retrospective review of historical records.
15991 | 1220648-2024-21814 | 20595552 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20595552&pc=OZD | 09/26/2022 | 11/03/2024 |  | 11/03/2024 |  |  | 09/26/2022 | 11/03/2024 |  |  |  |  | 11/03/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1327996 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient that needed automated impella controller (aic) to aic transfer. when transferred to nyu console, placement signal not reliable alarm appeared, no placement signal visible and impella defective alarm. motor current was within normal limit, patient noted hemodynamically stable. pump was unplugged and patient placed back into cooper aic. placement signal not reliable alarm appeared, but motor current normal with the patient being properly supported by 5.5. patient left on cooper aic. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the aic-pump issue was an eeprom data transmission corruption. this report is being filed as part of a retrospective review of historical records.
15992 | 1220648-2024-21815 | 20595560 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20595560&pc=OZD | 09/01/2022 | 11/03/2024 |  | 11/03/2024 |  |  | 09/01/2022 | 11/03/2024 |  |  |  |  | 11/03/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1279510 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had an automated impella controller (aic) pump connection failure at the case start. the aic was replaced as the backup controller was used for the case support. no harm or injury to the patient from the delay. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the pump detection issue was due to the rarely occurring epm crystal issue. the impellatronic pca does not have permanent damage, but this intermittent pump detection failure occurs infrequently and already has a sw mitigation in place. this report is being filed as part of a retrospective review of historical records.
15993 | 1220648-2024-21816 | 20595569 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20595569&pc=OZD | 08/30/2022 | 11/03/2024 |  | 11/03/2024 |  |  | 08/31/2022 | 11/03/2024 |  |  |  |  | 11/03/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1383536 |  | 0042-0000-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had an automated impella controller (aic) with controller error/alarms. the aic was replaced. no patient was involved. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the aic issue was contamination. there was an electrolyte leakage from the supercapacitor which caused damage to the pbm serial connection. this report is being filed as part of a retrospective review of historical records.
15994 | 1220648-2024-21817 | 20595585 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20595585&pc=OZD | 08/31/2022 | 11/03/2024 |  | 11/03/2024 |  |  | 08/31/2022 | 11/03/2024 |  |  |  |  | 11/03/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1427009 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported an automated impella controller (aic) failure with purge failure alarm occurring firstand 0.0 purge flow displaying on automated impella controller (aic). purge cassette was stopped per recommendations on aic. no patient harm reported. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the aic issue was contamination. this report is being filed as part of a retrospective review of historical records.
15995 | 1220648-2024-21818 | 20595595 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20595595&pc=OZD | 08/29/2022 | 11/03/2024 |  | 11/03/2024 |  |  | 08/29/2022 | 11/03/2024 |  |  |  |  | 11/03/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1261275 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | Battery Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was on an automated impella controller (aic) for mechanical circulatory support. it was reported that the device exhibited the following alarms: battery failure and controller error. no battery life was present on console. the device was replaced and set up was successful. no report of patient harm. | [5 empty] | The console (aic) was returned for evaluation. upon evaluation of the console, communication with optical bench was affected, at some point the console did not recognize pump being disconnected, which resulted in air in purge alarm not being cleared, but also in inability to shut down the console (aic won¿t allow shutdown if pump is connected). presumably in an attempt to shut down the console, and before the emergency shutdown went into effect, battery transport switch was engaged, which resulted in the ¿battery failure¿ alarm. therefore, all three reported failure modes were determined to have been related and are considered to be different symptoms of a single investigated failure mode - the cause of the aic issue was most likely a defective 4in1 board. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
15996 | 1220648-2024-21774 | 20594248 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20594248&pc=OZD | 11/02/2022 | 11/02/2024 |  | 11/02/2024 |  |  | 11/03/2022 | 11/02/2024 |  |  |  |  | 11/02/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER JP |  |  | ABIOMED, INC. |  |  | US |  | 1420740 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in aortic valvuloplasty was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the automated impella controller (aic) experienced a controller error. aic was replaced with another aic and the procedure was completed. there was no patient harm reported. | [5 empty] | The investigation into the controller error has been completed. the impella automated controller was returned for investigation. the data logs confirmed the alarm for a controller error. this was followed by several triggers of communication errors for the power battery manager 1 (pbm1) communication line. upon boot up, the console booted into a system self-check error, likely due to errors with the pbm1 communication line. while performing a visual inspection of the interior of the console, it was observed there was a leaking electrolyte surrounding capacitors c69 and c70 which had caused corrosion on the pbm. the cause of the controller error alarm is due to corrosion caused by leaking of electrolyte from c69 and c70 on the pbm. this report is being filed as part of a retrospective review of historical records.
15997 | 1220648-2024-21775 | 20594249 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20594249&pc=OZD | 10/20/2022 | 11/02/2024 |  | 11/02/2024 |  |  | 10/20/2022 | 11/02/2024 |  |  |  |  | 11/02/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1614416 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported an automated impella controller (aic) alarm indicating an unexpected shut down had occurred. the issue resolved after 30 seconds, but a new controller was used out of caution. there was no harm to the patient. | [5 empty] | E1: name of initial reporter is unknown. the investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the root cause of the unexpected controller shutdown was most likely due to an inability to open the dumper process (software-issue). the inability to open the dumper process was not determined due to the inability to reproduce. this report is being filed as part of a retrospective review of historical records.
15998 | 1220648-2024-21776 | 20594254 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20594254&pc=OZD | 11/02/2022 | 11/02/2024 |  | 11/02/2024 |  |  | 11/02/2022 | 11/02/2024 |  |  |  |  | 11/02/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1565140 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported an automated impella controller (aic) battery failure alarm was observed. a replacement aic was required, and the case resumed. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer for evaluation of the device. the cause of the issue was use related. the root cause of the failure was due to the battery switch not engaged. this report is being filed as part of a retrospective review of historical records.
15999 | 1220648-2024-21777 | 20594255 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20594255&pc=OZD | 11/02/2022 | 11/02/2024 |  | 11/02/2024 |  |  | 11/02/2022 | 11/02/2024 |  |  |  |  | 11/02/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1085761 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported an automated impella controller (aic) battery failure alarm was observed. the device was removed. there was no patient involvement. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the aic issue was a defective battery. this report is being filed as part of a retrospective review of historical records.
16000 | 1220648-2024-21778 | 20594256 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20594256&pc=OZD | 11/02/2022 | 11/02/2024 |  | 11/02/2024 |  |  | 11/02/2022 | 11/02/2024 |  |  |  |  | 11/02/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1122737 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that automated impella controller (aic) had intermittent alarms of ¿placement signal not reliable¿ notification since device insertion. aic soon after received yellow ¿controller error¿ alarm. console was replaced. no patient harm was reported. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the aic issue (fast-toggling transient loss of optical data) was most likely a defective 4in1 board. this report is being filed as part of a retrospective review of historical records.
16001 | 1220648-2024-21779 | 20594268 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20594268&pc=OZD | 11/01/2022 | 11/02/2024 |  | 11/02/2024 |  |  | 11/01/2022 | 11/02/2024 |  |  |  |  | 11/02/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1289301 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported for an automated impella controller (aic) the hospital biomed received the console with a note stating that there was a console malfunction on power up. there was no further information provided from the clinical staff. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported system self-check error issue has been completed. the product was returned, and the issue was confirmed. data analysis: aic logs showed sau_powermod1 communication errors. device analysis: pbm corrosion was identified on the top side of the pbm. fcn 71 had been implemented to address other pbm boards manufactured between march and november 2013 that do not yet show signs of damage due to leaking capacitors. pbm super cap is a pattern failure mode and usually occurs with pbms that have a sn between (b)(6) (inclusive). this pbms sn was (b)(6) which was outside of the date range. per the results of the clinical review and investigation, the root cause was determined to be due to a corroded pbm (super caps leak). this report is being filed as part of a retrospective review of historical records.
16002 | 1220648-2024-21780 | 20594273 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20594273&pc=OZD | 10/28/2022 | 11/02/2024 |  | 11/02/2024 |  |  | 10/28/2022 | 11/02/2024 |  |  |  |  | 11/02/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1684985 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that while the patient was on support, the console alarmed, unexpected shutdown. swapped out for new automated impella controller (aic). no reported harm to the patient. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the root cause of the unexpected controller shutdown was unable to be determined because it was not reproduced. this report is being filed as part of a retrospective review of historical records.
16003 | 1220648-2024-21781 | 20594281 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20594281&pc=OZD | 10/27/2022 | 11/02/2024 |  | 11/02/2024 |  |  | 10/27/2022 | 11/02/2024 |  |  |  |  | 11/02/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1084436 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that their automated impella controller (aic) displayed a battery failure alarm. a hard reset was attempted to troubleshoot the issue, but the failure remained. the console was removed from service and a loaner was provided to the customer site for future support as needed. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
16004 | 1220648-2024-21782 | 20594284 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20594284&pc=OZD | 10/28/2022 | 11/02/2024 |  | 11/02/2024 |  |  | 10/28/2022 | 11/02/2024 |  |  |  |  | 11/02/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1239159 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a controller error yellow alarm and continually beeping. no patient was involved. | [5 empty] | E1: name of the initial reporter is unknown the investigation into the controller error has been completed. the impella automated controller was returned for investigation. the data logs confirmed the reported alarm triggering issue. an impella defective alarm was issued with no pump connected. there were also issues properly shutting the console down as a result. the failure mode was reproduced on the returned console by booting up the console. after swapping the impellatronic pca, the failure was resolved. the root cause of the erroneous impella defective alarm and continuous beeping was due to a defective impellatronic pca, specifically the electromagnetic pulse circuitry. this report is being filed as part of a retrospective review of historical records.
16005 | 1220648-2024-21783 | 20594294 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20594294&pc=OZD | 10/25/2022 | 11/02/2024 |  | 11/02/2024 |  |  | 10/25/2022 | 11/02/2024 |  |  |  |  | 11/02/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1548488 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported an automated impella controller (aic) failure alarm. the console was removed due to this issue, and a new device was utilized. there was no report of patient harm or injury. | [5 empty] | E1: name of initial reporter is unknown. the investigation into the controller failure (red alarm) issue has been completed. the automated impella controller (aic) was returned for investigation. data analysis: no controller failure was observed in the logs. aic and rt logs align with the pattern failure mode, transient loss of optical data. aic logs confirmed the reported failure mode given there was a controller error alarm (opm comm stopped ¿ event set # 1043) triggered during the clinical procedure. the console was tested on an engineering lab bench. the failure mode was not reproduced while running an optical pump simulator. there was no issue with the console hardware. the root cause of the controller error was due to an aic software issue. this report is being filed as part of a retrospective review of historical records.
16006 | 1220648-2024-21784 | 20594369 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20594369&pc=OZD | 10/25/2022 | 11/02/2024 |  | 11/02/2024 |  |  | 10/25/2022 | 11/02/2024 |  |  |  |  | 11/02/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1319824 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Power Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. while preparing to insert, it was reported that the (aic) experienced a/c power issues. console not holding power when plugged into wall suspected failure with cord or plug. no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | The investigation into the a/c power issue has been completed. the automated impella controller (aic) was returned for investigation. the visual inspection of the returned plug for the ac power cord had a disconnected wire. the loose aic power cord wire caused the a/c power issues and inability to charge the batteries as the original incident description noted. this report is being filed as part of a retrospective review of historical records.
16007 | 1220648-2024-21785 | 20594371 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20594371&pc=OZD | 10/21/2022 | 11/02/2024 |  | 11/02/2024 |  |  | 10/21/2022 | 11/02/2024 |  |  |  |  | 11/02/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1194664 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported an automated impella controller (aic) battery failure during use of the device. the initial device was replaced with a new aic. there was no patient involvement. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the aic issue was a defective battery. this report is being filed as part of a retrospective review of historical records.
16008 | 1220648-2024-21786 | 20594386 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20594386&pc=OZD | 10/24/2022 | 11/02/2024 |  | 11/02/2024 |  |  | 10/24/2022 | 11/02/2024 |  |  |  |  | 11/02/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1136980 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a controller error alarm at the end of the case. additionally, it was reported the soft keys froze, the white plug was pulled to shut the pump off. there was no patient harm reported. | [5 empty] | The investigation into the controller error/kbd knob issue has been completed. the impella automated controller was returned for investigation. the data logs from the reported day of event are consistent with the complaint and show alarm #24 (controller error due to kbd comm loss) preceded by imc-kbd error (kbd is keyboard error) messages indicating communication issue. shortly after the alarm, the console registered multiple keypresses, but reportedly the software did not respond to the keypresses. the pump was unplugged, and the console successfully shut down (confirming that rotary push-button was operational, and software was responsive). the issue was not reproduced during testing of the returned product, and there was no damage to any components observed during inspection. due to the hazard associated with this malfunction presenting during active support (rather than during aic boot-up sequence), it is recommended that all components that might have contributed to this issue are considered suspect and are to be replaced. the cause of the issue was not determined since the issue could not be reproduced. this report is being filed as part of a retrospective review of historical records.
16009 | 1220648-2024-21787 | 20594387 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20594387&pc=OZD | 10/19/2022 | 11/02/2024 |  | 11/02/2024 |  |  | 10/19/2022 | 11/02/2024 |  |  |  |  | 11/02/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1188635 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a battery failure alarm was observed on the automated impella controller (aic). it was reported that the aic was removed as a result of the issue. there was no reported patient harm. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the aic issue was a battery cell failure. this report is being filed as part of a retrospective review of historical records.
16010 | 1220648-2024-21788 | 20594410 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20594410&pc=OZD | 10/18/2022 | 11/02/2024 |  | 11/02/2024 |  |  | 10/18/2022 | 11/02/2024 |  |  |  |  | 11/02/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1080190 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported battery failure alarm involving the automated impella controller (aic). the device was removed. there was no reported patient harm. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and therefore, an evaluation of the device was completed. the cause of the aic issue was a defective power supply. this report is being filed as part of a retrospective review of historical records.
16011 | 1220648-2024-21789 | 20594595 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20594595&pc=OZD | 10/18/2022 | 11/02/2024 |  | 11/02/2024 |  |  | 10/18/2022 | 11/02/2024 |  |  |  |  | 11/02/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1462815 |  | 0042-0010-EU |  | Yes |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Device Damaged by Another Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. it was reported that the automated impella controller had an issue with a part of the cable stuck in the automated impella controller and could not be removed. it was reported that the patient was stable. | [5 empty] | D9: the date of the device return for evaluation is unknown. the investigation for the reported issue has been completed. the impella device was received from the customer for evaluation. the cause of the issue was use related. the aic damage was caused by the pump torn off from aic. this report is being filed as part of a retrospective review of historical records.
16012 | 1220648-2024-21790 | 20594658 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20594658&pc=OZD | 09/21/2022 | 11/02/2024 |  | 11/02/2024 |  |  | 10/15/2022 | 11/02/2024 |  |  |  |  | 11/02/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1313339 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) was experiencing high purge pressures. this aic was replaced with another aic which resolved the high purge pressure issue. there was no patient harm reported. | [5 empty] | The investigation into the purge issue has been completed. the impella automated controller was returned for investigation. the data analysis of the logs from the reported day of event are consistent with the complaint and show that over time there was a gradual increase of purge pressure and decrease in purge flow (while purge cassettes 51051 and 90910 were used). however, it is very likely that what actually prompted the aic swap was the inability to continue purging due to the purge disk not being detected during cassette change (to 90526). analysis of the console log data one week prior to purge disk detection failure showed a rapid increase of purge flow (and decrease in purge pressure) while using purge cassette 918 which is consistent with purge leak. the inspection of the returned console revealed heavy contamination of purge insert and purge disk retainer, broken disk detect flag, and immobilized purge motor shaft (and compromised purge motor shaft gasket). it is most likely that purge fluid residue from the leak (crystallized dextrose) caused the purge flag to fail a week later, which prompted the console swap. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
16013 | 1220648-2024-21791 | 20594660 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20594660&pc=OZD | 10/13/2022 | 11/02/2024 |  | 11/02/2024 |  |  | 10/13/2022 | 11/02/2024 |  |  |  |  | 11/02/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1279510 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant had a pump automated impella controller (aic) connection and recognition failure. at the port where the impella catheter connects to console, the piece inside that recognizes catheter is loose and would not properly connect to catheter. it was reported that the procedure outcome was successful. no reported harm or injury to the patient. the aic was replaced per instructions for use recommendations. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the issue was use related (accidental damage to aic during prior use). this report is being filed as part of a retrospective review of historical records.
16014 | 1220648-2024-21792 | 20594670 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20594670&pc=OZD | 10/14/2022 | 11/02/2024 |  | 11/02/2024 |  |  | 10/14/2022 | 11/02/2024 |  |  |  |  | 11/02/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1554466 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that a battery failure was observed on the automated impella controller (aic). the device was removed. there was no reported patient harm. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the aic issue was a defective battery. this report is being filed as part of a retrospective review of historical records.
16015 | 1220648-2024-21793 | 20594672 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20594672&pc=OZD | 10/13/2022 | 11/02/2024 |  | 11/02/2024 |  |  | 10/13/2022 | 11/02/2024 |  |  |  |  | 11/02/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1618084 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) exhibited a battery failure. the resolution for this issue is unknown. no report of patient involvement. | [5 empty] | The console (aic) was returned for evaluation. upon observation of the console, the failure mode was reproduced. software showed a cell voltage decline in only one of the batteries 1's cells which occurred as the battery failure alarm triggered. therefore, the root cause was due to a defective battery. the root cause of the defective battery was due to single cell failure. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
16016 | 1220648-2024-21703 | 20586496 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20586496&pc=OZD | 12/09/2022 | 11/01/2024 |  | 11/01/2024 |  |  | 12/12/2022 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1554428 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 for mechanical circulatory support. when the automated impella controller (aic) was activated, it produced a battery failure alarm and could not be used. the console was replaced, and there was no reported patient harm. | [5 empty] | The device was returned and evaluated, and the investigation for the reported aic - battery failure (red alarm) has been completed. console interior connections were checked and verified. it was confirmed that persistent battery failure (transport connection blown/missing) was not due to a disengaged battery switch. when console was booted for the first time in japan, console alarms are consistent with what was seen in the field. when visually inspecting the batteries and pbm, it was observed that one of the batteries status leds were completely blank. the battery failure alarm was due to a defective battery 2 (decline of individual cell voltage). the root cause of the battery cell failure was undetermined. this report is being filed as part of a retrospective review of historical records.
16017 | 1220648-2024-21704 | 20586540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20586540&pc=OZD | 12/11/2022 | 11/01/2024 |  | 11/01/2024 |  |  | 12/11/2022 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1251624 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on an automated impella controller (aic) for mechanical circulatory support. complainant reported the aic experienced aic - controller failure (red alarm). no patient was involved. | [5 empty] | The investigation into the controller failure (red alarm) issue has been completed. the automated impella controller (aic) was returned for investigation. data analysis: in reviewing the imc logs the console outputted both the "ossiopsens set invalid_bad_pressure_sample_mask" and "process sample foropmin valid signal error: calculator placement signal" errors. then reviewing the 5s signals from the rt log at the time of these errors, the optical signals jump to xffff for an extended period of time (~5 minutes) which points to the pattern failure of "aic - loss of opm". device analysis: these failures were attempted to be reproduced but were unsuccessful. this known pattern failure, which is characterized by a temporary loss of optical data (placement, snr, contrast, lamp, gain) and triggering of a controller error alarm, has a low probability of reoccurrence. if needed, patient hemodynamics and impella position can be monitored with imaging. repair: perform a full functional check on the console and optical system. this report is being filed as part of a retrospective review of historical records.
16018 | 1220648-2024-21705 | 20586616 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20586616&pc=OZD | 12/27/2022 | 11/01/2024 |  | 11/01/2024 |  |  | 12/27/2022 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1380525 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Product Quality Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced a defect. the resolution for this issue is unknown. no report of patient injury or harm. | [5 empty] | The investigation has been completed. the console was sent back for further investigation. the root cause of the aic issue was contamination. this report is being filed as part of a retrospective review of historical records.
16019 | 1220648-2024-21706 | 20586627 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20586627&pc=OZD | 11/09/2022 | 11/01/2024 |  | 11/01/2024 |  |  | 11/09/2022 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1286792 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced system self-check errors. no patient was involved. | [5 empty] | The investigation has been completed. the console (b)(6) was sent back for further investigation. the root cause of the system check failure was due to the power battery manager corrosion. this report is being filed as part of a retrospective review of historical records.
16020 | 1220648-2024-21707 | 20586639 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20586639&pc=OZD | 11/12/2022 | 11/01/2024 |  | 11/01/2024 |  |  | 11/12/2022 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1445594 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  | White | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was on an automated impella controller (aic) for mechanical circulatory support. while on support, the automated impella controller (aic) experienced an unexpected shutdown, which was resolved through new console. no patient harm reported. | [5 empty] | The investigation has been completed. the console was sent back for further investigation. the root cause of the unexpected controller reboot was not determined due to the inability to reproduce. this report is being filed as part of a retrospective review of historical records.
16021 | 1220648-2024-21708 | 20586659 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20586659&pc=OZD | 11/24/2022 | 11/01/2024 |  | 11/01/2024 |  |  | 12/13/2022 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1126564 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) was experiencing a air in line purge issue. it was reported that the nurse unsuccessfully attempted to de-air the line twice. reportedly the cassette was changed with no success, then the console was switched out and the de-air process was successful. this aic was replaced with another aic resulting in no other issues with the aic. there was no patient harm reported due to the aic issue. | [5 empty] | The investigation into the purge issue/purge flag issue has been completed. the impella automated controller was returned for investigation. the data analysis of the logs from the reported day of event are partially consistent with the complaint and show that after almost 3 days of support using pump (b)(6), the console presented with purge-related alarms: air in purge system (404), purge system open (407), purge pressure low (431), and eventually controller failure due to purge pressure sensor defective (#414). contrary to the complaint problem description, the purge cassette (sn: (b)(6) has not been replaced with another purge cassette during troubleshooting on ic2188. finally, the purge cassette was removed from ic2188, and pump was disconnected from ic2188, to be used on backup console ic7767. logs from ic7767 confirm pump (b)(6) being connected and recognized as previously running, and also show that when the same purge cassette sn (b)(6) was used, the console presented with alarm ¿air in purge system¿ (#404) that was resolved after the purge cassette was replaced with another one (sn: (b)(6). after that, the support (using pump (b)(6) continued successfully. the long-term log and real time log data showed particular signals showing purge pressure, flow and motor ticks, which is consistent with sudden pressure loss (e.g. due to leak) and inability to build up pressure despite motor shaft advancing the purge cassette piston (which resulted in erroneously triggered alarm #414). the exact point of failure (leak) could not be identified due to the pump and purge cassette were not returned for analysis. visual inspection of the console showed evidence of a purge fluid leak, in the form of contamination and residue on various elements of the purge assembly (including inside the console). a separate investigation into pump (b)(6) (not returned for analysis) concluded that the root cause of the purge issue was unable to be determined. the console inspection and testing also revealed purge pressure sensor out-of-spec (unlikely to have contributed to the purge leak), and kinked impellatronic cable (unrelated to complaint). the cause of the issue was most likely use related. this report is being filed as part of a retrospective review of historical records.
16022 | 1220648-2024-21709 | 20586673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20586673&pc=OZD | 08/24/2022 | 11/01/2024 |  | 11/01/2024 |  |  | 08/24/2022 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1554466 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Female |  |  | White | Malfunction | No | Yes | Battery Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with high-risk percutaneous coronary intervention was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) experienced optical signal issues with no resolution specified and a battery failure red alarm with no resolution specified. there was no report of patient harm or injury. | [5 empty] | The investigation for the reported issue has been completed. aic logs showed that console ic8625 triggered a couple battery-related alarms after roughly six months: ltlogs showed that cell 4 of battery 1 was fully depleted which caused the battery to trigger an internal safety mechanism and internally lock up (no output). this resulted in the battery failure and communication failure alarms. battery 2 was showing signs of cell 1 voltage divergence but had not failed yet. console was tested. both batteries were replaced to resolve the battery failure alarm and was recorded to validate discharging/ charging. the root cause of the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
16023 | 1220648-2024-21710 | 20586675 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20586675&pc=OZD | 10/09/2024 | 12/20/2024 |  | 11/01/2024 |  |  | 10/09/2024 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025534536 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that impella cp had air in purge alarms that did not resolve despite 4 attempts of de-airing that was performed. the purge cassette was replaced, and the issue was resolved. it was further reported that the patient's leg was cold and the team was thinking of weaning and removing the impella. inotropes were reduced and the limb ischemia was better. patient was stable, and impella cp was successfully weaned after 4 days of support. | [5 empty] | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
16024 | 1220648-2024-21711 | 20586677 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20586677&pc=OZD | 10/09/2024 | 12/11/2024 |  | 11/01/2024 |  |  | 10/09/2024 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025543824 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported following impella 5.5 implant for 67-year-old male patient, a graft was needed to be placed for the implant. while tunneling a 10mm hemasheild platinum graft conduit, the surgeon accidentally punctured the subclavian vein attempting to pass the graft through the tunnel. cardiopulmonary bypass was initiated immediately as a result and a repair of the vein was performed post-procedure. impella 5.5 was placed via the graft and tunnel successfully. the clinical rep confirmed that the graft was needed for the impella placement. patient outcome at the end of support was 'survived'. | [5 empty] | The investigation for the vascular damage has been completed. data logs are not relevant to the complication and product was not returned for investigation. clinical details report that a surgical instrument punctured the subclavian vein while trying to tunnel the graft for impella placement. based on clinical details, the root cause of the vascular damage was determined to most likely be a use issue. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system and impella 5.5 with smartassist for use during section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿
16025 | 1220648-2024-21712 | 20586691 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20586691&pc=OZD | 03/26/2024 | 12/10/2024 |  | 11/01/2024 |  |  | 10/09/2024 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 50 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A publication was received from abiomed japan on 2024-10-09 entitled "case report: unmasking sustainable left ventricular recovery in chronic heart failure with axillary temporary mechanical circulatory support" in which an impella 5.5 pump support for a 50 year old female at mayo (florida) had impella support while being considered for transplant. there was an ae noted "an acute rise in lactate dehydrogenase (ldh) was seen on impella day 22, suggesting ongoing hemolysis despite systemic anticoagulation, a marker indicating an elevated risk of pump thrombosis." | [5 empty] | The investigation for the hemolysis has been completed. the pump was not returned for investigation. data logs were not retrieved due to missing pump and console information. according to the clinical report, the patient had concerns about hemolysis during the case and it was resolved after pump removal. the cause of the hemolysis issue was not determined since product was not returned and sufficient clinical information was not provided. added-e1 facility name corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ desai, a., sharma, s., luce, c., ruiz, j., & goswami, r. (2024). case report: unmasking sustainable left ventricular recovery in chronic heart failure with axillary temporary mechanical circulatory support. frontiers in cardiovascular medicine, 11. https://doi.org/10.3389/fcvm.2024.1407552.
16026 | 1220648-2024-21713 | 20586697 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20586697&pc=OZD | 10/09/2024 | 12/19/2024 |  | 11/01/2024 |  |  | 10/09/2024 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2025500084 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella cp support. while on support, the impella was pulled out of the left ventricle due to the patient's movement, and attempts to adjust the position were difficult. although the patient's hemodynamics did not change significantly, support was still required for the planned percutaneous coronary intervention, so a guidewire access port was used, impella was removed, a 16fr sheath was inserted in the same site, and clinical and professional skills assessment was inserted. the fixation ring was tightened when it was pulled out. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. based on clinical description, the root cause of positioning issue was most likely patient movement. corrections to the initial mfr report: d4 updated model number g1 mdr reporting contact email
16027 | 1220648-2024-21714 | 20586705 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20586705&pc=OZD | 10/09/2024 | 11/29/2024 |  | 11/01/2024 |  |  | 10/09/2024 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025542534 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and there was a hematoma at the access site. two units of blood products were given to the patient and support continued. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation into access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on manufacturer device report: 1220648-2024-21714. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16028 | 1220648-2024-21716 | 20586713 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20586713&pc=OZD | 02/16/2022 | 11/01/2024 |  | 11/01/2024 |  |  | 02/16/2022 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1548488 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during support, automated impella controller (aic) alarmed for an error. the pump and aic remained on for support awaiting care escalation. controller was replaced. there was no harm or injury noted to the patient. | [5 empty] | The investigation for the reported aic - controller error (yellow alarm) has been completed. the product was returned and the aic - controller error (yellow alarm) was confirmed. the cause of the issue was not determined since the failure was unable to be reproduced throughout testing. this report is being filed as part of a retrospective review of historical records.
16029 | 1220648-2024-21717 | 20586714 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20586714&pc=OZD | 10/09/2024 | 12/02/2024 |  | 11/01/2024 |  |  | 10/09/2024 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025481055 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the 66-year-old male patient on impella 5.5 support had to be shocked out of ventricular tachycardia. the patient is stable, and was successfully bridged to a heart transplant. | [5 empty] | The investigation for the reported ventricular fibrillation/ventricular tachycardia (vf/vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vf/vt could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16030 | 1220648-2024-21715 | 20586715 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20586715&pc=OZD | 10/09/2024 | 12/09/2024 |  | 11/01/2024 |  |  | 10/09/2024 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025544932 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  |  |  | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the 47-year-old male patient on impella 5.5 support experienced an ischemic stroke. ct scan showed the infarct stable. it was further reported that the patient knocked over the automated impella controller (aic) resulting in damage to the impella catheter plug. the white connector cable plug pins were broken in aic. impella disconnected alarm occurred. the impella was explanted. | [5 empty] | The investigation into pump stop and stroke/cva has been completed. the impella device was not returned for evaluation. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. the root cause of the pump stop is due to a use issue as the patient knocked over the automated impella controller (aic). instructions for use as follows: impella 5.5 with smartassist for use during cardiogenic shock. section: direct aortic insertion. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-21715. e4 should have been left blank. g1 reporting contact email. h6 code 4619 was reported incorrectly. new code was added to health effect - impact code. h10 complete instructions for use missing. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
16031 | 1220648-2024-21718 | 20586725 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20586725&pc=OZD | 10/09/2024 | 12/10/2024 |  | 11/01/2024 |  |  | 10/09/2024 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025519998 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | Arrhythmia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support there was high purge pressure noted. tissue plasma activator was initiated in the purge and the issue resolved. support continued. additionally, there was a decrease in impella flows, increased suction events and ectopy despite pump repositioning. decision was made to removed and replace axillary 5.5 impella. clot was noted on inflow and outflow of impella during removal. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation into low pump flow, thrombus, vf/vt/af/at, and high purge pressure has been completed. the pump was returned for investigation. no physical abnormalities were reported on the returned product. low pump flow and high purge pressure were not reproduced during testing. the pump was operated as intended. the data log shows that a 6-minute gradual rise was reported on the second day of pump use. the purge pressure was trending down but returned to the normal range after the given tissue plasminogen activator (tpa), according to the clinical and confirmed data log review. the cause of the low pump flow issue was most likely patient condition related, based on given clinical details. the cause of the high purge pressure issue was biomaterial, based on observed trend on data log review and provided given clinical details. also, the reported thrombus failure mode was removed since it was already addressed under the high purge pressure failure mode, as the patient's condition caused the formation of biomaterial. as the necessary information was not provided, the root cause of the vt/vf was not determined. d9 date device returned to manufacturer added. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-21718. e4 should have been left blank. g1 reporting contact email.
16032 | 1220648-2024-21719 | 20586731 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20586731&pc=OZD | 10/10/2024 | 12/11/2024 |  | 11/01/2024 |  |  | 10/10/2024 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025544065 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | Defective Device |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that multiple complications were encountered during impella 5.5 support. it was initially documented that an air in purge alarm appeared on the automated impella controller (aic) during patient support. the staff swapped out the purge cassette to resolve the issue. however, coinciding with the replacement, a new purge bag had to be ordered from the pharmacy resulting in a thirty to forty five minute period where no purge fluid was running through the system or to the pump. once the purge bag arrived and the purge lines were set-up, high purge pressure and purge system blocked alarms appeared on the aic. troubleshooting steps were discussed, but the staff ultimately declined use of tissue plasminogen activator due to the patient's existing disseminated intravascular coagulation condition. the pump then stopped unexpectedly and was explanted. there was no patient harm as a result of the failure. | [5 empty] | The investigation into high purge pressure and pump stop has been completed. the impella was returned cut at catheter and at the cannula. blood was present within the purge gap and in the purge line. data logs were reviewed and the logs showed purge pressure and purge flow dropped to 0 for approximately an hour during a bag change with immediate increase to high purge pressure after bag change was completed. additionally, motor current began trending up when purge values were at 0 before alarm high motor current pump stop occurred. as the high purge pressure was a result of no purge running in the pump for an extended period of time leading cascading events of blood ingress through the purge gap and into the motor causing the pump stop, the failure mode of "high purge pressure" was removed and investigated under "pump stop". the root cause of the pump stop was a use related issue as the long bag change caused blood ingress and high purge pressure leading to a pump stop. the following have been reported incorrectly on manufacturer device report 1220648-2024-21719. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: direct aortic insertion. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
16033 | 1220648-2024-21720 | 20586734 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20586734&pc=OZD | 10/10/2024 | 11/16/2024 |  | 11/01/2024 |  |  | 10/10/2024 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1464783 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller purge disc holder broke off. there was no harm to the patient. | [5 empty] | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. no data review is needed for this complaint because it was only reported that the retainer was broken. during device analysis the purge disc retainer was found to be broken. the root cause of the issue was broken purge disc retainer. e4 should have been left blank on manufacturer device report 1220648-2024-21720. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
16034 | 1220648-2024-21721 | 20586762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20586762&pc=PYX | 10/10/2024 | 11/27/2024 |  | 11/01/2024 |  |  | 10/10/2024 | 11/01/2024 |  |  | 11/30/2025 |  | 11/01/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025442262 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Difficult to Insert | Failure to Advance | [5 empty] | Bradycardia | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that complications were encountered during attempted delivery of an impella rp pump. upon initial insertion, resistance was met while attempting to advance the pump and the cannula kinked. the end of the peel away sheath subsequently split and the pump was pulled back for a second attempt. the patient began to decompensate and became hypotensive. the pump was again inserted, but the implant was unsuccessful and both the pump and introducer had to be removed. the patient went bradycardic and cardiopulmonary resuscitation was initiated until the patient stabilized. a new pump and introducer were then utilized for a successful implant. | [5 empty] | The investigation for the delivery issues, difficulty inserting, hypotension has been completed. the product was returned. the clinical details state that the cannula buckled just outside of the sheath causing the end of the sheath tip to split. the buckling of the cannula was thought to have been caused by the resistance felt in the left femoral vein as well as due to the patients anatomy. the sheath and cannula were evaluated and tip split and cannula kink were confirmed. the root cause of the introducer damage - mechanical is most likely due to the sheath tip split due to the patients vessels. the root cause of the delivery issues causing the kink in the cannula are due to the patients condition. the root cause of the hypotension could not be determined without additional information. the instructions for use for the impella rp system with automated impella controller states the following: section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
16035 | 1220648-2024-21722 | 20586817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20586817&pc=OZD | 10/09/2024 | 03/26/2025 |  | 11/01/2024 |  |  | 10/09/2024 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025524969 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 28 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported 28-year-old female patient was admitted in with cardiogenic shock and cardiomyopathy. the patient was attempted to be placed on impella cp and ecmo support. however, in the cath lab, prior to the percutaneous coronary intervention, the pump had optical signal issues and was therefore replaced with a new impella cp. support proceeded without harm or injury. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The impella device was discarded by the customer and therefore, ¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. e.1 added facility name as it was inadvertently omitted from the initial manufacturer device report number. g.1 revised reporting contact email since the initial manufacturer device report number was submitted in accordance with updated procedures. h.6 code 4114 was reported incorrectly on the initial manufacturer device report number. a new code was added to health effect - clinical code. due to the device being discarded, investigation findings and investigation conclusions codes were revised. h.10 revised additional manufacturer narrative since the initial manufacturer device report number was submitted due to the device being discarded.
16036 | 1220648-2024-21723 | 20586992 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20586992&pc=OZD | 10/07/2024 | 12/09/2024 |  | 11/01/2024 |  |  | 10/09/2024 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025548946 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant in japan reported that impella 5.5 pump was placed to support female patient admitted in for a high risk surgery for the a ventricular septal defect (vsd). the pump was placed and a left ventricular (lv) perforation was observed. the perforation was observed following pump repositioning. the patient survived and the surgical repair was done. the impella 5.5 was used for only 30 minutes of support, in addition to and ecmo and an intra-aortic balloon pump (iabp) support. | [5 empty] | The investigation into ventricle perforation has been completed. the impella device was not returned for evaluation. based on the clinical details provided, the root cause of the ventricle perforation is most likely due to use as it was perforated while attempting to adjust the position of the device. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-21723. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system and impella 5.5 with smartassist for use during section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
16037 | 1220648-2024-21724 | 20587201 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20587201&pc=OZD | 10/09/2024 | 12/05/2024 |  | 11/01/2024 |  |  | 10/09/2024 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025454965 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported delivery issues with an impella cp. the patient had bioprosthetic valve but no stenosis was observed. the doctor observed a tortuosity of the aortic arch that he believed made it difficult to advance impella into proper position. it was unsuccessful in inserting the first impella cp therefore another impella cp was used. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high-risk pci. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation into delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition since the implant was aborted due to patient's anatomy as tortuosity of the aortic arch was present.
16038 | 1220648-2024-21725 | 20587266 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20587266&pc=OZD | 10/01/2024 | 12/02/2024 |  | 11/01/2024 |  |  | 10/01/2024 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025534991 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Tachycardia | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a 79-year-old male patient on impella cp support was brought to the catheterization lab for femoral arterial sheath extracorporeal tubing bypass from left femoral artery to right femoral artery since there was no pulse on palpitation or doppler. the patient has iliac disease and severe peripheral artery disease. the nurse was still unable to find pulses on the right foot. the goal is to wean and remove the impella as soon as possible. additionally, the patient had episodes of ventricular tachycardia. it was further reported that the patient was transitioned to comfort care and expired on impella support. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer and therefore,an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into limb ischemia and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
16039 | 1220648-2024-21726 | 20587387 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20587387&pc=OZD | 10/09/2024 | 03/11/2025 |  | 11/01/2024 |  |  | 10/09/2024 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1824652 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller had a power cord damage issue. the cord got caught when trying to move the patient bed. there was no known patient harm. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - cabling issue has been completed. the console log showed an ac power was not connected. the console was returned for investigation. upon arrival the console underwent testing, which revealed an issue during visual inspection. two of the internal cables within the plug were found to be disconnected. further investigation involved removing the transparent cover of the plug. this revealed that both the green and white cables were disconnected. the issue was most likely the insufficient length of the wire covering, which was noticeably shorter than that of a properly functioning unit. the cables' outer sheath/jacket length was inadequate to prevent pinching at the plug, which the cord broke loose. root cause: the reported complaint was most likely due to an insufficient length of the wire outer sheath or covering (defective ac cord set). d2 common device name revised on manufacturer device report 1220648-2024-21726 in accordance with updated procedures. g1 reporting contact email revised on manufacturer device report 1220648-2024-21726 in accordance with updated procedures.
16040 | 1220648-2024-21727 | 20587426 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20587426&pc=OZD | 10/04/2024 | 12/05/2024 |  | 11/01/2024 |  |  | 10/04/2024 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025544836 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support and during support the patient had a few runs of ventricular tachycardia (vt). position was confirmed on echocardiography. the pump was slightly against the left ventricle wall, however the care team elected to not reposition and to keep the pump as is. the patient was on impella support for 190.66 hours before the patient expired. care was withdrawn due to the patient¿s lack of neurological status. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the impella 5.5 has been completed. the device nor sufficient clinical information was provided. data logs showed pump ran without issue during support. there were impella position in aorta and suction alarms triggered during support but resolved quickly. product was not returned for review. based on clinical description, the root cause of positioning issue was most likely use related. the root cause of the ventricular tachycardia (vt) could not be determined without additional information. corrections to the initial mfr report: g1 mdr reporting contact email.
16041 | 1220648-2024-21728 | 20587443 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20587443&pc=OZD | 12/07/2022 | 11/01/2024 |  | 11/01/2024 |  |  | 12/07/2022 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1129480 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported the automated impella controller (aic) experienced battery charging/other issues. no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
16042 | 1220648-2024-21729 | 20587504 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20587504&pc=OZD | 10/09/2024 | 12/10/2024 |  | 11/01/2024 |  |  | 10/09/2024 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025515164 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient had an impella 5.5 and after a difficult implant, the pump was inserted into the left ventricle and started. the pump immediately had a spike in the motor current and stopped itself to p0. the pump was gradually restarted at p2 to p5 with motor current higher than normal given p5 with flow saturation. impending pump failure was noted and the decision made to explant the pump and implant a new impella 5.5. there was no reported patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the pump stop and saturated flow has been completed. the product was returned and evaluated. the motor cable lines were all open. there were severe clamp marks found on the catheter and the catheter at the motor housing was slightly kinked. there was also biomaterial found on and around the impeller. the pump could not be run in the lab since the motor cable lines were open. log review showed an "impella stopped - motor current high" alarm after initial startup with a motor current spike. the pump was restarted and there were immediate saturated flows around 5.2 l/min. shortly after, the pump was turned back down to p-0 and then there was an alarm 119, indicating an electrical issue. it is likely that the motor cables lines were damaged by clamping during explant. the root cause of the pump stop was most likely biomaterial. the root cause of the saturated flow was most likely biomaterial.
16043 | 1220648-2024-21730 | 20587588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20587588&pc=OZD | 12/10/2022 | 11/01/2024 |  | 11/01/2024 |  |  | 12/10/2022 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  |  |  | 1319386 |  | 0042-0010 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 52 YR |  | Male |  |  | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. immediately after implanting the impella, the automated impella controller (aic) produced a "battery error 50%" message even though the device was plugged in. the console was replaced, and no patient harm was reported. | [5 empty] | The device was returned and evaluated, and the investigation for the reported aic - battery failure (red alarm) has been completed. console logs from the reported day of event are consistent with complaint and show battery-related alarms present upon console boot-up. same alarm presented a day prior to the reported. battery and power data embedded in lt log indicated that cell 4 of battery 1 was failing - this is a known pattern failure of cell imbalance. the cause of the aic (automated impella controller) issue was a defective battery. this report is being filed as part of a retrospective review of historical records.
16044 | 1220648-2024-21731 | 20587594 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20587594&pc=OZD | 10/10/2024 | 01/02/2025 |  | 11/01/2024 |  |  | 10/10/2024 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1817137 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Break | Detachment of Device or Device Component | Patient Device Interaction Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient kicked over the automated impella controller (aic) and the door that covers the power connection broke. the aic was switched and there was no harm to the patient. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - damage during therapy has been completed. during the visual inspection it is evident that the damages to the front panel optical connector and the pump connector cover were what the clinical staff meant to report. both were confirmed to be damaged. the root cause of this issue was use related. g1 reporting contact email revised on manufacturer device report 1220648-2024-21731 in accordance with updated procedures.
16045 | 1220648-2024-21733 | 20587757 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20587757&pc=OZD | 12/08/2022 | 11/01/2024 |  | 11/01/2024 |  |  | 12/08/2022 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  |  |  | 1147992 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced a battery failure (red alarm). no patient involvement was reported. | [5 empty] | The device was returned and evaluated, and the investigation for the reported aic - battery failure (red alarm) has been completed. console logs from the reported day of event are consistent with complaint and show multiple battery alarms upon console boot-up, indicating failure of battery 2. lt log data confirms that one cell of battery 2 had failed, which caused battery pack internal electronics to shut down that battery. this issue is a known pattern failure, caused by internal cell imbalance. the cause of the aic (automated impella controller) issue was a defective battery. this report is being filed as part of a retrospective review of historical records.
16046 | 1220648-2024-21734 | 20587780 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20587780&pc=OZD | 12/09/2022 | 11/01/2024 |  | 11/01/2024 |  |  | 12/09/2022 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1268745 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) showed a "system check failed - change console immediately" message during startup. the power was cycled, but the alarm remained. there was no reported patient involvement. | [5 empty] | The device was returned and evaluated, and the investigation for the reported system self-check error has been completed. the device logs are consistent with the reported failure mode. the logs show repetitive communication issues indicating errors in serial communication with pbm via channel 10. with the actual device, the system self-check failure was reproduced. the pbm was confirmed to be corrupted on both surfaces next to the super caps. the root cause of syscall error was due to a defective pbm. the root cause of the defective pbm was corrosion due to leaking super capacitors. this report is being filed as part of a retrospective review of historical records.
16047 | 1220648-2024-21732 | 20587800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20587800&pc=OZD | 07/06/2021 | 11/01/2024 |  | 11/01/2024 |  |  | 07/06/2021 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1314426 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 97 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The physician reported that the automated impella controller (aic) would not recognize when the pump was connected. it was reported the screen displayed "connect gray connectors arrow to arrow" when the pump was connected. this aic was replaced with another aic resulting in no issues with the replacement aic. there was no patient harm reported. | [5 empty] | The investigation into the pump shutdown/power on issue has been completed. the impella automated controller (aic) was returned for investigation. the data analysis of the logs indicated that immediately after the reported failure, it was determined that the pump recognition issues were due to a failure of the electrical power manager (epm) circuitry. the console was not sent for further investigation since this is a pattern failure, which resolves after power cycling the console and has a low chance of reoccurrence. software release 8.3 brought a change to reset the epm sau, which was intended to be a solution for ¿pump not recognized due to epm crystal issue¿. however, even after resetting the epm, the pump was not detected during the case start procedure. the cause of the failure to properly detect the pump was an issue with the crystal of the epm on the impellatronic board. this report is being filed as part of a retrospective review of historical records.
16048 | 1220648-2024-21735 | 20587839 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20587839&pc=OZD | 10/09/2024 | 12/03/2024 |  | 11/01/2024 |  |  | 10/09/2024 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025543759 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a hematoma was noted at the reposition sheath site when the patient was on impella cp support and arrived at the intensive care unit. the impella cp was repositioned. the site was re-sutured and angle matched. the outcome at time of explant was survived. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was most likely use since hemostasis was achieved by re-suturing and adjusting the angle.
16049 | 1220648-2024-21736 | 20587869 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20587869&pc=OZD | 12/09/2022 | 11/01/2024 |  | 11/01/2024 |  |  | 12/09/2022 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1288834 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the nurse was attempting to change the cassette, due to the hospital policy of 5 days per cassette, and the automated impella controller (aic) failed to progress through the steps. it was reported that the nurse called clinical support and was advised to change out the aic. the aic was replaced without issue and the patient reportedly remained stable throughout. additionally, a new purge cassette was used in the replacement aic. there was no patient harm reported. | [5 empty] | The investigation into the purge issue/purge flag issue has been completed. the impella automated controller was returned for investigation. the data analysis of the logs showed that during a purge cassette change, ic5090 began to alarm for "purge disc not detected". a physical and visual inspection of the purge assembly revealed that the purge disc flag had broken. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
16050 | 1220648-2024-21737 | 20587996 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20587996&pc=OZD | 12/08/2022 | 11/01/2024 |  | 11/01/2024 |  |  | 12/08/2022 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  |  |  | 1537012 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) produced a battery alarm. there was no reported patient involvement. | [5 empty] | The device was returned and evaluated, and the investigation for the reported aic - battery failure (red alarm) has been completed. lt log analysis of the battery data reveals that beginning before the complaint date of occurrence, cell 1 voltage of battery 1 had been declining for an extended period of time. it was confirmed that persistent battery failure (transport connection blown/missing) was not due to a disengaged battery switch. when console was booted for the first time in danvers, console alarms are consistent with what was seen in the field. console interior connections were checked and verified. when visually inspecting the batteries and pbm, it was observed that one of the batteries status leds were completely blank. the battery failure alarm was due to a defective battery 1 (decline of individual cell voltage). the root cause of the battery cell failure was undetermined. this report is being filed as part of a retrospective review of historical records.
16051 | 1220648-2024-21738 | 20588033 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20588033&pc=OZD | 10/10/2024 | 12/11/2024 |  | 11/01/2024 |  |  | 10/10/2024 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025553167 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during the impella cp insertion that the physician attempted multiple wires and catheters to cross the aortic valve but was unsuccessful. stenosis was suspected. the impella placement was aborted. the patient¿s outcome is unknown. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because the pump was to cross the valve due to stenosis and the insertion was aborted.
16052 | 1220648-2024-21739 | 20588034 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20588034&pc=PYX | 10/02/2024 | 12/06/2024 |  | 11/01/2024 |  |  | 10/02/2024 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1574757 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Obstruction of Flow | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) could possibly have an issue with the purge delivery. the aic will be returning for evaluation and there was no reported patient harm. | [5 empty] | The console was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation into automated impella controller (aic) - purge issue ¿ other has been completed. the cause of the issue was not determined since product was not returned.
16053 | 1220648-2024-21740 | 20588136 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20588136&pc=OZD | 12/03/2022 | 11/01/2024 |  | 11/01/2024 |  |  | 12/08/2022 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1200547 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced a battery failure (red alarm). there was no reported patient involvement. | [5 empty] | The device was returned and evaluated, and the investigation for the reported battery failure (red alarm) has been completed. lt log analysis of the battery data reveals that beginning before the complaint date of occurrence, cell 3 voltage of battery 1 had been declining for an extended period of time. it was confirmed that persistent battery failure (transport connection blown/missing) was not due to a disengaged battery switch. console interior connections were checked and verified. when visually inspecting the batteries and pbm, it was observed that one of the batteries status leds were completely blank. the battery failure alarm was due to a defective battery 1 (decline of individual cell voltage). the root cause of the battery cell failure was undetermined. this report is being filed as part of a retrospective review of historical records.
16054 | 1220648-2024-21742 | 20588250 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20588250&pc=OZD | 10/04/2024 | 12/05/2024 |  | 11/01/2024 |  |  | 10/04/2024 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025547461 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Atrial Fibrillation | Tachycardia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was presented with ventricular tachycardia (vt) storm with pulmonary hypertension and the decision was made to place the patient on impella cp support. it was reported that during the peel away sheath removal and insertion of a repo sheath, the pump came back near the valve. the doctor was unable to get it seated properly as like before. the doctor tried torque maneuvers to no avail. the doctor also tried getting a better position at bedside but pulled back across the aortic valve. the pump was explanted then reinserted and support was continued. the next day, the patient was critically ill and had a loss of pulsatility. positioning was confirmed, increase in volume and pressor support were noted. the patient was in and out of atrial fibrillation, vt, and paced rhythm with hypotension. the patient was shocked three times with minimal change. the following day, the family decided to withdraw care and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the impella cp has been completed. the device was not returned for evaluation. however, clinical details provided were sufficient to determine the most likely root cause. the root cause of the positioning issue was most likely use related since the pump pulled back during sheath removal. the root cause of the atrial fibrillation/ventricular tachycardia could not be determined without additional clinical details. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ corrections to the initial mfr report: g1-mdr reporting contact email.
16055 | 1220648-2024-21741 | 20588256 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20588256&pc=OZD | 12/07/2022 | 11/01/2024 |  | 11/01/2024 |  |  | 12/07/2022 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1133147 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella for mechanical circulatory support. during support, the automated impella controller (aic) showed a battery failure (red alarm). the console was replaced, and no patient harm was reported. | [5 empty] | The device was returned and evaluated, and the investigation for the reported aic - battery failure (red alarm) has been completed. ltlog analysis of the battery data reveals that beginning before the complaint date of occurrence, cell 2 voltage of battery 2 had been declining for an extended period of time. it was confirmed that persistent battery failure (transport connection blown/missing) was not due to a disengaged battery switch. when console was booted for the first time in lab, console alarms are consistent with what was seen in the field. console interior connections were checked and verified. when visually inspecting the batteries and pbm, it was observed that one of the batteries status leds were completely blank. the battery failure alarm was due to a defective battery 2 (decline of individual cell voltage). the root cause of the battery cell failure was undetermined. this report is being filed as part of a retrospective review of historical records.
16056 | 1220648-2024-21743 | 20588269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20588269&pc=OZD | 10/10/2024 | 12/11/2024 |  | 11/01/2024 |  |  | 10/10/2024 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025528698 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support when they went into ventricular tachycardia (vt). they were shocked and cardiopulmonary resuscitation was performed. the pump was weaned after the event. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation into vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-21743.
16057 | 1220648-2024-21744 | 20588550 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20588550&pc=OZD | 03/27/2023 | 12/02/2024 |  | 11/01/2024 |  |  | 10/03/2024 | 11/01/2024 |  |  | 11/30/2024 |  | 11/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024305092 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 58 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The clinical study protect iv reported that an impella cp was used to support percutaneous coronary intervention. post procedure in the intensive care unit, the patient was found to have a change in neurological status. the team reported that the patient suffered a cerebrovascular accident. the patient had a shunt placement monitored by neurology. decompressive craniectomy was offered but the family opted against surgery at this time and decided comfort care for the patient. the patient then expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." section: warnings, cautions & precautions: "to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers." | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined.
16058 | 1220648-2024-21745 | 20588670 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20588670&pc=OZD | 10/10/2024 | 12/11/2024 |  | 11/01/2024 |  |  | 10/10/2024 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025486085 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support, and during support a new right frontal lobe infarct was found. the pump was explanted shortly after. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email.
16059 | 1220648-2024-21746 | 20588683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20588683&pc=OZD | 10/05/2024 | 12/09/2024 |  | 11/01/2024 |  |  | 10/05/2024 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025523015 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp support for 76-year-old male patient, suction alarms appeared. there was flat motor current and reduced flows. the pump was repositioned but this did not resolve the issue. staff determined that the pump may be experiencing some obstruction. the pump was therefore removed. the patient was on impella support for 17.53 hours before the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | The investigation for the impella cp has been completed. data logs were reviewed which confirmed the failure mode and clinical narrative. logs showed pump started and run without issue for about 3.5 hours, then flow suddenly dropped to 0.5 l/min with motor current (mc) spikes that triggered auto suction response and suction alarms. this event remained to the rest of the case. the pump was returned for evaluation. visual inspection found a kink on can about 15 mm from outflow cage and cannula bonding site. no other issues were found on the rest of the pump. pump ran without issue under bench test. the root cause of low flow was improper positioning during patient movement resulting in the pump deep in the ventricle and a kink in the cannula. corrections to the initial mfr report: g1 mdr reporting contact email
16060 | 1220648-2024-21747 | 20588705 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20588705&pc=OZD | 10/13/2024 | 12/12/2024 |  | 11/01/2024 |  |  | 10/13/2024 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025533298 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support. while on support, labs were hemolyzed. education was provided on optimal pump position. blood units and fresh frozen plasma were given for hemoglobin 4. heparin was stopped. the family decided on comfort care for the patient. therefore, care was withdrawn and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the hemolysis has been completed. the clinical details stated that there were suction and positioning alarms occurring at the time of the hemolysis. the patient was low in volume and was also on an impella rp flex. data logs were reviewed which showed there were no indicative motor current spikes or a trend in the placement signal. due to lack of product returned, the root cause of the hemolysis could not be determined. b.2 revised date of patient death as it was previously entered incorrectly on initial manufacturer device report number 1220648-2024-21747. d.9 revised since the initial manufacturer device report number 1220648-2024-21747 was submitted as the pump with serial number (b)(6) was returned and not the pump in this report with serial number (b)(6). h.6 product was not returned was added since the initial manufacturer device report number 1220648-2024-21747 was submitted as the pump reported in this report was not returned as a different pump was returned. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2024-21747 was submitted.
16061 | 1220648-2024-21748 | 20588810 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20588810&pc=OZD | 10/11/2024 | 12/11/2024 |  | 11/01/2024 |  |  | 10/11/2024 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025547314 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Disability | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a cp support initiated for a high risk percutaneous coronary intervention patient. the cp was placed femorally and discussions were made for escalation to a 5.5 pump. the cp supported for 2 days until 5.5 placed. the limb was seen to be in need of a fasciotomy and the team days later made decision for the amputation, and an above knee amputation was done 4 days after explant. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: impella cp with smartassist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported limb ischemia has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the limb ischemia could not be determined. corrections to the initial mfr report:g1 mdr reporting contact email.
16062 | 1220648-2024-21750 | 20589054 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20589054&pc=OZD | 10/12/2024 | 12/24/2024 |  | 11/01/2024 |  |  | 10/12/2024 | 11/01/2024 |  |  | 05/31/2025 |  | 11/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024353800 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Female |  |  | White | Malfunction | No | Yes | Pumping Stopped | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella 5.5 support. while on support, a leak was noted from purge cassette. additionally, impella flow had changed to a saturated flow of 4.6l max and min also 4.6l. purge cassette was replaced. pump was explanted. the patient survived the event. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The pump was received for investigation. the purge leak and saturated pump flow failure modes were removed as a cascading event and addressed under the pump stop failure mode the cause of the pump stop was damage to the filter to tubing joint specifically at the filter neck. damage at the filter-to-tubing joint caused the purge leak, and drop in purge pressure led to blood ingress in the purge line and motor housing, which caused the saturated pump flow. instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: direct aortic insertion. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ g1 reporting contact email revised on manufacturer device report 1220648-2024-21750 in accordance with updated procedures.
16063 | 1220648-2024-21749 | 20589061 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20589061&pc=PYX | 10/04/2024 | 12/05/2024 |  | 11/01/2024 |  |  | 10/04/2024 | 11/01/2024 |  |  | 06/30/2025 |  | 11/01/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024371150 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that while loading and advancing the impella rp over the impella 0.027 wire, the medical doctor noticed the 0.027 wire had fallen too far back via fluoroscopy requiring further advancement. the impella rp was then slid back to the distal end of the 0.027 wire while maintaining the thread of the red lumen on the wire and back into the impella rp cannula. the scrub technician felt resistance while re-advancing the red lumen onto the 0.027 wire and into the impella rp cannula. on further inspection, the scrub noticed a ¿build up¿ in the outflow cage. the medical doctor was concerned about ¿build up¿ being introduced into the patient¿s body. the decision was made to exchange the impella rp. | [5 empty] | The investigation for the impella rp has been completed. the pump was returned for evaluation, upon visual inspection no abnormalities with the pump were noted. no guidewire or red lumen were returned for evaluation. clinical details noted resistance was felt when attempting to slide wire through red lumen and build-up was noted in outflow cage. the root cause of the delivery issue was most likely patient condition since the implant was aborted due to reported buildup on the outflow. | The pump was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella rp smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿
16064 | 1220648-2024-21751 | 20589170 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20589170&pc=PYX | 10/04/2024 | 10/21/2025 |  | 11/01/2024 |  |  | 10/04/2024 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1667361 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) received a controller error alarm when prepping and connecting to the impella. the aic was exchanged. | [5 empty] | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
16065 | 1220648-2024-21753 | 20589571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20589571&pc=PYX | 10/04/2024 | 12/06/2024 |  | 11/01/2024 |  |  | 10/04/2024 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025518034 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had impella rp flex support ongoing for a 58-year-old male patient admitted with cardiomyopathy and other systemic comorbidities. the team additionally had an impella 5.5 supporting. the pump was removed after 19 hours due to low flows. the pump had motor current and purge flow changes. the patient survived the explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the low pump flow has been completed. the pump was not returned for investigation. the data log shows that the motor current spiked while running at steady p-level, accompanied by a drop in pump flow. a 9-minute gradual purge pressure rise was observed during the low pump flow event with active high purge pressure alarm, and the placement signal spiked to 150. the cvp was positive and exhibited some spikes during the reported event. the data log trend indicates signs of biomaterial ingestion during the case run. according to the clinical report, high purge pressure and low pump flow were noted, along with a motor current spike. a clot was reported on the inlet/outlet cage of the explanted pump. the cause of the low pump flow issue was most likely biomaterial. the biomaterial interaction could also lead to the high purge pressure. the instructions for use for the impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller states the following: section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
16066 | 1220648-2024-21752 | 20589600 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20589600&pc=OZD | 12/07/2022 | 11/01/2024 |  | 11/01/2024 |  |  | 12/07/2022 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1393502 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. while on support, the aic experienced controller failure (red alarm). there was no report of patient harm or injury. | [5 empty] | The investigation into the controller failure (red alarm) issue has been completed. the automated impella controller (aic) was returned for investigation. data analysis: controller failure (purge pressure sensor (pps) defect ¿ event set #418) was confirmed in the aic logs. this alarm triggers when the voltage coming from the pps is out of range and was persistently occurring. the root cause of the controller failure alarm was due to a defective pps. before returning this console back to the customer, field service was instructed to replace the pps and perform a full functional check on the console and optical system. this report is being filed as part of a retrospective review of historical records.
16067 | 1220648-2024-21754 | 20589777 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20589777&pc=OZD | 10/10/2024 | 12/11/2024 |  | 11/01/2024 |  |  | 10/10/2024 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025547797 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the impella 5.5 device was not delivered due to cardiac anatomy prior to liberation from cardiopulmonary bypass after bioprosthetic aortic valve placement. the decision was made to abort the impella. there was no reported patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation into delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related, unable to deliver impella device due to cardiac anatomy as per the clinical description provided. the following have been reported incorrectly on manufacturer device report.
16068 | 1220648-2024-21755 | 20590012 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20590012&pc=OZD | 10/04/2024 | 12/09/2024 |  | 11/01/2024 |  |  | 10/04/2024 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025490139 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 84-year-old male patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. while on support, a slight hematoma noted in right groin, pressure held by bedside nurse and hematoma softened up and not growing in size. additionally, after repositioning sheath in place, it was noted there was sluggish flow in the right lower extremity. swelling was noted at the surgical site. decision was made to remove impella due to cold leg. care was withdrawn and patient eventually expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation into limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-21755.
16069 | 1220648-2024-21757 | 20590134 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20590134&pc=OZD | 12/06/2022 | 11/01/2024 |  | 11/01/2024 |  |  | 12/06/2022 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1093076 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. during support, the automated impella controller (aic) showed a battery failure (red alarm). the console was replaced to continue the procedure, and no patient harm was reported. | [5 empty] | The device was returned and evaluated, and the investigation for the reported battery failure (red alarm) has been completed. lt log analysis of the battery data reveals that beginning before the complaint date of occurrence, cell 1 voltage of battery 2 had been declining for an extended period of time. it was confirmed that persistent battery failure (transport connection blown/missing) was not due to a disengaged battery switch. when console was booted for the first time in lab, console alarms are consistent with what was seen in the field. the battery failure alarm was due to a defective battery 2 (decline of individual cell voltage). the root cause of the battery cell failure was undetermined. this report is being filed as part of a retrospective review of historical records.
16070 | 1220648-2024-21756 | 20590137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20590137&pc=OZD | 10/10/2024 | 12/27/2024 |  | 11/01/2024 |  |  | 10/10/2024 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025548325 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella cp placement, bleeding from the access site occurred. manual compression and injection of adrenaline around puncture site was done. the bleeding was reduced but was still persistence. the following day, the impella was removed. the patient survived the explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of access site bleeding was determined to be patient condition because the vessels are heavily calcified and vessels are completely covered with plaque. corrections to the initial mfr report: g1 mdr reporting contact email.
16071 | 1220648-2024-21758 | 20590295 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20590295&pc=OZD | 10/12/2024 | 12/03/2025 |  | 11/01/2024 |  |  | 10/12/2024 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1071329 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Display or Visual Feedback Problem | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp had issues with low flows and changing the aic resolved the flow issues. there was no harm to the patient. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
16072 | 1220648-2024-21759 | 20591404 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20591404&pc=OZD | 10/05/2024 | 12/05/2024 |  | 11/01/2024 |  |  | 10/05/2024 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025500753 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 63-year-old female patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. the pump was placed and an impella rp as well. the cp was supporting for 2 days when the pump was removed and replaced due to optical signal loss. the second pump was successfully placed. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation low pump flow has been completed. the impella device was not returned for evaluation. the cause of the low pump flow issue was most likely patient condition related to low blood volume, based on given clinical details and data log review. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-21759.
16073 | 1220648-2024-21760 | 20591509 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20591509&pc=OZD | 12/05/2022 | 11/01/2024 |  | 11/01/2024 |  |  | 12/05/2022 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1103522 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  |  | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that their automated impella controller (aic) shutdown unexpectedly and was followed by an unexpected controller shutdown alarm. the aic was removed and replaced with another aic as a result of the failure. there was no patient injury resulting from the noted issue. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the root cause of the unexpected controller shutdown (imminent shutdown) was not determined due to unable to reproduce. this report is being filed as part of a retrospective review of historical records.
16074 | 1220648-2024-21761 | 20592741 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20592741&pc=OZD | 12/02/2022 | 11/01/2024 |  | 11/01/2024 |  |  | 12/02/2022 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1603745 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on an automated impella controller (aic) for mechanical circulatory support. complainant reported the aic experienced controller failure (red alarm). no patient was involved. | [5 empty] | The investigation into the controller failure (red alarm) issue has been completed. the automated impella controller (aic) was returned for investigation. data analysis: there are two separate failures in the imc logs. attached log "imc (b)(6) transport switch" shows the initial battery failure alarm received after receiving the console back from pm. this alarm can trigger when the battery transport connection switch is disengaged, and the console is booted up on ac power. then attached log "imc (b)(6) pump overcurrent" shows that there were multiple pump stops from high motor current with a demo pump. device analysis: the error from the transport switch was able to be reproduced by disengaging battery switch. no other issues with the battery and power management system were observed. for the high motor current stops, the alarms were unable to be reproduced during investigation with pump (b)(6). since these errors occurred while using the demo pump 6/7, it is likely that the issues stemmed from the demo pump in use. this report is being filed as part of a retrospective review of historical records.
16075 | 1220648-2024-21762 | 20593030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20593030&pc=OZD | 11/11/2022 | 11/01/2024 |  | 11/01/2024 |  |  | 11/11/2022 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1274109 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that during preventive maintenance the automated impella controller (aic) was not displaying to proceed when starting a case after the purge cassette was inserted. this aic was replaced with another aic and the procedure was successfully completed. there was no patient involvement. | [5 empty] | E1: name of the initial reporter is unknown. the investigation into the purge issue has been completed. the impella automated controller was returned for investigation. the data analysis of the logs showed case start was initiated and the purge cassette and disc were inserted, however it can be seen that only the disc is inserted and detected. the case start wizard also does not progress due to this. additionally, a photo of the incident was taken, showing the cassette and disc inserted into the console with the case wizard halted. the returned console was opened and inspected with attention focused on the purge assembly. no visual issues were found. the cassette detection was tested again, via purgecmd vshmem, where it was seen that placing the cassette into the console would not trigger the registration. different cassettes were tested, however, testing resulted in the same error. the purge unit driver (pud) was swapped out with a known good pud, and cassette detection was functional again. purgecmd vshmem was run again showing successful detection. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
16076 | 1220648-2024-21763 | 20593745 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20593745&pc=OZD | 12/01/2022 | 11/01/2024 |  | 11/01/2024 |  |  | 12/01/2022 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1171908 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced a battery failure (red alarm). there was no reported patient harm. | [5 empty] | The device was returned and evaluated, and the investigation for the reported battery failure (red alarm) has been completed. aic logs confirmed the reported failure. there was a battery failure alarm and a battery 2 communication failure alarm due to low pack voltage. the lt power data from the complaint date showed a cell voltage decline in only one of battery 2¿s cells which occurred as the battery failure alarm triggered. the failure mode was reproduced on an engineering bench. the battery lcd indicator was blank (fully discharged). battery 2 was momentarily replaced to resolve the failure alarm. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
16077 | 1220648-2024-21764 | 20593757 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20593757&pc=OZD | 12/02/2022 | 11/01/2024 |  | 11/01/2024 |  |  | 12/02/2022 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1091705 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced a battery failure. no patient involvement was reported. | [5 empty] | The device was returned and evaluated, and the investigation for the reported battery failure (red alarm) has been completed. aic logs confirmed the reported failure. there was a battery failure alarm and a battery 2 communication failure alarm due to low pack voltage. additionally, a battery level low and battery failure (v<12v) was triggered on the same power on. the ltpowerdata from the complaint date showed a cell voltage decline in only one of battery 2¿s cells which occurred as the battery failure alarm triggered. the failure mode was reproduced on an engineering bench. the battery 2 pack voltage was measured to be 3.5 v rather than the expected 16 v, and the battery lcd indicator was blank (fully discharged). battery 2 was replaced to resolve the failure alarm. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
16078 | 1220648-2024-21765 | 20593763 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20593763&pc=OZD | 11/29/2022 | 11/01/2024 |  | 11/01/2024 |  |  | 11/29/2022 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1160667 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | Insufficient Information | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. while on support, the aic experienced controller failure (red alarm) that was resolved by replacing the aic. there was no report of patient harm or injury. | [5 empty] | The investigation into the controller failure (red alarm) issue has been completed. the automated impella controller (aic) was returned for investigation. data analysis: aic logs revealed that while the user was viewing the placement screen the console had unexpectedly rebooted while powered by ac. as the console booted back up, the impellatronic had to be reinitialized, and therefore, the pump was not running for roughly one minute. once the console booted back up, the previous p-level was resumed as designed. device analysis. the console was tested thoroughly on an engineering lab bench. the unexpected controller reboot was not reproduced. the epc supply from pbm was stable on both ac and battery power during a long-duration test (no interruptions). there were no intermittent electrical connections throughout the console that would result in an unexpected reboot. the root cause of the unexpected reboot could not be determined due to the inability to reproduce. before sending this console back to the customer, field service was instructed replaced the 4-in-1 (0042-1016/17) and compact flash (2025-0019) as a precaution. this report is being filed as part of a retrospective review of historical records.
16079 | 1220648-2024-21766 | 20593776 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20593776&pc=OZD | 12/02/2022 | 11/01/2024 |  | 11/01/2024 |  |  | 12/02/2022 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1240263 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. it was noted after preparing the pump, it was connected to the automated impella controller (aic). the aic then reported an optical sensor error. the staff disconnected the pump once and checked the sensor, but the problem persisted. a second aic was started. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the aic (automated impella controller) issue was a defective blue receptacle. this report is being filed as part of a retrospective review of historical records.
16080 | 1220648-2024-21767 | 20593781 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20593781&pc=OZD | 11/21/2022 | 11/01/2024 |  | 11/01/2024 |  |  | 11/21/2022 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1669290 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | Insufficient Information | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported the aic experienced controller failure (red alarm). no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | The investigation into the controller failure (red alarm)issue has been completed. the automated impella controller (aic) was returned for investigation. imc logs were reviewed and confirm only one of the reported failure modes. console was unexpectedly shutdown at the end of a patient case. once the console booted up, pmd serial communication issue entries were observed occasionally upon boot up. console recognized the unexpected shutdown by issuing the alarm #603 - unexpected controller shutdown. however, console logs had gaps which led to inability to confirm the reported controller failure alarm. the controller failure alarm was most likely occurred when imc logs stopped logging, therefore, there is no visibility from the existing logs. console was tested with known good pump and purge set for overnight case simulation. although the unexpected controller shutdown was not reproduced which supports the idea that the original issue was the inability to shut down the console properly, it was observed that the impellatronic cpld buzzer was beeping post overnight testing, indicating there was communication issue within impellatronic circuitry. this observed communication issue would lead to controller failure alarm as complaint reported. the controller failure, although unable to be confirmed via aic logs, was most likely due to a defective impellatronic (communication issues). the inability to properly shutdown the console was not determined due to inability to reproduce and confirm from the incomplete imc logs (due to corrupted compact flash cards). there was no clear evidence to prove any correlation between the corrupted compact flash cards and the inability to shut down the console. this report is being filed as part of a retrospective review of historical records.
16081 | 1220648-2024-21768 | 20593783 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20593783&pc=OZD | 05/12/2022 | 11/01/2024 |  | 11/01/2024 |  |  | 05/12/2022 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1651078 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  |  | White | Malfunction | No | Yes | Use of Device Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 for mechanical circulatory support. while on support, the automated impella controller (aic) experienced purge issues and an overheating battery, which were resolved by switching to a new console. no patient harm reported. | [5 empty] | The investigation has been completed. the console (b)(6) was sent back for further investigation. console was tested on pump and purge for extended period of time without issue. it is most likely that the issues seen in the field have occurred due to use issues related to the purge lines. there was no issue found during the investigation of the console. the failure mode could not be reproduced. this report is being filed as part of a retrospective review of historical records.
16082 | 1220648-2024-21769 | 20593800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20593800&pc=OZD | 03/20/2023 | 11/01/2024 |  | 11/01/2024 |  |  | 03/20/2023 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1261275 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced and optical sensor system error. alert. it was reported that a loaner device was sent to the customer. no report of patient involvement. | [5 empty] | E1: the name of the biomedical engineer is unknown. the investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the aic issue was a defective optical bench. this report is being filed as part of a retrospective review of historical records.
16083 | 1220648-2024-21770 | 20593810 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20593810&pc=OZD | 02/06/2023 | 11/01/2024 |  | 11/01/2024 |  |  | 02/06/2023 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1250484 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the aic experienced purge disc/flag issues. no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | The investigation into the purge disc/flag issues issue has been completed. the automated impella controller (aic) was returned for investigation. visual inspection of the purge assembly area confirmed a broken blue retainer. the cause of the issue was broken/missing purge disk retainer. this report is being filed as part of a retrospective review of historical records.
16084 | 1220648-2024-21771 | 20593819 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20593819&pc=OZD | 01/30/2023 | 11/01/2024 |  | 11/01/2024 |  |  | 01/30/2023 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1535515 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant had an automated impella controller in for a preventative maintenance (pm) and during the pm damage was observed. the burned power battery management (pbm) circuit board caused no harm or injury as there was no patient use at the time. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the root cause of the burnt capacitor was most likely a failure of the power battery management board. this report is being filed as part of a retrospective review of historical records.
16085 | 1220648-2024-21772 | 20593824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20593824&pc=OZD | 11/25/2022 | 11/01/2024 |  | 11/01/2024 |  |  | 11/25/2022 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1404666 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that their automated impella controller (aic) shutdown twice during patient support. the aic was removed and replaced with a new console to continue with patient support. there was no patient harm as a result of the reported failure. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the root cause of the unexpected controller shutdown was due to a software process failure. the root cause of the software process failure was not determined. this report is being filed as part of a retrospective review of historical records.
16086 | 1220648-2024-21773 | 20593827 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20593827&pc=OZD | 11/22/2022 | 11/01/2024 |  | 11/01/2024 |  |  | 11/22/2022 | 11/01/2024 |  |  |  |  | 11/01/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1116343 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Display or Visual Feedback Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. while priming, the automated impella controller (aic) produced a "failure to calibrate optical sensor" alarm, and there was a glitch with the aic display. the console was replaced with another, and there was no reported patient harm. | [5 empty] | The device was returned and evaluated, and the investigation for the reported optical signal issue and aic - display issue has been completed. imc logs confirmed the optical calibration failure upon case start. however, the reported screen glitch issue was unable to be confirmed. during testing, the device was confirmed to be within specifications, and the failure mode was unable to be reproduced. inspection on the optical fiber connection did not reveal any abnormality. the root cause of the optical sensor calibration issue was highly likely due to an air gap at the front panel optical connect because the user did not connect the pump all the way in. as for the other failure mode, the screen glitch was unable to be reproduced during case start simulation. long term camera monitoring for display artifacts did not reveal any screen malfunction. the root cause of the glitch screen was undetermined due to inability to reproduce. this report is being filed as part of a retrospective review of historical records.
16087 | 1220648-2024-21629 | 20577114 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20577114&pc=OZD | 10/02/2024 | 11/18/2024 |  | 10/31/2024 |  |  | 10/02/2024 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025526018 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Vascular Dissection | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was on impella 5.5 support. while on support, the patient experienced ventricular tachycardia (vt) and required being shocked 8 times. additionally, imaging showed an aortic dissection. care was withdrawn and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation vascular damage - great vessel and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during. section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
16088 | 1220648-2024-21630 | 20577146 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20577146&pc=OZD | 10/08/2024 | 12/09/2024 |  | 10/31/2024 |  |  | 10/08/2024 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1276714 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella support when the automated impella controller's (aic) screen was black and the pump turned off. the nurse found the aic screen was off. they pressed the silver knob and the screen came alive. the aic was restarted at the previous level like the aic to aic transfer screen. the aic and pump were restarted and support continued. there was no reported patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation for the console (aic) pump stop has been completed. console logs from the reported day of event are consistent with complaint and show that after approximately 5.5 hours of support, one of software processes (imcgui) became unresponsive, triggering watchdog mechanism that restarted the aic software (partial reset), as per design, to recover and restart the process. prior and during the restart, pump kept running, as it was supported by impellatronic board supporting pump autonomously. after the restart, pump was recognized as already connected, and then was auto-restarted at the prior p-level. during the software initialization, the pump was paused for about 25 seconds (time between reinitialization of pmd and pump start). ltlog and rtlog data indicates that other than during the partial reset, there were no interruptions of hemodynamic support. the console was returned for evaluation but the issue was not reproduced, and console operated within specification. when imcgui process was purposefully stopped, the system reacted in same fashion as in the field, and fully recovered after restart. the imcgui process is a part of aic software. the reason for the imcgui process to ¿crash¿ was not determined, although might have been related to a possible data corruption.
16089 | 1220648-2024-21631 | 20577164 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20577164&pc=OZD | 10/08/2024 | 12/05/2024 |  | 10/31/2024 |  |  | 10/08/2024 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025526759 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that during the impella 5.5 insertion there was difficulty advancing past innominate artery with buckling noted at motor housing and catheter junction. the impella was exchanged and there was no reported patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, difficulty advancing impella past innominate/subclavian as per the clinical description provided. corrections to the initial mfr report: g1 mdr reporting contact email.
16090 | 1220648-2024-21632 | 20577166 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20577166&pc=OZD | 10/08/2024 | 12/06/2024 |  | 10/31/2024 |  |  | 10/08/2024 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025521639 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 61-year-old female noted as high risk coronary intervention was implanted with an impella cp device for mechanical circulatory support. after impella implant, there was loss of pulses in the right lower extremity. right anterograde sheath was placed along with left retrograde sheath for fem-fem bypass. pulses were dopplerable and color and temperature returned. | [5 empty] | The investigation into limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-21632. e4 should have been left blank. g1 reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
16091 | 1220648-2024-21633 | 20577207 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20577207&pc=OZD | 10/07/2024 | 11/21/2024 |  | 10/31/2024 |  |  | 10/07/2024 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025516117 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during the impella cp placement after drapes were removed, impella in aorta alarm occurred. imaging under fluoroscopy confirmed outlet was in the descending aorta. the patient was re-prepped, and the doctor attempted several times to insert impella back across the aortic valve utilizing reaccess port without success. the impella cp was removed and insertion was aborted. there was no reported harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the impella cp has been completed. the device was not returned for evaluation. however, clinical details provided were sufficient to determine the most likely root cause. the cause of the positioning issue was most likely use related since the pump was pulled out while removing drapes from patient.
16092 | 1220648-2024-21634 | 20577317 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20577317&pc=OZD | 10/08/2024 | 12/11/2024 |  | 10/31/2024 |  |  | 10/08/2024 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025502467 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had an impella cp support ongoing for a 79-year-old male patient admitted in an acute myocardial infarction / cardiogenic shock (ami/cgs). the pump was placed and support was ongoing but there were signs of hemolysis. the team had labs and urine color change. the pump was troubleshot with fluids, p-level adjustment, and explanted after 32 hours. | [5 empty] | The investigation into hemolysis has been completed. the impella device was not returned for evaluation. the cause of the hemolysis issue was most likely patient condition related. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-21634. e4 should have been left blank. g1 reporting contact email | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16093 | 1220648-2024-21637 | 20577321 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20577321&pc=OZD | 10/08/2024 | 12/19/2024 |  | 10/31/2024 |  |  | 10/09/2024 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Delivered as Unsterile Product | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A complainant reported the insertion kit was opened and unsterile. no impella implantation and no patient harm was reported. patient and product information are unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the product damage-sterility issues has been completed. no product was returned for evaluation. clinical details noted insertion kit was opened and unsterile. the root cause of the packaging issues - sterility cannot be definitively determined.
16094 | 1220648-2024-21635 | 20577322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20577322&pc=OZD | 10/08/2024 | 12/11/2024 |  | 10/31/2024 |  |  | 10/08/2024 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025495347 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 90 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support due to acute decompensation. it was reported that the purge reservoir and filter sidearm cracked. several hours went by and air in purge system alarm occurred. bypass was successfully completed and the purge flow returned within normal range. at the end of support, the pump was successfully weaned and the patient survived the explant. | [5 empty] | The investigation into purge leak ¿ pump has been completed. no product was returned for analysis. insufficient data logs were provided. data logs provided ended on 10/04 while occurrence date was 10/8. the root cause of the purge leak - pump was not determined as no product or images were returned g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-21635. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
16095 | 1220648-2024-21636 | 20577323 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20577323&pc=OZD | 07/26/2024 | 12/03/2025 |  | 10/31/2024 |  |  | 10/09/2024 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1307027 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Female |  |  | White | Malfunction | No | Yes | Key or Button Unresponsive/not Working | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported that during impella cp support, the 69-year-old female patient, it was suddenly not possible to make changes on the automated impella controller (aic) via the soft buttons. nothing happened with pressing these buttons. the aic had to be changed and there was no patient harm reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Corrections to the initial mfr report: g1 mdr reporting contact email. d2 device name has been updated. h6 type of investigation code added.
16096 | 1220648-2024-21638 | 20577324 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20577324&pc=OZD | 05/12/2023 | 12/10/2024 |  | 10/31/2024 |  |  | 10/09/2024 | 10/31/2024 |  |  | 12/31/2024 |  | 10/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024325056 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 16 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | On 2024-10-09 abiomed japan forwarded a publication entitled catecholamine-induced cardiomyopathy secondary to pheochromocytoma rescued with percutaneous left ventricular assist device: novel application of the impella cp in a pediatric patient. the patient on impella cp support had a femoral artery dissection noted at the pump explant in (b)(6) 2023. the injury per publication was immediately recognized and repaired by vascular surgery with a good result and no complications to the affected extremity. | [5 empty] | The investigation into vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. citation: zumwalt, c. m., roybal, j., crittendon, i., lucas, v. s., & wells, d. a. (2024). catecholamine-induced cardiomyopathy secondary to pheochromocytoma rescued with percutaneous left ventricular assist device: novel application of the impella cp in a pediatric patient. ochsner journal, 24(3), 229¿232. https://doi.org/10.31486/toj.24.0028 the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-21638. e4 should have been left blank. g1 reporting contact email. g1 reporting contact fax number missing. h10: article citation missing | Refer to the literature attachment. the impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
16097 | 1220648-2024-21639 | 20577342 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20577342&pc=OZD | 10/07/2024 | 11/21/2024 |  | 10/31/2024 |  |  | 10/07/2024 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531275 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 77-year-old female patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. the impella cp pump was placed via single access. the pump allowed for percutaneous coronary intervention (pci) for the coronary artery disease (cad) and was explanted due to the hematoma at the access site. the hematoma then had manual pressure/compression. | [5 empty] | The investigation for the reported access site bleeding major has been completed. the impella device has not been returned for evaluation. however, as the necessary clinical information was not provided and the device was not returned, the root cause of the access site bleeding - major could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16098 | 1220648-2024-21640 | 20577389 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20577389&pc=OZD | 10/04/2024 | 12/13/2024 |  | 10/31/2024 |  |  | 10/08/2024 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025516568 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient developed lower limb ischemia following insertion of an impella cp pump for mechanical circulatory support. an antegrade sheath was subsequently placed and blood was sent from the extracorporeal membrane oxygenation circuit to resolve the condition. support was continued with the implanted pump until planned explant three days later. it was noted that there was no ischemia following the pump removal and the patient survived the explant. | [5 empty] | The investigation into limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-21640. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16099 | 1220648-2024-21641 | 20577434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20577434&pc=OZD | 06/01/2022 | 10/31/2024 |  | 10/31/2024 |  |  | 06/01/2022 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1443894 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced wan alarm message about no placement signal, which was resolved through new console. no patient harm was reported. | [5 empty] | The investigation has been completed. the console (ic7550) was sent back for further investigation. the root cause of the issue was not determined because issue was unable to be reproduced. this report is being filed as part of a retrospective review of historical records.
16100 | 1220648-2024-21642 | 20577472 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20577472&pc=OZD | 10/08/2024 | 12/11/2024 |  | 10/31/2024 |  |  | 10/08/2024 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532067 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 77-year-old female patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. during impella support, the patient had access site bleeding. two units of blood products were given to the patient, perclose were tightened, and angle matching was maintained. | [5 empty] | The investigation into access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-21642. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16101 | 1220648-2024-21643 | 20577538 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20577538&pc=OZD | 10/08/2024 | 12/05/2024 |  | 10/31/2024 |  |  | 10/08/2024 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1085772 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during prep of the automated impella controller that the he front panel was cracked. the controller was removed and replaced from service for any future cases. there was no patient involvement. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation into automated impella controller (aic) - damage before therapy has been completed. during the data analysis no evidence was found in the logs on the reported complaint occurrence date, as the reported failure mode was physical damage to the aic housing. most likely, the damage could have occurred before the therapy. the root cause for the front panel crack was not determined but may have been a use issue.
16102 | 1220648-2024-21644 | 20577949 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20577949&pc=OZD | 10/06/2024 | 12/18/2024 |  | 10/31/2024 |  |  | 10/09/2024 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532066 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemolysis | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 56-year-old male patient in an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed hemolysis. it is unclear if hemolysis was resolved. additionally, cerebral infarction and cerebral hemorrhage were found on computed tomography. no treatment was given and anticoagulant therapy was discontinued. the impella was removed. the patient survived the explant. | [5 empty] | The investigation into stroke/cva and hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-21644. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella cp with smartassist system section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
16103 | 1220648-2024-21645 | 20577963 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20577963&pc=OZD | 10/07/2024 | 12/10/2024 |  | 10/31/2024 |  |  | 10/08/2024 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025534074 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Rupture | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was further reported that the patient was eventually successfully weaned and survived the explant. | The complainant reported a 63-year-old male patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. while on support, oozing and bleeding was noted around the access site. manual pressure was applied and figure 8 suture was deployed on the site. additionally, according to echo, physician found rupture of mitral valve chordae and papillary muscle. |  |  |  |  | The investigation for the heart valve damage and the access site bleed has been completed. the pump was returned and evaluated. no abnormalities were found with the pump or repositioning sheath. logs were not applicable. clinical details state that the physician had difficulty correctly placing the impella during implant. the root cause of the heart valve damage was most likely improper positioning. the bleeding occurred after the switch to the repositioning sheath. it is unknown if best practices were followed such as forward suturing and angle matching. the bleeding resolved with manual pressure and a figure 8 suture. the root cause of the access site bleeding was most likely use related. added b5- patient outcome corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16104 | 1220648-2024-21646 | 20577966 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20577966&pc=PYX | 12/27/2022 | 10/31/2024 |  | 10/31/2024 |  |  | 12/27/2022 | 10/31/2024 |  |  | 12/31/2023 |  | 10/31/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023158235 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. the complainant had an impella rp that failed delivery despite all attempts. the pump was hung up on the previously placed extra-corporeal membrane oxygenation (ecmo) cannula. the attempted placements caused the rp to kink. it was removed and not replaced. no reported harm or injury. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the delivery issues was an interaction with another device. as noted in the impella ifu: impella smartassist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
16105 | 1220648-2024-21647 | 20578026 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20578026&pc=OZD | 10/10/2024 | 12/10/2024 |  | 10/31/2024 |  |  | 10/10/2024 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025554000 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the sterile sleeve separated from the toughy-lock when the scrub technician pulled on the sterile sleeve too hard. there was no harm to the patient. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the product damage has been completed. the product was returned and evaluated. the sterile sleeve was detached from the connector on the end closest to the motor housing. no defects were found with the product. the root cause of the sterile sleeve damage was use-related. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high-risk pci states the following: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
16106 | 1220648-2024-21648 | 20578408 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20578408&pc=OZD | 06/10/2021 | 10/31/2024 |  | 10/31/2024 |  |  | 06/10/2021 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1230742 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  |  | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella device for mechanical circulatory support. during support, the automated impella controller (aic) experienced a controller error yellow alarm that was resolved by switching to a backup console. there was no reported patient harm. | [5 empty] | The device was returned and evaluated, and the investigation for the reported controller error (yellow alarm) has been completed. imc logs show that the console lost communication with epm (on the impellatronic board) which resulted in a controller error. it also had a pbm related controller error due to low voltage. the failure was reproduced upon arrival. the console failed to bootup successfully and landed on the system self check failure screen (blue screen). replacing the impellatronic board fixed the failure mode. the cause of the yellow controller error was a defective impellatronic board. the cause of the defective impellatronic board is unknown. this report is being filed as part of a retrospective review of historical records.
16107 | 1220648-2024-21651 | 20578433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20578433&pc=OZD | 05/26/2021 | 10/31/2024 |  | 10/31/2024 |  |  | 05/26/2021 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1125336 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced a controller error yellow alarm. there was no patient involvement. | [5 empty] | The device was returned and evaluated, and the investigation for the reported controller error (yellow alarm) has been completed. console logs from the reported day of event are consistent with complaint and show a single occurrence of kbd communication error followed by alarm #24 (loss of communication with kbd). console was powered on again, the next day, but the alarm did not reoccur. after console was brought to (b)(6), the issue was initially not reproduced during several power-cycles. the kbd assembly and its ribbon cable were visually inspected, and no obvious signs of damage or degradation (e.g. corrosion) were observed. in the course of troubleshooting the console was set up to power-cycle 500 times (using power-cycling fixture), after which console logs were analyzed. analysis revealed 13 occurrences of kbd communication errors and 15 occurrences of alarm #24. after the 4-in-1 was replaced and test repeated, occurrences of kbd-related errors and alarms was reduced. the cause of the controller alarm upon boot-up was loss of communication between epc and kbd, most likely due to kbd assembly malfunction. this report is being filed as part of a retrospective review of historical records.
16108 | 1220648-2024-21650 | 20578462 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20578462&pc=OZD | 05/26/2021 | 10/31/2024 |  | 10/31/2024 |  |  | 05/27/2021 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1126837 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced a controller failure red alarm. there was no reported patient involvement. | [5 empty] | The device was returned and evaluated, and the investigation for the reported controller failure (red alarm) has been completed. rt logs show that all optical bench values railed to -16384. optical values stayed invalid for about 5 minutes, and the console displayed controller error #1045 (crc error from opm comm.). after that the values go back to normal and controller error goes away. the failure mode was not reproduced in field service. this report is being filed as part of a retrospective review of historical records.
16109 | 1220648-2024-21652 | 20578524 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20578524&pc=OZD | 05/25/2021 | 10/31/2024 |  | 10/31/2024 |  |  | 05/25/2021 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1088056 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced a controller error yellow alarm during preventive maintenance. there was no patient involvement. | [5 empty] | The device was returned and evaluated, and the investigation for the reported controller error (yellow alarm) has been completed. console logs from the reported day of event are consistent with complaint and show alarm # 1035 (opm communication stopped) upon console boot-up. after console was brought to the lab, the issue was reproduced. visual inspection of the console's internal components revealed broken cable that provides power from pbm to optical bench. this was due to the weight of mechanically unsupported ferrite, as the retaining clip was not securely attached to the optical bench carrier pcb. the cause of the controller alarm was loss of communication to optical bench due to broken cable that provides power to optical bench from pbm. broken cable was result of unsupported ferrite due to cable retainer not being installed securely. this report is being filed as part of a retrospective review of historical records.
16110 | 1220648-2024-21653 | 20578534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20578534&pc=OZD | 05/14/2021 | 10/31/2024 |  | 10/31/2024 |  |  | 05/14/2021 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1145437 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. while on support, the automated impella controller (aic) experienced a controller error yellow alarm with no resolution specified. there was no reported patient harm. | [5 empty] | The device was returned and evaluated, and the investigation for the reported controller error (yellow alarm) has been completed. the console was tested thoroughly on an engineering lab bench. this included connecting/ reading the eeprom of two different pump types more than ten times, while power cycling between connections. there were no impella defective alarms triggered throughout this testing, and further review of the imc logs revealed no underlying errors while reading. the impella defective failure mode was not reproduced under normal circumstances. however, an impella defective alarm was triggered by not connecting the pump connector fully during the read of the eeprom. the root cause of the impella defective alarm in the clinical setting was most likely due to the pump connector not fully seated at the blue receptacle. this report is being filed as part of a retrospective review of historical records.
16111 | 1220648-2024-21649 | 20578551 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20578551&pc=OZD | 06/02/2021 | 10/31/2024 |  | 10/31/2024 |  |  | 06/03/2021 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | ABIOMED, INC. |  |  | US |  | 1229947 |  | 0042-0000-EU |  | Yes |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced a controller error yellow alarm during check-in training. there was no reported patient involvement. | [5 empty] | During icu check - training, the console had a red controller error. no patient involved. the logs show controller error on the complaint date (2021-06-02) due to loss of communication with pmd (event #1022). there are numerous sau_motor_1 crc errors. the logs also show that during a case on 2021-05-24, the console had numerous sau_motor_1 errors and eventually ended up with a red controller error for dsp voltage too low (event #800) but it cleared within a minute. due to comm. errors with pmd, the console failed to detect pump disconnection and prevented the staff from shutting down the console. the failure (#1022) was reproduced in fs. replacing the impellatronics fixed the failure mode. before shipping the console back, fs was asked to replace impellatronic board and perform pm.
16112 | 1220648-2024-21654 | 20578577 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20578577&pc=OZD | 10/10/2024 | 12/23/2024 |  | 10/31/2024 |  |  | 10/10/2024 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025481287 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp purge fluid was leaking from the purge set and the pump connection. there was a crack on the impella side. the impella was successfully weaned and the patient survived the explant. | [5 empty] | The investigation for the purge leak has been completed. the pump and purge cassette were returned for investigation. a data log review was not required for this case. as per the clinical, a purge leak was observed from the pump side connection. a lengthwise crack on the side arm luer was found on the returned product. the purge was leaking from the luer connection during testing. no issue was found on returned purge cassette. the cause of the purge leak issue was sidearm yellow luer damage. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
16113 | 1220648-2024-21655 | 20578578 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20578578&pc=OZD | 05/19/2021 | 10/31/2024 |  | 10/31/2024 |  |  | 05/19/2021 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1422558 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced a controller failure red alarm. there was no reported patient involvement. | [5 empty] | The device was returned and evaluated, and the investigation for the reported controller error (yellow alarm) has been completed. console logs from the reported day of event are consistent with complaint and show alarm #1035 (opm communication loss) upon boot-up. after arriving in the lab, the issue was reproduced. during troubleshooting it's been determined that optical bench was drawing higher than normal current that from the pbm, which - in order to maintain constant power - was lowering voltage supplied to optical bench, which wasn't enough for the optical bench to operate (therefore communication stopped). the cause of the controller alarm upon bootup was loss of communication to optical bench, resulting from high current draw, due to defective ccd assembly on optical bench this report is being filed as part of a retrospective review of historical records.
16114 | 1220648-2024-21656 | 20578674 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20578674&pc=OZD | 10/10/2024 | 12/11/2024 |  | 10/31/2024 |  |  | 10/10/2024 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532988 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. the surgeon crossed aortic valve with the first impella, and went to advance further in the left ventricle but felt as though it had bent and saw the waveforms dampen. the decision was made to replace the impella. the patient survived the explant. | [5 empty] | The investigation into product damage (cannula kink) has been completed. the pump was returned for investigation. total pump run time was 8 minutes. cannula kink was noted near the outlet cage. no other physical abnormalities were observed on the returned product. pump was ran as expected during test in backet of water. data log shows the suction alarm and low pump flow after start of the case run without observed any trend of biomaterial interaction trend or positioning issue. as per the clinical, bent was felt during pump placement. the cause of the low pump flow issue was most likely kinked cannula. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-21656. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
16115 | 1220648-2024-21658 | 20579309 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20579309&pc=OZD | 10/10/2024 | 12/09/2024 |  | 10/31/2024 |  |  | 10/10/2024 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025503002 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella 5.5 support experienced ventricular tachycardia requiring cardioverting. additionally, the left arm, which is the impella side, has swelling. there is a lot of drainage from the impella 5.5 access site. it was managed by redressing the site. the patient remains on impella support. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email.
16116 | 1220648-2024-21657 | 20579367 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20579367&pc=OZD | 10/08/2024 | 12/10/2024 |  | 10/31/2024 |  |  | 10/08/2024 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025481042 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an elderly patient who had the impella cp placed via axillary after the 5.5 failed delivery. it was reported that the patient had mottled arm below elbow and believed to have thrombus, no pulses. the doctor elected to start dobutamine and case was aborted, the patient was taken to intensive care unit where she continued to decline, cardiopulmonary resuscitation initiated but ultimately the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation into limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-21657. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
16117 | 1220648-2024-21661 | 20582282 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20582282&pc=OZD | 02/06/2023 | 10/31/2024 |  | 10/31/2024 |  |  | 02/06/2023 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1477997 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella pump for mechanical circulatory support. while on support, the automated impella controller (aic) exhibited a controller error, the placement signal and left ventricle waveform disappeared from the screen. motor current and flows still were normal. no patient harm reported. | [5 empty] | The investigation has been completed. the console was sent back for further investigation. the root cause of the issue was not determined since failure could not be reproduced. this report is being filed as part of a retrospective review of historical records.
16118 | 1220648-2024-21664 | 20582305 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20582305&pc=OZD | 04/27/2022 | 10/31/2024 |  | 10/31/2024 |  |  | 04/27/2022 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | AUTOMATED IMPELLA CONTROLLER (AIC) |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1502178 |  | 0042-0010-UK |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Use of Device Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced a purge issue. no clinical details regarding resolution or patient involvement / condition were provided. | [5 empty] | The investigation has been completed. the console (b)(6) was sent back for further investigation. the console was tested in collaboration with power battery manager. the failure mode was reproduced. without a pressure disc connected, the console pressure sensor was reading around 1700 mm hg. replacing the pressure sensor resolved the issue. the root cause of the purge system priming issues was due to a defective purge pressure sensor. this report is being filed as part of a retrospective review of historical records.
16119 | 1220648-2024-21666 | 20582321 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20582321&pc=OZD | 01/22/2023 | 10/31/2024 |  | 10/31/2024 |  |  | 01/22/2023 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1129858 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  |  | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that their automated impella controller (aic) shutdown unexpectedly during impella support. the aic was subsequently removed and replaced with another console to resolve the reported failure. there was no patient harm resulting from the noted issue. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the aic issue was most likely a defective pbm board. this report is being filed as part of a retrospective review of historical records.
16120 | 1220648-2024-21668 | 20582344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20582344&pc=OZD | 01/18/2023 | 10/31/2024 |  | 10/31/2024 |  |  | 01/18/2023 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1332033 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Audible Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) exhibited no alarm sounds when the plug cable/suction alarms. the aic was replaced with another console and there were no further issues. no report of patient harm or injury. | [5 empty] | The investigation has been completed. the console was sent back for further investigation. the root cause of the aic issue was a loose connector that was unplugged. this report is being filed as part of a retrospective review of historical records.
16121 | 1220648-2024-21670 | 20582353 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20582353&pc=OZD | 09/12/2022 | 10/31/2024 |  | 10/31/2024 |  |  | 09/12/2022 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | Staff reported that when attempting to use the automated impella controller for educational purposes, it was found that the purge disc retainer was broken on the console. a backup controller was used. there was no patient involvement. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
16122 | 1220648-2024-21659 | 20582356 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20582356&pc=OZD | 03/01/2023 | 10/31/2024 |  | 10/31/2024 |  |  | 03/01/2023 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | XXXXXX |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that they had an automated impella controller (aic) freeze up on the startup screen when booting up. it was frozen for approximately 30-60 seconds before clearing and going to start screen. the console was replaced to resolve the issue. no patient impact or involvement. | [5 empty] | E1: the name of the initial reporter is unknown. the investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. console was continuously power cycled and then ran all weekend but was unable to reproduce the failure mode. the cause of the power on issues was due to failure of system check upon bootup because of timing for the initialization of the impellatronic. the console's soft-reboot mechanism had recovered the bad startup per design. the root cause of the initialization timing issue was not determined due to the inability to reproduce. this report is being filed as part of a retrospective review of historical records.
16123 | 1220648-2024-21660 | 20582361 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20582361&pc=OZD | 01/29/2023 | 10/31/2024 |  | 10/31/2024 |  |  | 01/29/2023 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1076069 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) exhibited a impella catheter defective alarm. the customer tried to reboot the device; however, this did not resolve the issue. no report of patient harm or injury. | [5 empty] | The investigation has been completed. the console was sent back for further investigation. the root cause of the aic issue was most likely a defective impellatronics board that prevented the eeprom from being read. this report is being filed as part of a retrospective review of historical records.
16124 | 1220648-2024-21667 | 20582378 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20582378&pc=OZD | 01/05/2023 | 10/31/2024 |  | 10/31/2024 |  |  | 01/05/2023 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1253749 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The user facility's biomedical engineer reported the during servicing the automated impella controller (aic), the ground cable did not fit into the strain relief. no patient involvement. | [5 empty] | The investigation has been completed. the console (ic4565) was sent back for further investigation. the root cause of the issue was an improperly sized cable. this report is being filed as part of a retrospective review of historical records.
16125 | 1220648-2024-21669 | 20582390 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20582390&pc=OZD | 01/11/2023 | 10/31/2024 |  | 10/31/2024 |  |  | 01/11/2023 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1100031 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the aic's purge clip was broken. a loaner device was sent to the customer. no report of patient harm or injury. | [5 empty] | The investigation into the purge disc/flag issues issue has been completed. the automated impella controller (aic) was returned for investigation. the device logs confirmed that there were multiple instances of "purge disc not detected alarm issued". upon visual inspection, the purge disc flag is missing/broken. this report is being filed as part of a retrospective review of historical records.
16126 | 1220648-2024-21671 | 20582393 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20582393&pc=OZD | 06/18/2023 | 10/31/2024 |  | 10/31/2024 |  |  | 06/18/2023 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1547850 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) experienced a controller error yellow alarm that self-resolved. no patient harm was reported. | [5 empty] | The investigation has been completed. the console (ic8510) was sent back for further investigation. this known pattern failure, which is characterized by a temporary loss of optical data (placement, snr, contrast, lamp, gain) and triggering of a controller error alarm, has a low probability of reoccurrence and was unable to be reproduced. the root cause of the issue was not determined since failure could not be reproduced. this report is being filed as part of a retrospective review of historical records.
16127 | 1220648-2024-21662 | 20582425 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20582425&pc=OZD | 02/06/2023 | 10/31/2024 |  | 10/31/2024 |  |  | 02/06/2023 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1502180 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) could not be switched on in battery mode. after plugging the device in the ac a burning smell was noticed. this occurred during preventative maintenance, no patient involvement. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. interior inspection of the console revealed that a capacitor c199 was damaged/ burned. also, both batteries were almost depleted according to the empty led bar on both batteries. after temporarily replacement of pbm and battery set, console was able to receive charging and operate without issue. given the sequence that console was not power up on battery mode (battery depleted already), followed by burning smell (capacitor failed) as ac power connected. the failure mode of console unable to power on was due to the depleted battery. the failure mode of pbm subcomponent failure was due to a defective pbm. this report is being filed as part of a retrospective review of historical records.
16128 | 1220648-2024-21663 | 20582433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20582433&pc=OZD | 01/27/2023 | 10/31/2024 |  | 10/31/2024 |  |  | 01/27/2023 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1465594 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Female |  |  |  | Malfunction | No | Yes | Defective Device | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella pump for mechanical circulatory support. it was reported that a controller error alarm (yellow) occurred on the automated impella controller (aic). the ao position waveform appeared and disappeared repeatedly, so the aic was replaced with another aic without any additional reported issues. no reported harm or injury to patient. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the issue was not determined since issue could not be reproduced. this report is being filed as part of a retrospective review of historical records.
16129 | 1220648-2024-21665 | 20582464 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20582464&pc=OZD | 01/12/2023 | 10/31/2024 |  | 10/31/2024 |  |  | 01/12/2023 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | ABIOMED, INC. |  |  | US |  | 1200541 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) showed a "system self-check failure" during startup. there was no reported patient involvement. | [5 empty] | The device was returned and evaluated, and the investigation for the reported system self-check error has been completed. aic logs confirmed the reported failure. instances of initialization failures for the 1st pbm serial access unit. no corrosion on pbm was identified. pbm replacement resolved subject failure mode. the root cause of the system self-check failure was due to defective pbm. this report is being filed as part of a retrospective review of historical records.
16130 | 1220648-2024-21672 | 20582564 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20582564&pc=OZD | 04/03/2023 | 10/31/2024 |  | 10/31/2024 |  |  | 04/03/2023 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | AUTOMATED IMPELLA CONTROLLER |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1673731 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced a controller error yellow alarm which was resolved through backup console. no patient was involved. | [5 empty] | The investigation has been completed. the console (ic9892) was sent back for further investigation. the root cause of the controller error alarm was due to a defective blue receptacle. this report is being filed as part of a retrospective review of historical records.
16131 | 1220648-2024-21674 | 20582654 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20582654&pc=OZD | 03/28/2023 | 10/31/2024 |  | 10/31/2024 |  |  | 03/28/2023 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | AUTOMATED IMPELLA CONTROLLER |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1287974 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) only recognized the impella pump after several attempts to connect it and gave the hint to check the light channel for contamination. no details are know regarding resolution, but the examinations could be completed successfully with the same aic. no patient was involved. | [5 empty] | The investigation has been completed. the console (b)(6) was sent back for further investigation. the imc logs are consistent with the reported issue. pump was inserted during case start and was detected successfully. however, the case was unable to start as an optical sensor system error during case start was observed. inspection on the optical fiber did not reveal any abnormality, 5s parameters and blub power are all within spec. console passed the case start wizard with known-good pump and purge. the root cause of the case start issue, optical system error, was due to user not fully inserted the pump into blue receptacle. this report is being filed as part of a retrospective review of historical records.
16132 | 1220648-2024-21675 | 20582656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20582656&pc=OZD | 05/23/2023 | 10/31/2024 |  | 10/31/2024 |  |  | 05/23/2023 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | OPTICAL AIC W/IMPELLA CONNECT |  |  | ABIOMED, INC. |  |  | US |  | 1680221 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant in japan had an aic replaced for a syscal error. there was no noted harm or injury from the aic (automated impella controller). | [5 empty] | The investigation for the reported aic - system self-check error has been completed. the root cause of the boot up issue, system self-check failure was due to a defective impellatronic board. clinical review: imc logs are consistent with the complaint. logs from the date of the complaint (may. 23rd 2023) revealed distinct communication issues with the electronic pressure measurement (epm) circuitry via serial port sau_ser6. 23/05/23 18:05:27 sau_ser6 cmd seq.101 timeout. reply err and sem_post. device analysis: upon start-up in abj lab, the console would issue system self check failed consistently. firmware check was performed, thus was able to confirm the eep communication issue with ¿n.a¿ in sau_sens_eep_1 hw revision and version. after swapping with a known-good impellatronic board, the console functioned as expected. the root cause of the boot up issue, system self-check failure was due to defective impellatronic board. before this console was returned to the customer, fs was instructed to, replaced the impellatronic pca (0042-1115), perform section 13-16 per preventive maintenance (0042-7309), perform a full functional check on the console and optical system (0042-7316). no corrective action is recommended at this time because the failure was determined to have a risk as low as practical (alap). failures of this type will be monitored for trends. this report is being filed as part of a retrospective review of historical records.
16133 | 1220648-2024-21673 | 20582681 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20582681&pc=OZD | 06/05/2023 | 10/31/2024 |  | 10/31/2024 |  |  | 06/05/2023 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1383042 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported automated impella controller (aic) screen was stuck on the abiomed screen during routine start up. it was reported that this was not associated with a case or patient encounter. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the root cause of the reported failure mode was unable to determine due to the failure mode was not reproduced. this report is being filed as part of a retrospective review of historical records.
16134 | 1220648-2024-21678 | 20582732 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20582732&pc=OZD | 03/06/2023 | 10/31/2024 |  | 10/31/2024 |  |  | 03/06/2023 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | AUTOMATED IMPELLA CONTROLLER |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1336911 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced notifications for ¿system failure check¿ appeared. powered back off and restarted aic, received same notification. confirmed battery transport switch on bottom of console was in the ¿on¿ position. no patient was involved. | [5 empty] | The investigation has been completed. the console ( (b)(6) ) was sent back for further investigation. the console reproduced the "system self check failure" in the lab. upon analyzing the pbm, there was residue around the super cap onto the communication line on the pcb which is indicative of the failure mode super cap corrosion the root cause of the aic issue was contamination of a communication line on the power battery manager. this report is being filed as part of a retrospective review of historical records.
16135 | 1220648-2024-21676 | 20582760 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20582760&pc=OZD | 05/15/2023 | 10/31/2024 |  | 10/31/2024 |  |  | 05/15/2023 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1122737 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that their automated impella controller (aic) failed to power on to the case start screen during lab education. coinciding with the power up failure, the aic was experiencing a reoccurring self-check failure. the console was subsequently removed from service. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the aic inability to boot up was a defective 4in1 board. this report is being filed as part of a retrospective review of historical records.
16136 | 1220648-2024-21681 | 20582780 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20582780&pc=OZD | 01/06/2023 | 10/31/2024 |  | 10/31/2024 |  |  | 01/06/2023 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1352449 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) had no sound during an alarm. no patient harm was reported. | [5 empty] | The investigation has been completed. the console was sent back for further investigation. the speaker cable was confirmed to be broken upon opening the console in field service. the root cause of the inaudible alarm was due to a broken speaker cable, which was likely due to a technician servicing error. this report is being filed as part of a retrospective review of historical records.
16137 | 1220648-2024-21677 | 20582805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20582805&pc=OZD | 02/22/2023 | 10/31/2024 |  | 10/31/2024 |  |  | 02/22/2023 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | AUTOMATED IMPELLA CONTROLLER |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1080190 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support, the automated impella controller (aic) experienced battery charging/other issues (battery temperature high alarm) that were "resolved quickly". no clinical details regarding resolution were provided. no patient harm reported. | [5 empty] | The investigation has been completed. the console (ic1105) was sent back for further investigation. the root cause is most likely due to power battery manager incorrectly reporting battery temperatures. this report is being filed as part of a retrospective review of historical records.
16138 | 1220648-2024-21679 | 20582819 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20582819&pc=OZD | 05/03/2023 | 10/31/2024 |  | 10/31/2024 |  |  | 05/03/2023 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1691800 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient with unknown race and ethnicity was implanted with impella device. during the start procedure when the impella device was plugged in to the automated impella controller (aic), there was an acoustic alarm and a black screen. it stayed like this. the aic and case went without complications. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the root cause of the issue was corrupt compact flash cards. this report is being filed as part of a retrospective review of historical records.
16139 | 1220648-2024-21680 | 20582877 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20582877&pc=OZD | 04/07/2023 | 10/31/2024 |  | 10/31/2024 |  |  | 04/07/2023 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1194348 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) was persistently showing a system self-check error. the console was removed from service, and there was no reported patient involvement. | [5 empty] | The device was returned and evaluated, and the investigation for the reported system self-check error has been completed. imc logs from the complaint date showed numerous powermod 1 communication issues which resulted in failure of system self-check. the pbm was confirmed to have corrosion on the communication traces next to the super caps. the root cause of syscall error was due to a defective pbm. the root cause of the defective pbm was corrosion due to leaking super capacitors. this report is being filed as part of a retrospective review of historical records.
16140 | 1220648-2024-21682 | 20582889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20582889&pc=OZD | 01/12/2023 | 10/31/2024 |  | 10/31/2024 |  |  | 01/12/2023 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1431282 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on an automated impella controller (aic) for mechanical circulatory support. while on support, the aic experienced a purge disc/flag issue that was resolved by switching to another console. no report of patient harm or injury. | [5 empty] | The investigation into the purge disc/flag issues issue has been completed. the automated impella controller (aic) was returned for investigation. when attempting to recreate the failure, the purge disc was unable to make complete contact resulting in a gap that caused the error of purge disc to validate sensor accuracy and perform a full functional check on the console and optical system. this report is being filed as part of a retrospective review of historical records.
16141 | 1220648-2024-21683 | 20582961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20582961&pc=OZD | 01/11/2023 | 10/31/2024 |  | 10/31/2024 |  |  | 01/11/2023 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1240266 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a battery failure red alarm that was resolved by switching to a backup console. no clinical details regarding patient involvement/condition were provided. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the aic issue was a defective pbm board. this report is being filed as part of a retrospective review of historical records.
16142 | 1220648-2024-21684 | 20582968 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20582968&pc=OZD | 01/11/2023 | 10/31/2024 |  | 10/31/2024 |  |  | 01/11/2023 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1563183 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a battery failure red alarm with no resolution specified. there was no patient involvement. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the failure mode was reproduced on an fs lab bench. the battery failure alarm was due to a defective battery. batt test command revealed cell imbalance with battery 1 cell 3. the root cause of the defective battery was due to single cell failure, a known pattern failure. this report is being filed as part of a retrospective review of historical records.
16143 | 1220648-2024-21685 | 20583285 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20583285&pc=OZD | 01/10/2023 | 10/31/2024 |  | 10/31/2024 |  |  | 01/10/2023 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1587787 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. the automated impella controller (aic) experienced a controller error alarm and was replaced. there was no patient harm reported. | [5 empty] | The investigation into the controller error issue has been completed. the impella automated controller was returned for investigation. the data analysis of the logs from the reported day of event are consistent with the complaint and show several events of ¿placement signal not reliable¿ and repeated alarms ¿controller error¿. this is a low probability event and lasts only a few minutes. if needed, patient hemodynamics and impella position can be monitored with imaging. the cause of the issue was unable to be determined since the controller error issue could not be reproduced. this report is being filed as part of a retrospective review of historical records.
16144 | 1220648-2024-21687 | 20583355 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20583355&pc=OZD | 01/10/2023 | 10/31/2024 |  | 10/31/2024 |  |  | 01/10/2023 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1199620 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | Insufficient Information | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the customer was placed onto impella cp support due to acute myocardial infarction / cardiogenic shock. while on support, the automated impella controller (aic) experienced a purge disc/flag issue that was resolved by changing the console. | [5 empty] | The investigation into the purge disc/flag issues issue has been completed. the automated impella controller (aic) was returned for investigation. data review: imc logs from the complaint date revealed that the console issued the purge disc not detected alarm several times. device analysis: the console was tested after arriving in field service (fs). the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken (missing at blue disc retainer). before returning this console back to the customer, fs was instructed to replace the purge flag, validate sensor accuracy via section 12 of the preventive maintenance procedure), and perform a full functional check on the console and optical system. this report is being filed as part of a retrospective review of historical records.
16145 | 1220648-2024-21686 | 20583391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20583391&pc=OZD | 01/10/2023 | 10/31/2024 |  | 10/31/2024 |  |  | 01/10/2023 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1141483 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | Insufficient Information | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the aic experienced a purge disc/flag issue. no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | The investigation into the purge disc/flag issues issue has been completed. the automated impella controller (aic) was returned for investigation. imc logs from the complaint date revealed that the console issued the purge disc not detected alarm several times. device analysis: the console was tested after arriving in field service (fs). the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken (missing at blue disc retainer). before returning this console back to the customer, fs was instructed to replace the purge flag, validate sensor accuracy via section 12 of the preventive maintenance procedure, and perform a full functional check on the console and optical system. this report is being filed as part of a retrospective review of historical records.
16146 | 1220648-2024-21688 | 20583498 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20583498&pc=OZD | 01/10/2023 | 10/31/2024 |  | 10/31/2024 |  |  | 01/10/2023 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1615855 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a battery failure red alarm. no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the failure mode was reproduced on an engineering lab bench. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure, a known pattern failure. this report is being filed as part of a retrospective review of historical records.
16147 | 1220648-2024-21689 | 20583602 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20583602&pc=OZD | 01/10/2023 | 10/31/2024 |  | 10/31/2024 |  |  | 01/10/2023 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1587053 |  | 0042-0040 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported patient was placed on impella 5.5 for hemodynamic support. while on support a placement signal drop and return was noted. unit swapped consoles as they were worried about failure. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the issue was not determined since issue could not be reproduced. this report is being filed as part of a retrospective review of historical records.
16148 | 1220648-2024-21690 | 20583623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20583623&pc=OZD | 01/09/2023 | 10/31/2024 |  | 10/31/2024 |  |  | 01/09/2023 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1129226 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | Insufficient Information | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the aic experienced purge disc/flag issues. no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | The investigation into the purge disc/flag issues issue has been completed. the automated impella controller (aic) was returned for investigation. aic logs revealed that the console could not detect the purge disc after the cassette was connected. there were no issues detecting the purge cassette. device analysis: the console was tested after arriving in field service (fs). the purge disc flag was confirmed to be broken. before returning this console back to the customer, fs was instructed to replace the purge disc retainer (due to hairline crack) /flag, validate sensor accuracy via section 12 of the preventive maintenance procedure, and perform a full functional check on the console and optical system. this report is being filed as part of a retrospective review of historical records.
16149 | 1220648-2024-21692 | 20583929 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20583929&pc=OZD | 12/31/2022 | 10/31/2024 |  | 10/31/2024 |  |  | 12/31/2022 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1494422 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Obstruction of Flow | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced multiple purge system alarms throughout the day. it was reported that unsuccessful troubleshooting steps were taken, including purge cassette change and tissue plasminogen activator added to the purge fluid. this aic was replaced with another aic and the procedure was successfully completed. there was no patient harm reported. | [5 empty] | The investigation into the purge issue has been completed. the impella automated controller was returned for investigation. the data analysis of the logs indicated occurring on 12/30/22, a sudden significant rise in purge pressure occurred. a change from a stable ~500 mmhg rose to over ~1000 mmhg. this triggered a persistent purge pressure high alarm. additionally, a persistent purge system blocked alarm was triggering on and off for the duration of the case on ic8192. the staff attempted to de-air the purge line and swapped out the purge cassette for a new one, however these troubleshooting methods were unsuccessful in clearing the persistent alarms that were triggering. after unsuccessful troubleshooting attempts, the staff performed an aic transfer to ic2041. the logs from constellation were reviewed for ic2041. it was observed that the staff used the same purge cassette (sn:(b)(6) when the transfer was done. it was also observed that there were no purge related alarms after swapping consoles. a physical inspection of the returned purge assembly was performed including the purge assembly as whole, purge disc retainer and flag, purge pressure sensor and communication ribbon cable. no issues were found with most of these items expect for the purge disc retainer. a small fracture was found on the left-hand side of the purge disc retainer. the screw posts of the purge disc retainer were inspected as well, but no further damage was observed. the returned aic ic8192 was set up on a pump and purge system in the burn-in room for ~12 hours. no issues occurred. ic8192 was set up on a similar system on an engineering lab bench. the console was run for an extended period of time while purge parameters were closely watched through telnet. no issues occurred. the ic8192 was able to successfully de-air and perform bolus and purge cassette changes during pump operation without issue. purge pressure performance was stable through all points of testing. a root cause for the persistent purge pressure high and purge system blocked alarms could not be determined due to the inability to reproduce the failures. this report is being filed as part of a retrospective review of historical records.
16150 | 1220648-2024-21691 | 20583974 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20583974&pc=OZD | 12/28/2022 | 10/31/2024 |  | 10/31/2024 |  |  | 12/28/2022 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1627246 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a battery failure red alarm. no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the battery failure alarm was due to a defective battery 1 (decline of individual cell voltage). the root cause of the battery cell failure was undetermined. this report is being filed as part of a retrospective review of historical records.
16151 | 1220648-2024-21693 | 20584036 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20584036&pc=OZD | 12/29/2022 | 10/31/2024 |  | 10/31/2024 |  |  | 12/29/2022 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1311567 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) experienced optical sensor error upon connection of the pump, which was resolved through replacing the device with a new console. no patient harm reported. | [5 empty] | The investigation has been completed. the console (ic5498) was sent back for further investigation. the root cause of the aic issue was a defective blue receptacle. this report is being filed as part of a retrospective review of historical records.
16152 | 1220648-2024-21694 | 20584083 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20584083&pc=OZD | 12/29/2022 | 10/31/2024 |  | 10/31/2024 |  |  | 12/29/2022 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | XXXXXXX |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced battery charging / other issues. no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. during testing in (b)(6), the issue was reproduced, and console showed ¿controller error¿ alarm. inspection of the console¿s components revealed misaligned communication cable between optical bench and 4-in-1 carrier. after the cable was re-seated, the issue was resolved and console was able to boot-up correctly. it is most likely that the cable was accidentally misaligned when the console¿s housing was being closed-up after the pm (5s parameters obtained during pm were within specification). the cause of the aic issue was a loose connector. this report is being filed as part of a retrospective review of historical records.
16153 | 1220648-2024-21695 | 20584448 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20584448&pc=OZD | 12/28/2022 | 10/31/2024 |  | 10/31/2024 |  |  | 12/28/2022 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1152332 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported patient was placed on impella cp with smart assist for hemodynamic support. while on support, the console used for the implant was a loaner and the patient was being transferred so the catheter was transferred to a different automated impella controller (aic). immediately the placement signal was identified as being unreliable. the catheter was then transferred back to the loaner console and the placement signal was fine. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the aic (automated impella controller) issue was a defective optical bench. this report is being filed as part of a retrospective review of historical records.
16154 | 1220648-2024-21696 | 20584799 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20584799&pc=OZD | 02/22/2023 | 10/31/2024 |  | 10/31/2024 |  |  | 02/22/2023 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1215712 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient came in as a segment elevation myocardial infarction (stemi) with severe hypotension, end-diastollic pressure (edp) 40. the decision was made to place the impella. when white cable was connected to the automated impella controller (aic), the aic was alarming "defective impella catheter: disconnect and replace impella catheter." the staff attempted multiple times to disconnect and reconnect white cable, and then top of aic yellow banner said "impella defective: do not use impella, replace impella." due to multiple attempts to re-plug white cable to clear alarm, critical nature of patient condition, and yellow banner saying, "do not use impella," decision made to open new impella. new impella connected, primed, and inserted without issue. the procedure outcome was unsuccessful, however, there was no reported harm or injury to the patient. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the console was tested with a number of service pumps as well as pump simulator, but the issue was not reproduced. the issue of rfid errors did not reoccur during normal operation of the console but could be reproduced by not installing purge cassette correctly. however, a defective purge cassette would also explain the issue, which could not be eliminated as a possible root cause due to the fact that the purge cassette was not returned for analysis. inspection of the console also revealed damage to the purge assembly component (purge disk detect flag and retainer), and their contamination with purge fluid residue. during functional testing of the console, it¿s been revealed that the pbm assembly was defective. the root cause is unable to be determined. this report is being filed as part of a retrospective review of historical records.
16155 | 1220648-2024-21697 | 20584801 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20584801&pc=OZD | 12/16/2022 | 10/31/2024 |  | 10/31/2024 |  |  | 12/16/2022 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1432278 |  | 0042-0010 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) showed an "emergency shutdown imminent" error message during a power charge test. however, after a couple of minutes of the alarm, the shutdown did not happen. there was no reported patient involvement. | [5 empty] | The device was returned and evaluated, and the investigation for the reported aic - battery charging/other issues has been completed. console logs from the reported day of event are consistent with complaint and show one instance where alarm #50 (¿emergency shutdown imminent¿) was not followed by controller shutdown. on another occasion, the console performed shutdown reportedly without the operator pressing the power button. in the course of troubleshooting the power button was replaced but issue was not resolved. after the pbm was replaced, however, the issue was resolved, proving that pbm assembly was defective. this report is being filed as part of a retrospective review of historical records.
16156 | 1220648-2024-21698 | 20584824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20584824&pc=OZD | 12/20/2022 | 10/31/2024 |  | 10/31/2024 |  |  | 12/20/2022 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1362195 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) showed a battery failure alarm when the console was powered on to check the date/time. there was no reported patient involvement. | [5 empty] | The device was returned and evaluated, and the investigation for the reported aic - battery failure (red alarm) has been completed. data logs confirmed battery failure alarms. subsequent boots following the complaint date also triggered these alarms. lt log analysis of the battery data reveals that beginning before the complaint date of occurrence, cell 4 voltage of battery 2 had been declining for an extended period of time. console interior connections were checked and verified. when visually inspecting the batteries and pbm, it was observed that one of the batteries status leds were completely blank. the battery failure alarm was due to a defective battery 2 (decline of individual cell voltage). the root cause of the battery cell failure was undetermined. this report is being filed as part of a retrospective review of historical records.
16157 | 1220648-2024-21699 | 20584830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20584830&pc=OZD | 12/13/2022 | 10/31/2024 |  | 10/31/2024 |  |  | 12/13/2022 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1184610 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) showed a battery failure red alarm. there was no reported patient involvement. | [5 empty] | The device was returned and evaluated, and the investigation for the reported aic - battery failure (red alarm) has been completed. aic logs confirmed the reported failure. there was a battery failure alarm and a battery level low alarm due to low pack voltage. the failure mode was reproduced on an engineering bench. the battery 1 lcd indicator was blank (fully discharged). batttest showed a missing ¿fu¿ for battery 1, indicating either battery is disengaged via switch, or it is not providing any voltage. see the attached batttest screenshot. the battery failure was due to a defective battery 1. the root cause of the defective battery 1 was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
16158 | 1220648-2024-21700 | 20584864 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20584864&pc=OZD | 12/03/2022 | 10/31/2024 |  | 10/31/2024 |  |  | 12/03/2022 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER JP |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER JP |  |  | ABIOMED, INC. |  |  | US |  | 1404666 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 30 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Use of Device Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was placed on an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) experienced a purge issue that was resolved by changing out the aic and using the same purge cassette. this aic was replaced with another aic and the procedure was successfully completed. there was no patient harm reported. | [5 empty] | The investigation into the purge issue/purge flag issue has been completed. the impella automated controller was returned for investigation. the data analysis of the logs from the complaint date revealed that the console issued the purge disc not detected alarm (event set #(b)(4)) several times. the returned console was tested. the issue with not properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be missing. the cause of the issue was a broken/missing purge flag. this report is being filed as part of a retrospective review of historical records.
16159 | 1220648-2024-21701 | 20584901 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20584901&pc=OZD | 12/15/2022 | 10/31/2024 |  | 10/31/2024 |  |  | 12/15/2022 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1207018 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Apparent Adverse Event | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that, upon transfer of the device down the hall, it fell onto the floor. an evaluation was requested to rule out any issues with the device. no patient involvement. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the console was inspected on an engineering bench. there was minimal damage to the console. the pressure sensor membrane foil had a dent and hairline crack in the retainer which was likely unrelated to the reported event. no other damage was observed (internal and external). this report is being filed as part of a retrospective review of historical records.
16160 | 1220648-2024-21702 | 20584908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20584908&pc=OZD | 12/13/2022 | 10/31/2024 |  | 10/31/2024 |  |  | 12/13/2022 | 10/31/2024 |  |  |  |  | 10/31/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1540253 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella 5.5 for mechanical circulatory support. the complainant reported that the console shut down for one minute during impella support. there was no patient harm or injury reported. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the console restart was a crash of the graphic component of console's software. this report is being filed as part of a retrospective review of historical records.
16161 | 1220648-2024-21475 | 20564062 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20564062&pc=OZD | 06/26/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 06/26/2023 | 10/30/2024 |  |  | 01/31/2025 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024310484 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 06/26/2023 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. impella cp was showing low flow shortly after implant. clot was suspected despite multiple attempts to reposition. upon removal, clot was noted in the inlet and outlet. thrombus was not noted in lv during lv gram. pt has history of hypercoagulability. impella cp successfully replaced for a new cp. the patient survived the procedure. | [5 empty] | The investigation into the reported issue has been completed. the returned pump was visually inspected. biomaterial observed in inlet cage and around impella too. no other abnormality found. root cause of the low pump flow was biomaterial due to patient condition related. this report is being filed as part of a retrospective review of historical records.
16162 | 1220648-2024-21476 | 20564067 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20564067&pc=OZD | 07/06/2022 | 10/30/2024 |  | 10/30/2024 |  |  | 07/06/2022 | 10/30/2024 |  |  | 11/30/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022157034 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Female |  |  |  | Malfunction | No | Yes | Difficult to Insert | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [37 empty] | A |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. there was difficulty inserting the impella cp pump through the introducer. the staff was unable to insert the impella which was cracked due to using force to insert. the pump was replaced. additional information noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
16163 | 1220648-2024-21477 | 20564069 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20564069&pc=OZD | 09/26/2022 | 10/30/2024 |  | 10/30/2024 |  |  | 09/26/2022 | 10/30/2024 |  |  | 03/31/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202337 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the complainant reported fecal matter found on the impella purge side arm. the pressure reservoir part of the impella purge was contaminated with feces. the pump was escalated and replaced. additional information noted care was withdrawn and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation into the reported issue has been completed. a visual inspection of the pump revealed fecal matter on the pressure reservoir. the cause of the product damage was use related. this report is being filed as part of a retrospective review of historical records.
16164 | 1220648-2024-21478 | 20564163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20564163&pc=OZD | 06/05/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 06/05/2023 | 10/30/2024 |  |  | 01/31/2025 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | GM |  | 2024332273 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Cardiac Arrest | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support for percutaneous coronary intervention. the delivery of the cp failed due to native anatomy and could not be placed. the patient expired during an emergent intubation and cardiopulmonary resuscitation was performed during the percutaneous coronary intervention. no allegations were made toward the impella as the cause of death. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine the root cause of the delivery issue was patient condition related to patient anatomy. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
16165 | 1220648-2024-21479 | 20565605 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20565605&pc=OZD | 10/03/2024 | 12/06/2024 |  | 10/30/2024 |  |  | 10/03/2024 | 10/30/2024 |  |  |  |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025505860 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 pump placed via the axillary to support a 59-year-old male patient admitted with cardiomyopathy and other systemic comorbidities and awaiting a heart transplant. the axillary site developed arm pain and was washed out in the or for the hematoma. the pump remained on despite this and the patient continues to await a transplant. | [5 empty] | The investigation into access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-21479. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16166 | 1220648-2024-21480 | 20565682 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20565682&pc=OZD | 10/03/2024 | 12/02/2024 |  | 10/30/2024 |  |  | 10/03/2024 | 10/30/2024 |  |  |  |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025567472 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support. while on support, the patient had ventricular tachycardia (vt) which required cardio-pulmonary resuscitation (cpr). impella was repositioned as thought too deep by the team. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation into vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
16167 | 1220648-2024-21481 | 20565965 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20565965&pc=OZD | 10/05/2024 | 08/22/2025 |  | 10/30/2024 |  |  | 10/05/2024 | 10/30/2024 |  |  |  |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025521676 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump placed to support a 79-year-old male patient admitted in an acute myocardial infarction/cardiogenic shock. the pump was placed but when in the intensive care unit, the pump came out of position and was therefore explanted and replaced. the pump support was for three days. the patient survived the explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The impella device was not received from the customer. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures. h.6. investigation conclusions revised on manufacturer device report in accordance with updated procedures. h.10. additional manufacture narrative revised on manufacturer device report in accordance with updated procedures.
16168 | 1220648-2024-21482 | 20566062 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20566062&pc=OZD | 10/06/2024 | 11/21/2024 |  | 10/30/2024 |  |  | 10/06/2024 | 10/30/2024 |  |  |  |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025551806 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  | White | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during an impella cp procedure access micro puncture was obtained and went in with perclose, but the perclose footplate got caught on something intravascular. it was hard to remove, but when it was removed, the tissue was stuck to the perclose. sheaths 5, 6, and 7 french were placed but bleeding continued around sheaths. an 11 french sheath was placed. bleeding was resolved but the contrast was injected and there was no flow down the leg. the impella sheath was placed and wired into the left ventricle but per vascular surgeon, the impella was removed. 14 french was pulled over the wire, and three perclose were inserted into the right femoral artery (rfa) and the rfa access site was closed. the patient pulses were dopplerable. | [5 empty] | The investigation for the impella cp has been completed. the device was not returned for evaluation. however, clinical details provided were sufficient to determine the most likely root cause. as the necessary clinical information was not provided, the root cause of the limb ischemia was not determined. the root cause of access site bleeding is determined to be use issue because the perclose failed and caught something intravascular causing the bleed. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16169 | 1220648-2024-21483 | 20566112 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20566112&pc=OZD | 10/04/2024 | 11/27/2024 |  | 10/30/2024 |  |  | 10/04/2024 | 10/30/2024 |  |  |  |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025562798 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had placed a impella cp to support a patient admitted in post partum with ventricular tachycardia arrest and ejection fraction of 5-10% in cardiogenic shock. the cp was placed femorally and the patient was transported to the tertiary center. the patient was seen to have signs of hemolysis with elevated pfhg. the limbs were sent to be clotting and pulses were lost. the pump was explanted and replaced by axillary 5.5 pump, and the surgeons performed a cutdown and limb fasciotomy. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation into limb ischemia and hemolysis has been completed. the impella device was not returned for evaluation. the root cause of hemolysis was most likely biomaterial ingestion. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. the following have been reported incorrectly on manufacturer device report 1220648-2024-21483. b3 date of the event. e4 should have been left blank. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
16170 | 1220648-2024-21484 | 20566126 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20566126&pc=OZD | 10/05/2024 | 12/18/2024 |  | 10/30/2024 |  |  | 10/05/2024 | 10/30/2024 |  |  |  |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025535080 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Difficult to Remove | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and the patient explanted the pump themselves due to transitional-syndrome. the patient was doing well after the explant. no blood products were required and no reported patient harm. | [5 empty] | The investigation for the positioning issues has been completed. no product was returned. log review showed a pump stop 115 alarm about 30 hours into support. it was not a true pump stop alarm as it was seen with impella disconnect event. there was also "impella position in ventricle" alarms shortly before the pump stop. the fm "removal issues" was updated to "positioning issues" because the patient explanted the pump, which ended support. the root cause of the positioning issue was use-related. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
16171 | 1220648-2024-21485 | 20566186 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20566186&pc=PYX | 02/20/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 02/20/2023 | 10/30/2024 |  |  | 08/31/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED |  |  | US |  | 2024289344 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with post-op left ventricular assist device (lvad) was implanted with an impella rp device for mechanical circulatory support. while on support, pump had low flows. the pump was explanted and replaced with protek duo. no patient harm was reported. | [5 empty] | The impella device was returned for evaluation. upon visual inspection, biomaterial was found wrapped around the impeller. the root cause of the low pump flow was determined to be ingested biomaterial. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
16172 | 1220648-2024-21486 | 20566212 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20566212&pc=PYX | 02/21/2023 | 10/03/2025 |  | 10/30/2024 |  |  | 02/21/2023 | 10/30/2024 |  |  | 11/30/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED |  |  | US |  | 2024286339 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. during impella support, an intermittent loss of placement signal occurred. patient support continued throughout the noted issue. there was no reported patient harm. | Additional information care was eventually withdrawn, and the patient expired. |  |  |  |  | The investigation for the reported optical issues has been completed. the impella device was not received from the customer nor data logs. therefore, the root cause of the optical issue optical signal issue was unable to be determined as no product or logs were returned for this investigation. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. all pertinent information available to abiomed has been submitted at this time. b1 adverse event updated. b2 death and death date updated. b5 updated to include death description. h1, h2 type of reportable event. h6 updated to include death coding corrections that were made from the initial manufacturer device report number 1220648-2024-21486. b7, other relevant history, including preexisting medical conditions d10, concomitant medical products and therapy dates g1 reporting contact email updated.
16173 | 1220648-2024-21487 | 20566299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20566299&pc=PYX | 02/23/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 02/23/2023 | 10/30/2024 |  |  | 05/31/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED |  |  | US |  | 2022103167 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with postcardiotomy cardiogenic shock (pccs)/low cardiac output syndrome (lcos) was implanted with an impella rp device for mechanical circulatory support. due to patient's anatomy and position of aortic root cannulation, impella rp & 5.5 could not be properly positioned. after several positioning attempts, case was pivoted to va extracorporeal membrane oxygenation (ecmo) and transfer to sunrise. procedure was aborted and no harm to the patent was reported. | [5 empty] | The investigation for the reported positioning issues has been completed. the impella device was not received from the customer. however, clinical details were sufficient to determine the root cause. the root cause of the positioning issue was patient condition related to the patient's anatomy. this report is being filed as part of a retrospective review of historical records.
16174 | 1220648-2024-21488 | 20566344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20566344&pc=OZD | 10/04/2024 | 03/27/2025 |  | 10/30/2024 |  |  | 10/07/2024 | 10/30/2024 |  |  |  |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025562790 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Renal Failure | [11 empty] | Required Intervention | Other | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported the patient was on impella cp support. while on support, stroke alert occurred, and computed tomography scan showed left parietal occipital and posterior temporal no hemorrhage. the patient was unable to move their right side however it was noted to have resolved but unknown as to how. the procedural outcome was the patient survived surgery. | Additional information received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry that the patient endured acute kidney injury with a "unknown (undetermined)" relationship to the impella device. per the report, creatinine was increased to 2.3 and patient oliguric. nephrology was consulted. continuous renal replacement therapy (crrt) initiated. it was reported that the patient/event has not recovered/not resolved. |  |  |  |  | The investigation for the renal failure and stroke has been completed. the retuned product was visually inspected and a cannula kink was observed which does not effect the results of the investigation. as this clinical information was not provided, the true root cause of the stroke could not be determined. the cause of the issue was not determined since the product was not returned and insufficient clinical details were provided. b.3 date of event was reported incorrectly on follow up 1 manufacturer device report number 1220648-2024-21488. it was reported correctly on the initial manufacturer device report number 1220648-2024-21488. d.4 revised serial number as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2024-21488. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2024-21488 was submitted in accordance with updated procedures. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The following information has been updated on manufacturer device report 1220648-2024-21488 to reflect additional information that was received. b2: outcomes attributed to adverse event updated. b3: date of additional event added. b5: updated to reflect additional information received. h6: health effect - clinical code and health effect - impact code updated to reflect. additional adverse event. investigation for stroke and renal failure is still pending a supplemental report will be submitted once investigation is complete. instructions for use: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-21488. e4: should have been left blank. g2: report source; study missing.
16175 | 1220648-2024-21490 | 20566456 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20566456&pc=OZD | 10/07/2024 | 06/26/2025 |  | 10/30/2024 |  |  | 10/07/2024 | 10/30/2024 |  |  |  |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025516550 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 78 YR |  | Female |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hematoma | Renal Failure | Respiratory Failure | [9 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | Additional information received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry. the patient that endured emergent exploration of right deltopectoral pocket hematoma and evacuation with a "unknown (undetermined)" relationship to the impella device used: "patient developed enlarging hematoma in the right deltopectoral pocket following insertion of impella 5.5. found to have slowly oozing muscle surface of the entire pocket. on further exploration the axillary artery to graft anastomosis appeared to be completely hemostatic. 1 additional suture was placed at the toes of the anastomosis for reinforcement. thereafter the entire muscle surface of the pocket was slowly and carefully cauterized paying attention to avoid the brachial plexus. thereafter the pocket was irrigated profusely with warm saline. following adequate hemostasis, 1 jackson pratt (jp) drain was placed and the wound was closed in running layers of vicryl's followed by skin staples. pressure dressing was placed on top of the wound. patient was brought back to the icu in critical and guarded condition.¿. the patient recovered with no sequelae. furthermore, the complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured cardiac arrest with a "unknown (undetermined)" relationship to the impella device used: ¿cardiac arrest at 6:39pm. initial rhythm appeared to be pulseless electrical activity (pea). chest compressions were started and received 1x dose of epi. on pulse check she was noted to be in pea with bradycardia concerning for 3rd degree av block. during the second round of compressions, she was noted to move her arm, so compressions were aborted, and she was found to have a pulse. transcutaneous pacing started at 1846 at 80bpm. she was intubated and started vasopressors.¿. the patient recovered with no sequelae. the complaint team also received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured respiratory failure with a "unknown (undetermined)" relationship to the impella device used: ¿intubated, started on multiple pressors and inotropes, and requiring crrt. chest x-ray (b)(6) 2024 with worsening interstitial and airspace opacities bl - concern for acute respiratory distress syndrome (ards). palliative meeting held on (b)(6) at 1330 with patient's family, palliative care physician, ccu team, and critical care team. patient's family collectively decided on dnr status on (b)(6).¿ it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | The complainant reported a 78-year-old female noted as high risk coronary intervention was implanted with an impella 5.5 device for mechanical circulatory support. during impella support, there was a hematoma at the access site. the patient went to the operating room for a washout and exploration. support continued. | Additional information was received via notification from abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured renal failure with a "possible" relationship to the impella device used. |  |  |  | The following have been reported incorrectly or missing from manufacturer device report 1220648-2024-21490. b2 changed to death, date of death is unknown. b5 updated with additional treatment information and adverse events. d6b according to the clinical information the impella pump was never explanted, thus date of explant is unknown. e4 should have been left blank. g2 report source; study was missing. h6 4641 was reported incorrectly. new code 4624 was added to health effect - impact code. h6 additional codes added health effect - clinical code and health effect - impact code to reflect adverse events mentioned in b5. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | B5-updated narrative with additional adverse event of renal failure. h6 clinical code 2041 added.
16176 | 1220648-2024-21492 | 20566472 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20566472&pc=OZD | 10/01/2024 | 11/24/2024 |  | 10/30/2024 |  |  | 10/01/2024 | 10/30/2024 |  |  |  |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025474209 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 pump placed to support a patient admitted in with respiratory distress, heart failure, and cardiogenic shock. it was reported that there were some concerns of hemolysis but no wishes to change course of care by the team. it was further reported that care was withdrawn and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation for the hemolysis has been completed. data logs were reviewed which showed generally pump ran without issue during support. there were multiple intermittent suction alarms triggered but resolved during the case. the pump was returned for evaluation. visual inspection found biomaterial inside the pump housing & wrapped around the impeller. no other issues were found on the rest of the pump. the root cause of hemolysis was biomaterial ingestion. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
16177 | 1220648-2024-21489 | 20566477 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20566477&pc=OZD | 10/07/2024 | 12/10/2024 |  | 10/30/2024 |  |  | 10/07/2024 | 10/30/2024 |  |  |  |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025547627 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hematoma | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support. a large clot burden was wrapped down the catheter length into descending aorta. once explanted, both the left femoral artery (impella) and left femoral vein ( venous ecmo cannula) were repaired. once vascular re-repair completed, it was noted the axillary site had a large new hematoma and over 200mls fresh blood in the jackson-pratt drain. the patient was taken to the operating room for evacuation of axillary hematoma, no acute bleeding sources found. the patient received 50mg protamine, 2units fresh frozen plasma and 1units platelet count. additionally, the patient suffered stroke and the pump was removed the next day. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to produring repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation for the stroke and the access site bleed has been completed. data logs are not relevant to the complication. returned product was investigated and no visual abnormalities were noted. clinical details state the hematoma formed shortly after insertion, but no acute source was found. the complication was said to be stable within 48 hours. due to lacking clinical details and no abnormalities being found with returned product, the root cause of the access site bleed could not be determined. the root cause of the stroke/cva could not be determined. corrections to the initial mfr report: g1-mdr reporting contact email.
16178 | 1220648-2024-21494 | 20566497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20566497&pc=OZD | 10/09/2024 | 12/10/2024 |  | 10/30/2024 |  |  | 10/09/2024 | 10/30/2024 |  |  |  |  | 10/30/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1222710 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The eu complainant had an automated impella controller (aic) alarm when transferring the patient from the intensive care unit to the cardiac catheterization lab. the patient was being transferred for pump explant, and the team did not replace the aic. no patient harm has been reported. | [5 empty] | The investigation for the console controller alarm-yellow alarm has been completed. data logs were reviewed which confirm that multiple controller error (pbm voltage out-of-spec) alarms were issued. the console was returned for evaluation. the failure was reproducible in the lab. internal circuitry of the console was inspected and found to be within specification. when the pbm was replaced with a known good one, the failure was resolved. the root cause for the controller error (pbm voltage out-of-spec) was most likely a defective pbm. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
16179 | 1220648-2024-21495 | 20566506 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20566506&pc=OZD | 10/09/2024 | 12/11/2024 |  | 10/30/2024 |  |  | 10/09/2024 | 10/30/2024 |  |  |  |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025548428 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | Obstruction/Occlusion | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that complications were encountered during attempted delivery of an impella 5.5 pump for mechanical circulatory support. it was stated that following a right axillary cutdown, the pump was unable to advance past calcification within the innominate. after multiple failed attempts, the implant was aborted. there was no patient harm. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported delivery issues has been completed. the impella device was not returned for evaluation. however, clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because of the calcification of the at innominate artery. despite multiple attempts impella insertion was aborted.
16180 | 1220648-2024-21491 | 20566514 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20566514&pc=OZD | 09/24/2024 | 12/10/2024 |  | 10/30/2024 |  |  | 10/07/2024 | 10/30/2024 |  |  |  |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025534253 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | We received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured right axillary hematoma with a "probable" relationship to the impella 5.5 device used: ¿(b)(6) 2024 : right axillary hematoma, pressure held, heparin paused.¿. the patient recovered with no sequelae. | [5 empty] | The investigation into access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report. e4 should have been left blank. g1 reporting contact email. g2 report source; study missing. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
16181 | 1220648-2024-21493 | 20566518 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20566518&pc=OZD | 10/08/2024 | 12/06/2024 |  | 10/30/2024 |  |  | 10/08/2024 | 10/30/2024 |  |  |  |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025536119 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and during support, the impella was pointed towards the mitral valve. staff attempted to reposition the pump, but it was caught in the papillary muscle. the pump was later explanted. the patient survived the event. | [5 empty] | The investigation for the positioning issue has been completed. the pump was not returned for investigation. data log review was not required for this case. according to the clinical details, the pump was found deep in the ventricle, towards the mitral valve after transfer to icu. the doctor was able to reposition it successfully. the cause of the positioning issue was most likely use related, based on given clinical details. corrections to the initial mfr report: g1 mdr reporting contact email | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
16182 | 1220648-2024-21496 | 20566531 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20566531&pc=OZD | 10/09/2024 | 12/11/2024 |  | 10/30/2024 |  |  | 10/09/2024 | 10/30/2024 |  |  |  |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025551219 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | Obstruction/Occlusion | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the physician attempted an impella cp placement after gaining arterial access via right groin, but severe peripheral arterial disease noted. the physician was unable to place peel away sheath and cp placement was aborted. the patient had an occluded iliac artery. the patient¿s outcome is unknown. | [5 empty] | The investigation into delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because the peel away sheath was unable to be inserted due to diseased condition of the vessels. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-21496. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
16183 | 1220648-2024-21497 | 20566537 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20566537&pc=OZD | 09/30/2024 | 11/19/2024 |  | 10/30/2024 |  |  | 09/30/2024 | 10/30/2024 |  |  |  |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025530605 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, distal pulses were absent bilaterally via doppler. distal lower extremity (le) found to be mottled. according to the physician, the patient¿s condition and history of peripheral arterial disease (pad) could have contributed to the ischemia. the patient continued to deteriorate leading to multisystem organ failure. the patient¿s family signed a do not resuscitate. the patient expired on support with four vasopressors/inotropes and ventilator support. escalation to an impella 5.5 or rp was not available at the facility, and patient transfer was not an option due to the patient¿s critical condition. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-21497: e4 should have been left blank. g1 reporting contact fax number missing.
16184 | 1220648-2024-21498 | 20566737 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20566737&pc=OZD | 10/06/2024 | 11/21/2024 |  | 10/30/2024 |  |  | 10/06/2024 | 10/30/2024 |  |  |  |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025524970 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Tachycardia | Ventricular Fibrillation | [10 empty] | Required Intervention | Life Threatening |  |  |  |  | VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 59-year-old male patient in an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during impella support, the impella cp was having impella in aorta alarm. the patient went into pulseless ventricular tachycardia, ventricular fibrillation and torsade. several rounds of cardiopulmonary resuscitation were done with return of spontanrous circulation. the patient was emergent cannulated on venoarterial extracorporeal membrane oxygenation. the impella was placed, verified, and flow turned to p8. no additional alarms on the impella were noted. flows are good. the patient is stable. | [5 empty] | The investigation for the impella cp has been completed. the device was not returned for evaluation. the necessary clinical information was not provided; therefore the root cause of the vt/vf was not determined. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
16185 | 1220648-2024-21499 | 20567147 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20567147&pc=PYX | 02/04/2022 | 10/30/2024 |  | 10/30/2024 |  |  | 02/04/2022 | 10/30/2024 |  |  | 12/31/2022 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022027410 |  | 004334 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. it was reported the physician encountered delivery issues during placement of the impella. reportedly the physician attempted to unsuccessfully advance the impella rp three times using multiple wires, however the physician could not pass the tricuspid valve despite torquing the catheter. the impella rp placement was aborted. there was no patient harm reported. | [5 empty] | The investigation into the delivery issue has been completed. the impella rp was returned for investigation. the returned product was visually inspected. evaluation of the returned product found no damage to the pump which would have contributed to a delivery issue. the root cause of the delivery issue was most likely patient condition due to diseased/small vessels. impella rp instructions for use for the related event are as follows: rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ this report is being filed as part of a retrospective review of historical records.
16186 | 1220648-2024-21500 | 20567191 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20567191&pc=PYX | 04/19/2022 | 10/30/2024 |  | 10/30/2024 |  |  | 04/19/2022 | 10/30/2024 |  |  | 07/31/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022122245 |  | 004334 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in right heart failure was implanted with an impella rp device for mechanical circulatory support. while on support, there was a high purge pressure that prompted the pc to be replaced. t he rp later stopped. the rp was able to be restarted and then when the pump was explanted due to the stop the rp was seen to have a kink. | [5 empty] | The investigation for the reported high purge pressure and pump stop issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the returned pump was visually inspected and a kink in the catheter was observed. a syringe was connected with the sidearm and a filter leak was reproduced. the catheter was cut open and blood was observed in the purge line. the cause of the high purge pressure was a kinked purge line. the cause of the pump stop was a kinked purge line resulting in high purge pressure and blood ingress. the cause of the purge leak was sidearm filter damage likely from insufficient adhesive at the edges. impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
16187 | 1220648-2024-21501 | 20567242 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20567242&pc=PYX | 12/28/2022 | 10/30/2024 |  | 10/30/2024 |  |  | 12/28/2022 | 10/30/2024 |  |  | 01/31/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023175909 |  | 004334 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support there was a low flow issue with the impella device. it was noted that the pump flow had been gradually dropping and was close to zero despite adjusting the p-level. the pump was subsequently removed as a result of the ongoing flow issue. there was no reported harm or injury to the patient. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. compliant returned pump visually inspected. biomaterials and slightly cannula kink observed. no broken blades found. root cause determined as biomaterial as the patient had several health conditions and had right side issue support. this report is being filed as part of a retrospective review of historical records.
16188 | 1220648-2024-21502 | 20567312 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20567312&pc=PYX | 02/19/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 02/19/2023 | 10/30/2024 |  |  | 08/31/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED |  |  | US |  | 2024289342 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. during support, the device exhibited a purge leak. it was reported that the purge pressure was low since insertion. no information regarding resolution is available. it was reported that the patient survived the procedure. | [5 empty] | The impella device was returned for evaluation. upon inspection, dyed purge fluid was run through the system. no leaks were reproduced on the sidearm of the pump. dyed purge fluid was noted on the housing of the returned cassette. the purge cassette was cut open and a leak was observed at the piston to cylinder interface. the root cause of the purge leak was determined to be a piston failure, and no issues were found with the pump. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
16189 | 1220648-2024-21503 | 20567320 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20567320&pc=PYX | 03/02/2022 | 10/30/2024 |  | 10/30/2024 |  |  | 03/02/2022 | 10/30/2024 |  |  | 07/31/2022 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 1488304 |  | 0046-0011 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 20 YR |  | Male |  |  | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an acute patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. the complainant reported pump was working properly, but suddenly low flow alarm. transthoracic echocardiography (tte) was performed, and no flow was provided by the pump. patient was brought to the operating room to remove the pump. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. no clinical details provided were sufficient to determine a root cause. this report is being filed as part of a retrospective review of historical records.
16190 | 1220648-2024-21504 | 20567373 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20567373&pc=OZD | 10/03/2024 | 12/06/2024 |  | 10/30/2024 |  |  | 10/04/2024 | 10/30/2024 |  |  |  |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025545601 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 52-year-old male noted as high risk coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, there were no distal pulses found in right root (impella side). fem to fem bypass was placed for distal perfusion to right leg. throughout the night, the patient began to complain of pain, leg became cold and then patient couldn¿t move leg. vascular was consulted and patient was emergently taken to the operating room to remove impella and perform thrombectomy. impella removed without issue. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation for the limb ischemia has been completed. the device was not returned for evaluation. however, clinical details provided were sufficient to determine the most likely root cause. the necessary clinical information was not provided, therefore the root cause of the limb ischemia was not determined.
16191 | 1220648-2024-21505 | 20567374 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20567374&pc=PYX | 02/19/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 02/19/2023 | 10/30/2024 |  |  | 08/31/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED |  |  | US |  | 2024289362 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella rp for mechanical circulatory support. during support, the device exhibited purge pressure issue with corresponding low flow. the device was explanted and replaced with a new impella rp. it was reported that the patient survived explant, no further information is available. | [5 empty] | The impella device was returned for evaluation. upon visual inspection, biomaterial was found wrapped around the impeller and in the purge gap. therefore, the cause of the low pump flow was determined to be ingested biomaterial. the root cause is unable to be determined. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
16192 | 1220648-2024-21507 | 20567414 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20567414&pc=PYX | 02/27/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 02/27/2023 | 10/30/2024 |  |  | 08/31/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | 23-13676 |  |  | ABIOMED |  |  | US |  | 2024289369 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella rp for mechanical circulatory support. during support, the device exhibited low flow and a variation in motor current, possible ingestion of biomaterial. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: the date of patient death is unknown. the impella device was returned for evaluation. the data logs were reviewed and show a motor current spike on the third day of support following the ramp up to p-8. during the motor current spike there is also a drop in speed, drop in flow and increase in placement signal. this is indicative of biomaterial ingestion. the product was returned, and a large thrombus was found wrapped around the impeller. the root cause of the thrombus is unable to be determined. this report is being filed as part of a retrospective review of historical records.
16193 | 1220648-2024-21506 | 20567416 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20567416&pc=OZD | 10/01/2024 | 12/04/2025 |  | 10/30/2024 |  |  | 10/01/2024 | 10/30/2024 |  |  |  |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1095481 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  |  |  | Death | Yes | Yes | Pumping Stopped | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported pump flow issues on the impella cp. the automated impella controller (aic) was returned for investigation of possible flow issues. the patient expired after the family decided to withdraw care after acute condition changes overnight. the relationship between the console and cause of death is unknown. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The product was returned for analysis. following the case, the console was turned on thirty minutes and immediately shut off. almost two hours after that, the console is booted up again and triggers the software watchdog due to a failing imcgui subprocess triggering a partial reboot by design. the console booted back up still triggering the software watchdog and did a partial reboot again. after two partial reboots, the system then does a full reboot after which the unexpected shutdown alarm triggers and the console boots up as expected. the console was power cycled 200 times but the imc gui subprocess failure was not reproduced. no pump stop events were identified during the case. the cause of the imc gui subprocess failing triggering the software watchdog could not be determined since it was not reproduced. d2 common device name updated. g1 reporting contact email updated.
16194 | 1220648-2024-21508 | 20567418 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20567418&pc=OZD | 10/04/2024 | 12/06/2024 |  | 10/30/2024 |  |  | 10/04/2024 | 10/30/2024 |  |  |  |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025509742 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 86-year-old male noted as high risk coronary intervention was implanted with an impella cp device for mechanical circulatory support. during impella implant, the guidewire kinked. a competitor's control wire 0.018 was used instead. there was no reported patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. | The investigation into product damage (guidewire damage - great vessel) has been completed. the impella device was not returned for evaluation. clinical details noted that wire kinked when attempting to access left ventricle. no damage was noted on guidewire before insertion. the root cause of the product damage (guidewire damage - great vessel) cannot be definitively determined as limited clinical details on patient's condition were provided and no product was returned to confirm damage. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-21508.
16195 | 1220648-2024-21509 | 20567439 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20567439&pc=PYX | 03/29/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 03/29/2023 | 10/30/2024 |  |  | 06/30/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED |  |  | US |  | 2023222773 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was going to be implanted with an impella rp device for mechanical circulatory support. during delivery, the impella rp advanced over the wire, and once in the right ventricle impella immediately stopped advancing. the physician created a stiff rail by adding backward tension on guidewire while also attempting to advance rp forward. the decision was made to remove the rp and wire and attempt to cross a second time. the buddy wire technique was utilized with a 0.035¿ amplatz super stiff wire. the physician also created clockwise torque; however, the rp was still unable to advance through the right ventricle. a third attempt to cross was made utilizing buddy wire technique, both clockwise and counterclockwise torque attempted but was unsuccessful. despite all efforts, the rp remained was stuck in the right ventricle and unable to advance forward. the right ventricle was severely dilated with minimal contractility. the decision was made to remove rp and support the patient with medications. no further information is available. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device was not received from the customer. however, clinical details were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition related, the impella was unable to pass through the vasculature due to patient's anatomy. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
16196 | 1220648-2024-21511 | 20567630 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20567630&pc=OZD | 10/04/2024 | 11/26/2024 |  | 10/30/2024 |  |  | 10/04/2024 | 10/30/2024 |  |  |  |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025533751 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a cp pump placed to support a patient admitted in with a st elevated myocardial infarction and other underlying cardiac and systemic comorbidities. the pump was malpositioned within the chordae/papillary. the attempted pump repositioning failed and the pump lost left ventricle positioning and was removed. the patient was stable post-explant, and no other device was placed. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the reported positioning issue has been completed. product was not returned for review. based on clinical description, the root cause of positioning issue was most likely use related.
16197 | 1220648-2024-21510 | 20567641 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20567641&pc=PYX | 03/30/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 03/30/2023 | 10/30/2024 |  |  | 06/30/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED |  |  | US |  | 2022117778 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Death | No | Yes | Positioning Failure | Pumping Stopped | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. it was reported that the patient became fully dependent on impella support. an echocardiogram was performed to ensure proper placement. the impella rp inlet cage was noted to be at the level of the hepatic vein. then, the patient started to have suction alarms on the lower levels of support. a chest x-ray demonstrated that the rp cannula had buckled, with the outlet cage no longer in the pulmonary artery but in the right ventricle instead. the patient continued to decompensate and spiral, until multiple rounds of cardio-pulmonary resuscitation (cpr) was initiated. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: date of patient death is unknown. the investigation for the reported positioning issue and pump stop has been completed. the impella device was not received from the customer. however, the data log was reviewed and that is when the pump stop was identified. suction with a motor current spike was observed upon activation and within 30 minutes of support, a high motor current pump stop was observed. 1 hour into support a shift in placement signal was observed likely correlating with the reported positioning issues. the root cause of the positioning issue was most likely patient movement, and the root cause of the pump stop was not determined since the product was not returned. the failure modes will be monitored and trended. as noted in the impella ifu: impella cp with smart assist system section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers. this report is being filed as part of a retrospective review of historical records.
16198 | 1220648-2024-21514 | 20567672 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20567672&pc=OZD | 10/06/2024 | 12/20/2024 |  | 10/30/2024 |  |  | 10/07/2024 | 10/30/2024 |  |  | 11/30/2025 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024434974 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp support ongoing for a patient admitted in acute myocardial infarction cardiogenic shock. the pump lost the placement signal but would return at certain p levels but the p level was not appropriate for the patient. the pump was explanted and extracorporeal membrane oxygenation replaced. the procedural outcome was the patient survived the explant. | [5 empty] | The impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the optical signal issue has been completed. the impella cp returned for evaluation. upon visual inspection, no biomaterial or kinks in the catheter were observed. optical parameters were checked in lab and all were within range. the pump passed the laser continuity test. data logs were reviewed and long term log showed placement signal not reliable (psnr) alarms being triggered at the beginning of the case. the alarms were resolved when the pump was turned down to lower p-levels. increases in placement signal occurring with higher p-levels. optical parameters during case were all stable except signal-to noise ratio (snr) was low when the placement signal was spiking and psnr alarms were triggered. snr was able to return to normal range at lower p-levels. mean flow was trending down throughout case. multiple suction alarms were noted in the second half of the long term logs. clinical details noted that psnr alarms were occurring at p-levels higher than p-6 and were able to be resolved at lower p-levels. however, lower p-levels were unable to support the patient. the root cause of the optical signal issue was most likely patient condition related. g.1 revised mdr reporting contact email since the initial manufacturer device report number 1220648-2024-21514 was submitted in accordance with updated procedures.
16199 | 1220648-2024-21512 | 20567698 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20567698&pc=PYX | 07/02/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 07/02/2023 | 10/30/2024 |  |  | 12/31/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024308931 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella rp explanted for low flows. at the time of explant, there was thrombus observed on the pump. the team replaced the impella device with protek duo. no harm or injury to patient reported. | [5 empty] | The investigation for the reported low pump flow issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the issue was biomaterial. this report is being filed as part of a retrospective review of historical records.
16200 | 1220648-2024-21513 | 20567719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20567719&pc=PYX | 07/06/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 07/06/2023 | 10/30/2024 |  |  | 12/31/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024308911 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. the physician was unable to place the device due to the patient's cardiac anatomy. no known patient harm or injury. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the cause of the issue was patient condition related. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿
16201 | 1220648-2024-21515 | 20567741 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20567741&pc=PYX | 07/13/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 07/13/2023 | 10/30/2024 |  |  | 05/31/2025 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024355132 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  |  | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with unknown ethnicity and cardiomyopathy was implanted with an impella rp for mechanical circulatory support. the physician was unable to position the impella rp. there was resistance in the right atrium. the decision was made to abort the procedure. the impella was not able to be implanted. no patient harm reported due to the issue. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the cause of the delivery issue was patient condition related due to patient anatomy. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿
16202 | 1220648-2024-21517 | 20567763 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20567763&pc=PYX | 05/29/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 05/29/2023 | 10/30/2024 |  |  | 01/31/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP EU |  |  | ABIOMED, INC. |  |  | US |  | 2023273805 |  | 0046-0011 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the physician was unable to deliver the device due to the patient's anatomy. impella support was aborted without any reported harm or injury to the patient. | [5 empty] | The investigation for the reported delivery issue has been completed. the root cause of the issue was patient condition related. as noted in the impella ifu: impella rp smartassist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves warning: pre-support evaluation axillary insertion of the impella 5.5 catheter direct aortic insertion use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings warning caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions warnings & cautions caution ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings warning: axillary insertion of the impella cp ® with smartassist ® catheter inserting the impella rp flex with smartassist® system catheter inserting the impella rp® with smartassist® system catheter ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ axillary insertion of the impella 5.5 catheter ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valveswarning: pre-support evaluation axillary insertion of the impella 5.5 catheter direct aortic insertion use of impella with transcatheter aortic valves warning caution: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella rp & rp flex rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: potential adverse events (united states ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
16203 | 1220648-2024-21516 | 20567821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20567821&pc=OZD | 10/05/2024 | 07/08/2025 |  | 10/30/2024 |  |  | 10/05/2024 | 10/30/2024 |  |  |  |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025534497 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 24 YR |  | Male |  |  | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Us complainant reported that a patient was on impella cp support. while on support, there were continuous suction events despite echocardiography showing ample volume. the doctor suspected clot/tissue entrainment and made the decision to replace the impella. the patient survived the event. | [5 empty] | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | Clinical details noted that pump was initially out of position but repositioning did not resolve low pump flows. patient had adequate volume on echocardiogram. the root cause of the low pump flow could not be definitively determined as clinical details note that patient had adequate volume and low pump flows continued even after pump was repositioned, additionally, no issue was found on the returned product. e.1 fax number was inadvertently omitted from the initial manufacturer device report number 1220648-2024-21516. g.1 revised reporting contact email from the initial manufacturer device report 1220648-2024-21516 in accordance with updated procedures. h.6 investigation findings and investigation conclusions were incorrectly entered on manufacturer device report number 1220648-2024-21516-1. health effect - impact code 4641 was also incorrectly entered on manufacturer device report number 1220648-2024-21516-1. additional type of investigation codes were added as they were inadvertently omitted from the first supplemental manufacturer device report number 1220648-2024-21516. h.10 investigation root cause for low pump flow was incorrectly entered on manufacturer device report number 1220648-2024-21516-1. | The investigation for the low pump flow and thrombus has been completed. the impella pump was returned for evaluation and visually inspected. there was no biomaterial, no cannula kink and no broken blade observed. pump run in water for more than 10 minutes at p-9 to reproduce the reported low pump flow issue. pump flow was 3.7 l/min and within specification limit as required. no issue observed. data logs were reviewed which showed suction alarms occurred. no other indication observed. the cause of the low pump flow issue most likely was patient condition due to suction alarms and rv failure. and patient was also on ecmo and potential lack of volume. the root cause of the thrombus could not be determined without additional clinical details.
16204 | 1220648-2024-21519 | 20567918 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20567918&pc=PYX | 06/13/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 06/13/2023 | 10/30/2024 |  |  | 02/28/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023204309 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. while on support, the physician noticed some clotting in the heart vessels which required flushing to clear the clot. after this there was a drop in flow and pressure of impella. the physician decided to use tissue plasminogen activator (tpa) to break any type of clotting blocking impella, and that helped resolve the issue. there was no harm to the patient as a result of this incident. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. data review confirmed the clinical narrative for low flow. data showed about 10 hours into the case, the flow dropped to 2.3 l/min but no alarms were triggered. the flow then increased gradually and back to normal range 1 hour later and remained to the rest of the case. based on the clinical description and data review, the root cause of low flow was most likely due to biomaterial ingestion. this report is being filed as part of a retrospective review of historical records.
16205 | 1220648-2024-21518 | 20567944 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20567944&pc=PYX | 05/29/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 05/29/2023 | 10/30/2024 |  |  | 02/28/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023199612 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella rp device for mechanical circulatory support. the catheter of the device was kinked/pinned from the torque applied in attempt to deliver the device. additionally, there was a crack in the 23 fr peel away sheath that was used for this insertion. this impella rp was removed and successfully was replaced with another impella rp. the procedural outcome was the patient survived. | [5 empty] | The investigation into the delivery issues/introducer damage - mechanical issue has been completed. the pump was not returned for investigation. the cause of the delivery issue was use related to torque insertion technique which resulted in catheter kink and sheath damage. the cause of the introducer damage issue was an introducer sheath split along the scorelines; this is attributed to the patient condition. as noted in the impella ifu: impella rp with smart assist system section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
16206 | 1220648-2024-21520 | 20567979 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20567979&pc=PYX | 06/28/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 06/28/2023 | 10/30/2024 |  |  | 10/31/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023204435 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella rp device for mechanical circulatory support. during implantation, the medical staff was unable to deliver impella rp due to the patient's anatomy. the physician decided to escalate to impella 5.5 for hemodynamic support. there was no harm to the patient as a result of this incident. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the issue was patient condition related based on provided clinical information. impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller instructions for use for the related event are as follows: section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings, ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
16207 | 1220648-2024-21521 | 20567995 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20567995&pc=PYX | 06/30/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 06/30/2023 | 10/30/2024 |  |  | 10/31/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023274781 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella rp device for mechanical circulatory support. while on support, the impella rp was started, then a high motor current alarm went off. the device stopped. the staff replaced the pump with a new pump. no known patient harm or injury. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the data logs were reviewed, and multiple motor current spike alarm were observed. alarm 13 high motor current pump stop was observed per log analysis. the cause of issue was not determined since product was not returned and based on inconclusive evidence per data log review. impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
16208 | 1220648-2024-21522 | 20568137 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20568137&pc=OZD | 10/04/2024 | 12/10/2024 |  | 10/30/2024 |  |  | 10/07/2024 | 10/30/2024 |  |  |  |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025548398 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient had an attempted impella 5.5 implant however, the impella 5.5 was unable to navigate through the vascular anatomy. decision made to abort 5.5 implant. the patient¿s outcome is unknown. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿. ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿. section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿. section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿. | The investigation for the impella 5.5 has been completed. the device was not returned for evaluation. however, clinical details provided were sufficient to determine the most likely root cause. the root cause of delivery issue is determined to be patient condition because the pump was unable to delivered due to tortuosity of the patient vessels.
16209 | 1220648-2024-21524 | 20568185 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20568185&pc=OZD | 10/08/2024 | 12/18/2024 |  | 10/30/2024 |  |  | 10/08/2024 | 10/30/2024 |  |  |  |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025538810 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support, the pump was stuck on the papillary muscle. several attempts were made to reposition and they were unsuccessful. the staff escalated to an impella 5.5. the patient survived the explant. | [5 empty] | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the reported positioning issue has been completed. the device was not returned for evaluation. however, clinical information provided is sufficient to determine the root cause. the cause of the positioning issue most likely was patient condition related due to anatomy of patient's heart (small aortic arch/ventricle) and no better positioning was possible. corrections to the initial mfr report: g1 mdr reporting contact email.
16210 | 1220648-2024-21523 | 20568186 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20568186&pc=PYX | 09/24/2021 | 10/30/2024 |  | 10/30/2024 |  |  | 09/24/2021 | 10/30/2024 |  |  | 02/28/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022065743 |  | 004334 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella rp for mechanical circulatory support. it was reported that despite two separate power cables and two separate consoles being used, the console did not recognize that the cable had been connected. the pump was exchanged and the automated impella controller recognized it. | [5 empty] | The investigation into the reported issue of "pump not detected" has been completed. the device was received for evaluation. evidence of twisting was observed. the cause of the pump unable to be recognized was bent connector pins from issues of improper technique. this report is being filed as part of a retrospective review of historical records.
16211 | 1220648-2024-21525 | 20568315 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20568315&pc=PYX | 11/14/2021 | 10/30/2024 |  | 10/30/2024 |  |  | 11/14/2021 | 10/30/2024 |  |  | 04/30/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022094485 |  | 004334 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the us complainant had a bipella patient. the cp was successfully placed, but due to anatomic limitation the rp was unable to be placed. in the repeated attempts the pump kinked. the procedure was aborted. no patient harm was reported. | [5 empty] | The investigation into the delivery issue has been completed. the impella rp was returned for investigation. the returned product was visually inspected, and no abnormalities were observed. the product did not fail to meet specifications. per review of the clinical details the cause of the delivery issue was most likely difficult patient anatomy. impella rp smartassist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
16212 | 1220648-2024-21526 | 20568320 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20568320&pc=OZD | 10/07/2024 | 12/05/2024 |  | 10/30/2024 |  |  | 10/07/2024 | 10/30/2024 |  |  |  |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025567130 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Hematuria | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and during support, there was positioning issues. the pump was shallow and the pigtail was pointed toward papillary muscle. the patient had hematuria in the foley, but labs were not hemolyzing. lactate dehydrogenase was rising. additionally, the patient had a hematoma. the pump was removed due to the positioning issues and the hematoma. the patient was escalated to an impella 5.5 and survived the event. | [5 empty] | The investigation for the impella cp has been completed. the device was not returned for evaluation. data logs were reviewed which showed pump ran without issue during support. no positioning alarm was triggered during support. product was not returned for review. based on clinical description and data review, the root cause of positioning issue was most likely use related. the root cause of access site bleeding was most likely patient condition. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16213 | 1220648-2024-21527 | 20568340 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20568340&pc=PYX | 05/10/2022 | 10/30/2024 |  | 10/30/2024 |  |  | 05/10/2022 | 10/30/2024 |  |  | 03/31/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022071461 |  | 004334 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella rp device for mechanical circulatory support. while on support, and trying to achieve hemostasis, the impella fell into the right ventricle (rv). multiple attempts were made to recross without wire, which were unsuccessful. the decision was made to place new pump and sheath. | [5 empty] | The impella device was returned, no issue was found with the sheath. the excessive bleeding was most likely due to improper usage. .impella rp with smart assist system with automated impella controller section: warnings and cautions instructions for use for the related event are as follows: ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound .¿ this report is being filed as part of a retrospective review of historical records.
16214 | 1220648-2024-21528 | 20568367 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20568367&pc=PYX | 06/24/2022 | 10/30/2024 |  | 10/30/2024 |  |  | 06/24/2022 | 10/30/2024 |  |  | 02/29/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023182437 |  | 004334 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella rp device for mechanical circulatory support. while on support, and upon successful placement of the device and support initiation, the ps read 180s/100s with flows between 0.0-0.5 on p6. a suction alarm occurred on p7. the decision was made to exchange the device for new impella rp. the pump was removed and replaced with no harm to the patient. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the data logs showed erratic waveforms in the placement signal, motor current, and pump flow upon starting the pump. soon after there were recurrent "suction" alarms with a loss in motor current pulsatility and low flows. there were no placement signal alarms. however, since the pump was not returned, the cause of the low pump flow was not determined. this report is being filed as part of a retrospective review of historical records.
16215 | 1220648-2024-21529 | 20568392 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20568392&pc=PYX | 07/13/2022 | 10/30/2024 |  | 10/30/2024 |  |  | 07/27/2022 | 10/30/2024 |  |  | 09/30/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023199504 |  | 0046-0035 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella rp device for mechanical circulatory support. two weeks after the explant of the impella, the team informed the field rep of a concern for hemolysis. no further information was provided. | [5 empty] | The investigation for the reported hemolysis issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the issue was most likely patient condition related since the hemolysis was not attributed to impella. this report is being filed as part of a retrospective review of historical records.
16216 | 1220648-2024-21530 | 20568462 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20568462&pc=OZD | 10/09/2024 | 12/31/2024 |  | 10/30/2024 |  |  | 10/09/2024 | 10/30/2024 |  |  |  |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025452570 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported the patient was having an impella 5.5 implanted and the implant was performed using dao, and it was explained to the patient that blind insertion carried the risk of left ventricular perforation. however, due to the space in the operating room, the catheter was inserted without fluoroscopy, which resulted in left ventricular perforation. the ventricle was patched and support was continued. the patient remained on impella support. | [5 empty] | The investigation for the ventricle perforation has been completed. data logs were not available for review and the product was not returned for investigation. during insertion of the impella 5.5 directly through the aorta, the physician decided not to use imaging guidance, which does not align with the recommendation. as a result, the left ventricle was perforated. based on clinical details, the root cause of the ventricle perforation was determined to most likely be a use issue. g.1 revised mdr reporting contact email since the initial manufacturer device report number 1220648-2024-21530 was submitted. | D.6b explant date was inadvertently omittted from follow-up 1 manufacturer device report number 1220648-2024-21530. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
16217 | 1220648-2024-21531 | 20568494 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20568494&pc=OZD | 10/07/2024 | 12/10/2024 |  | 10/30/2024 |  |  | 10/07/2024 | 10/30/2024 |  |  |  |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025539618 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a 53-year-old female patient was scheduled for impella 5.5 implant for mechanical circulatory support. during attempts to implant the device, it was noted that the pump was unable to advance past innominate due to the small vessel size. after multiple failed attempts, the implant was aborted and an impella cp pump was placed for ongoing support. there was no patient harm. | [5 empty] | The investigation for the delivery issue has been completed. the device was returned for evaluation. no issues were found in the returned product. clinical details provided were sufficient to determine the root cause. the root cause of the delivery issue was most likely patient condition since the implant was aborted due to patient's anatomy. added d9 device return date corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
16218 | 1220648-2024-21532 | 20568946 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20568946&pc=OZD | 10/07/2024 | 12/09/2024 |  | 10/30/2024 |  |  | 10/07/2024 | 10/30/2024 |  |  |  |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025547094 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that complications were encountered during attempted delivery of an impella 5.5 pump for mechanical circulatory support. following insertion, it was noted that narrowing at the innominate/aortic junction prevented advancement of the device. after multiple failed attempts, the implant was aborted and an impella cp was placed for support. there was no patient harm. | [5 empty] | The investigation into delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition related, impella unable to pass through innominate/aortic junction as per the clinical description provided. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-21532. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
16219 | 1220648-2024-21533 | 20569369 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20569369&pc=OZD | 10/09/2024 | 12/06/2024 |  | 10/30/2024 |  |  | 10/09/2024 | 10/30/2024 |  |  |  |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025550606 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 22 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support and during the implant, the patient had ventricular tachycardia multiple times. the staff performed cardiopulmonary resuscitation. support continued as a bridge to a transplant. the patient was taken to the operating room for a heart transplant. the impella 5.5 was explanted during the transplant procedure. the patient survived the explant. | [5 empty] | This report is being supplemented to provide additional information based on the legal manufacturer's final investigation. updated/corrected h6. the field service engineer that troubleshooted the issue confirmed that the malfunction was likely caused due to a loose cable connection. olympus will continue to monitor field performance for this device. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
16220 | 1220648-2024-21535 | 20569531 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20569531&pc=PYX | 08/14/2021 | 10/30/2024 |  | 10/30/2024 |  |  | 08/14/2021 | 10/30/2024 |  |  | 08/31/2022 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 1498341 |  | 004334 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  |  | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella rp for mechanical circulatory support. it was reported the pump was successfully placed however, the purge flow rate was trending down, and the purge pressure was trending up. reportedly the flow reached 2.1 and purge pressure greater than 1000. this impella rp was removed and successfully replaced with another impella rp and noted normal purge flow and pressure. there was no patient harm reported. | [5 empty] | The investigation into the high purge pressure issue has been completed. the impella rp was returned for investigation. the provided data log showed that purge pressure ranged from 900 to 1000 mmhg between 07:14 and 07:51 on 14aug2021. during this time, purge flow was approximately 2.5 milliliters per hour. no high purge pressure or purge flow decreased alarms were triggered. at pressure p-9, the pump struggled to maintain flow within the 3.9 to 4.4 liters per minute range. the evaluation of the returned pump found a suture like material wrapped around the impeller and purge gap. the pump was run without issue. of note, the purge cassette was not returned. the root cause of the high purge pressure and low purge flow was most likely the suture like material obstructing the purge gap due to user improper technique. impella rp instructions for use for the related event are as follows: impella rp with smart assist system section: using the automated impella controller with the impella rp with smart assist system catheter as noted in the impella ifu ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smart assist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smart assist system catheter whenever a rise in motor current is seen.¿ this report is being filed as part of a retrospective review of historical records.
16221 | 1220648-2024-21534 | 20569548 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20569548&pc=PYX | 08/03/2021 | 10/30/2024 |  | 10/30/2024 |  |  | 08/03/2021 | 10/30/2024 |  |  |  |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 1501414 |  | 004334 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella rp for mechanical circulatory support. it was reported that it took multiple attempts with the buddy wire to successfully place the impella rp. reportedly on the 3rd attempt the pump was successfully placed however the pump stopped approximately 5 to 10 minutes after placement, resulting in removal of the pump. the pump was noted to have a thrombus in the inflow upon removal. this impella rp was removed and successfully replaced with another impella rp. there was no patient harm reported. | [5 empty] | The investigation into the pump stop issue has been completed. the impella rp was returned for investigation. the data logs from the case were not returned. a visual inspection of the returned pump revealed the presence of significant biomaterial wrapped around the impeller. no other damage or abnormalities were found. in conclusion, per the results of the clinical review and visual inspection of the returned device was determined that the cause of the pump stop was ingested biomaterial. impella rp instructions for use for the related event are as follows: impella; rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: troubleshooting the purge system as noted in the impella ifu ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smart assist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smart assist system catheter whenever a rise in motor current is seen.¿. section: warnings & cautions: warnings. as noted in the impella ifu ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿. this report is being filed as part of a retrospective review of historical records.
16222 | 1220648-2024-21536 | 20569564 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20569564&pc=PYX | 11/03/2021 | 10/30/2024 |  | 10/30/2024 |  |  | 11/03/2021 | 10/30/2024 |  |  | 03/31/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022087526 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  |  | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant in the us had a rp support in which the flows were low. additional information noted care was withdrawn and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation into the reported issue has been completed. no device was received for evaluation. the cause of the low pump flow was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records.
16223 | 1220648-2024-21537 | 20569586 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20569586&pc=PYX | 01/04/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 01/04/2023 | 10/30/2024 |  |  | 04/30/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023208325 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp for mechanical circulatory support. the impella pump was unable to be delivered due to the patient's vasculature. impella support was aborted without any harm. | [5 empty] | The investigation for the reported delivery issue has been completed. no device was received for evaluation. the root cause of the delivery issue was patient condition related. as noted in the impella ifu: impella rp smartassist system with automated impella controller, section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. warning: pre-support evaluation, axillary insertion of the impella 5.5 catheter, direct aortic insertion, use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: warning caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions: warnings & cautions: caution: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: warning: axillary insertion of the impella cp ® with smartassist ® catheter, inserting the impella rp flex with smartassist® system catheter, inserting the impella rp® with smartassist® system catheter. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ axillary insertion of the impella 5.5 catheter. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves warning: pre-support evaluation: axillary insertion of the impella 5.5 catheter. direct aortic insertion: use of impella with transcatheter aortic valves. warning: caution: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella rp & rp flex rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller: section: potential adverse events (united states): ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings: ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
16224 | 1220648-2024-21538 | 20569608 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20569608&pc=PYX | 01/15/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 01/15/2023 | 10/30/2024 |  |  | 04/30/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023268846 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella rp for mechanical circulatory support. it was reported the pump experienced low pump. the patient remained on impella support and was successfully weaned. | [5 empty] | The investigation for low pump flow has been completed. the data logs confirm a low displayed flow. the motor current was consistent throughout the case showing that flows were adequate. the pump was rezeroed multiple times throughout the case. the pump was rezeroed while the patient was hypotensive and had volume issues. once the patient was given volume it is likely that the low flow condition resolved however the rezeroing of the sensor resulted in an offset in the dp and calculated flow. the root cause of the prolonged low flow was determined to be use related as the site rezeroed in an attempt to mitigate an issue unrelated to the functionality of the dp sensor. this report is being filed as part of a retrospective review of historical records.
16225 | 1220648-2024-21539 | 20569761 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20569761&pc=PYX | 01/15/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 01/15/2023 | 10/30/2024 |  |  | 06/30/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023221298 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp for mechanical circulatory support. the pump exhibited high purge pressure alarms and purge system blocked. the team observed a large clot burden in the time of explant. the patient remained on impella support and was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
16226 | 1220648-2024-21540 | 20569786 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20569786&pc=OZD | 03/31/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 03/31/2023 | 10/30/2024 |  |  | 02/29/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023182449 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  |  |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that during impella placement, suction alarms were received. it was stated that the pump was observed to have ingested a clot, and the device was explanted at bedside due to the issue. weaning was successful, and the procedural outcome was the patient survived. | [5 empty] | The investigation into the reported issue was completed. the device was not received for evaluation. based on the clinical details it was confirmed that the device ingested clot. data logs were reviewed and showed characteristics of ingestion. the root cause of the low pump flow most likely was biomaterial ingested. this report is being filed as part of a retrospective review of historical records.
16227 | 1220648-2024-21541 | 20569996 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20569996&pc=PYX | 02/01/2022 | 10/30/2024 |  | 10/30/2024 |  |  | 02/01/2022 | 10/30/2024 |  |  | 02/28/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022052379 |  | 004334 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was being implanted with an impella rp device for mechanical circulatory support. while prepping the device, the power cable would not plug into the rp catheter. upon further inspection, a prong on device was found to be bent and loose, and during the inspection the prong fell out of the device. attempted to still plug in power cable but upon connecting to aic, the device was not recognized. no other rp devices were available, therefore, the patient was transferred to a new hospital for support. | [5 empty] | The investigation for the reported product damage issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that it was confirmed that a pin was broken off on the rp side of the connection. the pump was plugged into the console and it was confirmed that the console was unable to recognize the pump. the root cause of the issue was determined to be a broken pin. the pin was most likely broken by connecting the pump to the patient cable at the incorrect alignment. this report is being filed as part of a retrospective review of historical records.
16228 | 1220648-2024-21542 | 20570158 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20570158&pc=PYX | 02/15/2022 | 10/30/2024 |  | 10/30/2024 |  |  | 02/15/2022 | 10/30/2024 |  |  | 12/31/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023158237 |  | 004334 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with unknown ethnicity in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. it was reported during preparation a purge fluid leak was observed from the purge reserve/filter. this impella rp was replaced with another impella rp resulting in no issues with the replacement impella rp. there was no patient harm reported. | [5 empty] | The investigation into the purge leak has been completed. the impella rp was returned for investigation. the returned product was visually inspected, and a filter leak was confirmed at the vents of the filter. the root cause of the purge system leak was determined to be sidearm filter damage, most likely as a result of ipa interaction. impella rp instructions for use for the related event are as follows: impella rp & rp flex impella; rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: cleaning as noted in the impella ifu ¿do not clean with or expose any part of the clear sidearm of the impella catheter (eg, infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir.¿ this report is being filed as part of a retrospective review of historical records.
16229 | 1220648-2024-21546 | 20570276 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20570276&pc=PYX | 07/28/2023 | 10/13/2025 |  | 10/30/2024 |  |  | 07/28/2023 | 10/30/2024 |  |  | 08/31/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023264147 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  |  |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. while on support the catheter kinked, causing a hole in the catheter. the patient was successfully weaned from rp support with an impella 5.5 still in place. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. g1 reporting contact email revised on manufacturer device report 1220648-2024-21546 in accordance with updated procedures. h5 selection was inadvertently omitted on the initial manufacturer device report number 1220648-2024-21546.
16230 | 1220648-2024-21543 | 20570280 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20570280&pc=PYX | 11/16/2022 | 10/30/2024 |  | 10/30/2024 |  |  | 11/16/2022 | 10/30/2024 |  |  | 02/28/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023204406 |  | 004334 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a placement signal not reliable alarm during use of the device. a decision was made to explant and replace the impella. no reported harm or injury to the patient. | [5 empty] | The investigation into the placement signal issue has been completed. the impella rp was returned for investigation. data logs showed that the "epm sensor defect" alarm (#701) triggered immediately upon initially plugging in the pump. the placement signal immediately spiked to 3000 upon connecting the pump, followed by the occurrence of the "placement signal not reliable" alarm. disconnecting and reconnecting the impella did not resolve the issues. the raw placement signal reading averaged just below 1000 during this time. no external damage or abnormalities were found on the returned pump. upon connecting the pump to an aic, the placement signal spiked out of range and the "placement signal not reliable" alarm reproduced. electrical resistance testing revealed open lines between motor phases 1 and 2 and 2 and 3, as well as between s+/s- phases. there was no evidence of any use-related issue causing damage to the wiring. therefore, it was concluded that the cause of the placement signal issue was an open electrical line. this report is being filed as part of a retrospective review of historical records.
16231 | 1220648-2024-21545 | 20570348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20570348&pc=PYX | 01/06/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 01/06/2023 | 10/30/2024 |  |  | 04/30/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023268856 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 25 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella rp for mechanical circulatory support. during a cassette change, the impella stopped. attempted to restart per protocol without success. the surgeon elected to remove the device without replacement due to successful weaning. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the data logs were returned and confirmed the high motor current pump stop. the pump stop occurred during a purge cassette change and was unable to be restarted. the pump was returned, and it had been rinsed before return. no biomaterial was found in the pump. electrical testing was performed, and all motor phases were within specification. the pump was then sent out for ct scan. the ct scans showed the retainer in the bearing had worn and multiple struts were deformed/broken causing the pump to stop. the root cause of the high motor current pump stop was determined to be bearing wear. impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
16232 | 1220648-2024-21548 | 20570353 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20570353&pc=PYX | 03/31/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 03/31/2023 | 10/30/2024 |  |  | 11/30/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023281754 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp for mechanical circulatory support. the impella pump was unable to be delivered due to the patient's vasculature. impella support was aborted without any harm to patient. | [5 empty] | The investigation into the delivery issue has been completed. the root cause of the issue was most likely patient condition related based on provided clinical information. as noted in the impella ifu: impella rp smartassist system with automated impella controller. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. warning: pre-support evaluation. axillary insertion of the impella 5.5 catheter. direct aortic insertion. use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿. section: warnings & cautions: warnings. warning caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿. section: warnings & cautions: cautions. warnings & cautions. caution. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿. section: warnings & cautions: warnings. warning: axillary insertion of the impella cp ® with smartassist ® catheter . inserting the impella rp flex with smartassist® system catheter. inserting the impella rp® with smartassist® system catheter. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿. axillary insertion of the impella 5.5 catheter. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valveswarning: pre-support evaluation. axillary insertion of the impella 5.5 catheter. direct aortic insertion. use of impella with transcatheter aortic valves. warning. caution: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿. ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella rp & rp flex rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: potential adverse events (united states). section: potential adverse events (united states ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings. ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿. section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿. this report is being filed as part of a retrospective review of historical records.
16233 | 1220648-2024-21544 | 20570418 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20570418&pc=OZD | 09/30/2024 | 11/27/2024 |  | 10/30/2024 |  |  | 09/30/2024 | 10/30/2024 |  |  | 07/31/2025 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024370781 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a 56-year-old female patient in a cardiogenic shock (cgs) was implanted with an impella 5.5 device for mechanical circulatory support. while on support, bleeding was noted at the access site. the patient received blood products and pressure dressing was applied. bleeding was also noted from the chest tubes. suspected blood in abdomen. care was withdrawn. patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the access site bleed and vascular damage has been completed. no product was returned for evaluation. clinical details noted that patient was bleeding from multiple sites while on support. additionally, bleeding at access site was managed with blood products and pressure dressings. the root cause of the access site bleeding - major was most likely patient condition related as bleeding was occurring at multiple sites. the failure mode of "vascular damage - peripheral vessel" was removed as no explicit mention of vessel damage occurred at the access site and since the pump was inserted via axillary artery, the abdominal bleeding is not related to the impella. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16234 | 1220648-2024-21547 | 20570436 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20570436&pc=PYX | 03/31/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 03/31/2023 | 10/30/2024 |  |  | 08/31/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022132502 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient of unknown race was implanted with an impella rp for mechanical circulatory support. it was reported that kinks in the driveline were observed preventing the impella device to advance. impella support was aborted without any harm to patient. | [5 empty] | The investigation for delivery issue has been completed. the root cause of the delivery issue was most likely due to patient condition related as they observed catheter kink because of patient anatomy. as noted in the impella instructions for use: impella rp smart assist system with automated impella controller section: warnings and cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings and cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella rp smart assist system with automated impella controller: section: warnings and cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings and cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ impella rp smartassist system with automated impella controller. section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ this report is being filed as part of a retrospective review of historical records.
16235 | 1220648-2024-21549 | 20570461 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20570461&pc=PYX | 07/01/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 07/01/2023 | 10/30/2024 |  |  | 04/30/2025 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024337974 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient needing an impella rp for mechanical circulatory support. it was reported that during implantation, the physician had trouble delivering the device and was unable to advance any further. the procedure was aborted as the device was unable to pass through the patient's vasculature. no further information was provided. | [5 empty] | The investigation for the reported delivery issue has been completed. however, as the necessary clinical information was not provided and the device was not returned, the root cause of the delivery issue could not be determined. as noted in the impella ifu: impella & rp smartassist system with automated impella controller. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿. section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿. ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿. this report is being filed as part of a retrospective review of historical records.
16236 | 1220648-2024-21550 | 20570485 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20570485&pc=OZD | 09/30/2024 | 11/19/2024 |  | 10/30/2024 |  |  | 09/30/2024 | 10/30/2024 |  |  |  |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532113 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a 69-year-old male patient was on impella cp support due to left main disease with 100% occlusion. while on support, oozing was noted at the groin. the groin area was exposed and upon examination, the repositioning sheath blue hub was retracted from the skin tract by about 1.5 inches and was sutured down. the vascular surgeon removed the sutures and then inserted the repositioning sheath appropriately and angel matched. bleeding stopped, was noted when the repositioning sheath was moved around. the vascular surgeon performed a surgical cutdown and placed a suture around the arteriotomy over the sheath. the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation into access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was most likely use issue since the bleed was resolved by angle matching and sutures. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16237 | 1220648-2024-21552 | 20570600 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20570600&pc=OZD | 09/30/2024 | 10/30/2024 |  | 10/30/2024 |  |  | 09/30/2024 | 10/30/2024 |  |  |  |  | 10/30/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1465870 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant also reported the automated impella controller (aic) (b)(6) (subject device) displayed the purge menu options, and the screen froze. screen would not advance to get to placement screen. console was replaced, and normal flows resumed. information pertaining to devices in use during the event, but reported elsewhere. us complainant reported the patient was on impella cp support due to non-st elevated myocardial infarction (nstemi). while on support, purge alarms were noted on screen. purge cassette was changed. pump functioned appropriately, and the patient was supported. additionally, the impella became displaced. impella was repositioned at bedside with echocardiography. pericardial fluid seen on echocardiography with right ventricle compression. the patient decompensated in holding department while waiting on transfer. started on neo in addition to levo that was started post pump insertion. the patient returned to catheterization lab. impella repositioned, 83cm at sheath. the complainant also reported the automated impella controller (aic) (b)(6) displayed the purge menu options, and the screen froze. screen would not advance to get to placement screen. console was replaced. the patient expired during transit. the likely cause of death is unrelated to the subject device. | [5 empty] | The impella device was returned and evaluated. there was no issue found with the console. the device functioned/operated to specification. this is one of three medical device reports associated with the event. refer to initial medical device reports for automated impella controller serial number (b)(6), automated impella controller serial number (b)(6) and impella pump serial number (b)(6) with date of report: october 30, 2024.
16238 | 1220648-2024-21553 | 20570619 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20570619&pc=PYX | 01/28/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 01/28/2023 | 10/30/2024 |  |  | 02/28/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023188570 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella rp for mechanical circulatory support. the impella pump was unable to be delivered due to the catheter subsequently kinked. the implant rp was aborted and another rp was subsequently placed for patient support. there was no reported harm or injury to the patient. | [5 empty] | The investigation for the reported delivery issue has been completed. complaint devices not returned for investigation base on the clinical data the kink occurred while attempted to pass the rp because the guidewire was bent and resistance was met at the turn in the aorta. the root cause was most likely patient condition related. impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller instructions for use for the related event are as follows: section: warnings and cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings and cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
16239 | 1220648-2024-21551 | 20570622 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20570622&pc=OZD | 09/30/2024 | 10/30/2024 |  | 10/30/2024 |  |  | 09/30/2024 | 10/30/2024 |  |  |  |  | 10/30/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1347481 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant also reported the automated impella controller (aic) (b)(6) (subject device) displayed the purge menu options, and the screen froze. screen would not advance to get to placement screen. console was replaced. information pertaining to other devices in the event, but reported elsewhere. us complainant reported the patient was on impella cp support due to non-st elevated myocardial infarction (nstemi). while on support, purge alarms were noted on screen. purge cassette was changed. pump functioned appropriately, and the patient was supported. additionally, the impella became displaced. impella was repositioned at bedside with echocardiography. pericardial fluid seen on echocardiography with right ventricle compression. the patient decompensated in holding department while waiting on transfer. started on neo in addition to levo that was started post pump insertion. the patient returned to catheterization lab. impella repositioned, 83cm at sheath. the complainant also reported the automated impella controller (aic) (b)(6) displayed the purge menu options, and the screen froze. screen would not advance to get to placement screen. console was replaced, and normal flows resumed. the patient expired during transit. the likely cause of death is unrelated to the subject device. | [5 empty] | The impella device was returned and evaluated. there was no issue found with the console. the device functioned/operated to specification. this is one of three medical device reports associated with the event. refer to initial medical device reports for automated impella controller serial number (b)(6), automated impella controller serial number (b)(6) and impella pump serial number (b)(6) with date of report: october 30, 2024.
16240 | 1220648-2024-21554 | 20570754 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20570754&pc=PYX | 01/22/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 01/22/2023 | 10/30/2024 |  |  | 02/29/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023182494 |  | 004334 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp for mechanical circulatory support. it was reported the pump there were high purge pressure alarms and the pump stopped and was restarted multiple times. the case wrap up indicates the patient remained on impella support and was successfully weaned. | [5 empty] | The investigation for pump stop has been completed. the root cause of the high motor current pump was most likely ingested biomaterial. as noted in the impella instructions for use: impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella rp smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella rp smart assist system with automated impella controller section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers. this report is being filed as part of a retrospective review of historical records.
16241 | 1220648-2024-21555 | 20570800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20570800&pc=PYX | 01/23/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 01/25/2023 | 10/30/2024 |  |  | 03/31/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP EU |  |  | ABIOMED, INC. |  |  | US |  | 2022073989 |  | 0046-0011 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella rp for mechanical circulatory support. it was reported after 31 days the pump experienced pump stop. the case wrap up indicates the patient remained on impella support and was successfully weaned | [5 empty] | The investigation for pump stop has been completed. the data logs confirm a high motor current pump stop on the 31st day of support. the event occurs right after a brief stop in purge flow and air in purge alarms. however, the purge pressure returned to normal values after the pump stop. the root cause of the pump stop was unable to be determined as no product was returned for evaluation. as noted in the impella instructions for use: impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella rp smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella rp smart assist system with automated impella controller section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers. this report is being filed as part of a retrospective review of historical records.
16242 | 1220648-2024-21556 | 20570809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20570809&pc=PYX | 01/29/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 01/30/2023 | 10/30/2024 |  |  | 09/30/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023266263 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella rp for mechanical circulatory support. during support, placement signal not reliable alarm occurred. the case wrap up indicates the patient remained on impella support and was successfully weaned. | [5 empty] | The investigation for optical signal issue has been completed. the root cause of the optical signal issue was most likely biomaterial. this report is being filed as part of a retrospective review of historical records.
16243 | 1220648-2024-21557 | 20570880 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20570880&pc=PYX | 01/29/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 01/29/2023 | 10/30/2024 |  |  | 02/28/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP US |  |  | ABIOMED, INC. |  |  | US |  | 2022064702 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella rp for mechanical circulatory support. it was reported that kinks were observed preventing the impella device to deliver. impella support was aborted without any harm or injury to the patient. | [5 empty] | The investigation for the reported delivery issue has been completed. the root cause of the delivery issue was patient condition related based on the provided clinical information. as noted in the impella ifu: impella rp smartassist system with automated impella controller. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. warning: pre-support evaluation. axillary insertion of the impella 5.5 catheter. direct aortic insertion. use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. warning caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions. warnings & cautions. caution. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. warning: axillary insertion of the impella cp ® with smartassist ® catheter . inserting the impella rp flex with smartassist® system catheter. inserting the impella rp® with smartassist® system catheter. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ axillary insertion of the impella 5.5 catheter. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valveswarning: pre-support evaluation. axillary insertion of the impella 5.5 catheter. direct aortic insertion. use of impella with transcatheter aortic valves. warning. caution: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella rp & rp flex rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: potential adverse events (united states). section: potential adverse events (united states. ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings. ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
16244 | 1220648-2024-21558 | 20570921 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20570921&pc=OZD | 09/30/2024 | 11/18/2024 |  | 10/30/2024 |  |  | 09/30/2024 | 10/30/2024 |  |  |  |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025496339 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp (subject device) support due to non-st elevated myocardial infarction (nstemi). while on support, purge alarms were noted on screen. purge cassette was changed. pump functioned appropriately, and the patient was supported. additionally, the impella became displaced. impella was repositioned at bedside with echocardiography. pericardial fluid seen on echocardiography with right ventricle compression. the patient decompensated in holding department while waiting on transfer. started on neo in addition to levo that was started post pump insertion. the patient returned to catheterization lab. impella repositioned, 83cm at sheath. the complainant also reported the automated impella controller (aic) (b)(6) displayed the purge menu options, and the screen froze. screen would not advance to get to placement screen. console was replaced. the complainant also reported the automated impella controller (aic) (b)(6) displayed the purge menu options, and the screen froze. screen would not advance to get to placement screen. console was replaced, and normal flows resumed. the patient expired during transit. the likely cause of death is unrelated to the subject device. | [5 empty] | The investigation into positioning issues and ventricle perforation has been completed. positioning issues: while the patient was being moved to a new bed, the pump became displaced. neither data logs nor returned product are relevant to the investigation. based on the clinical information provided, the root cause of the positioning issues was determined to be patient movement. ventricle perforation: after transferring the patient to a new bed, the pump was displaced and pericardial fluid and rv compression was then seen on echo. data logs are not relevant to the investigation. returned product was investigated and there were no visual abnormalities. based on clinical details provided, the root cause of the ventricle perforation was determined to most likely be patient movement. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this is one of three medical device reports associated with the event. refer to initial medical device reports for automated impella controller serial number (b)(6), automated impella controller serial number (b)(6) and impella pump serial number (b)(6) with date of report: on (b)(6) 2024.
16245 | 1220648-2024-21559 | 20570944 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20570944&pc=PYX | 01/30/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 01/30/2023 | 10/30/2024 |  |  | 08/31/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022132495 |  | 004334 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. t he complainant reported that the implanted impella rp pump did not sit well in the patient and subsequently shifted during sheath removal. as a result, the decision was made to remove the pump and utilize a new impella rp for ongoing support. upon explant, it was further noted that there was a clot within the outlet of the pump. there was no harm to the patient. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. clinical details provided were sufficient to determine a root cause. the root cause of the positioning issue was most likely use related as the pump came out of position while the introducer was being peeled away. this report is being filed as part of a retrospective review of historical records.
16246 | 1220648-2024-21560 | 20570958 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20570958&pc=PYX | 02/04/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 02/04/2023 | 10/30/2024 |  |  |  |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022153823 |  | 0046-3196 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. impella rp insertion aborted due to patient anatomy. coronary artery bypass graft procedure performed without impella support. there was no reported harm or injury to the patient. | [5 empty] | H4; serial number, expiration date, udi, h4; manufacturing date unavailable. the impella device was not received from the customer, therefore, investigation of the device was not possible. the root cause of the delivery issue was determined to be patient condition per the clinical details. this report is being filed as part of a retrospective review of historical records.
16247 | 1220648-2024-21561 | 20571065 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571065&pc=PYX | 02/13/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 02/13/2023 | 10/30/2024 |  |  | 03/31/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2022074728 |  | 004334 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. while on support the impella moved into the right ventricle after the patient was turned. upon removal of the rp, a large clot was noted when the device was removed from the peel away sheath. there was no reported harm or injury to the patient. | [5 empty] | The investigation for the reported thrombus has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, information provided was sufficient to determine a root cause. root cause of the issue was patient movement overnight due to improper technique. this report is being filed as part of a retrospective review of historical records.
16248 | 1220648-2024-21562 | 20571139 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571139&pc=OZD | 04/13/2022 | 10/30/2024 |  | 10/30/2024 |  |  | 04/13/2022 | 10/30/2024 |  |  | 12/31/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023181001 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported during exchanging the peel away for a repositioning sheath the impella popped out of the ventricle, and when trying to cross the aortic valve the impella kinked. this impella placement was aborted and a new replacement impella was placed. | [5 empty] | The investigation into the positioning issue has been completed. the impella pump was returned for investigation. per clinical review while exchanging the introducer with the repositioning sheath, the impella was inadvertently pulled out of the left ventricle. it was noted that the introducer was not fully removed from the body before it was peeled away. the clinicians attempted to reinsert the pump into the left ventricle, but the cannula kinked, preventing successful re-access. the data logs showed the "impella position in aorta" suddenly triggered, accompanied by aortic placement signal waveforms. a cannula kink was observed near the outflow cage of the returned pump. otherwise, no damage or abnormalities were found. after reviewing the clinical information, it was determined that the cause of the positioning issues (and the subsequent cannula kink) was use related. improper technique was used when removing the introducer and the impella is not indicated for wireless re-access per ifu 0048-9007. the cause of the positioning issue was due to improper technique by the end user. this report is being filed as part of a retrospective review of historical records.
16249 | 1220648-2024-21563 | 20571155 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571155&pc=PYX | 06/14/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 06/14/2023 | 10/30/2024 |  |  | 12/31/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024296426 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. while on support the pump was found to be kinked. the pump was replaced as a result of the noted issue. there was no reported harm or injury to the patient. | [5 empty] | The investigation for the reported product damage (cannula kink) has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, the root cause was most likely use related due to improper technique where imaging was not performed prior to impella placement. this report is being filed as part of a retrospective review of historical records.
16250 | 1220648-2024-21564 | 20571224 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571224&pc=OZD | 09/30/2024 | 11/26/2024 |  | 10/30/2024 |  |  | 09/30/2024 | 10/30/2024 |  |  |  |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025545136 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a 61-year-old male patient in an acute myocardial infarction / cardiogenic shock (ami/cgs) was implanted with an impella cp device for mechanical circulatory support. the patient had ventricular tachycardia (vt) throughout support. pump position was confirmed and the flows remained at 3.3 l/min during vt. it is unknown if the impella caused or contributed to the patient's vt. the patient was on impella support for 213.43 hours before the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined. corrections to the initial mfr report: g1 mfr site name and address. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16251 | 1220648-2024-21566 | 20571230 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571230&pc=OZD | 10/01/2024 | 12/04/2024 |  | 10/30/2024 |  |  | 10/01/2024 | 10/30/2024 |  |  |  |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025534371 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Death | Yes | No | Increase in Pressure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The eu complainant had a cp pump inserted for cardiogenic shock / st elevated myocardial infarction patient. the pump had high purge pressure. they attempted to troubleshoot for kinking and other causes of high purge pressure. the team made the choice to intervene with tissue plasminogen activator / lysis. the patient expired while on support. the patient had needed resuscitation due to deep cardiogenic shock with high dose of catecholamines, lack of volume, and lactate of 12.9. | [5 empty] | The investigation into high purge pressure has been completed. the impella device was not returned for evaluation. as the clinical details mention tissue plasminogen activator (tpa) addition resolved the issue, the root cause of the high purge pressure was most likely biomaterial the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-21566: e4 should have been left blank. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed.
16252 | 1220648-2024-21565 | 20571233 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571233&pc=OZD | 10/01/2024 | 12/12/2024 |  | 10/30/2024 |  |  | 10/01/2024 | 10/30/2024 |  |  |  |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025535444 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male |  |  | White | Death | Yes | No | Fluid/Blood Leak | Material Puncture/Hole | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella cp support. while on support, purge fluid was noted leaking from the red impella plug. pericardial effusion was noted during high-risk percutaneous coronary intervention (hrpci) and was deemed not to be impella-related. patient expired within three hours of admission to unit. | [5 empty] | The investigation into purge leak has been completed. impella cp was returned for evaluation. upon visual inspection, blood was present at seams of red handle. when opened, blood was present in the red handle. a hole at the 85cm mark was also present. data logs were reviewed and a downtrend in motor current at end of lt log with a simultaneous decrease in pump flows while running at p-6. clinical details noted purge fluid was leaking from red handle during support. as it was likely blood leaking from the red handle in the field, the failure mode of "purge leak - pump" was changed to "product damage (hole in catheter)". the root cause of the product damage (hole in catheter) was use related. the following have been reported incorrectly on manufacturer device report 1220648-2024-21565. e4 should have been left blank. g1 reporting contact email. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: section: cleaning. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
16253 | 1220648-2024-21567 | 20571238 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571238&pc=OZD | 11/25/2021 | 02/13/2025 |  | 10/30/2024 |  |  | 11/25/2021 | 10/30/2024 |  |  | 08/31/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022133350 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup |  |  |  | 11/25/2021 |  |  |  |  | The complainant reported a patient noted with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. while on support, pump stopped multiple times. device was replaced with another impella cp pump. no patient harm was reported. | [5 empty] | The investigation was completed. an impella connect screenshot received from the field confirmed elevated motor current and high purge pressure about 3 hours prior to the pump stop. neither the product nor the data logs were returned. after reviewing the clinical information, it was determined that the cause of the pump stop was most likely blood ingress as a result of high purge pressure. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
16254 | 1220648-2024-21568 | 20571242 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571242&pc=OZD | 07/16/2022 | 10/30/2024 |  | 10/30/2024 |  |  | 07/16/2022 | 10/30/2024 |  |  | 02/29/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023197509 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | Increase in Pressure | Deformation Due to Compressive Stress |  |  |  | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that during impella cp support, there was alarming on the aic for low flow and blocked lumen. the decision was made to replace the pump and guidewire. the wire was seen to be damaged at it's tip. the product was replaced but, due to the exchange a hematoma developed at the access site. the hematoma was troubleshot effectively without the need for rbcs. staff later noted impella slightly deep at 4.5 cm. urine has transitioned to cy from red/cherry. the patient remained on 3 of dobutamine, and .2 of nipride. no additional information provided. | [5 empty] | The investigation for the reported issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the root cause of the low pump flow and the purge system blocked alarm was due to biomaterial. however, the cause of the damaged guidewire was not determined since the guidewire was not returned. this report is being filed as part of a retrospective review of historical records.
16255 | 1220648-2024-21569 | 20571252 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571252&pc=OZD | 11/28/2021 | 10/30/2024 |  | 10/30/2024 |  |  | 11/28/2021 | 10/30/2024 |  |  | 09/30/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022136695 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support. while on support, impella cp ingested a clot. device was replaced with another impella cp pump. patient survived the explant procedure but subsequently expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The device was returned and evaluated, and the investigation for the reported thrombus has been completed. evaluation of the returned pump found biomaterial around the impeller. the root cause of the low flow was ingested biomaterial. instructions for use for the related event are as follows: impella cp with smart assist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smart assist system section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
16256 | 1220648-2024-21570 | 20571264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571264&pc=OZD | 07/26/2022 | 10/30/2024 |  | 10/30/2024 |  |  | 07/26/2022 | 10/30/2024 |  |  | 04/30/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218589 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during support, the device had a suction alarm that was unresolved by repositioning or adding fluids. the pump was explanted and replaced, and there was no reported harm to the patient. | [5 empty] | The investigation for the reported low pump flow has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the likely root cause of the low pump flow was due to biomaterial, most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
16257 | 1220648-2024-21571 | 20571285 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571285&pc=OZD | 07/27/2022 | 10/30/2024 |  | 10/30/2024 |  |  | 07/27/2022 | 10/30/2024 |  |  | 12/31/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023175835 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an unknown primary indication was implanted with an impella cp for mechanical circulatory support. during support, the pump failed to get flows over 2 l/min and was eventually showing suction at setting p2. the team suspected that a clot was causing the issue, so the pump was removed and replaced with a second device. there was no reported harm to the patient. | [5 empty] | The investigation for the reported low pump flow has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the data logs were returned and confirm the low flow and suction alarms. multiple spikes in motor current can also be seen. the root cause of the low flow was most likely biomaterial ingestion as the clinical report mentions biomaterial being observed, and the multiple motor current spikes in the logs suggest biomaterial interacting with the impeller. this report is being filed as part of a retrospective review of historical records.
16258 | 1220648-2024-21572 | 20571287 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571287&pc=OZD | 10/22/2021 | 10/30/2024 |  | 10/30/2024 |  |  | 10/17/2024 | 10/30/2024 |  |  | 08/31/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022127439 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Cardiac Arrest | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was needing an impella cp device for mechanical circulatory support. upon insertion of the impella it would not cross the femoral bifurcation due to calcification. unfortunately the patient coded and return of spontaneous circulation (rosc) could not be achieved and the patient expired. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The investigation for the reported delivery issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the root cause of the delivery issues was determined to be patient condition. this report is being filed as part of a retrospective review of historical records.
16259 | 1220648-2024-21573 | 20571329 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571329&pc=OZD | 09/20/2021 | 05/28/2025 |  | 10/30/2024 |  |  | 09/20/2021 | 10/30/2024 |  |  | 02/28/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022050191 |  | 0048-0034-JP |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the complainant reported bleeding at the puncture site. the staff noted that there was arterial bleeding and a blood transfusion was required. the impella was removed and extracorporeal membrane oxygenation was left in place. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | The relationship between impella and the access site bleeding was unable to be determined as no product, logs, or clinical details were provided for this investigation. this report is being filed as part of a retrospective review of historical records.
16260 | 1220648-2024-21574 | 20571336 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571336&pc=OZD | 09/09/2021 | 10/30/2024 |  | 10/30/2024 |  |  | 10/17/2024 | 10/30/2024 |  |  | 05/31/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022090831 |  | 0048-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  |  |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that there was a delivery issues due to kinking of the impella catheter. the patient was noted to have femoral artery tortuosity. impella support was aborted without any harm due to this issue. additional information was provided that noted the patient expired; however, this was caused by a bleed in the opposing femoral artery with no relationship to the impella. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). no patient harm related to the device was identified. | [5 empty] | The investigation for the reported delivery issue (cannula kink) has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the delivery issues was patient condition, specifically tortuous vasculature. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿. ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿. this report is being filed as part of a retrospective review of historical records.
16261 | 1220648-2024-21575 | 20571352 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571352&pc=OZD | 08/28/2022 | 10/30/2024 |  | 10/30/2024 |  |  | 08/28/2022 | 10/30/2024 |  |  | 04/30/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023224434 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support there was access site bleeding noted. upon removal of the peel away sheath, bleeding was noted around the impella site, with manual pressure applied. the impella was weaned and explanted and the site closed with preclose x3 using post closure technique. it was reported that the patient survived the procedure. | [5 empty] | The investigation for the reported bleeding issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the bleeding was most likely acts were 280 seconds which were above recommendation. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
16262 | 1220648-2024-21576 | 20571360 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571360&pc=OZD | 12/09/2022 | 10/30/2024 |  | 10/30/2024 |  |  | 12/09/2022 | 10/30/2024 |  |  | 08/31/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023269843 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, suction alarms were received, and the pump speed level would not run above p2. pump positioning was confirmed as adequate through fluoroscopy. it was stated that due to a lesion in the right coronary artery, it was possible that the suction alarms were caused by right ventricle failure. as a result of the patient¿s age and comorbidities, the patient was not eligible for escalation of care to the impella rp. the patient¿s blood pressure was more stable, and the decision was made to remove the impella cp and medically manage the patient. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
16263 | 1220648-2024-21577 | 20571368 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571368&pc=OZD | 10/12/2022 | 10/30/2024 |  | 10/30/2024 |  |  | 08/09/2024 | 10/30/2024 |  |  | 05/31/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023219738 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Other | Life Threatening | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient died after a subarachnoid hemorrhage. it was stated that the cause of death is difficult to find due to coagulation control. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
16264 | 1220648-2024-21578 | 20571386 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571386&pc=OZD | 03/30/2022 | 03/13/2025 |  | 10/30/2024 |  |  | 03/29/2022 | 10/30/2024 |  |  | 12/31/2022 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022039779 |  | 0048-0034-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced hemolysis and an access site bleed requiring a blood transfusion. no further intervention was reported, but the patient survived their procedures. | [5 empty] | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. as noted in the impella ifu: access site bleeding: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ hemolysis: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records.
16265 | 1220648-2024-21579 | 20571394 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571394&pc=OZD | 12/21/2022 | 10/02/2025 |  | 10/30/2024 |  |  | 12/21/2022 | 10/30/2024 |  |  | 04/30/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023226129 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the pump had low flows and was replaced after troubleshooting measures had failed. no lasting harm or injury to patient was reported. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | This report is being filed as part of a retrospective review of historical records. the investigation of low pump flow has been completed. the impella device was returned for evaluation. data logs show the flows initially as expected. after the physician reinserted the pump, the flows were seen to be around 3.0 l/min max at p8, which is slightly below expected. the pump was inspected in the lab and no abnormalities were found, and the low pump flow was not reproduced. thus, the low pump flow is likely due to patient condition. g1 reporting contact email revised on manufacturer device report 1220648-2024-21579 in accordance with updated procedures.
16266 | 1220648-2024-21580 | 20571402 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571402&pc=OZD | 01/04/2022 | 10/30/2024 |  | 10/30/2024 |  |  | 01/04/2022 | 10/30/2024 |  |  | 09/30/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022146997 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Death | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that an impella cp failed to deliver due to native anatomy limitations. the iliac was intervened upon by plain old balloon angioplasty (poba) but, the pump failed delivery, and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | B2: date of patient death is unknown. the investigation for the reported delivery issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the root cause of the delivery issue was the patient's vasculature. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high-risk pci section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. this report is being filed as part of a retrospective review of historical records.
16267 | 1220648-2024-21581 | 20571413 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571413&pc=OZD | 08/31/2021 | 10/30/2024 |  | 10/30/2024 |  |  | 10/17/2024 | 10/30/2024 |  |  | 05/31/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022089492 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. it was reported that when attempting to use the 14fr peel away sheath, the guidewire was noted to be deformed. additionally, during priming there was a purge cassette leak, the purge cassette was replaced. additional information was provided that noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported issues has been completed. the impella device has not been returned for evaluation. the cause of the product damaged (guidewire damage-peripheral vessels) issue was not determined as no product was returned for analysis. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
16268 | 1220648-2024-21584 | 20571487 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571487&pc=OZD | 06/03/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 06/03/2023 | 10/30/2024 |  |  | 01/31/2025 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2024328669 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  |  |  | Death | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Ventricular Fibrillation | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. during delivery, the long impella sheath was placed in the right femoral artery. then the patient began to decompensate again leading to ventricular tachycardia, multiple shocks were administered and return of spontaneous circulation was achieved. the aortic valve was crossed with the pigtail, and the impella advanced over the 018 wire. the impella pigtail crossed into the left ventricle, but the physician was unable to advance the cannula. the impella was removed, and a new 018 wire was placed in the left ventricle. the impella again advanced over the wire, however the physician was unable to cross the aortic valve with the impella cannula. the impella was once again removed. the patient received multiple rounds of cardio-pulmonary resuscitation with multiple shocks and recovery during this entire process. the patient was also intubated with a temporary pacer and multiple drips. the patient did not survive the procedure. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported delivery issue and ventricular fibrillation has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were sufficient to determine a root cause. the root cause of the delivery issue was determined to be patient condition as the pump was unable to cross the valve and multiple rounds of cpr were performed while attempting to place the device. as noted in the impella ifu: impella cp with smart assist system. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. this report is being filed as part of a retrospective review of historical records.
16269 | 1220648-2024-21587 | 20571496 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571496&pc=OZD | 05/20/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 05/20/2023 | 10/30/2024 |  |  | 12/31/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024321390 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that their impella cp pump experienced multiple pump stops due to high motor current during impella support. additionally patient experienced bleeding and blood products were given. the pump was beyond indication date and was subsequently explanted because of the failures. patient is stable post explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) : ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
16270 | 1220648-2024-21582 | 20571503 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571503&pc=OZD | 07/19/2021 | 10/30/2024 |  | 10/30/2024 |  |  | 07/19/2021 | 10/30/2024 |  |  | 02/28/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022064786 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 62 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock /low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. the complainant reported that due to low heart function, the impella device was introduced at the time of heart-lung machine withdrawal. the purge pressure became 1000 or more and the purge liquid became 1 or less, a red alarm sounded, the purge pressure and purge flow rate gradually increased, therefore a replacement is required. the catheter was successfully replaced. | [5 empty] | The investigation for the reported high purge pressure issue has been completed. the impella device was received from the customer. the high purge pressure and low purge flow was unable to be reproduced. therefore, the root cause of the high purge pressure was unable to be determined. this report is being filed as part of a retrospective review of historical records.
16271 | 1220648-2024-21585 | 20571513 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571513&pc=OZD | 09/17/2022 | 07/09/2025 |  | 10/30/2024 |  |  | 09/17/2022 | 10/30/2024 |  |  | 02/29/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023238343 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. although the impella was placed without issue, optimal flows could not be achieved. the device was repositioned multiple times with no success. during repositioning, the pump was pulled back across the valve, leading to the decision to start over. upon inspection, the pump was found to be twisted and the pigtail bent, prompting the decision to open a new pump. ultimately, the device was explanted and replaced with a new impella cp device. | [5 empty] | The impella device was not received from the customer. no product was returned. data logs showed that recurrent "suction" alarms (including low pulsatility alarms) triggered soon after the initiation of support. the motor current had reduced pulsatility throughout, and pump flow never exceeded an average of 2 l/min. notably, the placement signal waveform trended downwards throughout support. after reviewing the clinical information, it was determined that the cause of the low pump flow was most likely patient condition related. this report is being filed as part of a retrospective review of historical records. | The cause of the positioning issue was not determined since product was not returned. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures 1220648-2024-21585. h10 the results of positioning issues was inadvertently omitted on the initial manufacturer device report number 1220648-2024-21585. h6 type of investigation 4112 was inadvertently omitted on the initial manufacturer device report number 1220648-2024-21585. h6 investigation conclusions 50 was erroneously entered on the initial manufacturer device report number 1220648-2024-21585. h6 investigation findings 114 was inadvertently omitted on the initial manufacturer device report number 1220648-2024-21585.
16272 | 1220648-2024-21586 | 20571521 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571521&pc=OZD | 01/11/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 02/10/2023 | 10/30/2024 |  |  | 08/31/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023268305 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was implanted with an impella cp as an escalation of support from an intra-aortic balloon pump (iabp). prior to the patient's transfer to north shore hospital, the physician was unable to acquire distal pulses via doppler on right lower extremity after the impella placement. it was noted that the patient was on levophed. on arrival to transferred facility, the physician was able to obtain doppler pulses to right lower extremity without any form of intervention to the limb. the patient was made a do not resuscitate and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer and it was determined that the use of the device cannot conclusively be associated with the outcome. the patient's peri-procedural condition was critical. | [5 empty] | The investigation for the reported limb ischemia has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the limb ischemia is most likely due to patient condition since the patient was on levophed. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
16273 | 1220648-2024-21588 | 20571522 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571522&pc=OZD | 07/19/2021 | 10/30/2024 |  | 10/30/2024 |  |  | 07/19/2021 | 10/30/2024 |  |  | 05/31/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022084108 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella cp for mechanical circulatory support. there were lower than expected flows during impella support. the pump was removed and replaced without further issue or any harm to the patient. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was received from the customer and therefore, an evaluation of the device confirmed biomaterial. the root cause of the low pump flow was ingested biomaterial. this report is being filed as part of a retrospective review of historical records.
16274 | 1220648-2024-21583 | 20571529 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571529&pc=OZD | 05/30/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 05/30/2023 | 10/30/2024 |  |  | 02/28/2025 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024331564 |  | 1000080 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. unable to achieve hemostasis at impella cp access site in the right groin. the patient was taken to remove the impella, double perclose deployed. oozing through perclose. pressure held manually. no new impella inserted. the patient survived. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the access site bleeding issue was not determined since product was not returned and with insufficient clinical information. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
16275 | 1220648-2024-21590 | 20571535 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571535&pc=OZD | 04/11/2023 | 02/27/2025 |  | 10/30/2024 |  |  | 04/11/2023 | 10/30/2024 |  |  | 10/31/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024285455 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported a failure to prime the device. as a result, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported priming issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that there is no problem found for the priming issue. dextrose buildup was found on the returned purge cassette. however, if there was a blockage in the purge cassette during case start then hpp would have been found. additionally, the pump ran without issue in the lab with a good purge cassette. this report is being filed as part of a retrospective review of historical records.
16276 | 1220648-2024-21589 | 20571546 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571546&pc=OZD | 06/12/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 06/12/2023 | 10/30/2024 |  |  | 04/30/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218526 |  | 0048-0047 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported that an impella cp was unable to advance through the patient's vasculature due to an existing dissection and calcium in the common iliac. the implant was subsequently aborted as a result of the noted difficulty. no patient harm was reported due to the issue. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. root cause of the delivery issue was patient condition related as due to poor vascular anatomy,
16277 | 1220648-2024-21591 | 20571551 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571551&pc=OZD | 07/26/2021 | 05/23/2025 |  | 10/30/2024 |  |  | 10/16/2024 | 10/30/2024 |  |  | 04/30/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022079763 |  | 0048-0024 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  |  | Asian | Death | Yes | Yes | Pumping Problem | [6 empty] | Sepsis | [12 empty] | Other | Required Intervention | Life Threatening | Death | [21 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support there were concerns that the pump was not working. additional information received indicated the pump was not working due to septic shock caused by sepsis. the pump was replaced. however, despite replacing the pump, the hemodynamics after that became unstable and the patient developed multiple organ failure (mof), resulting in death. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
16278 | 1220648-2024-21592 | 20571553 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571553&pc=OZD | 04/21/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 04/21/2023 | 10/30/2024 |  |  | 09/30/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2023279345 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced access site bleeding, and the physician decided to explant the device. the bleeding was managed by administering 2 percloses. it was later noted that the patient went into vasoplegic shock, white blood cell count was (b)(4) and was likely in mixed picture shock which led to his death hours later. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The investigation for the reported access site bleeding has been completed. the product was returned. per the results of the clinical review and investigation: the returned pump was inspected, and no abnormalities were observed on the pump or the repositioning sheath. a leak test was performed and no water leaked from the valve on the repositioning unit. the cause of the access site bleeding was most likely patient condition related and related to the patient's vessel conditions as well as obesity. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
16279 | 1220648-2024-21593 | 20571561 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571561&pc=OZD | 05/08/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 05/08/2023 | 10/30/2024 |  |  | 01/31/2025 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024311042 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while advancing the pump under fluoroscopy through the sheath into the aorta attending physician noticed a kink on the cannula and made the decision to replace it with a new one due to safety concerns. | [5 empty] | No product was returned for evaluation. the cause of the delivery issue was most likely due to the catheter kink based on the clinical information provided and that patient was unable to stay still. this report is being filed as part of a retrospective review of historical records.
16280 | 1220648-2024-21594 | 20571572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571572&pc=OZD | 11/07/2022 | 10/30/2024 |  | 10/30/2024 |  |  | 11/07/2022 | 10/30/2024 |  |  | 05/31/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023218728 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  |  | Death | Yes | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, and the impella pump was in the left ventricle, there was no flow rate and a flat motor curve. echocardiogram could only reveal a poor image so positioning could not be confirmed. it was reported that the patient¿s abdomen was bulging, and the flow rate did not improve with reposition. the device was removed, and the care was withdrawn and the patient expired after explant of the device. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that returned complaint pump visually inspected. no kink observed. no broken blades observed. no biomaterial or foreign materials observed. impella flow rate looks normal and optical sensor was functioning properly. log checked. high motor current events observed and some position in ventricle alarms. the cause of the issue was improper positioning of the device. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
16281 | 1220648-2024-21595 | 20571611 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571611&pc=OZD | 04/15/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 04/15/2023 | 10/30/2024 |  |  | 03/31/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191389 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported the patient¿s leg was cold to the touch and the toe was discolored. it was stated that bleeding was identified at the impella access site. heparin administration was reduced to manage this complication, and it was reported that the problem was resolved. additional information notes that care was withdrawn, and the procedural outcome was the patient expired at explant. it was noted that the patient death was not groin or bleeding related, but the decision to withdraw care was due to multimorbid patient with beginning multi organ failure. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. based on clinical description, the root cause of limb ischemia was most likely due to patient condition and the root cause of the access site bleeding was due to high patient activated clotting time (act) values. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
16282 | 1220648-2024-21596 | 20571727 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571727&pc=OZD | 01/20/2023 | 03/10/2025 |  | 10/30/2024 |  |  | 01/20/2023 | 10/30/2024 |  |  | 09/30/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023273036 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the impella cp pump pulled back into the aorta during patient support. the physician was unable to re-cross the valve and the decision was made to explant the pump. support continued with an intra-aortic balloon pump (iabp) and extracorporeal membrane oxygenation (ecmo). it was reported that the patient survived the procedure. | [5 empty] | The impella device evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported positioning issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
16283 | 1220648-2024-21597 | 20571734 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571734&pc=OZD | 11/10/2022 | 10/30/2024 |  | 10/30/2024 |  |  | 11/10/2022 | 10/30/2024 |  |  | 07/31/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023259423 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the pump exhibited low flows. the customer tried to resolve the issue, however, the alarms remained for low flow and flat motor current. the patient was successfully weaned from the impella and the device explanted. | [5 empty] | The investigation for the reported low pump flow and pump stop has been completed. the impella device was not received from the customer; however, the data logs were returned. upon review of the logs and clinical information, it was determined the cause of the low pump flow and pump stop was mostly likely ingested biomaterial. instructions for use for the related event are as follows: using the automated impella controller with the impella catheter. ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
16284 | 1220648-2024-21598 | 20571818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571818&pc=OZD | 11/14/2022 | 10/30/2024 |  | 10/30/2024 |  |  | 11/14/2022 | 10/30/2024 |  |  | 07/31/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022114100 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was attempted to be implanted with an impella cp device for mechanical circulatory support. the team attempted to implant the pump via the femoral artery, but stenotic occlusion of the aorta prevented the pump from advancing. this pump was removed, and a second impella device was placed via axillary insertion. there was no reported harm to the patient. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was patient anatomy/condition. as noted in the impella ifu: delivery issues: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
16285 | 1220648-2024-21599 | 20571851 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571851&pc=OZD | 01/27/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 01/27/2023 | 10/30/2024 |  |  | 09/30/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023281060 |  | 0048-0047 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the complainant reported a high motor current alarm that resolved with no intervention. support proceeded unimpacted and it was reported that there was no harm to the patient. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the pump stop was not determined since neither product nor data logs were returned. as noted in the impella ifu: impella cp with smart assist system. section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
16286 | 1220648-2024-21600 | 20571901 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20571901&pc=OZD | 07/28/2021 | 10/30/2024 |  | 10/30/2024 |  |  | 10/17/2024 | 10/30/2024 |  |  | 02/28/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022067878 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support there was a decrease in flow rate; however, there was no problem with the catheter position. the patient was under ecpella management, but there was pulse pressure on the a line. an echocardiogram indicated the left ventricle (lv) was stretched, and was determined not to be assisted by the impella. the impella was removed and a piece of tissue was found in the cannula. additional information was provided that noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported low pump flow issues has been completed. the impella device has not been returned for evaluation. however, clinical details were sufficient to determine the root cause. the root cause of the low pump flow was most likely ingested biomaterial. this report is being filed as part of a retrospective review of historical records.
16287 | 1220648-2024-21601 | 20572039 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20572039&pc=OZD | 08/04/2021 | 10/30/2024 |  | 10/30/2024 |  |  | 10/17/2024 | 10/30/2024 |  |  | 05/31/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022090894 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the impella cp had low pump flows that prompted pump replacement. the 2nd pump was placed and percutaneous coronary intervention (pci) was performed. during the pci the patient required multiple rounds of cardiopulmonary resuscitation and despite efforts, the patient died. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported low pump flow issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the low pump flow was ingested biomaterial. this report is being filed as part of a retrospective review of historical records.
16288 | 1220648-2024-21602 | 20572086 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20572086&pc=OZD | 10/15/2022 | 10/30/2024 |  | 10/30/2024 |  |  | 10/15/2022 | 10/30/2024 |  |  | 02/29/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023202740 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. on day 20 of support, a sudden pump stop occurred. normal troubleshooting attempts to restart the device was unsuccessful. the impella cp was explanted and replaced to address the issue. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of pump stop. the device was not received for evaluation. the root cause of the pump stop was unable to be determined since the product and data logs were not returned. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: troubleshooting the purge system unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
16289 | 1220648-2024-21603 | 20572268 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20572268&pc=OZD | 08/26/2021 | 03/13/2025 |  | 10/30/2024 |  |  | 08/26/2021 | 10/30/2024 |  |  | 06/30/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022107669 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  |  | White | Death | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, a kinked cannula was identified. the physician chose to remove and replace the impella, and there was no reported harm to the patient. survival outcome at explant noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | This complaint has been identified as part of a retrospective review. the cause of the cannula kink was not determined as no clinical information was provided. | The impella device was not returned by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
16290 | 1220648-2024-21604 | 20572452 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20572452&pc=OZD | 10/25/2022 | 10/30/2024 |  | 10/30/2024 |  |  | 10/25/2022 | 10/30/2024 |  |  | 02/29/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023182156 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  | White | Malfunction | No | Yes | Delivered as Unsterile Product | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the physician intended to insert the pump without the use of a guidewire to expedite the procedure. during the insertion, the patient required brief cardiopulmonary resuscitation (cpr). in the haste of the situation, the pump fell off the table and became contaminated. consequently, the decision was made to explant the device and replace it with a new impella cp device. the patient survived the procedure. | [5 empty] | The investigation has been completed for the reported issue of packaging issues- sterility. the device was received for evaluation. the root cause of the packaging issues- sterility was use related since the pump was dropped and became non-sterile. this report is being filed as part of a retrospective review of historical records.
16291 | 1220648-2024-21605 | 20573115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20573115&pc=PYX | 05/07/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 05/07/2023 | 10/30/2024 |  |  | 01/31/2025 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024310605 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella rp device for mechanical circulatory support. while on support, that the pump stopped and would not restart. the pump was explanted as a result of the noted failure. there was no patient harm. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the product was returned with the catheter cut. the motor and cannula were not returned, so limited analysis could be completed. the data logs were reviewed and show a high motor current pump stop on day 22 of support. the pump stop is sudden with slight irregularities in mc in the minute leading up to it. the cause of the high motor current pump stop was unable to be determined as insufficient product was returned for analysis. impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
16292 | 1220648-2024-21606 | 20575247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20575247&pc=PYX | 12/15/2021 | 10/30/2024 |  | 10/30/2024 |  |  | 12/15/2021 | 10/30/2024 |  |  | 03/31/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022070507 |  | 004334 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  |  | Death | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the rp catheter kinked at multiple sites when inserting over 0.027. additional information noted that the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported issue has been completed. visual inspection of the returned pump revealed two significant kinks along the catheter. the kinks were located near the motor housing and were about 10 cm apart from each other. no other damage or abnormalities were found. after reviewing the clinical information, it was determined that the cause of the delivery issues was catheter kinking due to patient condition, specifically the condition of the patient's pulmonary arteries. this file is being submitted as part of a retrospective review of historical records after which medical safety has updated the report type.
16293 | 1220648-2024-21608 | 20575329 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20575329&pc=PYX | 04/25/2022 | 10/30/2024 |  | 10/30/2024 |  |  | 04/25/2022 | 10/30/2024 |  |  | 02/29/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023193594 |  | 0046-0011 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  |  | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support there was low purge flow with high purge pressure during use of the device. additional information was provided that noted care was withdrawn and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported high purge pressure has been completed. however, as the necessary clinical information was not provided and the device was not returned, the root cause of the issue could not be determined. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
16294 | 1220648-2024-21609 | 20575365 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20575365&pc=PYX | 06/10/2022 | 10/30/2024 |  | 10/30/2024 |  |  | 06/10/2022 | 10/30/2024 |  |  | 03/31/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023204616 |  | 004334 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Hispanic |  | Death | Yes | Yes | Positioning Failure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. while on support the device lost optimal position. after multiple attempts were made to reposition the device, the decision was made to replace the device with another impella rp device. additional information was provided that noted care was withdrawn and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported positioning issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, information provided was sufficient to determine a root cause. the root cause of the delivery issues was patient movement. this report is being filed as part of a retrospective review of historical records.
16295 | 1220648-2024-21610 | 20575366 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20575366&pc=PYX | 12/15/2021 | 10/30/2024 |  | 10/30/2024 |  |  | 12/15/2021 | 10/30/2024 |  |  | 01/31/2022 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 1468332 |  | 004334 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Hemodynamic instability | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella rp for mechanical circulatory support. multiple rp catheters kinked during insertion and failed delivery. the procedure outcome is the patient passed away and there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The returned pump was visually inspected and multiple kinks in the catheter were confirmed. the root cause of the reported delivery issues was determined to be the observed catheter kinks. refer to es2020-080 for further information. based on the provided information, the catheter kinks were most likely caused by excessive force during pump insertion. as noted in the impella ifu: impella rp smartassist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
16296 | 1220648-2024-21611 | 20575376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20575376&pc=PYX | 01/17/2022 | 10/30/2024 |  | 10/30/2024 |  |  | 01/17/2022 | 10/30/2024 |  |  | 11/30/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022147048 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission |  |  |  |  | 01/17/2022 |  |  |  |  | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. after multiple attempts made to advance the rp, the cannula would buckle and would not advance past the tricuspid valve. md noted a heavier/stiffer wire would help tremendously. additional information noted that care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. this report is being filed as part of a retrospective review of historical records. | [5 empty] | The investigation into the reported issue has been completed. the device was not received for evaluation; however, clinical information was sufficient to determine the root cause. the cause of the cannula kink during insertion was most likely small vessels since the patient was on inotropes/vasopressors.
16297 | 1220648-2024-21607 | 20575378 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20575378&pc=PYX | 01/14/2022 | 10/30/2024 |  | 10/30/2024 |  |  | 01/14/2022 | 10/30/2024 |  |  | 01/31/2022 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 1468728 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. additional information noted the patient expired. the pump reportedly stopped. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation into the reported issue has been completed. neither the product nor the data logs were returned. therefore, the cause of the pump stop was not determined. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high-risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
16298 | 1220648-2024-21612 | 20575417 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20575417&pc=PYX | 06/17/2022 | 10/30/2024 |  | 10/30/2024 |  |  | 06/17/2022 | 10/30/2024 |  |  | 03/31/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022074902 |  | 0046-0011 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | Yes | Pumping Stopped | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. while on support there was a pump stop when the physician increased the p-level above p2. the pump was repositioned. additional information was provided that noted care was withdrawn and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | Date of patient death (b2) and explant date (d6) are unknown. the investigation for the reported pump stop has been completed. however, as the necessary clinical information was not provided and the device was not returned, the root cause of the pump stop could not be determined. as noted in the impella ifu: impella rp with smart assist system. section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
16299 | 1220648-2024-21613 | 20575471 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20575471&pc=PYX | 12/19/2021 | 10/30/2024 |  | 10/30/2024 |  |  | 12/19/2021 | 10/30/2024 |  |  | 10/31/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2022143582 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Decrease in Pressure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 12/19/2021 |  |  |  |  | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. few hours after insertion low purge pressure was observed. multiple attempts were made by adjusting the fluid to address the issue, to no success. the team did not replace the pump due to the family withdrew care. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer. the root cause of the low purge was unable to be determined as no product or data was returned for analysis. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella rp with smartassist system section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
16300 | 1220648-2024-21614 | 20575473 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20575473&pc=PYX | 04/01/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 04/20/2023 | 10/30/2024 |  |  | 01/31/2025 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024310633 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella rp for mechanical circulatory support. during impella placement, the impell pump was unable to be delivered due to the patient's vasculature. the procedural outcome was the patient expired on support. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The root cause of the delivery issue was most likely patient condition, as the pump was able to be inserted on a second attempt after a flow directed catheter was reinserted to redirect the wire into the rpa. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿
16301 | 1220648-2024-21615 | 20575484 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20575484&pc=PYX | 04/15/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 04/15/2023 | 10/30/2024 |  |  | 01/31/2025 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024305329 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission |  |  |  |  | 04/15/2023 |  |  |  |  | The complainant reported a patient was implanted with an impella rp post cardiac surgery. the complainant reported that during impella placement, the device was unable to achieve a flow greater than 1.6 liters per minute. it was reported that the pump was not replaced due to this issue and the case continued. the procedural outcome was the patient survived explant, but eventually expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The data logs were returned and a high motor current spike, speed deviation, and drop in flow were observed around the same time the clinical details specify that vaecmo was added. the product was returned and biomaterial was confirmed to be around the impeller. the root cause of the low flow was determined to be ingested biomaterial. the biomaterial was likely released from the dialysis cath and into the patient when vaecmo was inserted. the biomaterial was able to form as the patient had sub therapeutic anticoagulation for the majority of the case. this report is being filed as part of a retrospective review of historical records.
16302 | 1220648-2024-21616 | 20575501 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20575501&pc=PYX | 01/01/2022 | 10/30/2024 |  | 10/30/2024 |  |  | 01/01/2022 | 10/30/2024 |  |  | 02/28/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022064744 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device as additional bipella support for mechanical circulatory support. upon insertion of the rp, the pump initially detected a purge pressure high alarm followed by a purge system blocked alarm. appropriate trouble shooting steps were performed (i.e. switching to a new purge cassette) however the problem persisted indicating the purge lumen of the rp was blocked. the physician insisted the pump be replaced since the patient was unstable. the replacement impella rp pump was successfully placed with the same cassette and had no issues. | [5 empty] | An investigation into the high purge pressure issue has been completed. the impella rp was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were insufficient to determine the root cause. additionally, neither the product nor data logs were returned for analysis. the cause of the high purge pressure was not determined since neither the product, nor the data logs were returned. impella rp instructions for use for the related event are as follows: impella rp with smart assist system. section: using the automated impella controller with the impella rp with smart assist system catheter as noted in the impella ifu ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smart assist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smart assist system catheter whenever a rise in motor current is seen.¿ this report is being filed as part of a retrospective review of historical records.
16303 | 1220648-2024-21617 | 20575503 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20575503&pc=PYX | 02/01/2022 | 10/30/2024 |  | 10/30/2024 |  |  | 02/01/2022 | 10/30/2024 |  |  | 02/28/2022 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 1469832 |  | 004334 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. while on support, device experienced delivery issues where the device could not past the pulmonic valve, despite multiple attempts. the pump was removed and a second rp pump was used. there was no harm to the patient. | [5 empty] | The investigation into the delivery issue has been completed. the impella rp was returned for investigation. the returned product was visually inspected and a kink in the catheter was confirmed adjacent to the motor housing. the root cause of the reported delivery issues was determined to be the observed catheter kink. based on the provided information, the catheter kink was most likely caused by excessive force during pump insertion due to improper technique by the end user. as noted in the impella ifu: impella rp smartassist system with automated impella controller. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
16304 | 1220648-2024-21618 | 20575510 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20575510&pc=PYX | 05/11/2023 | 10/27/2025 |  | 10/30/2024 |  |  | 05/11/2023 | 10/30/2024 |  |  | 08/31/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024297024 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Hispanic |  | Death | Yes | Yes | Restricted Flow rate | Failure to Sense | Device Sensing Problem |  |  |  |  | Unspecified Vascular Problem | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with right heart failure was implanted with an impella rp. the device was placed and supported the patient for 6 days with reported pump flow and placement signal issue. the team then withdrew care, and the patient expired. no allegation that the rp caused or contributed to the death, but there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was returned and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the pump was returned for investigation. the pump passed the laser continuity test and optical bench test. all electrical values were within spec. there were no catheter or cannula kinks, no biomaterial on sensor, or blood ingress in the purge line. some biomaterial was found around the impeller blades however, due to which the pump was unable to be run in the lab as it triggered motor current high alarms. data logs were returned for analysis. data log analysis confirmed periods of negative central venous pressure (cvp0. the optical parameters were checked in the logs and the sensor appears to be fully functional. the clinical details mention that the placement signal issues begin when extra-corporeal membrane oxygenation is initiated. this likely causes a local suction condition that applies negative pressure to the sensor. a motor current spike was seen towards the end with rise in purge pressures which might point towards biomaterial ingestion, however, sustained low pump flows for the same p-level were seen earlier in the case, which aligned with the cvp dropping to negative values, pointing towards patient condition when considering that the patient was on other support devices simultaneously based on clinical data. no suction alarms were triggered. placement signal was stable initially, then started trending upwards, and became unstable as cvp dropped to negative values. no placement signal not reliable alarms were triggered. optical signal issue: the root cause of the optical signal issue was most likely patient condition based on data log analysis, product evaluation, and clinical details. low pump flow: the root cause of the low pump flow was most likely patient condition based on data log analysis, product evaluation, and clinical details. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated to include placement signal coding b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated.
16305 | 1220648-2024-21619 | 20575547 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20575547&pc=PYX | 07/13/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 07/13/2023 | 10/30/2024 |  |  | 02/29/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023184786 |  | 004334 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  |  |  | Death | Yes | Yes | Increase in Pressure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that they encountered a purge pressure high/purge pressure blocked alarm during impella support. the purge cassette was replaced to initially resolve the alarms; however, the issue returned a short time after. two mg tpa and 50 cc of water were hung as purge fluid and the alarms were again resolved. additional information noted the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation into the reported issue has been completed. no device was received for evaluation. the cause of the high purge pressure issue was most likely biomaterial based on provided clinical information and data log review. as noted in the impella ifu: impella rp with smartassist system section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ this report is being filed as part of a retrospective review of historical records.
16306 | 1220648-2024-21620 | 20575560 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20575560&pc=PYX | 02/16/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 02/16/2023 | 10/30/2024 |  |  | 12/31/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED |  |  | US |  | 2024295468 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  |  | Death | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient was very unstable, already had a cp in place. however, the patient acutely decompensated and was brought to ccl. md stated that we will try to place an rp however the patient's heart was not contracting well, and he knew it may be very difficult to place. when attempting to place device the rp could not cross. pump kept getting stuck in the right ventricle and unable to advance forward. after several attempts, the decision was made to abort the insertion. additional information the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation of the reported issue was completed. no device was received for evaluation; however, clinical information was sufficient to determine the root cause of the issue. the root cause of the delivery issue was determined to be patient condition as the patient's heart was not contracting well making the insertion past the pv difficult. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: potential adverse events (united states): ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings: ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
16307 | 1220648-2024-21621 | 20575570 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20575570&pc=PYX | 10/31/2022 | 10/30/2024 |  | 10/30/2024 |  |  | 10/31/2022 | 10/30/2024 |  |  | 04/30/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023252915 |  | 0046-0011 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Increase in Pressure | [6 empty] | Renal Failure | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. complainant reported the bipella ran in the hospital without heparin. this led to the rp pump showing the error message purge flow low alarm. switched to nabic as purge solution. pf increased again. due to the multiple organ failure, the physicians and family decided to withdraw from therapy. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | Investigation into the reported issue has been completed. no data logs were sent. visual inspection of the returned pump revealed the presence of biomaterial on and around the impella as well as within the purge gap. no other abnormalities were found. therefore, it was determined that the cause of the high purge pressure was biomaterial occluding the purge gap. the lack of heparin in the purge most likely led to the biomaterial buildup. as noted in the impella ifu: impella rp with smartassist system section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ this report is being filed as part of a retrospective review of historical records.
16308 | 1220648-2024-21622 | 20575583 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20575583&pc=PYX | 01/24/2023 | 05/28/2025 |  | 10/30/2024 |  |  | 01/24/2023 | 10/30/2024 |  |  | 07/31/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023276898 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the impella rp flex experienced low flows 30 minutes after pump start. no systemic anticoagulation was used. heparin was being used in the purge. patient arrested and expired on the second day of support. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. | No product was returned for evaluation. the data logs were evaluated and 30 minutes after case start a motor current spike, speed spike, and dp sensor increase/flow decrease was observed. this pattern is indicative of biomaterial ingestion. low flow followed this event. the root cause of the low flow was most likely biomaterial ingestion.
16309 | 1220648-2024-21623 | 20575601 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20575601&pc=OZD | 09/28/2021 | 10/30/2024 |  | 10/30/2024 |  |  | 09/28/2021 | 10/30/2024 |  |  | 05/31/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022088570 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella cp device for mechanical circulatory support. alarm on screen "impella stopped retrograde flow" without completing any priming or insertion of pump. console restarted & second attempt made, attempted to go into "case start" and then manually priming pump through purge menu, third attempt again unsuccessful with new aic. case aborted and a new pump opened and successfully inserted without difficulty. there was no reported harm or injury to the patient. | [5 empty] | The investigation for the reported case start issue has been completed. the product was returned and the issue was confirmed. the root cause of the case start issue was most likely the pump being started prematurely prior to placement. this report is being filed as part of a retrospective review of historical records.
16310 | 1220648-2024-21624 | 20575604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20575604&pc=PYX | 08/05/2022 | 10/30/2024 |  | 10/30/2024 |  |  | 08/05/2022 | 10/30/2024 |  |  | 12/31/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022154349 |  | 004334 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 25 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. impella rp placement was attempted. pt had a large rv and insertion was attempted over a 0.35¿ guidewire. guidewire was removed and pump start attempted. motor current high alarm immediately occurred, and pump was unable to be started. the rp was removed and a second rp was attempted to be placed however decision was made to stop all efforts due to prolonged coding time. the act was unknown. additional information noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | Investigation into the reported issue has been completed. the logs were reviewed and confirm the pump was unable to start and produced high motor current pump stops. the product was returned, and no biomaterial or kinks were found. the pump ran without issue and had adequate flows. the root cause of the pump being unable to start was unable to be determined as the issue was unable to be reproduced. this report is being filed as part of a retrospective review of historical records.
16311 | 1220648-2024-21625 | 20575623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20575623&pc=PYX | 03/02/2022 | 10/30/2024 |  | 10/30/2024 |  |  | 03/02/2022 | 10/30/2024 |  |  | 07/31/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022131581 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Other | Death | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported a pump stop during use of the device. a micro-puncture wire that was left in the patient was suspected as being ingested into the impella. additional information noted the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | Investigation of the reported issue has been completed. the product was returned and scratches and idents were found around the hub and the inflow cage. the data logs confirm that a sudden high motor current pump stop occurred. the clinical details suggest that the foreign wire was ingested into the rpsa. the root cause of the pump stop was determined to be interaction with the micro puncture guidewire. as noted in the impella ifu: impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
16312 | 1220648-2024-21626 | 20575625 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20575625&pc=OZD | 02/06/2023 | 10/30/2024 |  | 10/30/2024 |  |  | 02/06/2023 | 10/30/2024 |  |  | 11/30/2024 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024287208 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an acute patient undergoing cardiac surgery was implanted with an impella cp device for mechanical circulatory support. while inserting the impella, the md was unable to implant the device due to the patient's anatomy. there was no patient harm reported. | [5 empty] | The investigation into the reported issue was completed. no device was received for evaluation; however, clinical information was sufficient to determine the root cause of the issue. the root cause of the delivery issue was patient condition related to patient anatomy. as noted in the impella ifu: impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿. section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿. ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿. section: warnings & cautions: warnings. ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿. section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿. this report is being filed as part of a retrospective review of historical records.
16313 | 1220648-2024-21627 | 20575635 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20575635&pc=PYX | 03/05/2022 | 10/30/2024 |  | 10/30/2024 |  |  | 03/05/2022 | 10/30/2024 |  |  | 03/31/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2022081523 |  | 004334 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a failed rp delivery. support with this device was aborted. there was no harm to the patient resulting from this issue. | [5 empty] | Device was returned: - the cause of the failure mode was patient condition related. instructions for use for the related event are as follows: section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
16314 | 1220648-2024-21628 | 20575651 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20575651&pc=PYX | 07/14/2022 | 10/30/2024 |  | 10/30/2024 |  |  | 07/14/2022 | 10/30/2024 |  |  | 09/30/2023 |  | 10/30/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023193947 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella rp device for mechanical circulatory support. during insertion, the physician was unable to place the rp into the appropriate position. the pump would not fully advance due to native anatomy. additional information was provided that noted the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, information provided was sufficient to determine a root cause. the root cause of the delivery issues was patient condition related. as noted in the impella ifu: impella rp with smart assist system. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ this report is being filed as part of a retrospective review of historical records.
16315 | 1220648-2024-21389 | 20555065 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20555065&pc=OZD | 10/03/2024 | 12/09/2024 |  | 10/29/2024 |  |  | 10/04/2024 | 10/29/2024 |  |  |  |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025537478 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that when trying to reposition the impella 5.5, left ventricle perforation occurred. the impella was removed, and the ventricle was repaired. the patient survived the explant. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation into ventricle perforation has been completed. the device was not returned for evaluation. however, clinical details provided were sufficient to determine the most likely root cause. the root cause of the ventricle perforation was most likely use-related as evidenced by clinical detail of failure occurring during repositioning. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-21389.
16316 | 1220648-2024-21390 | 20555119 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20555119&pc=OZD | 10/05/2024 | 11/27/2024 |  | 10/29/2024 |  |  | 10/05/2024 | 10/29/2024 |  |  |  |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025561564 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient developed a hematoma at their access site following impella cp implant. it was noted that multiple attempts were necessary to place antegrade sheaths at the ecpella sites immediately prior to the formation of the hematoma. manual pressure was applied to the sites and three units of blood were given to treat the condition. the impella site was reported to be stable and support with the pump continued. the patient was successfully weaned from the impella four days later and the patient survived the explant. | [5 empty] | The investigation into access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was most likely patient condition since hematoma was observed on multiple sites. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16317 | 1220648-2024-21391 | 20555224 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20555224&pc=OZD | 10/05/2024 | 12/09/2024 |  | 10/29/2024 |  |  | 10/05/2024 | 10/29/2024 |  |  |  |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025551711 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, the patient had hemolysis. the staff reduced the p level and monitored. it is unknown if the hemolysis resolved. additionally, the patient had a retroperitoneal bleed the next day which caused a five gram drop requiring multiple blood products. when initially inserted, the peel away sheath was left in place. the care team thought that the bleed came from patient moving around a lot causing no flow to impella extremity and bleed. the impella was removed at bedside and peel away sheath was removed in catheterization lab for a controlled environment. at the end of case, the patient had no pulses in bilateral extremities and vascular surgery was consulted for intervention. the procedural outcome was the patient survived the explant. | [5 empty] | The investigation into vascular damage - peripheral vessel, limb ischemia, and hemolysis has been completed. returned complaint device visually inspected. no biomaterial observed. no broken blade or sharp edge observed. data log reviewed: many suction alarms occurred. no other indication observed confirm the cause of reported hemolysis issue. dropping the cvp from 8 to 6 at the time of the hemolysis issue occurred mostly related to patient condition with many suction alarms too. the cause of the hemolysis issue most likely was patient condition related. the cause of the vascular damage - peripheral vessel cannot be determined since insufficient clinical data provided. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-21391. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
16318 | 1220648-2024-21392 | 20555923 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20555923&pc=OZD | 10/07/2024 | 11/21/2024 |  | 10/29/2024 |  |  | 10/07/2024 | 10/29/2024 |  |  |  |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532409 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump placed for a patient admitted in with acute myocardial infarction/cardiogenic shock. the pump was placed for three days when the pump had low/no flows. the team observed this after the pump was repositioned at the bedside. the pump was explanted and replaced by a new impella cp and support proceeded. the patient survived the explant. | [5 empty] | The investigation for the impella cp has been completed. the device was not returned for evaluation. however, clinical details provided were sufficient to determine the most likely root cause. the cause of low pump flow issue most likely was improper positioning due to improper technique. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
16319 | 1220648-2024-21393 | 20556067 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20556067&pc=OZD | 09/02/2024 | 12/10/2024 |  | 10/29/2024 |  |  | 10/07/2024 | 10/29/2024 |  |  |  |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025527400 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured second degree atrioventricular (av) block, mobitz type 1 with a possible relationship to the impella device used. in post catheterization recovery period, bradycardia develop with telemetry evidence of av block second degree mobitz type 1 with subsequent mild hypotension. a coincidental hematoma at impella sheath insertion site was discovered at this time as well. dopamine started, sheath removed & manual pressure held for 70 minutes. rhythm and blood pressure were stabilized with dopamine just prior to sheath removal. dopamine was weaned off after hemostasis was achieved. av block was present at the time of disconnect. it was reported that the patient/event has not recovered/not resolved. the patient survived the explant. | [5 empty] | The investigation for the arrhythmia has been completed. neither the device nor clinical details were returned for evaluation. therefore, the root cause of the arrhythmia could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
16320 | 1220648-2024-21394 | 20556162 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20556162&pc=OZD | 10/01/2024 | 12/02/2024 |  | 10/29/2024 |  |  | 10/01/2024 | 10/29/2024 |  |  |  |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025549837 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella cp to support a 38-year-old male patient admitted in for a high risk coronary intervention (hrpci). the pump was placed but, the pump flows were low and so the team replaced the pump. there was a presumed ingested clot. | [5 empty] | The impella device was received for evaluation. upon investigation completion, a supplemental mdr will be filed. | The investigation into low pump flow has been completed. the data logs confirmed the failure mode & clinical narrative. logs showed after pump started, motor current spiked followed low flow that triggered auto suction response & suction alarms. this event remained to the rest of the case. visual inspection found biomaterial inside the pump housing and wrapped around the impeller. no other issues were found on the rest of the pump. the root cause of the low flow was biomaterial ingestion.
16321 | 1220648-2024-21395 | 20556292 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20556292&pc=OZD | 10/07/2024 | 12/06/2024 |  | 10/29/2024 |  |  | 10/07/2024 | 10/29/2024 |  |  |  |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025542821 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that impella 5.5 was not able to be delivered despite several troubleshooting attempts. the patient had tortuous anatomy. the 5.5 was removed and impella cp was place successfully. there was no harm to the patient. | [5 empty] | The investigation into delivery issues has been completed. the impella device was not returned for evaluation. the root cause of the delivery issue was most likely patient condition since the implant was aborted due to patient's anatomy as tortuosity of the aortic arch was present. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-21395. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
16322 | 1220648-2024-21396 | 20556399 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20556399&pc=OZD | 10/06/2024 | 08/22/2025 |  | 10/29/2024 |  |  | 10/06/2024 | 10/29/2024 |  |  |  |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025486962 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump for the support of a surgical patient. the procedure was high risk. the pump was supporting during a neurological event. the cerebral infarct and causal link to the impella could not be made. neurologically the patient is still not fully recovered however the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | G1 reporting contact email revised on manufacturer device report in accordance with updated procedures. h6 investigation conclusions revised on manufacturer device report in accordance with updated procedures.
16323 | 1220648-2024-21398 | 20556412 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20556412&pc=OZD | 10/07/2024 | 12/09/2024 |  | 10/29/2024 |  |  | 10/07/2024 | 10/29/2024 |  |  |  |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025524381 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that during preparation for impella support, the tech noted that the 7 fr companion sheath was compressed in the middle and bent at the end. the sheath was replaced and the procedure was completed successfully. there was no noted patient harm. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation into introducer damage ¿ mechanical has been completed. the cause of the issue was not determined since product was not returned. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-21398.
16324 | 1220648-2024-21397 | 20556426 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20556426&pc=OZD | 10/06/2024 | 12/02/2024 |  | 10/29/2024 |  |  | 10/06/2024 | 10/29/2024 |  |  |  |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025547073 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support. while on support, the patient had runs of ventricular tachycardias (vt) that require shock. pressors and inotropes remained off and the patient remained on anti-arrythmic drops for continued episodes of vt. additionally it was reported that there was a thrombus noted. an echocardiogram was done, pump appears in good position, no suction or positioning alarms. treatment of the thrombus is unknown however the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can led to systemic embolization, serious injury, or death.". in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation for the reported ventricular tachycardia (vt) and thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of vt and thrombus could not be determined as the device was not returned nor sufficient clinical details.
16325 | 1220648-2024-21400 | 20556432 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20556432&pc=OZD | 10/07/2024 | 12/09/2024 |  | 10/29/2024 |  |  | 10/07/2024 | 10/29/2024 |  |  |  |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025529012 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient developed a pocket hematoma at their axillary site during impella 5.5 support. systemic anticoagulation was put on temporary hold to assess / manage the condition. support remains ongoing with the implanted pump therefore explant date and patient outcome are unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation into access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-21400. d6b missing. e4 should have been left blank. g1 reporting contact email.
16326 | 1220648-2024-21399 | 20556433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20556433&pc=OZD | 10/07/2024 | 12/23/2024 |  | 10/29/2024 |  |  | 10/07/2024 | 10/29/2024 |  |  |  |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025518979 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A complainant reported the patient was on impella cp support. while on support, oozing was noted at the access site. femostop applied at 40mmhg above the access site. bleeding in the groin stopped, but hemoglobin (hb) was at 5 requiring a transfusion. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined.
16327 | 1220648-2024-21401 | 20556439 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20556439&pc=OZD | 10/07/2024 | 12/10/2024 |  | 10/29/2024 |  |  | 10/07/2024 | 10/29/2024 |  |  |  |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025534985 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a patient was on impella cp support. while on support, droplets of blood were noted around impella catheter while inside peel-away sheath. the patient had large hematomas at both groin sites. right femoral artery site, which was accessed but not used for impella, was perclosed x 1; hematoma pushed out with manual pressure. left femoral artery site, which was the impella access site, with large hematoma as well, manual pressure applied to site. the concern for further bleeding and patient brought back to cath lab to assess for potential retro-peritoneal bleeding with angiogram of peripherals. in meantime, patient received four units packed red blood cells, three liter normal saline, one bottle albumin, and two fresh frozen plasma. the decision was made to remove the impella and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation into introducer damage ¿ mechanical has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was most likely patient condition related, there was multiple site bleeding observed as per the clinical description provided. the following have been reported incorrectly on manufacturer device report 1220648-2024-21401. e4 should have been left blank. g1 reporting contact email. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16328 | 1220648-2024-21402 | 20556443 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20556443&pc=OZD | 10/08/2024 | 12/10/2024 |  | 10/29/2024 |  |  | 10/08/2024 | 10/29/2024 |  |  |  |  | 10/29/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1209418 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that when the automated impella controller (aic) was turned on, an impella defective alarm occurred. the aic was removed. no patient involvement. | [5 empty] | The investigation for the console "pump not detected" issue has been completed. data logs were reviewed which confirm an impella defective alarm was issued on the date of occurrence. an intermittent impella defective alarm was observed during investigation in the lab without a pump being attached to the console, complicating using isolation for troubleshooting. the pump was returned for evaluation. inspection of internal circuitry revealed no anomaly. despite power supply being adequate, the impellatronics board is suspected and will be replaced as a likely defective component. the root cause for the impellatronic defective alarm is most likely a defective impellatronics board. corrections to the initial mfr report: g1 mdr reporting contact email. h6 impact code 925 removed. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
16329 | 1220648-2024-21403 | 20556455 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20556455&pc=OZD | 09/25/2024 | 12/11/2024 |  | 10/29/2024 |  |  | 10/08/2024 | 10/29/2024 |  |  |  |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025545418 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Hemolysis | Mitral Perforation | Mitral Valve Insufficiency/ Regurgitation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support and during support, the doctor expressed concern that the impella placement potentially caused severe mitral regurgitation related to mitral valve tear, but they stated that the team was not sure of the etiology. the patient also had hemolysis while on support. additionally, there were suction alarms. the pump was explanted. the patient survived the event. | [5 empty] | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: impella cp with smartassist catheter insertion ¿impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve¿ section: use of echocardiography for positioning of the impella catheter impella catheter inlet to the aorta ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation for the reported low pump flow, heart valve damage, positioning issue, and hemolysis has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the low pump flow, heart valve damage, positioning issue, and hemolysis could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email.
16330 | 1035166-2024-00119 | 20556574 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20556574&pc=DYB | 10/01/2024 | 10/29/2024 |  | 10/29/2024 |  |  | 10/03/2024 | 10/29/2024 |  |  | 05/01/2024 |  | 10/29/2024 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3006 |  |  | OSCOR INC. |  |  | US |  | DOR9061174 |  | 0052-3006 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [5 empty] | 74 YR |  | Female |  |  | White | Malfunction | No | Yes | Fluid/Blood Leak | Material Integrity Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Patient's condition when presented to the physician was cardiomyopathy; patient is a 74-year-old, white female. during the procedure when the physician was inserting the introducer into the 8mm graft; the sheath started leaking, and it was found that the tip was split. the device was not used in a tortuous path when event occurred. the clinical representative does not know the brand name, model or lot number of the second introducer that was used to successfully complete the procedure. there was no medical intervention, no additional surgery, and no procedure delay. additional devices used during the procedure are an impella 5.5 with smartassist s2 set, us (1000100) lot number 2025547602, serial number (b)(6). no patient injury reported. the device is not available for evaluation. | [5 empty] | No product was provided for evaluation. procedure was completed successfully with another device. there was no medical intervention, no additional surgery, and no procedure delay. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. without the return of the actual device in question for evaluation, oscor inc. is unable to determine the exact cause for this incident. at this time, it is not possible to assign a definitive root cause for the event as reported. the complaint is non-verifiable as the product was not returned for evaluation. in conclusion, based on the information available at this time it cannot be confirmed that the product failed to meet requirements. the event will be re-evaluated if the device is returned for evaluation or additional information becomes available. oscor will continue to monitor this event type and risk. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
16331 | 1220648-2024-21404 | 20556778 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20556778&pc=OZD | 10/08/2024 | 12/11/2024 |  | 10/29/2024 |  |  | 10/08/2024 | 10/29/2024 |  |  |  |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025517797 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | Material Twisted/Bent |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRA CORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported the patient expired after impella cp explant. per medical safety officer review use of the device cannot conclusively be associated with the outcome. | The complainant reported that during the impella cp insertion the cannula got kinked under the motor and it was unable to be straighten out. flow was not able to be turned up without getting suction. the cp was eventually explanted and exchanged for an impella 5.5 while on extra corporeal membrane oxygenation support. the patient survived the event. |  |  |  |  | The investigation into low pump flow and product damage (cannula kink) has been completed. the pump was not returned for investigation. the data log shows low pump flow from the start of the case with an active suction alarm. no abnormalities were reported regarding biomaterial interaction or positioning issues. according to the clinical report, a cannula kink was noted under imaging, and the doctor was unable to straighten it out. the doctor decided to replace the pump. the cause of the low pump flow issue was most likely kinked cannula. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-21404. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
16332 | 1220648-2024-21405 | 20556786 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20556786&pc=OZD | 10/07/2024 | 12/09/2024 |  | 10/29/2024 |  |  | 10/07/2024 | 10/29/2024 |  |  |  |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025490940 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump placed to support a patient admitted in with a st elevation myocardial infarction. there was an access site bleed the day after insertion. the team troubleshot the issue with stitching, injecting lidocaine, and two units of red blood cells. these medical interventions and manual hold remedied the issue. the impella cp was exhanged for an impella 5.5 the next day. the patient survived the explant. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation into access site bleeding has been completed. the device was not returned for evaluation. however, clinical details provided were sufficient to determine the most likely root cause. the root cause of access site bleeding is determined to be use issue because the bleeding was resolved by placing a two stitches around the sheath. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-21405.
16333 | 1220648-2024-21406 | 20557163 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20557163&pc=OZD | 10/08/2024 | 12/11/2024 |  | 10/29/2024 |  |  | 10/08/2024 | 10/29/2024 |  |  |  |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025520942 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, the patient developed hematoma. manual pressure was held and purse string suture was placed around the repositioning sheath. heparin was paused. hematoma was resolved with manual pressure, and the patient survived the event. | [5 empty] | The investigation for the reported access site bleed has been completed. the device was not returned for evaluation. however, clinical information was sufficient to determine the root cause. the root cause of access site bleeding is determined to be use issue because the hemostasis was achieved by placing a purse string around the repo sheath. corrections to the initial mfr report: g1 mdr reporting contact email. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16334 | 1220648-2024-21407 | 20557221 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20557221&pc=OZD | 06/18/2023 | 10/29/2024 |  | 10/29/2024 |  |  | 06/18/2023 | 10/29/2024 |  |  | 12/31/2024 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024302457 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | No | Yes | Restricted Flow rate | [6 empty] | Appropriate Term / Code Not Available | [37 empty] | A |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support there was low pump flow. due to constant suction alarms and multiple repositioning attempts, the doctor questioned catheter integrity. the patient was taken to the cath lab to exchange for a new catheter. additional information was provided that noted the patient expired. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The investigation for the reported low pump flow has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical information provided was sufficient to determine a root cause. the root cause of the low pump flow issue was most likely due to the patient's condition. this report is being filed as part of a retrospective review of historical records.
16335 | 1220648-2024-21409 | 20557230 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20557230&pc=OZD | 10/14/2024 | 09/26/2025 |  | 10/29/2024 |  |  | 10/14/2024 | 10/29/2024 |  |  |  |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025542877 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Tachycardia | [12 empty] | Required Intervention | Other |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported the patient was on impella 5.5 support. the patient had runs of ventricular tachycardia. the pump was explanted four days later and the patient survived the explant. | The complainant also reported the patient was impella 5.5 support experienced a placement signal not reliable alarm, and the position was adjusted. |  |  |  |  | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation for the reported ventricular tachycardia (vt) has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the vt could not be determined. corrections to the initial mfr report: g1 mdr reporting contact email. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. there was no problem found regarding the pump for the optical signal issue. the optical signal issue is captured against the console in 24-37994-2. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include placement signal issue description. d10 concomitant medical products and therapy dates.
16336 | 1220648-2024-21410 | 20557302 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20557302&pc=OZD | 07/06/2023 | 10/29/2024 |  | 10/29/2024 |  |  | 07/06/2023 | 10/29/2024 |  |  | 03/31/2025 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024347902 |  | 1000080 |  | No |  |  | Not Available |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock needing an impella cp device for mechanical circulatory support. it was reported that during implantation there was a failed delivery due to a calcified common iliac artery. the patient's blood pressure and heart rate dropped. the impella was removed and flushed. decision was made to proceed with balloon aortic valvuloplasty (bav). patient began to arrest while balloon was being positioned. cardiopulmonary resuscitation in progress and the bav was performed. the impella was then able to be implanted in the left ventricle. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, information provided was sufficient to determine a root cause. the root cause of the delivery issues was due to patient condition. this report is being filed as part of a retrospective review of historical records.
16337 | 1220648-2024-21411 | 20557343 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20557343&pc=OZD | 10/08/2024 | 12/06/2024 |  | 10/29/2024 |  |  | 10/08/2024 | 10/29/2024 |  |  |  |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531234 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a cp pump placed to support a patient admitted in for coronary artery bypass graft, valve repair, atriclip, and maze procedure. when the cp was placed, the patient had ventricular tachycardia and was shocked 2x. pump positioning was also reviewed. the pump remains on for support. | [5 empty] | The investigation into vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6b added g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-21411. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
16338 | 1220648-2024-21408 | 20557372 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20557372&pc=OZD | 06/15/2023 | 10/29/2024 |  | 10/29/2024 |  |  | 06/15/2023 | 10/29/2024 |  |  | 12/31/2024 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024321341 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support there were low pump flows and continuous suction alarms. the staff was unable to optimize support and increase above p4 without suction alarms present. the pump position appeared to be optimal under fluoroscopy and transthoracic echocardiogram (tte). intravenous fluid was administered despite adequate preload and no signs of right ventricle failure on tte were noted. the impella cp device was removed and replaced with a new pump. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported low pump flow has been completed. however, as the necessary clinical information was not provided and the device was not returned, the root cause of the issue could not be determined. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
16339 | 1220648-2024-21412 | 20557511 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20557511&pc=PYX | 08/17/2021 | 10/29/2024 |  | 10/29/2024 |  |  | 08/17/2021 | 10/29/2024 |  |  | 03/31/2023 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP US PUMP SET |  |  | ABIOMED, INC. |  |  | US |  | 2022081556 |  | 004334 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted with cardiomyopathy was implanted with an impella rp device for mechanical circulatory support. while on support, pump had no flow and placement signal numbers went extremely high. the p level was decreased and then increased back to the previous p level. the flows resumed and the patient was maintained on support. no harm to the patient was reported. | [5 empty] | The investigation for the reported low pump flow has been completed. the product was returned, and the low pump flow was confirmed. the root cause of the low pump flow was determined to be ingested biomaterial. this report is being filed as part of a retrospective review of historical records.
16340 | 1220648-2024-21413 | 20558034 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20558034&pc=PYX | 09/09/2022 | 10/29/2024 |  | 10/29/2024 |  |  | 09/09/2022 | 10/29/2024 |  |  | 06/30/2024 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023226918 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device as part of bipella support for mechanical circulatory support. the rp had low flows after the swan was explanted from the patient. the rp was replaced and the patient remains on support. there was no reported harm or injury to the patient. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the low pump flow issue was due to patient anatomy, most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
16341 | 1220648-2024-21414 | 20558412 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20558412&pc=PYX | 11/10/2022 | 10/29/2024 |  | 10/29/2024 |  |  | 11/10/2022 | 10/29/2024 |  |  | 07/31/2024 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023248915 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that during insertion the 0.027 wire was inserted through the inlet portion of the pump and possible damage to the optical sensor was done. staff was informed and about best practices along with not touching the shiny portions of the pump during insertion. pump will be sent back to make sure there was no issue with optical sensor. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the possible product damage was most likely related to handling use during prep. this report is being filed as part of a retrospective review of historical records.
16342 | 1220648-2024-21415 | 20558517 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20558517&pc=OZD | 10/08/2024 | 12/05/2024 |  | 10/29/2024 |  |  | 10/08/2024 | 10/29/2024 |  |  |  |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025539581 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support the patient's plasma free hemoglobin was elevated. the staff attempted to reposition the pump many times. however, the pump was deep and the patient's left ventricle was enlarged. the pump was explanted. the patient¿s outcome is unknown. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation into hemolysis and positioning issues has been completed. the data logs confirmed pump was in improper position corresponded with clinical description that triggered alarms impella position unknown. no other issues were found on the logs. the product was returned in damaged condition. visual inspection found no issues on the pump. unable to conduct hemolysis test on damaged product. the root cause of hemolysis was most likely due to pump was not in the proper position. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-21415.
16343 | 1220648-2024-21416 | 20558649 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20558649&pc=OZD | 10/08/2024 | 11/27/2024 |  | 10/29/2024 |  |  | 10/08/2024 | 10/29/2024 |  |  |  |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025546548 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support. while on support, the patient had oozing around the impella cp access site. two units of blood were given. manual pressure was held for 30 minutes. femostop was applied. two days later, the pump was successfully weaned and the patient survived the explant. | [5 empty] | The investigation into access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16344 | 1220648-2024-21418 | 20558715 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20558715&pc=PYX | 05/06/2023 | 10/29/2024 |  | 10/29/2024 |  |  | 05/06/2023 | 10/29/2024 |  |  | 07/31/2024 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023276902 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp for mechanical circulatory support. it was reported the pump experienced low pump flow, resulting in the pump being explanted. a clot was subsequently noted on the pigtail of the removed device. the device was explanted, and the patient survived the procedure. | [5 empty] | The investigation for the reported low pump flow/thrombus has been completed. the pump was returned and biomaterial was found wrapped around the impeller. the root cause of the low pump flow/thrombus was determined to be biomaterial ingestion likely due to the lack of anticoagulation. this report is being filed as part of a retrospective review of historical records.
16345 | 1220648-2024-21419 | 20558767 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20558767&pc=OZD | 10/08/2024 | 12/02/2024 |  | 10/29/2024 |  |  | 10/08/2024 | 10/29/2024 |  |  |  |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025551765 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Life Threatening | Required Intervention | [23 empty] | A |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that one day following impella 5.5 implant, the patient began to experience left sided weakness. a scan of the head was obtained and it was concluded that the patient had experienced an ischemic stroke. low dose heparin orders were set to continue and the patient stabilized as support continued. the patient remains on impella support. | [5 empty] | The investigation for the reported stroke has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the stroke could not be determined. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
16346 | 1220648-2024-21417 | 20558867 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20558867&pc=PYX | 04/28/2023 | 10/29/2024 |  | 10/29/2024 |  |  | 04/28/2023 | 10/29/2024 |  |  | 12/31/2024 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024320113 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp for mechanical circulatory support. it was reported the pump experienced low pump, resulting in this pump being replaced with another impella rp. the device was explanted, and the patient survived the procedure. | [5 empty] | The investigation for the reported low pump flow has been completed. the suction alarm and motor current spike indicate that biomaterial was likely ingested and had impacted the motor and caused subsequent low pump flow. the root cause of the low pump flow was most likely biomaterial based on the log analysis. this report is being filed as part of a retrospective review of historical records.
16347 | 1220648-2024-21420 | 20559066 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20559066&pc=PYX | 05/09/2023 | 10/29/2024 |  | 10/29/2024 |  |  | 05/09/2023 | 10/29/2024 |  |  | 09/30/2024 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023270974 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp for mechanical circulatory support. it was reported the pump was inadvertently pulled back across the pulmonic valve while the mattress sutures were being tightened around the repositioning sheath. attempts to re-advance the pump across the valve were unsuccessful and the pump was replaced for continued support. patient survived the procedure. | [5 empty] | The investigation for the reported positioning issue has been completed. no product was returned for evaluation. the root cause of the positioning issue was most likely use related based on the provided clinical information. this report is being filed as part of a retrospective review of historical records.
16348 | 1220648-2024-21423 | 20559158 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20559158&pc=OZD | 09/29/2024 | 11/30/2025 |  | 10/29/2024 |  |  | 09/29/2024 | 10/29/2024 |  |  |  |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1087212 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported that upon transfer of the automated impella controller (aic), the impella 5.5 experienced impella undetectable error. staff tried to connect the catheter back to the original aic and the catheter still was undetectable. the issue persisted despite switching consoles. the patient survived the event. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
16349 | 1220648-2024-21424 | 20559194 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20559194&pc=PYX | 04/14/2023 | 10/29/2024 |  | 10/29/2024 |  |  | 04/14/2023 | 10/29/2024 |  |  | 08/31/2024 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024289360 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella rp device for mechanical circulatory support. it was reported the pump experienced low pump, resulting in this pump being replaced with another impella rp. the device was explanted, and the procedural outcome was the patient survived. | [5 empty] | The investigation for low pump flow has been completed. the data logs were reviewed and an ingestion event was identified on the 7th day of support. a motor current spike and speed deviation were observed without a p-level change, this suggests impact to the impeller. the placement signal increases after this event and corresponding calculated flow drops. the pump was returned and ingested biomaterial was found in the cannula and around the impeller. the root cause of the low pump flow was determined to be ingested biomaterial. the lack of systemic anticoagulation likely contributed to the thrombus ingestion. this report is being filed as part of a retrospective review of historical records.
16350 | 1220648-2024-21421 | 20559250 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20559250&pc=OZD | 10/08/2024 | 12/02/2024 |  | 10/29/2024 |  |  | 10/08/2024 | 10/29/2024 |  |  |  |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025569816 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp placed to support a patient admitted in for a coronary artery bypass graft, for underlying acute myocardial infarction/cardiogenic shock. the pump was placed but the limb became ischemic. the artery was occluded. the impella cp was explanted and extracorporeal membrane oxygenation was placed. the patient survived the explant. | [5 empty] | The investigation into limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16351 | 1220648-2024-21426 | 20559370 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20559370&pc=PYX | 04/19/2023 | 10/29/2024 |  | 10/29/2024 |  |  | 04/19/2023 | 10/29/2024 |  |  | 12/31/2024 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024308899 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient with post-op left ventricular assist device (lvad) was implanted with an impella rp device for mechanical circulatory support. it was reported the pump experienced low pump flow, resulting in pump being explanted. it was observed that fibrous material had encapsulated the pump motor. the device was explanted, and the patient survived the procedure. | [5 empty] | The investigation for the reported low pump flow has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. base on the clinical data that upon removal large chunk of fibrous material had encapsulated motor, and data logs reviewed and showed drop in flow at high p-level. the root cause of the low pump flow most likely was biomaterial ingestion. this report is being filed as part of a retrospective review of historical records.
16352 | 1220648-2024-21422 | 20559459 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20559459&pc=PYX | 04/13/2023 | 10/29/2024 |  | 10/29/2024 |  |  | 04/13/2023 | 10/29/2024 |  |  | 01/31/2025 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024305485 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella rp device for mechanical circulatory support. it was reported the pump experienced low pump flows due to possible clot ingestion, resulting in this pump being replaced with another impella rp. the device was explanted, and the procedural outcome was the patient survived. | [5 empty] | The investigation for low pump flow has been completed. the logs show a motor current spike, speed deviation without a p-level change, increase in pap, increase in purge pressure ad decrease in purge flow on the 24th day of support. motor current remains elevated and flow is low until the removal of the pump. the product was returned and biomaterial was found wrapped around the impeller. the root cause of the low pump flow was determined to be ingested biomaterial. this report is being filed as part of a retrospective review of historical records.
16353 | 1220648-2024-21425 | 20559548 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20559548&pc=PYX | 04/14/2023 | 10/29/2024 |  | 10/29/2024 |  |  | 04/14/2023 | 10/29/2024 |  |  | 01/31/2025 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024311902 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. it was reported during an attempt to place the impella pump, the device could not be delivered due to the patient's vasculature. it was stated that the process of impella implantation as aborted. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to patient condition. as noted in the impella ifu: impella rp smartassist system with automated impella controller, section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. warning: pre-support evaluation, axillary insertion of the impella 5.5 catheter, direct aortic insertion, use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: warning caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions, warnings & cautions, caution: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. warning: axillary insertion of the impella cp ® with smartassist ® catheter. inserting the impella rp flex with smartassist® system catheter. inserting the impella rp® with smartassist® system catheter. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ axillary insertion of the impella 5.5 catheter ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion. section: use of impella with transcatheter aortic valveswarning: pre-support evaluation, axillary insertion of the impella 5.5 catheter, direct aortic insertion, use of impella with transcatheter aortic valves. warning: caution: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella rp & rp flex rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: potential adverse events (united states), section: potential adverse events (united states. ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings: ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
16354 | 1220648-2024-21427 | 20559633 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20559633&pc=PYX | 04/20/2023 | 10/29/2024 |  | 10/29/2024 |  |  | 04/20/2023 | 10/29/2024 |  |  | 08/31/2024 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024291202 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient noted with postcardiotomy cardiogenic shock (pccs)/low cardiac output syndrome (lcos) was implanted with an impella rp device for mechanical circulatory support. it was reported the pump experienced low pump flow, resulting in pump being explanted. the device was explanted, and the patient survived the procedure. | [5 empty] | The investigation for the reported low pump flow has been completed. placement signal increased and corresponding calculated flow dropped. low pump flow was persistent after this event and lasted the rest of the case. the product was returned and biomaterial was found wrapped around the base of the impeller. the root cause of the low pump flow was determined to be ingested biomaterial. this report is being filed as part of a retrospective review of historical records.
16355 | 1220648-2024-21429 | 20559792 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20559792&pc=PYX | 04/27/2023 | 10/29/2024 |  | 10/29/2024 |  |  | 04/27/2023 | 10/29/2024 |  |  | 08/31/2024 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024291192 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp for mechanical circulatory support. the pump exhibited high purge pressure alarms followed by low purge pressure alarms. these resolved autonomously and the patient remained on support and was successfully weaned. no patient harm was reported. | [5 empty] | The investigation for the reported high purge pressure has been completed. there are also suction alarms noted during the first half of the case. the root cause of the high purge pressure was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records.
16356 | 1220648-2024-21428 | 20559851 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20559851&pc=PYX | 04/24/2023 | 10/29/2024 |  | 10/29/2024 |  |  | 04/24/2023 | 10/29/2024 |  |  | 01/31/2025 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024305345 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  |  | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The us complainant had placed a rp flex for support after having a heartmate 3 on the month prior. the rp flex was placed after some difficulty achieving appropriate position. the pump stopped after just under 17 days of support. no harm or injury from the pump stop. the rp is not a full support pump. | [5 empty] | The investigation for the reported pump stop has been completed. the data logs confirm a high motor current pump stop on the 17th day of support. motor current and speed irregularities were noted in the minute leading up to the pump stop. the product was returned and sent out no irregularities were found. the pump was sent out to ct scan and wear was seen to the metal retainer in the bearing. the root cause of the pump stop was determined to be bearing wear. "as noted in the impella instructions for use: impella rp with smartassist system: section: using the automated impella controller with the impella catheter. ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella rp smart assist system with automated impella controller. section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿" this report is being filed as part of a retrospective review of historical records.
16357 | 1220648-2024-21430 | 20559949 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20559949&pc=OZD | 07/14/2021 | 10/29/2024 |  | 10/29/2024 |  |  | 07/14/2021 | 10/29/2024 |  |  | 02/28/2023 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP PUMP SET, JAPAN |  |  | ABIOMED. INC. |  |  | US |  | 2022052342 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  |  | Asian | Death | Yes | Yes | Pumping Stopped | [6 empty] | Arrhythmia | Hematoma | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Life Threatening | Required Intervention | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support, the patient experienced access site bleeding and a hematoma, a blood transfusion was administered. also, the patient experienced arrythmia, the resolution for this issue is unknown. additionally, the pump stopped and was restarted, it was reported that after the pump stop the patient's creatinine level increased. the patient died during support; it was reported that the device was working normally at the time of death. it was noted that reason for the patient's death is unknown, it could have been acute exacerbation of idiopathic pulmonary fibrosis or dilated cardiomyopathy exacerbation. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported pump stop has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the pump stop and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. as noted in the impella ifu: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
16358 | 1220648-2024-21431 | 20560385 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20560385&pc=OZD | 01/13/2022 | 09/22/2025 |  | 10/29/2024 |  |  | 01/13/2022 | 10/29/2024 |  |  | 08/31/2023 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2022127547 |  | 0048-0034-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support, the patient had dark brown urine indicating hemolysis, the resolution for this issue is unknown. survival outcome at explant noted the patient expired. the cause of the patient's death was noted as multiple organ failure due to sepsis. it was reported that the impella cp worked as intended. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. | [5 empty] | The investigation for the reported hemolysis has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the hemolysis and its relation to impella were not determined since the product, data logs, and sufficient clinical information were not provided. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as a part of retrospective review. due to limited clinical information and lack of available data/product, the root cause of this investigation is not determined. d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2024-21431. g1 reporting contact email revised on manufacturer device report 1220648-2024-21431 in accordance with updated procedures. h6 health effect - impact code 1802 was erroneously entered on the initial manufacturer device report number 1220648-2024-21431.
16359 | 1220648-2024-21433 | 20561431 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20561431&pc=PYX | 06/12/2023 | 10/29/2024 |  | 10/29/2024 |  |  | 06/12/2023 | 10/29/2024 |  |  | 08/31/2024 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024297005 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp for mechanical circulatory support. while on support, the pump experienced low pump, resulting in the pump being explanted and replaced. a clot was subsequently noted on the outlet cage of the removed device. the device was explanted, and the procedural outcome was the patient survived. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that a cannula kink near the inlet was observed. a big clot observed in the outlet cage. it looks like the cannula kink occurred in transportation as no indication regarding the kink even after explanation. the data log reviewed showed motor current spikes but didn't show last few days of the case. the root cause of the low pump flow was biomaterial ingestion due to patient condition related. this report is being filed as part of a retrospective review of historical records.
16360 | 1220648-2024-21432 | 20561520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20561520&pc=OZD | 06/12/2023 | 10/29/2024 |  | 10/29/2024 |  |  | 06/12/2023 | 10/29/2024 |  |  |  |  | 10/29/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1577697 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) purge disc holder had snapped off, though not during an active case. there was no reported patient involvement. | [5 empty] | The device was returned and evaluated, and the investigation for the reported purge disc/flag issues has been completed. the device logs were reviewed but not relevant to the complaint. visual inspection of the purge assembly area confirmed a ripped/ missing blue retainer. the failure mode was component failure due to broken/ missing purge retainer. this report is being filed as part of a retrospective review of historical records.
16361 | 1220648-2024-21435 | 20561582 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20561582&pc=PYX | 12/22/2022 | 10/29/2024 |  | 10/29/2024 |  |  | 12/22/2022 | 10/29/2024 |  |  | 05/31/2024 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023213695 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with right heart failure was being implanted with an impella rp device for mechanical circulatory support. during implantation, the impella rp pump was attempted to be delivered several times. the pigtail did cross into the pulmonary artery (pa) briefly but came back. after that, the physician had no success in getting the pump into position due to the size of the right ventricle (rv). hemodynamics were stable, the case was aborted, and the patient was stable. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the delivery issue was patient condition related and related to the size of the right ventricle. impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller instructions for use for the related event are as follows: section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
16362 | 1220648-2024-21434 | 20561619 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20561619&pc=PYX | 05/21/2023 | 10/29/2024 |  | 10/29/2024 |  |  | 05/21/2023 | 10/29/2024 |  |  | 04/30/2025 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024334326 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. while on support, the patient experienced bleeding at access site, and femotop, manual pressure, and suture was done to address the issue. the case wrap up indicates the patient remained on impella support and was successfully weaned. | [5 empty] | The investigation for the reported access site bleeding issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the access site bleeding was most likely lack of vessel recoil because the bleeding occurred after repo unit insertion and was resolved with manual pressure. impella rp flex with smartassist instructions for use & clinical reference manual instructions for use for the related event are as follows: instructions for use for the related event are as follows: section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ section: warnings and cautions: ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records.
16363 | 1220648-2024-21438 | 20561724 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20561724&pc=OZD | 06/02/2023 | 10/29/2024 |  | 10/29/2024 |  |  | 06/02/2023 | 10/29/2024 |  |  |  |  | 10/29/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1308456 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a purge motor issue. no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | E1; name of initial reporter is unknown. the investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. there was no issue found during the case. the device functioned/operated to specification. this report is being filed as part of a retrospective review of historical records.
16364 | 1220648-2024-21444 | 20561795 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20561795&pc=PYX | 04/17/2022 | 10/29/2024 |  | 10/29/2024 |  |  | 04/17/2022 | 10/29/2024 |  |  | 05/31/2023 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022103146 |  | 004334 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella rp device for mechanical circulatory support. while on support, a pump stop occurred after the team had removed the heparin several hours earlier. the pump stop prompted the team to explant the pump. the pump was replaced and support proceeded on new pump. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that a large clot was found around the impeller. thus, the root cause of the pump stop is biomaterial. impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
16365 | 1220648-2024-21439 | 20561833 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20561833&pc=PYX | 02/08/2023 | 10/29/2024 |  | 10/29/2024 |  |  | 02/08/2023 | 10/29/2024 |  |  | 08/31/2024 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024289359 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Male |  |  | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with right heart failure was implanted with an impella rp device for mechanical circulatory support. while on support, there was a low pump flow. the device was not explanted as a result and there was no harm to the patient. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the data logs were reviewed and confirmed a motor current spike, speed deviation without a change in p-level, slight decrease in flow, and slight increase in placement signal. the product was returned and biomaterial was found wrapped around the impeller. the root cause of the low pump flow was determined to be ingested biomaterial. this report is being filed as part of a retrospective review of historical records.
16366 | 1220648-2024-21449 | 20561838 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20561838&pc=PYX | 02/19/2023 | 10/29/2024 |  | 10/29/2024 |  |  | 02/19/2023 | 10/29/2024 |  |  | 08/31/2024 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED |  |  | US |  | 2023260142 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, impella rp dislodged out of position when peel away sheath was attempted to be removed. physician was unable to reposition the device. therefore, rp catheter was removed, and another rp was inserted as replacement. no patient harm was reported. | [5 empty] | The investigation for the reported positioning issues has been completed. the impella device was not received from the customer. however, clinical details were sufficient to determine the root cause. the root cause of the positioning issue was determined to be use related as the pump was pulled back out of position while the peel away sheath was being removed. this report is being filed as part of a retrospective review of historical records.
16367 | 1220648-2024-21437 | 20561839 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20561839&pc=OZD | 06/02/2023 | 10/29/2024 |  | 10/29/2024 |  |  | 06/02/2023 | 10/29/2024 |  |  |  |  | 10/29/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1619563 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced now as not recognizing catheter and stated ¿make sure gray to gray connection is secure¿, which was resolved by changing the console. no patient was involved. | [5 empty] | The investigation has been completed. the console (ic9307) was sent back for further investigation. the console was tested to attempt to recreate the failure mode but this was unsuccessful. the console passed all testing during the investigation leading to the cause of this failure mode to be most likely due to the pattern failure epm crystal on the impellatronic board. the pump detection issue was most likely due to the rarely occurring epm crystal issue. the impellatronic pca does not have permanent damage, but this intermittent pump detection failure occurs infrequently and already has a sw mitigation in place. this report is being filed as part of a retrospective review of historical records.
16368 | 1220648-2024-21450 | 20561850 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20561850&pc=PYX | 12/18/2022 | 10/29/2024 |  | 10/29/2024 |  |  | 12/18/2022 | 10/29/2024 |  |  | 04/30/2023 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023268853 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella rp for mechanical circulatory support. no initial complaint by the customer. the pump was returned and found to have low flows. there was not any noted harm or injury during the support from the low flows. no further information is available. | [5 empty] | The investigation into the low pump flow issue has been completed. the data logs were evaluated and a general trend down in motor current and pump flow was seen throughout the case. an evaluation of the physical device noted a kink in the cannula. during follow up with the physician, it was noted that the kink was not present on imaging and was likely a result of shipping the device back to abiomed. the kink was not suspected to be the root cause of the low flow. no biomaterial or thrombus was found on the returned product. the returned pump was run in a basin of water and produced adequate flows. the root cause of the low flow was determined to most likely be due to patient condition as the patient had a packed chest. this report is being filed as part of a retrospective review of historical records.
16369 | 1220648-2024-21451 | 20561868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20561868&pc=PYX | 12/20/2022 | 10/29/2024 |  | 10/29/2024 |  |  | 12/20/2022 | 10/29/2024 |  |  | 01/31/2024 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023173520 |  | 004334 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. during support, and right after insertion the pump stopped. the device was explanted and replaced. no reported harm or injury to the patient. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was received from the customer and therefore, an evaluation of the device was completed. the returned pump was visually inspected, and biomaterial was observed on the impeller. the pump was run at p9 in a basin of water and pump flow were within specification at 4.2 l/min. the cause of the issue was biomaterial. the root cause of the problem is unable to be determined. as noted in the impella ifu: impella rp with smart assist system. section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
16370 | 1220648-2024-21436 | 20561995 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20561995&pc=PYX | 12/25/2022 | 10/29/2024 |  | 10/29/2024 |  |  | 12/25/2022 | 10/29/2024 |  |  | 04/30/2023 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023268852 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with right heart failure was implanted with an impella rp device for mechanical circulatory support. while on support, the impella rp was failing to achieve appropriate flows. the pump was removed, and the right ventricle [rv] was monitored for potential further mechanical support. there was no harm to the patient. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the logs were returned and a motor current spike, increase in placement signal, and a drop in speed independent of a p-level change were all seen in the same event which is indicative of clot ingestion. the product was returned and ingested thrombus was found wrapped around the impeller and in the purge gap. the root cause of the low flows was determined to be biomaterial. this report is being filed as part of a retrospective review of historical records.
16371 | 1220648-2024-21452 | 20562029 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20562029&pc=PYX | 02/17/2023 | 10/29/2024 |  | 10/29/2024 |  |  | 02/17/2023 | 10/29/2024 |  |  | 09/30/2024 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED |  |  | US |  | 2023270987 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was going to be implanted with an impella rp device for mechanical circulatory support. during insertion, the right ventricle was grossly dilated and large. the impella would not advance out of the pulmonary artery. eventually very high resistance/frictional force noted to advance the catheter through the 23 fr peelaway sheath. however, the device never reached a satisfactory position. several attempts were made at placement with varying wires. all proved to be unsuccessful. eventually, it became difficult to even advance the impella through the peelaway sheath. the physician decided to abort the implant and explanted the device. no further information is available. | [5 empty] | The impella device was returned for evaluation. upon visual inspection, a kink was observed in the introducer along with tip damage which is consistent with the difficult delivery. the root cause of the delivery issue was determined to be patient condition as this was a salvage case and the pump was unable to be inserted due to the patient's anatomy. as noted in the impella ifu: impella rp with smart assist system section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ this report is being filed as part of a retrospective review of historical records.
16372 | 1220648-2024-21440 | 20562071 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20562071&pc=PYX | 09/14/2021 | 10/29/2024 |  | 10/29/2024 |  |  | 09/14/2021 | 10/29/2024 |  |  |  |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 1485147 |  | 004334 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was implanted with an impella rp for mechanical circulatory support. there was a kink in the catheter shaft of the pump, another pump was used, and there was no report of patient harm requiring medical intervention. | [5 empty] | The investigation into the product damage issue has been completed. the impella pump was returned for investigation. a kink in the catheter shaft was confirmed during product review. thus, the root cause of the delivery issues is due to a kink in the catheter. it was determined from the clinical details and the returned product inspection, to be the observed cannula kink resulting from improper technique by the end user. this report is being filed as part of a retrospective review of historical records.
16373 | 1220648-2024-21441 | 20562083 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20562083&pc=PYX | 09/17/2021 | 10/29/2024 |  | 10/29/2024 |  |  | 09/17/2021 | 10/29/2024 |  |  | 09/30/2022 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP US |  |  | ABIOMED, INC. |  |  | US |  | 1503745 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient post cardiotomy cardiogenic shock/low cardiac output syndrome was implanted with an impella rp for mechanical circulatory support. the patient's anatomy caused a kink/bend in the inlet of the pump, another device was used, and there was no report of patient harm requiring medical intervention. | [5 empty] | The product was not returned for investigation. after reviewing the clinical information, it was determined that the cause of the delivery issues was patient condition, specifically patient vessel anatomy. this report is being filed as part of a retrospective review of historical records.
16374 | 1220648-2024-21442 | 20562086 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20562086&pc=PYX | 02/14/2023 | 09/26/2025 |  | 10/29/2024 |  |  | 02/14/2023 | 10/29/2024 |  |  | 08/31/2024 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024289357 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 26 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the rp flex was implanted after the patient's continued decompensation during placement of the impella 5.5. placement signal not reliable alarm was observed during insertion/initial placement. the lines in the rij were removed and the placement signal alarms subsided. after removal of the lines, it was noted that the rp flex was positioned below the pulmonic valve. therefore, the device was repositioned. the pump was successfully weaned and explanted. | The complainant reported a patient was implanted with an impella rp device for mechanical circulatory support. while on support, the rp flex came back below the pulmonic valve. there were no issues after the peel away was removed and repositioning sheath inserted. the pump came back after two additional lines from rij were removed. the patient was successfully weaned from rv support, and remained on impella 5.5 with plans to slowly wean as tolerated. |  |  |  |  | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the product was returned and the dp sensor and optical sensor were tested and found to be functioning without issue. the pump was run in a bucket of water and placement signal remained stable. the data logs were reviewed and show placement signal not reliable on ramp up and elevated pap and cvp. the root cause of the placement signal issue was determined to be use related as the dp and optical sensor were both likely responding to interaction with additional 2 lines in the rij. the issue resolved once the lines were removed. based on the clinical details the root cause of the positioning issues was determined to be use related as the pump was pulled out of position while removing the other lines in the rij. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical signal issue as part of a retrospective review of optical signal issues in accordance with fda recommendation. | The investigation for the reported positioning issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the product was returned and the dp sensor and optical sensor were tested and found to be functioning without issue. the pump was run in a bucket of water and placement signal remained stable. the data logs were reviewed and show psnr on ramp up and elevated pap and cvp. the root cause of the placement signal issue was determined to be use related as the dp and optical sensor were both likely responding to interaction with additional 2 lines in the rij. the issue resolved once the lines were removed. based on the clinical details the root cause of the positioning issues was determined to be use related as the pump was pulled out of position while removing the other lines in the rij. this report is being filed as part of a retrospective review of historical records.
16375 | 1220648-2024-21443 | 20562091 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20562091&pc=PYX | 04/14/2022 | 10/29/2024 |  | 10/29/2024 |  |  | 04/14/2022 | 10/29/2024 |  |  | 09/30/2022 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 1502967 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was being implanted with an impella rp device for mechanical circulatory support. during implantation, because of the patient¿s anatomy, the physician was unable to get the pump up into the pulmonary artery (pa) through direct approach and the outlet was damaged by the physician. a second pump was prepped, primed and attempted with a wire and still unable to get into the pa. the physician opted for veno-arterial extracorporeal membrane oxygenation (va ecmo) and aborted the impella rp implantation. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the issue was patient condition related. impella cp with smartassist system & impella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller instructions for use for the related event are as follows: section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
16376 | 1220648-2024-21445 | 20562098 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20562098&pc=PYX | 12/03/2022 | 10/29/2024 |  | 10/29/2024 |  |  | 12/03/2022 | 10/29/2024 |  |  | 07/31/2024 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023250967 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was going to be implanted with an impella rp device for mechanical circulatory support. the impella rp was unable to be delivered to the right ventricle [rv] for right sided support. the rp insertion was attempted by multiple troubleshooting measures inclusive of wire exchanges. the rp was not delivered and there was no harm or injury reported. the delivery was aborted. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were sufficient to determine a cause. the root cause of the delivery issue was unable to be determined since the product, data logs, and sufficient clinical information were not provided. as noted in the impella ifu: impella rp with smart assist system: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. this report is being filed as part of a retrospective review of historical records.
16377 | 1220648-2024-21446 | 20562101 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20562101&pc=PYX | 12/08/2022 | 10/29/2024 |  | 10/29/2024 |  |  | 12/08/2022 | 10/29/2024 |  |  | 08/31/2024 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023250939 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Death | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was going to be implanted with an impella rp device for mechanical circulatory support. it was reported that during delivery of the device the introducer and catheter kinked, and the device was explanted. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was possible. clinical details provided were sufficient to determine a root cause. the root cause of the introducer issue was a sheath kink most likely due to patient condition as the patient's vasculature was reportedly difficult to navigate. as noted in the impella ifu: impella rp with smart assist system section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. this report is being filed as part of a retrospective review of historical records.
16378 | 1220648-2024-21454 | 20562113 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20562113&pc=OZD | 05/10/2023 | 10/29/2024 |  | 10/29/2024 |  |  | 05/10/2023 | 10/29/2024 |  |  |  |  | 10/29/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1400770 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Damaged Prior to Use | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the grey door on the automated impella controller (aic) was broken. no patient was involved. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the root cause of the aic damage ¿ broken optical connector cover was unable to determine. this report is being filed as part of a retrospective review of historical records.
16379 | 1220648-2024-21448 | 20562129 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20562129&pc=PYX | 12/09/2022 | 10/29/2024 |  | 10/29/2024 |  |  | 12/09/2022 | 10/29/2024 |  |  | 09/30/2024 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023266232 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a low flow issue with an impella rp device. as a result of the low pump flow and coinciding increased motor current, the rp pump was explanted from the patient. patient support continued with venoarterial extracorporeal membrane oxygenation (va ecmo). there was no reported harm to the patient. | [5 empty] | The investigation for the reported low flow has been completed. the device was not returned. data logs review showed several purge pressure alarm occurred and sudden drop of motor current. clinical details reported possibility of clot burden, and the patient began to deteriorate, usage of inotropes and pressor was increased. root cause most likely biomaterial ingestion. this report is being filed as part of a retrospective review of historical records.
16380 | 1220648-2024-21453 | 20562199 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20562199&pc=PYX | 02/18/2023 | 10/29/2024 |  | 10/29/2024 |  |  | 02/18/2023 | 10/29/2024 |  |  | 05/31/2024 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED |  |  | US |  | 2023212219 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult or Delayed Positioning | Failure to Sense | Deformation Due to Compressive Stress |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella rp for mechanical circulatory support. after the device was implanted, the staff noticed clot on the pig tail/wire and removed the pump. the pump was ran on p2 in saline for couple seconds to clear any clot and the wire was contaminated by scrub. the staff used wire cutters to cut contaminated end, but too rigid for pump to advance over. everything was removed, fluoroscopy was provided and then it was noted that the impella rp sheath was kinked. the sheath was exchanged on amplatz super stiff without issue. the device was reinserted with some difficulty passing through tricuspid valve. then the device exhibited a placement signal lost message and bleeding was noticed at the access site. the device was successfully implanted. no resolution specified for the access site bleeding or placement signal lost message. it was reported that the patient survived the procedure. | [5 empty] | The investigation for the reported thrombus, optical signal, access site bleeding, introducer issue has been completed. the impella device was not received from the customer. however, clinical details were sufficient to determine the root cause. the cause of the thrombus issue was patient condition related. the cause of the introducer issue was a introducer sheath kink. the cause of the placement signal issue was not determined since product was not returned. the cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical details were provided. as noted in the impella ifu: impella rp with smart assist system section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ this report is being filed as part of a retrospective review of historical records.
16381 | 1220648-2024-21447 | 20562241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20562241&pc=OZD | 05/27/2023 | 10/29/2024 |  | 10/29/2024 |  |  | 05/27/2023 | 10/29/2024 |  |  |  |  | 10/29/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1186352 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a purge disc/flag issue, which was resolved by transferring to a different aic. no clinical details regarding patient condition were provided. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the issue was broken/ missing purge disk retainer. this report is being filed as part of a retrospective review of historical records.
16382 | 1220648-2024-21455 | 20562792 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20562792&pc=OZD | 03/23/2023 | 10/29/2024 |  | 10/29/2024 |  |  | 03/23/2023 | 10/29/2024 |  |  |  |  | 10/29/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1132455 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was placed onto impella cp support for high-risk percutaneous coronary intervention. while on support, the automated impella controller (aic) experienced a controller error yellow alarm that was resolved by changing out the controller. no patient harm was reported. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. however, the issue was not reproduced. monitoring of kbd communication during several aic start-up cycles did not reveal any irregularities. nevertheless, due to the age of components that could be responsible for the communication failure (kbd assembly, ribbon cable), it had been recommended that all the above components be replaced to minimize chance of aic alarm re-occurring. the cause of the controller error was most likely due to a defective kbd/display. this report is being filed as part of a retrospective review of historical records.
16383 | 1220648-2024-21456 | 20563319 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20563319&pc=OZD | 03/22/2023 | 10/29/2024 |  | 10/29/2024 |  |  | 03/22/2023 | 10/29/2024 |  |  |  |  | 10/29/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1102031 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a purge issue that was resolved by swapping out the console. no clinical details regarding patient condition were provided. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. despite exhaustive efforts, the failure was unable to be reproduced during investigation. it is likely that an electro static discharge (esd) event on the purge assembly communication line could have caused the momentary purge fluid miscalculation. the cause of the aic issue was a defective purge assembly misreporting the purge fluid volume. this report is being filed as part of a retrospective review of historical records.
16384 | 1220648-2024-21457 | 20563424 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20563424&pc=OZD | 03/15/2023 | 10/29/2024 |  | 10/29/2024 |  |  | 03/15/2023 | 10/29/2024 |  |  |  |  | 10/29/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1693473 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was placed onto impella 5.5 support due to cardiomyopathy. while on support, the automated impella controller (aic) experienced a purge issue. that was resolved by switching to a backup controller. no patient harm was reported due to the issue . | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the blue receptacle had a gap between itself and where it usually is placed which would have caused the purge assembly to be unable to detect purge fluid during operation as well as other errors stemming from an unaligned purge system. the cause of the issue was a broken blue receptacle. this report is being filed as part of a retrospective review of historical records.
16385 | 1220648-2024-21458 | 20563523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20563523&pc=OZD | 03/02/2023 | 10/29/2024 |  | 10/29/2024 |  |  | 03/02/2023 | 10/29/2024 |  |  |  |  | 10/29/2024 |  | AUTOMATED IMPELLA CONTROLLER |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1702808 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) console was showing an unexpected shut down message and began shutting down with loud beeping . there was no patient involved. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the complaint was able to be repeatedly reproduced during the investigation with the watchdog triggering a restart due to failed imcgui communication additionally, the kbd communication intermittently failed. known good 4 in 1 and compact flashes were swapped into ic10305. (the original 4 in 1 was the new design whereas the known good is the original design. cables were also swapped for the new connector configuration.) the compact flash cards on a known good console did not trigger the failure mode. the cause of the issue was defective components. this report is being filed as part of a retrospective review of historical records.
16386 | 1220648-2024-21459 | 20563534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20563534&pc=OZD | 07/06/2023 | 10/29/2024 |  | 10/29/2024 |  |  | 07/07/2023 | 10/29/2024 |  |  |  |  | 10/29/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1273761 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient of unknown age/race/ethnicity noted as high-risk percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. during support, the automated impella controller (aic) showed "purge disk not recognized." there was no reported patient harm. | [5 empty] | The device was returned and evaluated, and the investigation for the reported purge disc/flag issues has been completed. the device logs from the reported complaint date revealed that the console issued the purge disc not detected alarm. the console was tested and the issue with properly detecting the purge pressure disc was reproduced, the purge flag was confirmed to be broken. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
16387 | 1220648-2024-21460 | 20563629 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20563629&pc=OZD | 02/23/2023 | 10/29/2024 |  | 10/29/2024 |  |  | 02/23/2023 | 10/29/2024 |  |  |  |  | 10/29/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1371242 |  | 0042-0010-EU |  | No |  |  | No |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) exhibited a battery fault message prior to a planned procedure. the staff were informed on how to clear the error message. no report of patient harm or injury. | [5 empty] | The investigation has been completed. the consoles did not need to be sent back for investigation due to the error being resolved in the field. however, the complaint was able to be reproduced eliciting the same response to what the customer saw in the field. the cause of the issue was the transport switch not being flipped after receipt. this report is being filed as part of a retrospective review of historical records.
16388 | 1220648-2024-21461 | 20563637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20563637&pc=OZD | 02/14/2023 | 10/29/2024 |  | 10/29/2024 |  |  | 02/14/2023 | 10/29/2024 |  |  |  |  | 10/29/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1100963 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) exhibited a self-check failed message post software update. no patient involvement. | [5 empty] | The investigation has been completed. the console ( (b)(6) ) was sent back for further investigation. the root cause of the aic issue was a software issue upgrading. this report is being filed as part of a retrospective review of historical records.
16389 | 1220648-2024-21462 | 20563644 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20563644&pc=OZD | 03/09/2022 | 10/29/2024 |  | 10/29/2024 |  |  | 03/09/2022 | 10/29/2024 |  |  | 06/30/2023 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022108229 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to be implanted with an impella cp device for mechanical circulatory support. the physician was unable to advance the device past the aorta-iliac bifurcation and the procedure was aborted. additional information was provided that noted the patient expired. there was no reported relation of the impella with the patient's outcome. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was patient anatomy, most likely due to patient condition. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high-risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings. ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
16390 | 1220648-2024-21463 | 20563647 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20563647&pc=OZD | 03/10/2022 | 10/29/2024 |  | 10/29/2024 |  |  | 03/10/2022 | 10/29/2024 |  |  | 11/30/2023 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023166528 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during implantation, while exchanging the 14 fr introducer for the repositioning sheath, the cannula became kinked. attempts to correct the kink were unsuccessful, so the pump was removed and replaced. there was no reported harm or injury to the patient. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The investigation for the reported cannula kink has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the cannula kink was a use issue, as the impella moved upon exchange of the 14fr for the gwru. this report is being filed as part of a retrospective review of historical records.
16391 | 1220648-2024-21464 | 20563655 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20563655&pc=OZD | 03/20/2022 | 03/27/2025 |  | 10/29/2024 |  |  | 03/20/2022 | 10/29/2024 |  |  | 03/31/2023 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1476557 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic |  | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support there were lower than expected flows. during removal of the peel-away introducer, the impella shifted under mitral structure. the physician attempted to reposition the impella; however, was not able to get the pump back across the aortic valve but was able to then cross the valve and was in excellent position. after repositioning, there was suction at p-2, and the cannula appeared to have a kink in it. the pump was removed and replaced with no harm to the patient. additional information was provided that noted the patient expired. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | The investigation for the reported positioning and product damage issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that a kink in the cannula was confirmed, no other damage was seen. the cause of the positioning issues was use related, as the impella migrated as the gwru was inserted. this report is being filed as part of a retrospective review of historical records.
16392 | 1220648-2024-21465 | 20563684 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20563684&pc=OZD | 08/04/2021 | 10/29/2024 |  | 10/29/2024 |  |  | 08/04/2021 | 10/29/2024 |  |  | 06/30/2023 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022094221 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  |  | Death | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the device experienced suction issues indicating low flow, the staff repositioned the device to help resolve this issue. it was reported that flows increased. additionally, the device exhibited an optical signal issue the optical sensor showed erratic waveforms and sudden, large changes in value. an echocardiogram was performed and verified that the device was in the correct position. the device also exhibited a ¿impella in ventricle¿ alarm when impella was not in the ventricle (confirmed by echocardiogram) and the motor current was pulsatile. the staff decided to disable the placement monitoring. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported optical signal and low flow issue has been completed. the device log was reviewed and confirmed that there were multiple impella position unknown and impella position in the ventricle alarms triggered. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were not sufficient to determine a root cause. the cause of the optical signal and low flow issue were not determined since the product, data logs, and sufficient clinical information were not provided. this report is being filed as part of a retrospective review of historical records.
16393 | 1220648-2024-21466 | 20563697 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20563697&pc=OZD | 07/01/2022 | 03/27/2025 |  | 10/29/2024 |  |  | 07/01/2022 | 10/29/2024 |  |  | 12/31/2023 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023178654 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Fluid/Blood Leak | Increase in Pressure | Pumping Stopped |  |  |  |  | Thrombosis/Thrombus | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with post cardiotomy cardiogenic shock / low cardiac output syndrome was on an automated impella controller (aic) for mechanical circulatory support. during support, the device exhibited several leaking purge cassettes. the staff replaced several purge cassettes; however, they kept getting errors of no fluid detected. purge fluid was exiting the yellow hub. all troubleshooting done and unable to resolve. temporarily protected the impella using pressure back and normal iv tubing on purge system. also, low purge flow/high purge pressure was noted, the staff changed the purge, and the purge flow increased, then the pump stopped. the pump stopped several times and was able to be restarted, however the issues did not resolve. a stat echocardiogram was performed and noted an apical thrombus and another thrombus on the pump above the inflow. the team stated that the only option was to keep the pump in place as removing would eject clot and cause a possible stroke. it was reported that the patient died on support after cardio-pulmonary resuscitation. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported issues has been completed. the field long term log was received and reviewed. a steady increase in purge pressure was observed 5 hours into the run, then high purge pressure alarms which then temporarily resolved but then a sudden spike in purge pressure was observed with purge line click-on not detected alarms and then a pump stop alarm. no product was returned. only a data stick was returned with event logs. the cause of the high purge pressure and resulting pump stop was most likely biomaterial. the root cause of the biomaterial was most likely patient condition. this report is being filed as part of a retrospective review of historical records. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
16394 | 1220648-2024-21467 | 20563789 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20563789&pc=OZD | 09/20/2022 | 05/23/2025 |  | 10/29/2024 |  |  | 09/21/2022 | 10/29/2024 |  |  | 05/31/2024 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | GM |  | 2023218178 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  | White | Injury | Yes | Yes | Positioning Failure | Device Sensing Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced access site bleeding with heparin lowered from the purge. there were optical signal issues, and positioning issues with the pump. the impella was repositioned; however, the placement signal was not clear. in addition, the motor current was not pulsatile, and the color doppler showed confusion above and below the valve. these issues were resolved by removing the pump. the patient ultimately expired, use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This complaint has been identified as part of a retrospective review. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined.
16395 | 1220648-2024-21468 | 20563907 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20563907&pc=OZD | 07/28/2022 | 10/29/2024 |  | 10/29/2024 |  |  | 07/28/2022 | 10/29/2024 |  |  | 04/30/2024 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023219493 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, patient experienced access site bleeding. upon further checking, a hematoma was noted. hematoma was then expressed out by the peripheral artery (pa) and the site was redressed. the following day, the patient started to bleed from other sites. the heparin was discontinued as a result. care was later withdrawn and patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smartassist system (access site bleeding) section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist system (vascular dissection) section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
16396 | 1220648-2024-21469 | 20563919 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20563919&pc=OZD | 01/14/2022 | 10/06/2025 |  | 10/29/2024 |  |  | 03/29/2022 | 10/29/2024 |  |  | 08/31/2023 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022117275 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the patient had bloody sputum and bleeding including hematoma. during transfer to the icu, the patient woke up and moved significantly. to keep him at rest, he was intubated after entering the icu. the position of the impella was reevaluated and appeared to be a little deep and was adjusted. blood transfusion and suction were given. additional information was provided that noted the patient suffocated due to bloody sputum and experienced a sudden cardiac death. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | B3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2024-21469. g1 reporting contact email revised on manufacturer device report 1220648-2024-21469 in accordance with updated procedures. h6 investigation conclusions 11 was erroneously entered on the initial manufacturer device report number 1220648-2024-21469.
16397 | 1220648-2024-21470 | 20563946 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20563946&pc=OZD | 06/24/2022 | 10/29/2024 |  | 10/29/2024 |  |  | 06/24/2022 | 10/29/2024 |  |  | 02/29/2024 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023201656 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Infusion or Flow Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. the impella cp had a saturated pump flow. the decision was made by the patient's family to withdraw care. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation of the reported issue has been completed. the returned pump was visually inspected and biomaterial was observed on the impeller. the cause of the issue was biomaterial. this report is being filed as part of a retrospective review of historical records.
16398 | 1220648-2024-21471 | 20563981 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20563981&pc=OZD | 06/16/2023 | 10/29/2024 |  | 10/29/2024 |  |  | 06/16/2023 | 10/29/2024 |  |  | 03/31/2025 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | IMPELLA CP |  |  | US |  | 2024342332 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | Appropriate Term / Code Not Available | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the pump was placed, started, and then stopped. the pump was replaced after three restart attempts. additional information was provided that noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported pump stop has been completed. however, as the necessary clinical information was not provided and the device was not returned, the root cause of the pump stop could not be determined. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
16399 | 1220648-2024-21472 | 20563982 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20563982&pc=OZD | 08/19/2021 | 10/29/2024 |  | 10/29/2024 |  |  | 08/19/2021 | 10/29/2024 |  |  | 12/31/2022 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022039802 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | Asian | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Appropriate Term / Code Not Available | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump flow rate was lower than the previous day. it was decided to replace the device; however, the patient expired due to multi-organ failure. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported issues has been completed. the impella device has not been returned for evaluation. clinical details were insufficient to determine the root cause. the cause of the low pump flow was not determined. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
16400 | 1220648-2024-21473 | 20564012 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20564012&pc=OZD | 07/19/2023 | 10/29/2024 |  | 10/29/2024 |  |  | 07/19/2023 | 10/29/2024 |  |  | 03/31/2025 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024342943 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Cardiac Arrest | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention needing an impella cp device for mechanical circulatory support. the physician attempted to implant the impella but could not be advanced as the patient was presented with aortic calcification. the doctor implanted an intra-aortic balloon pump instead; however, the patient was tachycardic and experienced ventricular tachycardia. the patient continued to code and eventually expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
16401 | 1220648-2024-21474 | 20564030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20564030&pc=OZD | 07/14/2022 | 10/29/2024 |  | 10/29/2024 |  |  | 07/14/2022 | 10/29/2024 |  |  | 02/29/2024 |  | 10/29/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200301 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hypovolemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the device immediately developed low flows and suction. the device was replaced with a second pump, which also had low flows, so an echocardiogram was performed and showed a compressed left ventricle and hypovolemia. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The device was returned and evaluated, and the investigation for the reported low pump flow has been completed. the reported low pump flow was unable to be reproduced. the low pump flow was most likely due to poor patient condition. this report is being filed as part of a retrospective review of historical records.
16402 | 1220648-2024-21346 | 20541719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541719&pc=PYX | 03/07/2023 | 10/27/2024 |  | 10/28/2024 |  |  | 03/07/2023 | 10/28/2024 |  |  | 11/30/2024 |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023281755 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. while on support, the rp was unable to advance into an adequate position during delivery due to resistance met in the patient's vasculature. the implant was aborted and an impella 5.5 was subsequently placed for patient support. there was no reported harm or injury to the patient. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to patient condition. as noted in the impella ifu: impella rp smartassist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves warning: pre-support evaluation axillary insertion of the impella 5.5 catheter direct aortic insertion use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: warning caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions: warnings & cautions caution ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings warning: axillary insertion of the impella cp ® with smartassist ® catheter inserting the impella rp flex with smartassist® system catheter inserting the impella rp® with smartassist® system catheter ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ axillary insertion of the impella 5.5 catheter ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valveswarning: pre-support evaluation axillary insertion of the impella 5.5 catheter direct aortic insertion use of impella with transcatheter aortic valves warning caution: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella rp & rp flex rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: potential adverse events (united states) section: potential adverse events (united states ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
16403 | 1220648-2024-21347 | 20541723 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541723&pc=PYX | 03/08/2023 | 10/27/2024 |  | 10/28/2024 |  |  | 03/08/2023 | 10/28/2024 |  |  | 08/31/2024 |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024296981 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. while on support there were noted signs of hemolysis. the pfhg levels were elevated. the patient was successfully weaned from the impella and the device explanted. | [5 empty] | The investigation for the reported hemolysis has been completed. only the cannula portion of the pump was returned. the product was examined and biomaterial was found. no other abnormalities were found. the data logs were reviewed and showed a motor current spike, speed deviation, and drop in flow which is indicative of ingestion, however the ingestion event was after the reported hemolysis began. the biomaterial is unlikely to be the root cause of the hemolysis. the pump was soldered to a new handle and run through a hemolysis test. the pump passed hemolysis testing with an mih of 6.64. this confirmed that a product malfunction was not the cause of the hemolysis. the root cause of the hemolysis was most likely patient condition as the patient had other devices in place. this report is being filed as part of a retrospective review of historical records.
16404 | 1220648-2024-21344 | 20541724 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541724&pc=PYX | 08/27/2022 | 10/27/2024 |  | 10/28/2024 |  |  | 08/27/2022 | 10/28/2024 |  |  | 12/31/2023 |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022156756 |  | 004334 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Sense | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient undergoing cardiac surgery was implanted with an impella rp device for mechanical circulatory support. the impella rp had placement signal issues. the device was explanted and they physician decided to place the proteck duo. at the pump explant thrombosis was observed. | [5 empty] | The logs show spikes in motor current followed by an elevated placement signal, elevated motor current, and decreased flows. the product was returned and a large thrombus was found in the inlet cage interacting with the back of the differential pressure sensor. the pump was unable to start until the clot was ejected. once ejected the pump ran without issue, the placement signal was stable and without issue. the root cause of the placement signal issue was determined to be ingested biomaterial. this report is being filed as part of a retrospective review of historical records.
16405 | 1220648-2024-21345 | 20541726 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541726&pc=PYX | 02/27/2023 | 10/27/2024 |  | 10/28/2024 |  |  | 02/27/2023 | 10/28/2024 |  |  | 07/31/2024 |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023276893 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. while on support there was low pump flow. attempts to resolve the issue by reseating the white cable only marginally improved the problem. the patient was ultimately escalated to a protek duo. no harm to the patient resulting from low pump flows. | [5 empty] | The investigation for the reported low pump flow has been completed. the product was returned and the low pump flow was confirmed. the root cause of the low flow was determined to be due to biomaterial ingestion. this report is being filed as part of a retrospective review of historical records.
16406 | 1220648-2024-21348 | 20541734 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541734&pc=PYX | 03/11/2023 | 10/27/2024 |  | 10/28/2024 |  |  | 03/11/2023 | 10/28/2024 |  |  | 07/31/2024 |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023281625 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Consumer | Health Professional | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. while on support there was lower than expected pump flow. the pump was removed and support continued via another manufacture's device (protek). there was no reported patient harm. | [5 empty] | The investigation for the reported low pump flow has been completed. the product was returned and the low pump flow was confirmed. the root cause of the sustained low pump flow was determined to be ingested biomaterial. the source of the ingested biomaterial was most likely the venous ecmo cannula which was removed shortly after starting impella rp flex support and around the time of ingestion. this report is being filed as part of a retrospective review of historical records.
16407 | 1220648-2024-21349 | 20541737 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541737&pc=PYX | 03/20/2023 | 10/27/2024 |  | 10/28/2024 |  |  | 03/20/2023 | 10/28/2024 |  |  | 01/31/2025 |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024305320 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. while on support there was elevated motor current and decreased pump flow post initiation. the pump was removed and a clot ingestion was noted. support continued via a protek duo device. there was no patient harm reported. | [5 empty] | The investigation for the reported low pump flow has been completed. the product was returned and the low pump flow was confirmed. the root cause of the low flow was determined to be ingested biomaterial, which likely originated form preexisting cannulation/lines. this report is being filed as part of a retrospective review of historical records.
16408 | 1220648-2024-21350 | 20544095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20544095&pc=OZD | 10/04/2024 | 12/05/2024 |  | 10/28/2024 |  |  | 10/04/2024 | 10/28/2024 |  |  |  |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025517071 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 pump being delivered and the 5.5 pump failed to deliver. there were multiple attempts made to deliver the 5.5. pump but it would not advance beyond the distal innominate. after multiple attempts with buddy wire, the decision was made to remove it and implant a impella cp. there was no harm to the patient. | [5 empty] | The investigation for the impella 5.5 has been completed. the device was returned for evaluation. clinical details provided were sufficient to determine the root cause. the root cause of delivery issue is determined to be patient condition because the pump was unable to advance beyond the innominate artery. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
16409 | 1220648-2024-21351 | 20544210 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20544210&pc=OZD | 07/11/2022 | 10/08/2025 |  | 10/28/2024 |  |  | 07/11/2022 | 10/28/2024 |  |  | 02/29/2024 |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191091 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Life Threatening | Death | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had poor conditions and massive bleeding due to liver injury, resulting in disseminated intravascular coagulation (dic), a do not attempt resuscitation (dnar) policy was taken. the family made the choice to make the patient a dnr and the patient expired. use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records. | G1 reporting contact email revised on manufacturer device report 1220648-2024-21351 in accordance with updated procedures. h6 investigation conclusions code was erroneously entered on the initial manufacturer device report number 1220648-2024-21351. upon further review there was heparin in the purge solution and in accordance with updated procedures, the case was re-evaluated as death. the following fields were changed from the initial manufacturer device report number 1220648-2024-21351. b2 death and date of death. h1 type of report revised to death. h6 health effect - impact code 1802 included.
16410 | 1220648-2024-21352 | 20544351 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20544351&pc=OZD | 10/07/2024 | 11/27/2024 |  | 10/28/2024 |  |  | 10/07/2024 | 10/28/2024 |  |  |  |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025536060 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had placed a cp for the support of a patient admitted in acute myocardial infarction / cardiogenic shock, but due to persistent bleeding, it was explanted after just over 4 hours. the team had attempted to remedy with manual hold. the patient survived. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device has been discarded by the customer, therefore, investigation of the device is not possible. should any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16411 | 1035166-2024-00115 | 20544660 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20544660&pc=DYB | 10/10/2024 | 12/31/2024 |  | 10/28/2024 |  |  | 10/14/2024 | 10/28/2024 |  |  |  |  | 10/28/2024 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3021 |  |  | OSCOR INC. |  |  | US |  | DP-20999 |  | 0052-3021 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Distributor | [5 empty] | NA |  | Unknown |  |  |  | Injury | Yes | No | Material Integrity Problem | [6 empty] | Arrhythmia | Cardiac Arrest | Low Blood Pressure/ Hypotension | [10 empty] | Life Threatening | Hospitalization | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | We have been notified of a complaint involving an oscor 23 fr introducer (material #0052-3021, lot # dp-20999). the complainant reported that complications were encountered during attempted delivery of an impella rp pump. upon initial insertion, resistance was met while attempting to advance the pump and the cannula kinked. the end of the peel away sheath (23 fr introducer) subsequently split, and the pump was pulled back for a second attempt. the patient began to decompensate and became hypotensive. the pump was again inserted, but the implant was unsuccessful, and both the pump and introducer had to be removed. the patient went bradycardic and cardiopulmonary resuscitation was initiated until the patient stabilized. a new pump and introducer were then utilized for a successful implant. regulatory reporting required for the report of a damaged introducer, delivery issues, and patient decompensation during delivery complications resulting in intervention. the patient was admitted for acute respiratory failure, arrythmia, and hypotension.cpr was initiated when the patient became bradycardic during delayed implant.the physician believed that the left sided access combined with the patient's enlarged heart caused too much resistance and resulted in the impella kinking and the sheath splitting. | (b)(6) 2024: customer reported the following information: there was no extension or delay in surgical procedure resulting from the failure. |  |  |  |  | Conclusion not yet available, evaluation in process. a follow-up report will be submitted as soon as the investigation is complete. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | The following sections were updated in follow-up 1. b4, b5, g3, g6, h2, h6 & h11. no product was returned to oscor inc. for evaluation; however, a complaint notification was provided. three attempts were made to obtain the device, but the device has not returned for evaluation. without the return of the actual device in question for evaluation, oscor inc. is unable to determine the exact cause for this incident. a review of the manufacturing processes indicates the production operators are instructed to 100% visually inspect for any obvious defect prior to shipment. at this time, it is not possible to assign a definitive root cause for the event as reported. the complaint is non-verifiable as the product was not returned for evaluation. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
16412 | 1220648-2024-21353 | 20544762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20544762&pc=OZD | 10/07/2024 | 12/09/2024 |  | 10/28/2024 |  |  | 10/07/2024 | 10/28/2024 |  |  |  |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025527744 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 27 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported high purge pressure alarm occurred on impella cp. no kinks were found. the impella was reduced to p0 and restarted at p9 while motor current was monitored, and blockage was not cleared. tissue plasminogen activator was added to the purge. after approximately 30 minutes, flows and pressures returned to normal limits. the patient survived the event. | [5 empty] | The investigation into high purge pressure has been completed. the device was not returned for evaluation. however, clinical details provided were sufficient to determine the most likely root cause. the root cause of the high purge pressure was most likely biomaterial. g1 reporting contact email was reported incorrectly on manufacturer device report 1220648-2024-21353. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed.
16413 | 1220648-2024-21354 | 20545184 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20545184&pc=OZD | 10/07/2024 | 12/10/2024 |  | 10/28/2024 |  |  | 10/07/2024 | 10/28/2024 |  |  |  |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025539620 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient had an attempted impella 5.5 implant, however the pump would not advance impella past innominate. multiple attempts were made and the implant was aborted. the patient¿s outcome is unknown. | [5 empty] | The investigation into delivery issues has been completed. the device was returned for evaluation. clinical details provided were sufficient to determine the most likely root cause. the root cause of the delivery issue was most likely patient condition related. the following have been reported incorrectly on manufacturer device report 1220648-2024-21354: g1 reporting contact email. h6 code 4627 was reported incorrectly. new code was added to health effect - impact code. | The impella device has been returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use state the following: section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
16414 | 1220648-2024-21355 | 20545200 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20545200&pc=OZD | 10/07/2024 | 11/27/2024 |  | 10/28/2024 |  |  | 10/07/2024 | 10/28/2024 |  |  |  |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025549920 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The us complainant had a cp pump placed to support a patient admitted in with cardiomyopathy and cardiogenic shock. the patient had support for 4 days until wean and explant. during the support there was bleeding/persistent ooze that was treated by red blood cell infusion. the patient survived the event. | [5 empty] | The investigation for the reported access site bleed has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the access site bleed could not be determined. | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16415 | 1220648-2024-21356 | 20545927 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20545927&pc=PYX | 04/03/2022 | 10/28/2024 |  | 10/28/2024 |  |  | 04/03/2022 | 10/28/2024 |  |  | 01/31/2024 |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023173465 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. there was a delivery/guidewire issue. difficult anatomy resulted in lost position. positioning was re-gained via product trouble-shooting options. decision made by physician to pull a new impella when friction was noted at wire-pigtail connection. the wire and pigtail had a "friction". no patient harm was reported. | [5 empty] | Investigation into the reported delivery issue has been completed. no device was received for evaluation; however, clinical information was sufficient to determine the root cause of the issue. the cause of the issue was patient condition related. this report is being filed as part of a retrospective review of historical records.
16416 | 1220648-2024-21357 | 20546423 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20546423&pc=PYX | 04/12/2022 | 10/28/2024 |  | 10/28/2024 |  |  | 04/12/2022 | 10/28/2024 |  |  | 05/31/2023 |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  |  |  | 2022103196 |  | 004334 |  | Device was returned to manufacturer |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. there was a failure to deliver the rp due to anatomic limitations. the procedure was aborted. no patient harm was reported. | [5 empty] | The investigation into the reported delivery issue has been completed. no device was received for evaluation; however, the clinical information was sufficient to determine the root cause of the issue. the cause of the issue was patient condition related. as noted in the impella ifu: impella rp smartassist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves. warning: pre-support evaluation axillary insertion of the impella 5.5 catheter direct aortic insertion use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings warning caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions warnings & cautions caution ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings warning: axillary insertion of the impella cp ® with smartassist ® catheter inserting the impella rp flex with smartassist® system catheter inserting the impella rp® with smartassist® system catheter. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ axillary insertion of the impella 5.5 catheter ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves warning: pre-support evaluation axillary insertion of the impella 5.5 catheter direct aortic insertion use of impella with transcatheter aortic valves warning caution: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella rp & rp flex rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: potential adverse events (united states) section: potential adverse events (united states ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
16417 | 1220648-2024-21358 | 20546609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20546609&pc=PYX | 03/28/2023 | 10/28/2024 |  | 10/28/2024 |  |  | 03/28/2023 | 10/28/2024 |  |  | 08/31/2024 |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023260151 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella rp device for mechanical circulatory support. it was reported that during impella placement, high purge pressure was observed along with a purge system alarm. a leak at the purge luer lock was noted. the medical staff performed a purge sidearm bypass due to this issue. the patient remained on impella support and was successfully weaned. | [5 empty] | The investigation for the reported high purge pressure and purge leak has been completed. the returned product was visually inspected and purge bypass was observed but the detached sidearm was not returned. the outflow was inspected and biomaterial was observed blocking the purge gap. the cause of the high purge pressure was most likely biomaterial. the cause of the purge leak was not determined since product was not returned. impella rp with smartassist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (eg, infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir.¿ this report is being filed as part of a retrospective review of historical records.
16418 | 1220648-2024-21360 | 20547256 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20547256&pc=PYX | 03/23/2023 | 10/28/2024 |  | 10/28/2024 |  |  | 03/23/2023 | 10/28/2024 |  |  | 12/31/2024 |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024295639 |  | 0046-0035 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock needing an impella rp device for mechanical circulatory support. it was reported that during an attempt to place the pump, the device could not be delivered due to the patient's vasculature. the procedure was aborted. there was no reported harm or injury to the patient. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
16419 | 1220648-2024-21359 | 20547401 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20547401&pc=PYX | 03/21/2023 | 10/28/2024 |  | 10/28/2024 |  |  | 03/21/2023 | 10/28/2024 |  |  | 12/31/2024 |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024305821 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock needing an impella rp device for mechanical circulatory support. during insertion the rp was unable to be delivered due to the patient's anatomy. plans were to provide the patient with bipella support; however, since the rp could not be delivered, the patient remained on cp support alone. there was no reported harm or injury to the patient. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to patient condition. this report is being filed as part of a retrospective review of historical records.
16420 | 1220648-2024-21362 | 20547923 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20547923&pc=OZD | 07/17/2022 | 10/28/2024 |  | 10/28/2024 |  |  | 07/17/2022 | 10/28/2024 |  |  | 02/29/2024 |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023198978 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Injury | No | Yes | Detachment of Device or Device Component | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. after the impella was implanted, the physicians noticed blood leaking from the red handle at the connection between the patient cable and the red handle. the device was explanted and replaced with a new device. survival outcome at explant was that the patient survived. | [5 empty] | Investigation into the reported issue has been completed. the device was received for evaluation. a visual inspection of the pump revealed a hole in the catheter between the 82 and 83 cm marks. the cause of the product damage was blood ingress due to a hole in the catheter resulting from improper use. this report is being filed as part of a retrospective review of historical records.
16421 | 1220648-2024-21363 | 20548142 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20548142&pc=OZD | 02/27/2023 | 10/28/2024 |  | 10/28/2024 |  |  | 02/27/2023 | 10/28/2024 |  |  | 09/30/2024 |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023275804 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. low flows/suction were noted due to a kinked cannula. the team made the choice to remove the impella. it was not replaced due to anatomy contributing to the kink. the patient was placed on an intra-aortic balloon pump. additional information noted that the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation into the reported issue has been no completed. no device was received for evaluation. data logs showed that there was low flow and non-pulsatile motor current, accompanied by suction alarms, almost immediately after pump activation. the alarms do not resolve before the pump was explanted. after reviewing the clinical information, it was determined that the cause of the low pump flow was most likely a kinked cannula due to the patient's anatomy. this report is being filed as part of a retrospective review of historical records.
16422 | 1220648-2024-21365 | 20548257 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20548257&pc=OZD | 04/25/2023 | 10/13/2025 |  | 10/28/2024 |  |  | 04/25/2023 | 10/28/2024 |  |  | 11/30/2024 |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024306791 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Tachycardia | Vascular Dissection | [11 empty] | Other | Death | Life Threatening | Required Intervention | [21 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support the patient experienced ventricular tachycardia (vt) and pulmonary hemorrhage. the patient received atrial lead placement given recurrence of vt. upon return to icu, she experienced a hypoxia arrest secondary to pulmonary hemorrhage. the patient experienced another episode of vt and pacing was increased from 70 to 100. given recurrent vt and pulmonary hemorrhage, decision was made between family and medical team to transition towards comfort measures only. additional information was provided that noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system. section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The root cause of the vascular damage peripheral vessel could not be determined as insufficient clinical details were provided. the root cause of the vf/vt/af/at could not be determined as insufficient clinical details were provided. the cause of low pump flow issue was biomaterial. b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2024-21365. b2 death was inadvertently omitted on the initial manufacturer device report number 1220648-2024-21365. g1 reporting contact email revised on manufacturer device report 1220648-2024-21365 in accordance with updated procedures. h6 medical device problem code 2993 was erroneously entered on the initial manufacturer device report number 1220648-2024-21365.
16423 | 1220648-2024-21364 | 20549706 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20549706&pc=OZD | 03/14/2023 | 10/28/2024 |  | 10/28/2024 |  |  | 03/14/2023 | 10/28/2024 |  |  | 09/30/2024 |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023276825 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced an access site bleed that was treated by sheath advancement and suture. additionally, during the support, the aic alarm failed which prompted the cp to stop. the team replaced the aic on the pump patient who was pump dependent. the patient was supported appropriately after the aic was replaced, per ifu recommendations for the use a backup aic, until the patient expired after being made a do not resuscitate. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation into the reported issue has been completed. no device was received for evaluation. the root cause of the access site bleeding was use related and related to the improper fixation of the pump. as noted in the impella ifu: impella cp with smartassist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states) . ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. this report is being filed as part of a retrospective review of historical records.
16424 | 1220648-2024-21361 | 20550071 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20550071&pc=OZD | 05/01/2023 | 10/28/2024 |  | 10/28/2024 |  |  | 05/01/2023 | 10/28/2024 |  |  | 09/30/2024 |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023275797 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support there was a low flow issue. it was noted that the measured impella flow went down to 0.0. the support level was increased in an attempt to troubleshoot the issue, however, the rate remained 0.0. as a result of the ongoing issue, the decision was made to replace the device with another impella cp pump. additional information was provided that noted the patient expired. medical safety review of the event for the patient's demise noted use of the impella device cannot conclusively be associated with the [death] outcome. the patient's peri-procedural condition was critical. | [5 empty] | The investigation for the reported low pump flow has been completed. the product was returned and the low pump flow was confirmed. the cause of the low pump flow was ingested biomaterial. this report is being filed as part of a retrospective review of historical records.
16425 | 1220648-2024-21371 | 20550911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20550911&pc=OZD | 10/06/2024 | 12/10/2024 |  | 10/28/2024 |  |  | 10/06/2024 | 10/28/2024 |  |  |  |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025541820 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella cp support. while on support, the patient experienced atrial fibrillation (af) with rapid ventricular response (rvr). patient was given loading dose of amio and continued drip. additionally, the patient had large, bloody bowel movement. a decision was made to explant impella to stop heparin and better manage gastrointestinal (gi) bleed. additionally, the doctor performed internal balloon tamponade of impella sheath arteriotomy using a 6 x 40 balloon. he inflated and left for 10 minutes and dropped to assess. he did this for approximately 35 minutes. he then placed a 7.0 x 50 viabahn covered stent. the patient survived the event. | [5 empty] | The investigation for the gastrointestinal bleed, vascular damage, and atrial fibrillation (af) has been completed. product was not returned and logs are not applicable. clinical details were not sufficient to determine the roo cause. therefore, the root cause of the cause of the gi bleed, vascular damage, and af could not be determined since insufficient clinical details were provided. corrections to the initial mfr report: g1 mdr reporting contact email | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use state the following: section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
16426 | 1220648-2024-21372 | 20551277 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20551277&pc=OZD | 02/21/2023 | 10/28/2024 |  | 10/28/2024 |  |  | 02/21/2023 | 10/28/2024 |  |  | 10/31/2024 |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2024288377 |  | 0048-0045 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump was inadvertently dislodged by the staff. as the pump was already scheduled to be removed later that day, the decision was made to skip troubleshooting, and the device was explanted. additional information was provided that noted the family withdrew care and the patient subsequently expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported positioning issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, the root cause of the positioning issue was use issue as the impella was dislodged earlier. this report is being filed as part of a retrospective review of historical records.
16427 | 1220648-2024-21366 | 20551282 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20551282&pc=OZD | 03/29/2022 | 10/28/2024 |  | 10/28/2024 |  |  | 03/29/2022 | 10/28/2024 |  |  | 12/31/2023 |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023180094 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Hispanic |  | Death | Yes | Yes | Restricted Flow rate | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the pump appeared to be in the left ventricle in proper position; however, flows would not go over 0.2 in p2 or on auto. multiple attempts were made to reposition without improvement in flows. the pump was replaced and achieved adequate flows immediately. additional information was provided that noted the patient expired. there was no reported relation of the impella with the patient's outcome. | [5 empty] | The returned product was visually inspected and a kink in the cannula was observed. the long introducer was used in the case indicating a diseased vessel condition causing the kink but that was unable to be confirmed. the cause of the issue was a kinked cannula. this report is being filed as part of a retrospective review of historical records.
16428 | 1220648-2024-21369 | 20551982 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20551982&pc=PYX | 03/27/2023 | 10/28/2024 |  | 10/28/2024 |  |  | 03/27/2023 | 10/28/2024 |  |  | 08/31/2024 |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | RP FLEX W SMART ASSIST SET, US |  |  | ABIOMED, INC. |  |  | US |  | 2024291184 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella rp device for mechanical circulatory support. during impella placement, the pump flows were observed to be low. it was reported that after researching the device history, it was noted the impella pump had ingested a clot the prior afternoon. the device was explanted and replaced with protek. the patient continued on extracorporeal membrane oxygenation support. | [5 empty] | The investigation for the reported low pump flow issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the data logs were returned and confirm a motor current spike, speed deviation, elevation in placement signal, and decrease in flow on the third day of support. this is characteristic of biomaterial ingestion. the product was retuned, and biomaterial was found wrapped around the impeller. the root cause of the low pump flow was determined to be ingested biomaterial. impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. this report is being filed as part of a retrospective review of historical records.
16429 | 1220648-2024-21367 | 20552011 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20552011&pc=OZD | 04/06/2022 | 10/13/2025 |  | 10/28/2024 |  |  | 04/06/2022 | 10/28/2024 |  |  | 01/31/2024 |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023185047 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Cardiac Arrest | Ventricular Fibrillation | [11 empty] | Required Intervention | Life Threatening | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock needing an impella cp device for mechanical circulatory support. during implantation, the patient coded during the process requiring cardiopulmonary resuscitation and four more shocks, return of spontaneous circulation was obtained. in addition, the complainant reported high purge pressure during patient support. additional information was provided that noted the patient expired. there was no reported relation of the impella with the patient's outcome. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | H6 : added "3243-stress problem identified" to coding to account for kink. | High pressure purge: some evidence of biomaterial was seen in the purge gap. the pump was run in the lab and the high purge pressure was not reproduced, with purge pressure at around 750 mmhg. the data logs how a sudden spike in purge pressure indicative of a potential kink in the line, thus the root cause is most likely a purge line kink. vt/vf: this complaint has been identified as part of a retrospective review. due to limited clinical information and lack of available data/product the root cause of this patient coding investigation is not determined. b2 selection death was inadvertently omitted on the initial manufacturer device report number 1220648-2024-21367. g1 reporting contact email revised on manufacturer device report 1220648-2024-21367 in accordance with updated procedures. h1 type of report was erroneously selected on the initial manufacturer device report number 1220648-2024-21367. h6 type of investigation code 4121 and 4112 was inadvertently omitted on the initial manufacturer device report number1220648-2024-21367-1. h10 investigation results was inadvertently omitted on the initial manufacturer device report number 1220648-2024-21367-1.
16430 | 1220648-2024-21370 | 20552322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20552322&pc=OZD | 10/06/2024 | 12/09/2024 |  | 10/28/2024 |  |  | 10/06/2024 | 10/28/2024 |  |  |  |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025539557 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that impella 5.5 was not inserted due to inadequate vessel size. axillary was too small, and 5.5 would no longer advance. the doctor decided to insert an impella cp instead due to anatomy issues. the patient remained stable and no reported patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation into delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because the impella 5.5 was unable to be inserted due to small vessels of the patient and cp was inserted successfully. the following have been reported incorrectly or missing from manufacturer device report 1220648-2024-21370: e4 should have been left blank. g1 reporting contact email.
16431 | 1220648-2024-21380 | 20552533 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20552533&pc=OZD | 12/08/2022 | 10/28/2024 |  | 10/28/2024 |  |  | 12/08/2022 | 10/28/2024 |  |  | 06/30/2024 |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023231727 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Death | Required Intervention | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient was to be weaned off for recovery of his heart function, however the patient expired of sudden septic shock. the physician's comment was the death was related to the impella due to the possibility of infection due to a foreign object had been inserted in the body. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist instructions for use for the related event are as follows: section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
16432 | 1220648-2024-21373 | 20552704 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20552704&pc=PYX | 10/07/2024 | 12/10/2024 |  | 10/28/2024 |  |  | 10/07/2024 | 10/28/2024 |  |  |  |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025540956 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Renal Failure | [11 empty] | Required Intervention | [5 empty] | CONTINUOUS RENAL REPLACEMENT THERAPY | EXTRA CORPOREAL MEMBRANE OXYGENATION | [17 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | Us complainant reported that a patient was on impella rp flex support due to refractory right ventricular failure (rvf). while on support, the patient developed hemolysis. additionally, the patient had renal failure which required continuous renal replacement therapy (crrt). the patient¿s outcome is unknown. | [5 empty] | The impella device has not been received from the customer, therefore, investigation of the device has not been possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation for the hemolysis and the renal failure has been completed. the patient was on va ecmo and crrt. the data logs show that immediately upon activation of the pump there is a small motor current spike followed by a dip in flows and increase in purge pressure. there is a spike in cvp signaling there are pressures on the dp sensor from likely biomaterial. about 10 hours later there is a step up in motor current and a change in flow with no p level change. they are in suction conditions the whole time. about 19 hours later, there is a bag change followed by an increase in motor current, and decrease in flow with no p-level change. the device was returned and there was biomaterial wrapped around the impeller. there were no other abnormalities found. the pump was hemolysis tested and the mih was 6.14. the root cause of the hemolysis is most likely due to biomaterial. the root cause of the renal failure is most likely due to biomaterial. the instructions for use for the impella rp system with automated impella controller & impella rp flex with smart assist system with automated impella controller instructions for use & clinical reference manual: section: potential adverse events (united states): ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ added d9-device return to mfr box checked. corrections to the initial mfr report: g1 mdr reporting contact email.
16433 | 1220648-2024-21376 | 20552856 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20552856&pc=PYX | 11/11/2022 | 10/28/2024 |  | 10/28/2024 |  |  | 11/11/2022 | 10/28/2024 |  |  | 06/30/2024 |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023232128 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was being implanted with an impella rp device for mechanical circulatory support. during implantation, the impella rp was inserted without issue. the wire was removed, and the pump turned on at p2 and immediately the pump stopped and alarmed pump failure due to high motor current. they attempted to restart the pump multiple times with no success. the pump was removed and replaced with new rp and started without issue. | [5 empty] | The investigation for the reported pump did not start issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that data logs showed the "impella stopped - motor current high" alarm triggered immediately upon starting the pump, accompanied by a sudden motor current spike to over 1600. the alarm and motor current spike recurred with every subsequent restart attempt. no damage or abnormalities were found on the returned pump. when the impella was connected to an aic, it started and ran with no issues or relevant alarms. therefore, the root cause of the pump's failure to start was not determined since the issue could not be reproduced. this report is being filed as part of a retrospective review of historical records.
16434 | 1220648-2024-21378 | 20552907 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20552907&pc=OZD | 11/04/2022 | 10/28/2024 |  | 10/28/2024 |  |  | 11/04/2022 | 10/28/2024 |  |  | 05/31/2024 |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2023224893 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Death | No | Yes | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [5 empty] | VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support, the patient showed shock vital signs and dobutamine was discontinued in order to reduce the number of infusions required for hemodynamic stabilization and computer tomography (ct) imaging. after completing the ct and returning to the intensive care unit (icu), the impella's flow rate became insufficient and blood pressure dropped significantly, and a venoarterial extracorporeal membrane oxygenation (va-ecmo) was inserted. it was reported that the patient died and was discharged without improvement in cardiac function. per a note by the attending physician, "there is a high possibility that the flow rate of the impella has decreased due to the collapse of the left ventricle and the rapid deterioration of cardiac function". no further information was provided. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported product damage (cannula kink) has been completed. the product was returned, and the low flow was confirmed. data logs showed there was a simultaneous, temporary drop in motor current, motor speed, and pump flow upon the sudden occurrence of the "suction" alarm about 1 hour into support. there was no change in p-level at this time. subsequently, there were frequent "suction" alarms throughout the remainder of support, accompanied by a loss in motor current pulsatility and a gradual increase (saturation) of flow despite no change in p-level. per the results of the clinical review and investigation: visual inspection of the returned pump revealed a significant mass of biomaterial on and around the outflow cage and the impeller. no other abnormalities were found. in conclusion, it was determined that the cause of the low pump flow was ingested biomaterial. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
16435 | 1220648-2024-21368 | 20552995 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20552995&pc=OZD | 01/18/2023 | 10/28/2024 |  | 10/28/2024 |  |  | 01/18/2023 | 10/28/2024 |  |  | 09/30/2024 |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023277281 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [11 empty] | Life Threatening | Other | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support along with extracorporeal membrane oxygenation (ecmo). while on support, the patient began vomiting followed by facial droop and acute cognitive changes. the team concluded that the patient was having a stroke. the patient was then taken for a computed tomography of the head that revealed a thrombus and then taken to surgery for a thrombectomy. in addition, it was reported that the thrombus could have come from the patient's atrial fibrillation and that systemic anticoagulation has not been on since the implant due to bleeding from the ecmo cannulas. additional information was provided that noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
16436 | 1220648-2024-21383 | 20553080 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20553080&pc=OZD | 05/29/2019 | 11/11/2024 |  | 10/28/2024 |  |  | 10/01/2024 | 10/28/2024 |  |  | 11/30/2019 |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.0 |  |  | OZD |  |  | IMPELLA 5.0 |  |  | ABIOMED, INC. |  |  | US |  | 1317245 |  | 005066-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 48 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Intraoperative Cardiac Valve Injury | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella 5.0 pump delivered to support a patient admitted in with cardiomyopathy and bridging to final care/measures. the impella 5.0 was placed while transplant conversations were ongoing. the support with the impella 5.0 and extracorporeal membrane oxygenation was successful for 14 days when explanted. at the time of explant, the team made an observation of aortic valve damage per a publication written by the medical team. the explant took place for placement of the heartmate ii. the valve was repaired by a park's stitch. the patient survived the event. | [5 empty] | Literature citation: a case of extracorporeal circulation where impella was converted to an implantable ventricular assist device using the mics approach extracorporeal circulation technology 2023; vol.50. no4,458- (0-1-5) h.10: literature citation was inadvertently omitted on the initial manufacturer device report number 1220648-2024-21383 and was added. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
16437 | 1220648-2024-21384 | 20553129 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20553129&pc=PYX | 04/12/2023 | 10/28/2024 |  | 10/28/2024 |  |  | 04/12/2023 | 10/28/2024 |  |  | 06/30/2023 |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022109594 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella rp device for mechanical circulatory support. it was reported the pump stopped. after multiple attempts, the pump was able to be restarted, and the case continued. the device was explanted, and the procedural outcome was the patient survived. | [5 empty] | The investigation for pump stop has been completed. no device was received for evaluation. the root cause of the issue was most likely improper positioning of the device resulting in an interaction with the patient. as noted in the impella instructions for use: impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella rp smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella rp smart assist system with automated impella controller section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
16438 | 1220648-2024-21387 | 20553193 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20553193&pc=OZD | 10/03/2024 | 12/06/2024 |  | 10/28/2024 |  |  | 10/03/2024 | 10/28/2024 |  |  |  |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025525311 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient had attempted to be impella cp implant and it was aborted due to anatomy issues. the impella cp was prepped and inserted into the artery, but it would not advance. the implant was aborted and the patient remained on extracorporeal membrane oxygenation. no patient harm was reported. | [5 empty] | The investigation into delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue is determined to be patient condition because of patient anatomy. h6 code 4627 was reported incorrectly on manufacturer device report 1220648-2024-21387. new code was added to health effect - impact code. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
16439 | 1220648-2024-21377 | 20553194 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20553194&pc=PXY | 12/06/2022 | 10/28/2024 |  | 10/28/2024 |  |  | 12/07/2022 | 10/28/2024 |  |  | 09/30/2024 |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PXY |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023270446 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. while on support, the pump stopped several times, and the team removed and replaced the rp pump. support proceeded with no harm or injury. | [5 empty] | The investigation for the reported pump stop issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that a visual inspection of the pump did not reveal any abnormalities. the pump was attached to an aic and run at p9 and the high motor current pump stop did not reproduce. the pump was also run at p1 and p2 and the pump stop does not reproduce at these low p levels either. the log confirms the high motor current pump stops in the field. the log with the rp speed pulse filter applied shows several motor current spikes during the short pump run, suggesting that there was biomaterial interaction with the impeller. the root cause of the pump stop was most likely biomaterial. impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella rp with smartassist system instructions for use for the related event are as follows: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
16440 | 1220648-2024-21386 | 20553203 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20553203&pc=PYX | 10/03/2024 | 11/11/2024 |  | 10/28/2024 |  |  | 10/03/2024 | 10/28/2024 |  |  | 10/31/2025 |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024421944 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella rp flex support and while on support, the patient had red urine and the pump position was checked. a computed tomography scan was taken showing a clot near the distal outlet cannula in the left pulmonary atresia. the impella rp flex was explanted and a thrombectomy was performed. the patient survived the explant. | [5 empty] | H.6 code 4641 was reported incorrectly on the initial manufacturer device report number 1220648-2024-21386. a new code was added to health effect - impact code. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed.
16441 | 1220648-2024-21388 | 20553240 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20553240&pc=OZD | 10/03/2024 | 11/24/2024 |  | 10/28/2024 |  |  | 10/03/2024 | 10/28/2024 |  |  |  |  | 10/28/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1242768 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Unintended Application Program Shut Down | [6 empty] | Low Cardiac Output | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) shut down completely. the aic was powered down and restarted. the patient¿s mean arterial pressure dropped to 50s during this time as a result of the aic stopping but was otherwise stable. the plan was in place already to escalate to impella 5.5 that day. the patient survived the explant. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | Additional information is listed in g3, g6, h3, h6, h10. the device was returned for evaluation. the console was tested for analog outputs with pump and without pump connected and with test cavity. a thorough visual inspection of the internal components was also performed. the root cause of the unexpected controller shutdown was not determined due to the inability to reproduce.
16442 | 1220648-2024-21374 | 20553269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20553269&pc=OZD | 04/17/2023 | 10/28/2024 |  | 10/28/2024 |  |  | 04/17/2023 | 10/28/2024 |  |  | 06/30/2023 |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 380 14F LT CMR SET (US) |  |  | ABIOMED, INC. |  |  | US |  | 2022099806 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support there was a low flow issue and a suspected clot. the pump was replaced as a result of the noted issue and it was discovered that there was a clot on the outlet cage of the device. additional information was provided that noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported low pump flow and thrombus has been completed. the product was not returned for review but based on clinical description and data review, the root cause of low flow was most likely due to biomaterial ingestion and the root cause of the thrombus was most likely due to patient condition. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
16443 | 1220648-2024-21379 | 20553274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20553274&pc=PYX | 04/11/2023 | 10/28/2024 |  | 10/28/2024 |  |  | 04/11/2023 | 10/28/2024 |  |  | 11/30/2024 |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024286683 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella rp device for mechanical circulatory support. it was reported the pump experienced low pump flows, resulting in this pump being replaced with another impella rp. the patient remained on impella support and was successfully weaned. no patient harm was reported. | [5 empty] | The investigation for the reported low pump flow has been completed. the pump was sent back for further investigation. the returned pump was visually inspected and biomaterial was observed in the purge gap. the root cause of the issue was most likely biomaterial. this report is being filed as part of a retrospective review of historical records.
16444 | 1220648-2024-21381 | 20553278 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20553278&pc=OZD | 03/23/2022 | 10/28/2024 |  | 10/28/2024 |  |  | 03/23/2022 | 10/28/2024 |  |  | 07/31/2023 |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022108798 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced bleeding due to liver injury and expired during support with no issue or causal relationship to the impella device. use of the device cannot conclusively be associated with the outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. this report is being filed as part of a retrospective review of historical records.
16445 | 1220648-2024-21375 | 20553321 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20553321&pc=OZD | 12/05/2022 | 10/28/2024 |  | 10/28/2024 |  |  | 12/05/2022 | 10/28/2024 |  |  | 08/31/2024 |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023266142 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention |  |  |  |  | VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. during support, the patient experienced access site bleeding. the overnight surgeon placed purse sutures, and the pressure held. it was noted that the patient was also placed with a venoarterial extracorporeal membrane oxygenation (vaecmo) that is not an abiomed device and there was oozing from the ecmo cannulas. heparin was withheld. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. this was attributed to the patient's neuro status. | [5 empty] | B2: date of patient death is unknown the investigation for the reported access site bleeding has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the clinical details provided were sufficient to determine a root cause. the root cause is attributed to the patient condition because the patient had several medical and health conditions that were noted prior to the impella implant. the device was implanted according to the instruction for use recommendations. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
16446 | 1220648-2024-21385 | 20553386 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20553386&pc=OZD | 10/01/2024 | 11/25/2024 |  | 10/28/2024 |  |  | 10/03/2024 | 10/28/2024 |  |  |  |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025486421 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support, the patient had hemolysis. plasma free hemoglobin (pfhgb) was increased to 100 and creatine increased from 1.5 - 3.0. the pump position was confirmed and it was not in the papillary. tissue plasminogen activator was added to the purge (no purge issues) and the pfhgb dropped to 40. the patient remains on support. | [5 empty] | The investigation for the reported hemolysis has been completed. the device nor sufficient clinical information was returned for evaluation. therefore, the root cause of the hemolysis could not be determined. | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
16447 | 1220648-2024-21382 | 20553440 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20553440&pc=OZD | 05/25/2023 | 10/28/2024 |  | 10/28/2024 |  |  | 05/31/2023 | 10/28/2024 |  |  | 10/31/2024 |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023277252 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during implantation, the guidewire tip broke while attempting to insert the impella device. a new guidewire was used with no harm to the patient. the survival outcome at explant noted the patient expired due to bleeding from the ecmo cannulation site. use of the device cannot conclusively be associated with the outcome. no patient harm related to the device was identified. | [5 empty] | The investigation for the reported delivery issues issue has been completed. the impella device was discarded by the customer therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the guidewire damage was most likely use related shaping the guide wire since there was no issue observed out-of-box and the damage occurred prior to impella insertion. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
16448 | MW5161762 | 20558299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20558299&pc=OZD | 10/17/2024 | 10/21/2024 |  | 10/28/2024 |  |  |  | 10/28/2024 |  |  |  |  | 10/28/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | ABIOMED IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. | [8 empty] | Yes | [6 empty] | Health Professional |  |  | Voluntary report | Yes | Risk Manager | [6 empty] | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | [7 empty] | Hemorrhage/Blood Loss/Bleeding | Foreign Body In Patient | [11 empty] | Other | Life Threatening | Required Intervention | [26 empty] | Initial submission | [9 empty] | After the placement of the impella, the catheter tip was not in proper position. after multiple attempts to reposition the impella, the md proceeded to remove the device. upon removal the md noticed the tip of the impella was still in the patient and verified with imaging. the md attempted to remove the tip using different approaches. during that time, the patient remained hemodynamically stable but lost a significant amount of blood. pt received 1 unit of prbcs. manual pressure was held throughout the case to assist with hemostasis. after unsuccessful retrieval of the tip, cts was contacted for assistance. it was discussed that the patient needed emergent surgery to repair cfa as well as retrieve the impella tip. a iabp was placed in the rfa in the interim. patient was emergently transferred to the or and tip was removed.
16449 | 1220648-2024-21250 | 20541232 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541232&pc=PYX | 02/19/2023 | 10/27/2024 |  | 10/27/2024 |  |  | 02/19/2023 | 10/27/2024 |  |  | 07/31/2024 |  | 10/27/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023232085 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella rp device for mechanical circulatory support due to pulmonary hypertension. while on support, there was a placement signal not reliable alarm that occurred. white power cable was confirmed to be pressed in all the way. the pump was attempted to be unplugged and then plugged back in. waveforms were restored to baseline for about four hours and then the placement signal unreliable alarm happened again. the decision was made to not try to unplug the power cable again and to continue with the current situation. the patient was seeing improvement since the implant so the decision was made to leave the impella rp as is. there was no patient harm as a result of the failure. | [5 empty] | The investigation for the reported optical signal issue has been completed. the product was not returned. the returned data logs were reviewed and optical signal metrics and cvp were stable. the dp signal showed drift prior to a motor current spike indicating a malfunction of the dp sensor. the cause of the placement signal issue was most likely sensor drift. the product not returned to evaluate the cause of the drift. this report is being filed as part of a retrospective review of historical records.
16450 | 1220648-2024-21251 | 20541241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541241&pc=PYX | 03/10/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 03/10/2022 | 10/27/2024 |  |  | 02/28/2023 |  | 10/27/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022065838 |  | 004334 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was to be implanted with an impella rp device for mechanical circulatory support. there was a failed delivery due to poor visibility. a second device was attempted, but the procedure was ultimately aborted. there was no harm to the patient. | [5 empty] | The investigation into the reported delivery issue has been completed. clinical information has been reviewed and the pump and purge cassette were received for evaluation. the cause of the delivery issue was use related since the delivery was aborted due to poor visibility from the c-arm. as noted in the impella instructions for use: impella rp smart assist system with automated impella controller section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: potential adverse events (united states): ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ this report is being filed as part of a retrospective review of historical records.
16451 | 1220648-2024-21252 | 20541259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541259&pc=OZD | 02/03/2023 | 10/27/2024 |  | 10/27/2024 |  |  | 02/03/2023 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA OPTICAL CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1471015 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | Insufficient Information | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the aic experienced purge disc/flag issues. no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | The investigation into the purge disc/flag issues issue has been completed. the automated impella controller (aic) was returned for investigation. imc logs from the complaint date supported the customer's complaint that the purge cassette was not seen as present in the console. the cause of the issue was a missing purge flag. this report is being filed as part of a retrospective review of historical records.
16452 | 1220648-2024-21253 | 20541275 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541275&pc=OZD | 10/02/2024 | 12/06/2024 |  | 10/27/2024 |  |  | 10/02/2024 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025546137 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp placed for support of a patient admitted with cardiomyopathy, heart failure, and shortness of breath. the access site had an ooze that was remedied by the placement of extra suture, red blood cells infusion and albumin. two days later, the site continued bleeding and the decision was to explant the impella to minimize risk. the patient remained stable and the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation for the access site bleed has been completed. the returned product was visually inspected and no abnormalities were found. the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical details provided.
16453 | 1220648-2024-21254 | 20541276 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541276&pc=OZD | 02/03/2023 | 10/27/2024 |  | 10/27/2024 |  |  | 02/03/2023 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1219723 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was placed onto impella cp support for high risk percutaneous coronary intervention. while on support, the pump experienced a purge issue and controller failure red alarm that were resolved by replacing the aic (automated impella controller). there was no report of patient harm or injury. | [5 empty] | The investigation into the purge disc/flag issues issue has been completed. the automated impella controller (aic) was returned for investigation. imc logs from the complaint date revealed that the console issued the air in purge alarm several times. additionally, the console showed alarm #414 for 43 seconds for purge pressure sensor defective. rt logs show that the purge pressure was adequately maintained before these alarms occurred but then did not recover. device analysis: the console was tested after arriving in field service. the failure mode was unable to be reproduced during investigation. a fluid leak could have caused the failures seen causing the fluctuation in purge pressure. because the device was able to run the purge cassette successfully after switching, it is unlikely to be an error with the cassette itself. the cause of the aic issue could be a defective purge assembly caused by a fluid leak. this report is being filed as part of a retrospective review of historical records.
16454 | 1220648-2024-21255 | 20541289 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541289&pc=OZD | 02/02/2023 | 10/27/2024 |  | 10/27/2024 |  |  | 02/03/2023 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1141734 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient noted as high-risk percutaneous coronary intervention was implanted with an impella device for mechanical circulatory support. when starting the automated impella controller (aic), it said the pump was not recognized by the aic. the device was restarted and recognized the pump at first, but then the pump did not begin functioning. a second pump was used and again was not recognized on the first attempt, but began to function properly after a restart. there was no reported patient harm. | [5 empty] | The device was returned and evaluated, and the investigation for the reported pump issue has been completed. as seen in the device log: case start: pump not detected - power cycle epm" is recorded in the logs during start-up. this is an indication of the epm crystal failure mode. the console was returned to field service but the failure mode was unable to be reproduced which is frequent for this failure mode. the pump detection issue was due to the rarely occurring epm crystal issue. this report is being filed as part of a retrospective review of historical records.
16455 | 1220648-2024-21256 | 20541292 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541292&pc=OZD | 01/31/2023 | 10/27/2024 |  | 10/27/2024 |  |  | 01/31/2023 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1613085 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the aic experienced a purge disc/flag issue. no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | The investigation into the purge disc/flag issues issue has been completed. the automated impella controller (aic) was returned for investigation. the logs for ic showed the purge cassette inserted but no purge disk detected because the purge disc retainer had completely broken off. device analysis: the purge disc retainer was broken which would have caused the inability to complete the case wizard. the cause of the issue was a broken purge disc retainer. this report is being filed as part of a retrospective review of historical records.
16456 | 1220648-2024-21258 | 20541299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541299&pc=OZD | 02/04/2023 | 10/27/2024 |  | 10/27/2024 |  |  | 02/14/2023 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1276243 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella device for mechanical circulatory support. during support, the automated impella controller (aic) produced a brief pump stop message which quickly resolved. the aic was replaced as a preventative measure, and there was no reported patient harm. | [5 empty] | Data logs show that an "impella stopped (rpm deviations" alarm triggered and lasted for 2 seconds before resolving and all performance metrics went to 0. the alarm resolved on its own and the pump restarted and performed with no issues. however, the cause of the pump stop was not determined since the product was not returned. this report is being filed as part of a retrospective review of historical records.
16457 | 1220648-2024-21257 | 20541300 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541300&pc=OZD | 10/02/2024 | 11/25/2024 |  | 10/27/2024 |  |  | 10/02/2024 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025517655 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella cp pump for support and the patient was taken to the operating room (or) for coronary artery bypass graft. while in the or, the pump motor current spiked and the pump stopped and would not re-start after multiple attempts to do so. it was working fine until the last plegia and the medical doctor stated he had placed another cross clamp across aorta and may have caught the motor but was not sure. the pump would not re-start and was removed at end of case. no harm to the patient was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿. section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿. impella cp with smartassist for use during cardiogenic shock. section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿. section: impella stopped. ¿if the impella catheter has stopped suddenly: try to restart the catheter at previous p-level. if the impella does not restart, try to restart at p-2. if the impella does not restart or stops again, wait 1 minute and try to restart again. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿. | The investigation for the pump stop has been completed. data logs were investigated, and the reported pump stop was confirmed when alarm 13 triggered. it occurred just after the team disabled surgical mode and attempted to start the pump flow back up, as they had done several times before successfully. there are also clear attempts to continue to restart the pump, as alarm 13 triggers 17 more times, and alarm 115 triggers 3 times. there were no abnormalities leading up to the complication. returned product was investigated and the pump passed electrical testing on the motor cables. however, the outflow cage had clear deformation. when trying to rotate the impeller past deformed struts, it was not able to pass at all or not without excessive force. this is supported by the report that the physician believes they may have clamped down on the pump just before the pump stop. based on clinical details and returned product, it was determined that the root cause of the pump stop was a damaged outflow cage.
16458 | 1220648-2024-21259 | 20541301 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541301&pc=OZD | 02/08/2023 | 10/27/2024 |  | 10/27/2024 |  |  | 02/08/2023 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1465870 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. while on support, the pump experienced purge disc/flag issues, which were resolved by replacing the aic. there was no report of patient harm or injury. | [5 empty] | The investigation into the purge disc/flag issues issue has been completed. the automated impella controller (aic) was returned for investigation. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
16459 | 1220648-2024-21260 | 20541302 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541302&pc=OZD | 02/06/2023 | 10/27/2024 |  | 10/27/2024 |  |  | 02/06/2023 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1424180 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Female |  |  | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant in the eu had an aic (automated impella controller) failure. the purge disc did not recognize and detect. the team replaced the aic and support proceeded without harm or injury from delay. ifu was followed and backup controller used. | [5 empty] | The investigation into the purge disc/flag issues issue has been completed. the automated impella controller (aic) was returned for investigation. imc logs recorded multiple purge disc not detected alarm prior to the case. every time the purge alarm would be resolved only when purge cassette was removed. visual inspection of the purge assembly area found no issue. a known-good pc and purge disk was inserted without detection issue. the failure mode was unable to be reproduced. using purge commend to check the status of purge system confirmed the pc and disk was recognized. the failure mode was unable to be determined due to unable to reproduce. before returning this console back to the customer, field service was instructed to: replace the purge retainer and purge flag as precaution. this report is being filed as part of a retrospective review of historical records.
16460 | 1220648-2024-21261 | 20541303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541303&pc=OZD | 10/02/2024 | 12/02/2024 |  | 10/27/2024 |  |  | 10/02/2024 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025474727 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during explant of the impella cp, a clot was detected as well as on trans-thoracic echocardiography. the staff confirmed that only 1/3 of the grafts was still patent due to thrombus. an impella 5.5 was elected to be implanted despite obvious clot. the patient is stable and survived the event. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. warnings and cautions: warnings: pre-support evaluation. impella cp with smart assist catheter insertion. axillary insertion of the impella cp with smart assist catheter. use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation for the reported thrombus has been completed. neither the device nor sufficient clinical information was returned for evaluation. the root cause of thrombus could not be determined as the device was not returned nor sufficient clinical details.
16461 | 1220648-2024-21262 | 20541315 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541315&pc=OZD | 12/01/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 12/01/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1330939 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The biomedical engineer reported that the automated impella controller (aic) experienced a controller error yellow alarm. a loaner was sent to the customer to initiate return of this device. there was no patient harm reported. | [5 empty] | E1: name of the initial reporter is unknown the investigation into the controller error has been completed. the impella automated controller was returned for investigation. the console logs from the reported day of event were not available due to file corruption, however, the last available logs from 2022-11-30 had multiple entries indicating sau_powermod_1 communication issues. the log entries are consistent with power battery manager circuit board one communication failures due to corrosion. a visual inspection of the console confirmed the theory, as corrosion was found next to the supercaps c69, c70. the cause of the aic issue was contamination. this report is being filed as part of a retrospective review of historical records.
16462 | 1220648-2024-21263 | 20541317 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541317&pc=OZD | 11/27/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 11/30/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1664272 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The biomedical engineer reported that the automated impella controller (aic) experienced a controller error. a loaner was sent to the customer to initiate return of this device. there was no patient harm reported. | [5 empty] | E1: name of the initial reporter is unknown. the device was returned for evaluation. the console logs from the reported day of event were mostly consistent with the complaint and show that although there were no entries on 2022-11-30, there were two controller error alarms on 2022-11-27, due to opm communication cdc invalid (1045), that were resolved within few minutes. combined with data from long term and real time logs, the investigation was able to determine that the reported issue was a known pattern failure caused by a temporary loss of optical placement signal. there¿s no need for repairs if the condition self-corrects. this is a low probability event and lasts only a few minutes. the cause of the issue was not determined since the issue was not reproduced. this report is being filed as part of a retrospective review of historical records.
16463 | 1220648-2024-21266 | 20541318 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541318&pc=OZD | 11/14/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 01/04/2023 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1675422 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant, the night nurse, reported that the automated impella controller (aic) experienced a controller error. the day shift changed out the aic. the staff reported the patient was extremely non-compliant and had been caught unplugging iv¿s and ¿messing with¿ the purge side arm. the patient had also been refusing to wear bedside monitors for hemodynamics and dressing changes. there was no patient harm reported. | [5 empty] | The investigation into the controller error/loss of opm data has been completed. the impella automated controller was returned for investigation. the controller error and loss of placement signal is due to an optical bench fw communication protocol anomaly. the placement signal not reliable alarm on ic2188 (backup console) were likely due to issues writing to the impella 5.5 on console (b)(6) before it was disconnected. these eeprom writing issues were not reproduced during testing of (b)(6), and the root cause was not determined. this report is being filed as part of a retrospective review of historical records.
16464 | 1220648-2024-21264 | 20541319 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541319&pc=OZD | 11/25/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 11/28/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1337720 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The biomedical engineer reported that the automated impella controller (aic) experienced a controller error. the aic was a loaner that had a note attached indicated the controller error. no patient involvement reported. | [5 empty] | E1: name of the initial reporter is unknown. the investigation into the controller error/loss of opm data has been completed. the impella automated controller was returned for investigation. long term log data and real time log data allowed the investigator to determine that the reported issues (manifested by controller alarms) is a known pattern failure caused by a temporary loss of optical placement signal. there¿s no need for repairs if the condition self-corrects. this is a low probability event and lasts only a few minutes. the cause of the issue was not determined since the issue could not be reproduced. this report is being filed as part of a retrospective review of historical records.
16465 | 1220648-2024-21265 | 20541320 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541320&pc=OZD | 11/22/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 11/28/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1399060 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a controller error. there was no patient harm reported. | [5 empty] | The device was returned for evaluation. the investigation into the controller error/loss of opm data has been completed. the impella automated controller was not returned for investigation. the customer sent back the incorrect console and this was received on (b)(6) 2022. based on the cloud logs, the investigator was able to identify that pump 394490 had been used on console ic6896, and there was a controller error alarm on (b)(6) 2022 (although no evidence of a backup console being used). after analyzing logs from console ic6896, it was determined that the reported issue is a known pattern failure caused by a temporary loss of optical placement signal. there¿s no need to send the console in for repairs if the condition self-corrects. this is a low probability event and lasts only a few minutes. if needed, monitoring of the patient¿s hemodynamics and impella position can be accomplished through imaging. console ic6896 has not been returned to abiomed, but console ic1479 (which happened to be due for pm) had been sent in its place. the cause of the issue was not determined since the product was not returned the issue could not be reproduced. this report is being filed as part of a retrospective review of historical records.
16466 | 1220648-2024-21267 | 20541321 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541321&pc=OZD | 11/21/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 11/21/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1144654 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a controller error. the device was removed from use and a loaner was sent to the customer to initiate return of this device. there was no patient harm reported. | [5 empty] | E1: name of the initial reporter is unknown. the investigation into the controller error has been completed. the impella automated controller was returned for investigation. the data logs from the reported day of event are not consistent with the complaint as there is no ¿controller override error¿ and console has not been used past 2022-10-20. the reported controller errors and alarms were caused by purge assembly failure (piston blocked, pud communication issues), due to the purge motor being immobilized, causing excessive current draw from the power battery manager (pbm), which caused ¿pbm voltage out of spec¿ being reported. the cause of the aic issue was contamination. (purge fluid leak). this report is being filed as part of a retrospective review of historical records.
16467 | 1220648-2024-21268 | 20541322 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541322&pc=OZD | 11/18/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 11/18/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1603745 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported the aic experienced purge disc/flag issues. the resolution for this issue is unknown. no report of patient involvement. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the issue was a defective component. the purge disc retainer broke off. this report is being filed as part of a retrospective review of historical records.
16468 | 1220648-2024-21270 | 20541323 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541323&pc=OZD | 11/20/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 11/20/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1116077 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported the patient was placed on impella 5.5 for hemodynamic support. while on support, the aic experienced purge disc/flag issues that were resolved by replacing the aic. no report of patient injury or harm. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the failure mode of purge disk detection issue was due to broken/ missing purge flag. this report is being filed as part of a retrospective review of historical records.
16469 | 1220648-2024-21269 | 20541324 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541324&pc=OZD | 11/17/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 11/17/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1318441 |  | 0042-0000-CA |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported the aic experienced a purge disc/flag issues. the device was sent to abiomed for repair. this issue occurred during a procedure, no report of patient harm or injury. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. it is most likely that console¿s inability to detect purge cassette or purge disk were caused by intermittent malfunction of the board due to loss of power, which in turn was caused by corrosion on due to the supercaps leaking. this report is being filed as part of a retrospective review of historical records.
16470 | 1220648-2024-21271 | 20541325 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541325&pc=OZD | 11/19/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 11/19/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1075036 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported patient was placed on impella cp for hemodynamic support. while on support, the aic experienced purge disc/flag issues that were resolved by replacing the aic. no report of patient harm or injury. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the issue was a defective component. a physical inspection of the purge assembly was performed, and it was observed that the purge disc flag had broken off. additionally, the purge retainer was found to have micro-fractures. this report is being filed as part of a retrospective review of historical records.
16471 | 1220648-2024-21272 | 20541326 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541326&pc=OZD | 11/17/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 11/17/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2023216840 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that aic error message: purge disc not detected. the complainant was able to successfully exchange for a new console. no patient harm or injury noted. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause of the issue could not be determined since issue could not be reproduced. this report is being filed as part of a retrospective review of historical records.
16472 | 1220648-2024-21273 | 20541327 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541327&pc=OZD | 11/11/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 11/11/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1205049 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a controller error. the device was removed from use and a loaner was sent to the customer to initiate return of this device. there was no patient harm reported. | [5 empty] | E1: name of the initial reporter is unknown. the investigation into the controller error has been completed. the impella automated controller was returned for investigation. the data logs from the reported day of event are consistent with the complaint and show controller error alarm #24 due to loss of communication with the assembly. during initial testing, the alarm was not reproduced, and all aic boot-ups were successful. however, after more thorough testing that included monitoring of the communication between the controller (at42qt1060) and 4-in-1 carrier microcontroller (msc1210), we were able to observe false keyboard readings being interpreted as key presses, and resulting in errors very similar to those observed in the field. it is our conclusion that the kbd glass assembly was defective. the cause of the aic issue was a defective kbd controller assembly. this report is being filed as part of a retrospective review of historical records.
16473 | 1220648-2024-21275 | 20541329 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541329&pc=OZD | 10/21/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 10/21/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1145540 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a broken purge/disc flag retainer was observed on the automated impella controller (aic). the customer did not want a loaner device. no report of patient involvement. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. console purge assembly was inspected, with special attention to the purge disc retainer and purge disc flag. no issues were found with either. the purge assembly was disassembled, and the purge pressure sensor was inspected, as well as the ribbon cable. no issues were found. this report is being filed as part of a retrospective review of historical records.
16474 | 1220648-2024-21274 | 20541330 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541330&pc=OZD | 10/25/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 10/25/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1171906 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) exhibited purge disk failure. the initial device was replaced with a new aic. no patient involvement, this issue was discovered during an in-service. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. visual inspection of the purge assembly area confirmed a tilted/ broken blue retainer. the issue with detecting purge disc was due to defective purge retainer. this report is being filed as part of a retrospective review of historical records.
16475 | 1220648-2024-21276 | 20541331 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541331&pc=OZD | 10/22/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 10/22/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 2023192131 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. the complainant reported the purge cassette was not detected on the automated impella controller (aic). no report of patient harm or injury. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. console purge assembly was inspected. upon opening the purge assembly door, it was confirmed the purge disc retainer had broken/snapped off of the purge assembly. this report is being filed as part of a retrospective review of historical records.
16476 | 1220648-2024-21277 | 20541332 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541332&pc=OZD | 10/19/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 10/19/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1106952 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a purge disc issue on the automated impella controller (aic). no report of patient involvement. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. the purge retainer was confirmed to be broken. this report is being filed as part of a retrospective review of historical records.
16477 | 1220648-2024-21278 | 20541334 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541334&pc=OZD | 10/20/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 10/20/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1126026 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. during support, the device alarmed purge disk not detected. the staff removed and reinserted purge disk, however the alarm did not resolve. the initial aic was replaced with a new device and the case continued. no report of patient harm or injury. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. the cause of the issue was a defective component due to a broken container. this report is being filed as part of a retrospective review of historical records.
16478 | 1220648-2024-21279 | 20541335 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541335&pc=OZD | 10/12/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 10/17/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1313340 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Obstruction of Flow | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the automated impella controller (aic) was experiencing high purge pressures during support with an impella 5.5 pump. the original aic was replaced with another aic and the high purge pressure was resolved. it was noted that the patient was stable, no patient harm or injury reported. | [5 empty] | The investigation into the purge issue has been completed. the impella automated controller was returned for investigation. the data analysis of the logs from the reported day of event show that over time there was a gradual increase of purge pressure and decrease in purge flow. the returned console was tested and confirmed to be within specification. will continue to monitor and trend. this report is being filed as part of a retrospective review of historical records.
16479 | 1220648-2024-21280 | 20541336 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541336&pc=OZD | 10/11/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 10/11/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | XXXXXX |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) exhibited a battery failure message. a loaner device was sent to the customer. no report of patient involvement. | [5 empty] | The console (aic) was returned for evaluation. upon observation of the console, the failure mode was reproduced on an engineering bench. the battery 2 lcd indicator was blank (fully discharged). there was a cell voltage imbalance for battery 2. therefore, battery 2 was defective and the root cause is single cell failure. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
16480 | 1220648-2024-21281 | 20541348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541348&pc=OZD | 10/05/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 10/05/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1146851 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | Other Health Care Professional | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) exhibited a battery failure. a loaner was shipped to the customer, no report of patient involvement. | [5 empty] | The console (aic) was returned for evaluation. upon observation of the device, it was confirmed that persistent battery failure (transport connection blown/missing) was not due to a disengaged battery switch. when visually inspecting the batteries and pbm, it was observed that one of the batteries status leds were completely blank. the root cause of the aic issue was a defective battery. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
16481 | 1220648-2024-21282 | 20541369 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541369&pc=OZD | 10/11/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 10/11/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1500693 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that during preventative maintenance the automated impella controller (aic) exhibited battery failure. the customer did not request a replacement device and there was no patient involvement. | [5 empty] | The console (aic) was returned for evaluation. upon evaluation of the console, the aic was connected through the ac power and turned on; however, the reported failure mode alarm (battery failure) occurred, and the complaint was reproduced. the original batteries were replaced with a set of good batteries. once the controller was booted all battery related alarms were cleared. the root cause of the battery failure alarm was due to depleted batteries. batteries were depleted over time because they were not routinely charged by the customer. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
16482 | 1220648-2024-21286 | 20541421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541421&pc=OZD | 07/02/2023 | 10/27/2024 |  | 10/27/2024 |  |  | 07/02/2023 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1488654 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) purge disk holder was broken and the black pin was visibly out and lying under blue purge disc holder. the holder lifted away from the console. there was no reported patient involvement. | [5 empty] | The investigation for the reported purge disc/flag issues has been completed. the device logs were reviewed and confirmed that after purge cassette insertion, purge fluid not detected error was present. visual inspection of the purge assembly confirmed a broken/detached blue retainer. minimal residue in the vicinity of the purge assembly. the failure cause was a broken/missing purge retainer. this report is being filed as part of a retrospective review of historical records.
16483 | 1220648-2024-21287 | 20541427 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541427&pc=OZD | 09/23/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 09/23/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1363891 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that during a software update, the automated impella controller (aic) exhibited a battery failure alarm. attempts to turn power button off and back on didn¿t resolve the problem. the aic had been plugged in to a red outlet but only registered 50% battery. no patient involvement. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, it was confirmed the reported battery failure was not due to a disengaged battery switch. once the console was booted the console alarms were consistent with what the reported failure issue. therefore, the root cause of the battery failure was due to a defective battery. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
16484 | 1220648-2024-21290 | 20541428 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541428&pc=OZD | 09/12/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 09/12/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1302712 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that upon startup of the automated impella controller (aic), the unit displayed a "battery failure" message despite being plugged into ac power. while attempting to troubleshoot the issue, the unit subsequently displayed a "battery comm. failure" message. the device was removed and replaced with a new console. there was no patient harm resulting from the reported issue. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the aic issue was a defective battery. this report is being filed as part of a retrospective review of historical records.
16485 | 1220648-2024-21284 | 20541429 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541429&pc=OZD | 09/30/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 09/30/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1181669 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Company representation | [5 empty] | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) exhibited a battery failure alarm. when plugging in the console to ac power battery the percentage goes to fifty percent and the fuse for battery number 2 isn't visible. when unplugging the console battery, the percentage goes to one hundred percent, but battery failure alarm keeps going off. no indication of patient involvement, harm, or injury. | [5 empty] | The console (aic) was returned for evaluation. upon evaluation of the console, the reported issue was reproduced. after opening the console, it was determined that a loose screw was located on top of the power battery board (pbm), shorting its electronics. unfortunately, removing the screw did not resolve the issue as the pbm assembly already sustained electrical damage. therefore, a replacement pbm, battery set, and pbm bracket were applied and a functional check; as well as an electrical safety test was performed on the console. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
16486 | 1220648-2024-21283 | 20541430 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541430&pc=OZD | 09/21/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 09/21/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1154602 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The biomedical engineer reported that the automated impella controller (aic) was experiencing a purge cassette error requiring repair. a loaner device was sent to the customer. there was no patient involvement, harm, or injury reported. | [5 empty] | The investigation into the purge issue has been completed. the impella automated controller was returned for investigation. the data analysis of the logs from the complaint date revealed that the console issued the purge disc not detected alarm. the returned console was tested, the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
16487 | 1220648-2024-21288 | 20541432 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541432&pc=OZD | 09/15/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 09/15/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1085222 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) exhibited battery failure. a loaner was sent to the customer. no patient involvement was reported. | [5 empty] | The console (aic) was returned for evaluation. upon evaluation of the console, it was confirmed the persistent battery failure was not due to a disengaged battery switch. once the console was booted the alarms were consistent with the reported failure mode. visual inspection of the batteries revealed one of the batteries leds was completely blank. therefore, the root cause of the battery failure alarm was due to a defective battery 2. the failure modes will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
16488 | 1220648-2024-21285 | 20541433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541433&pc=OZD | 09/28/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 09/28/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER JP |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER JP |  |  | ABIOMED, INC. |  |  | US |  | 1325609 |  | 0042-0000-JP |  | No |  |  | No |  | No | No | Other |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the automated impella controller's (aic) purge motor was not properly applying torque. no report of patient involvement, harm, or injury. | [5 empty] | E1: name of the initial reporter is unknown. the investigation into the purge issue has been completed. the impella automated controller was not returned to the manufacturer for investigation. the data logs were not available for analysis. the japan field service representative reported that when attempting to test the torque specifications of the purge motor during preventive maintenance the motor was not applying torque to specifications. the field service engineer replaced the purge motor and reported that after replacement no issues were observed with motor torque. the motor passed preventive maintenance specifications after replacement. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
16489 | 1220648-2024-21291 | 20541434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541434&pc=OZD | 09/08/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 09/08/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1423376 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 64 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was on an automated impella controller (aic) for mechanical circulatory support. the automated impella controller (aic) exhibited an aic not recognizing the disk error. a replacement device was sent to the customer. it was reported that the procedure was successful' no harm or injury to the patient. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, the issue with properly detecting the purge pressure disc was reproduced, and the purge flag was confirmed to be broken (missing at blue disc retainer). the purge flag was replaced, and a functional check was performed on the console and optical system. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
16490 | 1220648-2024-21292 | 20541436 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541436&pc=OZD | 09/07/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 09/07/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | XXXXXX |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a battery failure, controller error (yellow alarm) and a backup controller was utilized. no reported harm or injury reported. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the aic issue was a defective battery. this report is being filed as part of a retrospective review of historical records.
16491 | 1220648-2024-21289 | 20541437 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541437&pc=OZD | 09/13/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 09/13/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1162036 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) exhibited a battery failure. a loaner was sent to the customer, no report of patient involvement. | [5 empty] | The console (aic) was returned for evaluation. upon evaluation of the console, the persistent battery failure was not due to a disengaged battery switch. once the console was booted, the alarm was consistent with the reported failure alarm. after visually inspecting the batteries and pbm, it was observed one of the batteries leds was completely blank. therefore, the battery failure alarm was due to a defective battery. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
16492 | 1220648-2024-21293 | 20541440 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541440&pc=OZD | 07/03/2023 | 10/27/2024 |  | 10/27/2024 |  |  | 07/03/2023 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1122390 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was placed onto impella cp support due to acute myocardial infarction/cardiogenic shock. while on support, the automated impella controller (aic) experienced a controller error yellow alarm that was resolved by transferring to a new console. the patient survived the procedure. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer. the logs from the complaint date showed a controller error, which confirms the reported failure. when the controller error occurs, the real time log showed the optical signal dropped to ¿-16384,¿ which points to the failure mode transient loss of opm data. the console was tested on an engineering lab bench. the failure mode was not reproduced by running the console with the optical pump. there was no issue with the console hardware. the root cause for the controller error was not determined due to the inability to reproduce. this report is being filed as part of a retrospective review of historical records.
16493 | 1220648-2024-21294 | 20541505 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541505&pc=PYX | 07/28/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 07/28/2022 | 10/27/2024 |  |  | 01/31/2024 |  | 10/27/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023173472 |  | 004334 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the physician was unable to deliver the device due to the patient's anatomy. impella support was aborted without any harm. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received for evaluation; however, the clinical information was sufficient to determine the root cause of the issue. the cause of the issue was patient condition related. as noted in the impella ifu: impella rp smartassist system with automated impella controller: section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. warning: pre-support evaluation: axillary insertion of the impella 5.5 catheter: direct aortic insertion: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: warning caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions: warnings & cautions: caution: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: warning: axillary insertion of the impella cp ® with smartassist ® catheter. inserting the impella rp flex with smartassist® system catheter. inserting the impella rp® with smartassist® system catheter. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ axillary insertion of the impella 5.5 catheter ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves. warning: pre-support evaluation, axillary insertion of the impella 5.5 catheter, direct aortic insertion, use of impella with transcatheter aortic valves. warning: caution: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella rp & rp flex rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller: section: potential adverse events (united states): section: potential adverse events (united states: ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings: ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records. this report is being filed as part of a retrospective review of historical records.
16494 | 1220648-2024-21295 | 20541507 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541507&pc=OZD | 07/12/2023 | 11/17/2025 |  | 10/27/2024 |  |  | 07/12/2023 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1080530 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella 5.5 support kicked the console over and cracked it. the automated impella controller was swapped for a new one. the patient survived the procedure. | [5 empty] | The impella console was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records.
16495 | 1220648-2024-21296 | 20541508 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541508&pc=PYX | 07/28/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 07/28/2022 | 10/27/2024 |  |  | 04/30/2024 |  | 10/27/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023208273 |  | 0046-0035 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the physician was unable to deliver the device due to the patient's anatomy. impella support was aborted without any harm. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received for evaluation; however, the clinical information was sufficient to determine the root cause of the issue. the cause of the issue was patient condition related. as noted in the impella ifu: impella rp smartassist system with automated impella controller section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. warning: pre-support evaluation axillary insertion of the impella 5.5 catheter direct aortic insertion use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings warning caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions warnings & cautions caution ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings warning: axillary insertion of the impella cp ® with smartassist ® catheter inserting the impella rp flex with smartassist® system catheter inserting the impella rp® with smartassist® system catheter ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ axillary insertion of the impella 5.5 catheter ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valveswarning: pre-support evaluation axillary insertion of the impella 5.5 catheter direct aortic insertion use of impella with transcatheter aortic valves warning caution: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella rp & rp flex rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: potential adverse events (united states) section: potential adverse events (united states ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
16496 | 1220648-2024-21297 | 20541511 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541511&pc=PYX | 07/29/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 07/29/2022 | 10/27/2024 |  |  | 03/31/2024 |  | 10/27/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023225328 |  | 0046-0011 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 22 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during the support low pump flow was seen due to clot formation. the pump was removed and replaced. no patient harm was reported. | [5 empty] | The investigation for the reported issue of low pump flow has been completed. the impella device was received for evaluation. a visual inspection of the pump did not reveal any issues. no kinks or biomaterial were found. the failure mode could not be re-created in the lab and no logs were returned. based on the provided clinical information, the root cause of the low pump flow was most likely improper positioning of the device. this report is being filed as part of a retrospective review of historical records.
16497 | 1220648-2024-21298 | 20541512 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541512&pc=PYX | 09/18/2021 | 10/27/2024 |  | 10/27/2024 |  |  | 09/18/2021 | 10/27/2024 |  |  | 01/31/2023 |  | 10/27/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022031788 |  | 004334 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. it was reported that the impella flows were decreasing during support. the clinician noted that the low pump flows may have been due to a thrombus in the catheter. the impella was removed, and manual pressure was held for forty minutes resulting in the patient stabilizing without the support of the pump. there was no patient harm reported. | [5 empty] | The investigation into the low pump flow has been completed. impella rp was not received from the customer. the site reported that pump flows decreased from 3 liters per minute (lpm) to 2 lpm, and placement signal was captured as 100s/90s millimeters of mercury following the placement of a new swan catheter. the root cause of the reported low pump flow/placement signal issue was unable to be determined through this investigation as no product or data logs were provided. this report is being filed as part of a retrospective review of historical records.
16498 | 1220648-2024-21299 | 20541513 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541513&pc=OZD | 02/01/2023 | 10/27/2024 |  | 10/27/2024 |  |  | 02/01/2023 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1451015 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a purge disc/flag issue. no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | The investigation into the purge disc/flag issues issue has been completed. the automated impella controller (aic) was returned for investigation. the logs for ic7663 showed the purge cassette inserted but no purge disk detected. the blue purge retainer was cracked and the purge flag was missing that would have caused the inability to complete the case wizard and have the purge pressure disc recognized. root cause: the cause of the issue was a broken blue receptacle. this report is being filed as part of a retrospective review of historical records.
16499 | 1220648-2024-21301 | 20541519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541519&pc=OZD | 06/14/2023 | 10/27/2024 |  | 10/27/2024 |  |  | 06/14/2023 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1371242 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Nurse | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that the controller showed yellow and red battery failure alarm. this was a servicing issue. no patient case was involved. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the automated impella controller logs confirmed the reported failure. there was a battery failure alarm and a battery 2 communication failure alarm due to loss of communication to the battery. though there is no long term powerdata from the complaint date, it is most likely a cell imbalance with battery 2, which caused the battery pack internal electronics to shut down the battery 2. the console was tested on a lab bench. the failure mode was reproduced while running the automated impella controller (aic) without alternating current power. the battery failure and the communication issues was also confirmed by performing a batttest in the terminal window and by visual inspection of the battery liquid crystal display screen. the failure mode was reproduced on an engineering lab bench. the battery failure alarm was due to a defective battery. the root cause of the defective battery was most likely due to cell imbalance. the cause of the aic issue was a defective battery. this report is being filed as part of a retrospective review of historical records.
16500 | 1220648-2024-21300 | 20541520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541520&pc=OZD | 01/30/2023 | 10/27/2024 |  | 10/27/2024 |  |  | 01/31/2023 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1316939 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  | Asian | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock on an automated impella controller (aic) for mechanical circulatory support. while on support, the aic experienced purge disc/flag issues that were resolved by replacing the aic. | [5 empty] | The investigation into the purge disc/flag issues issue has been completed. the automated impella controller (aic) was returned for investigation. the complaint specifies a "purge pressure transmitter failure alarm" but the alarms seen in the imc logs from the complaint date revealed that the console issued the purge disc not detected alarm several times and was unable to complete case wizard. device analysis: the console was tested after arriving in field service. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be missing. root cause: the cause of the issue was a broken/missing purge flag. this report is being filed as part of a retrospective review of historical records.
16501 | 1220648-2024-21302 | 20541521 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541521&pc=OZD | 01/30/2023 | 10/27/2024 |  | 10/27/2024 |  |  | 01/30/2023 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1122789 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient noted as high risk percutaneous coronary intervention was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the aic experienced a purge disc/flag issue. no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | The investigation into the purge disc/flag issues issue has been completed. the automated impella controller (aic) was returned for investigation. imc logs from the complaint date revealed that the console issued the purge disc not detected alarm several times and was unable to complete case wizard. while the complaint says that the problem was the inability to recognize the purge cassette, the purge cassette was able to be seen by the console but the purge disc was unable to be detected. device analysis: the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be missing. root cause: the cause of the issue was a broken/missing purge flag. this report is being filed as part of a retrospective review of historical records.
16502 | 1220648-2024-21303 | 20541522 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541522&pc=OZD | 12/22/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 12/22/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1164687 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced a pump stop during testing. there was no patient involvement. | [5 empty] | The device was returned and evaluated, and the investigation for the reported pump stop has been completed. imc logs show overnight the pump was tried and resulted in a pump stop. the data log confirmed the occurrence of the pump stop with the demonstration/testing pump, but this pump also failed with another console. the initial pump stop was due to an electrical open circuit causing a malfunctioning pump. this console was tested, but the pump stop failure was unable to be reproduced. speaking with the rep, it is likely that the faulty demo pump caused the intermittent pump stops. this report is being filed as part of a retrospective review of historical records.
16503 | 1220648-2024-21304 | 20541523 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541523&pc=PYX | 12/28/2021 | 10/27/2024 |  | 10/27/2024 |  |  | 01/24/2022 | 10/27/2024 |  |  | 09/30/2023 |  | 10/27/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022137161 |  | 004334 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported by the clinical educator that during the preparation of the impella rp for use the console was showing the pump as undetected, despite the white cable being plugged in and connected to the pump. the console was switched out and the white plug unplugged and plugged back in, however the message "impella device defective" was generated. a new impella rp was opened and successfully used. there was no patient harm reported. | [5 empty] | The investigation into the pump not detected has been completed. the impella rp was returned for investigation. the returned product was visually inspected and broken and twisted pins were observed in the blue handle gray connector. the pump was connected to a console and was unable to be detected. the cause of the defective pump issue was broken pump plug pins due to improper technique by the end user. this report is being filed as part of a retrospective review of historical records.
16504 | 1220648-2024-21305 | 20541527 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541527&pc=OZD | 06/13/2023 | 10/27/2024 |  | 10/27/2024 |  |  | 06/13/2023 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1494422 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient undergoing coronary artery bypass was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the automated impella controller experienced purge disc/flag issues that were resolved by changing the consoles. the patient survived the procedure. | [5 empty] | The investigation for the reported purge disc/flag issues has been completed. the logs showed repeated inability for the purge to go home indicative of a malfunction in the purge system. the purge disc retainer was completely cracked and the purge flag was missing when the console was returned for investigation. the cause of the issue was a broken purge disc retainer and missing purge flag. this report is being filed as part of a retrospective review of historical records.
16505 | 1220648-2024-21306 | 20541528 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541528&pc=PYX | 01/12/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 01/12/2022 | 10/27/2024 |  |  | 01/31/2023 |  | 10/27/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2022035367 |  | 004334 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support. it was reported during placement of the pump the physician had difficulty advancing the pump. reportedly when the physician pulled the catheter back a kink was observed close to the motor, additional attempts to advance the pump past the sheath on reentry were unsuccessful. this pump placement was aborted, and a replacement pump was successfully placed. there was no patient harm reported. | [5 empty] | The investigation into the product damage (cannula kink) has been completed. the impella rp was returned for investigation. the returned product was visually inspected and damage to the outflow cage could be seen, and two significant kinks were also found in the catheter of the device. the root cause of the damaged outflow cage was most likely due to patient condition as the physician was unable to advance the pump.the pump was observed to have a deformed outflow cage but did not result in a pump failure. the cause of the product damage was a catheter kink. the root cause of the damaged outflow cage was most likely due to patient condition as the physician was unable to advance the pump past the pv. this report is being filed as part of a retrospective review of historical records.
16506 | 1220648-2024-21307 | 20541532 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541532&pc=OZD | 05/30/2023 | 10/27/2024 |  | 10/27/2024 |  |  | 05/30/2023 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1162413 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a battery failure ¿ red alarm, which was resolved by replacing the console. no patient was involved. | [5 empty] | The investigation for the reported aic - battery failure (red alarm) issues has been completed. the controller device has been returned for evaluation. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
16507 | 1220648-2024-21308 | 20541536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541536&pc=OZD | 05/15/2023 | 10/27/2024 |  | 10/27/2024 |  |  | 05/15/2023 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1356952 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a battery failure red alarm. no patient was involved. | [5 empty] | The investigation for the reported aic - battery failure (red alarm) issues has been completed. the controller device has been returned for evaluation. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
16508 | 1220648-2024-21309 | 20541537 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541537&pc=OZD | 06/13/2023 | 10/27/2024 |  | 10/27/2024 |  |  | 06/13/2023 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1273766 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the automated impella controller produced a purge alarm and the purge disk mount was found to be broken. no resolution was specified. the patient successfully was weaned from the unit and survived the explant. | [5 empty] | The investigation for the reported purge disc/flag issues has been completed. the logs showed during the case, a purge disc not detected error while the pump and purge cassette were inserted. this alarm persisted despite efforts to reinsert the purge cassette. additionally, there were piston blocked errors in the logs. the purge disc retainer was completely cracked and the purge flag was missing when the console was returned for investigation. additionally, there was a glucose ingress. the cause of the issue was a broken purge disc retainer and missing purge flag. this report is being filed as part of a retrospective review of historical records.
16509 | 1220648-2024-21310 | 20541538 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541538&pc=OZD | 05/08/2023 | 10/27/2024 |  | 10/27/2024 |  |  | 05/08/2023 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1288948 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a battery failure red alarm. no patient was involved. | [5 empty] | The investigation for the reported aic - battery failure (red alarm) issues has been completed. the controller device has been returned for evaluation. clinical details were sufficient to determine the root cause. the battery failure was due to a defective battery 2. the root cause of the defective battery 2 was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
16510 | 1220648-2024-21311 | 20541540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541540&pc=OZD | 05/09/2023 | 10/27/2024 |  | 10/27/2024 |  |  | 05/09/2023 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1400770 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was placed onto impella cp support due to acute myocardial infarction / cardiogenic shock. while on support, the automated impella controller (aic) experienced a battery failure red alarm that resolved on its own. no patient harm was reported. | [5 empty] | The investigation for the reported aic - battery failure (red alarm) issues has been completed. the controller device has been returned for evaluation. clinical details were sufficient to determine the root cause. complaint cause is most likely due to pbm incorrectly reporting battery temperatures. this report is being filed as part of a retrospective review of historical records.
16511 | 1220648-2024-21312 | 20541541 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541541&pc=OZD | 03/14/2023 | 10/27/2024 |  | 10/27/2024 |  |  | 03/14/2023 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1299745 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a battery failure red alarm. no patient was involved. | [5 empty] | The investigation for the reported aic - battery failure (red alarm) issues has been completed. the controller device has been returned for evaluation. clinical details were sufficient to determine the root cause. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
16512 | 1220648-2024-21313 | 20541542 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541542&pc=OZD | 12/13/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 12/13/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1384825 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a purge disc/flag issue. no patient was involved. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the issue was a broken retainer. this report is being filed as part of a retrospective review of historical records.
16513 | 1220648-2024-21314 | 20541543 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541543&pc=OZD | 12/16/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 12/16/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1468305 |  | 0042-0010 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was placed onto impella support for high-risk percutaneous coronary intervention. while on support, the automated impella controller (aic) experienced purge disc/flag issues that were resolved by switching out the aic. no patient harm was reported. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has not been returned for evaluation. the cause of the issue was not determined since product was not returned. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
16514 | 1220648-2024-21315 | 20541544 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541544&pc=OZD | 12/27/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 12/27/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1071915 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was placed onto impella 5.5 support due to cardiomyopathy. while on support, the automated impella controller (aic) experienced purge disc/flag issues with no resolution specified. no patient harm was reported. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. upon visual inspection, the purge disc flag was missing/broken. the cause of the issue was a broken/missing purge flag. this report is being filed as part of a retrospective review of historical records.
16515 | 1220648-2024-21316 | 20541545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541545&pc=OZD | 12/13/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 12/13/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1576159 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a purge disc/flag issue that was resolved by swapping out the console. no patient was involved. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the issue was a broken purge retainer. this report is being filed as part of a retrospective review of historical records.
16516 | 1220648-2024-21317 | 20541551 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541551&pc=OZD | 12/07/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 12/07/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1319820 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automated impella controller (aic) experienced purge disc/flag issues. no patient was involved. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
16517 | 1220648-2024-21318 | 20541552 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541552&pc=OZD | 11/30/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 11/30/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1431278 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported, the automated impella controller (aic) experienced purge disc/flag issues. no patient was involved. | [5 empty] | The investigation for the reported aic-purge disc/flag issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
16518 | 1220648-2024-21319 | 20541553 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541553&pc=OZD | 12/07/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 12/07/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1308008 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility biomedical engineer reported the automated impella controller (aic) experienced purge disc/flag issues. no patient was involved. | [5 empty] | The investigation for the reported aic - purge disc/flag issues has been completed. the impella device has been returned for evaluation. clinical details were sufficient to determine the root cause. the cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
16519 | 1220648-2024-21320 | 20541554 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541554&pc=OZD | 03/27/2023 | 10/27/2024 |  | 10/27/2024 |  |  | 03/27/2023 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1107441 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller console had a blank screen. a new console was used to address the issue. no patient was involved. | [5 empty] | The investigation for the reported issue has been completed. the logs are consistent with the complaint reported, and another failure mode was discovered. the patient case started on march 23 late at night. an impella 5.5 pump was implemented without issue for 14 hours until noon the next day. suddenly, a ¿placement signal not reliable" alarm presented, with an invalid raw optical placement signal as real time logs shown. one minute after the placement signal alarm, the console recognized a running pump disconnection. due to the mechanism/criteria of the pump disconnection, it was likely due to a partially disconnected pump connector at the automated impella controller (aic) blue receptacle, causing an unreliable placement signal. this pump was likely fully electrically disconnected moments later. later, the pump was pushed fully in, and the console recognized the pump connection. once the pump eeprom was read, the ¿pump running time¿ parameter was not updated/read correctly. the pump was able to run for almost another hour until the similar ¿running pump disconnection¿ condition occurred with the same invalid placement signal. afterwards, the pump was reconnected. however, this time the same pump was unable to successfully read the eeprom flow characteristic curves 2 even with 3 retries. simultaneously, the console¿s software (sw) processes stopped heartbeat, which triggered a build-in remedy to perform a sw restart on the console with a ¿core dump¿ information in logs upon next boot up. although, the console rebooted without issue, eeprom reading for calibration data of the pump repeatedly failed until the console was exchanged for a backup. the console was received for investigation and repair. the console was tested with a known good pump for at least 300 power cycles of impella 5.5 pump connection and disconnection with usage data updated for each cycle. however, no eeprom writing or reading issue was reproduced during testing. the reported failure mode ¿ unexpected console reboot was due to a sw restart to mitigate the sw processes crash. the root cause of eeprom reading issue was unable to be determined due to unable to be reproduced. this report is being filed as part of a retrospective review of historical records.
16520 | 1220648-2024-21321 | 20541556 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541556&pc=OZD | 04/06/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 04/06/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1503160 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. the automated impella controller (aic) experienced a controller error yellow alarm, which was resolved by changing the console. no patient harm reported. | [5 empty] | The investigation has been completed. there was no issue found during the case. the alarm presented was result of unusually low ambient pressure and console's reaction was as designed. this report is being filed as part of a retrospective review of historical records.
16521 | 1220648-2024-21322 | 20541557 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541557&pc=OZD | 03/25/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 03/25/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1210876 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support, the automated impella controller (aic) experienced a purge disc/flag issue that was resolved by changing the console. | [5 empty] | The investigation has been completed. the console (ic3990) was sent back for further investigation. the purge disc retainer and flag were confirmed to be broken. the root cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
16522 | 1220648-2024-21324 | 20541558 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541558&pc=OZD | 03/07/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 03/07/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1420348 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. while on support, the automated impella controller (aic) experienced a purge disc/flag issue that was resolved by changing the console. no patient harm reported. | [5 empty] | The investigation has been completed. the console (ic7300) was sent back for further investigation. the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. the root cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
16523 | 1220648-2024-21323 | 20541559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541559&pc=OZD | 03/19/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 03/19/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1250486 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced purge disc/flag issues, which was resolved through backup console. no patient was involved. | [5 empty] | The investigation has been completed. the purge disc retainer was confirmed to be broken which likely effected the sensor¿s ability to detect the purge disc. the root cause of the purge component detection issues was due to a broken purge disc retainer. this report is being filed as part of a retrospective review of historical records.
16524 | 1220648-2024-21325 | 20541560 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541560&pc=OZD | 02/23/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 02/23/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1151535 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced purge disc/flag issues. no clinical details regarding resolution or patient involvement/condition were provided. no patient was involved. | [5 empty] | The investigation has been completed. the console was sent back for further investigation. visual inspection of the purge assembly area confirmed a ripped/ missing blue retainer. the root cause of the aic issue was a defective component. this report is being filed as part of a retrospective review of historical records.
16525 | 1220648-2024-21326 | 20541563 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541563&pc=OZD | 02/15/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 02/15/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1389732 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller failed to recognize a purge cassette when the purge disc was clipped into place. no clinical details regarding resolution or patient involvement/condition were provided. no patient was involved. | [5 empty] | The investigation has been completed. the console was sent back for further investigation. the purge disk detect flag was missing. the root cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
16526 | 1220648-2024-21327 | 20541564 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541564&pc=OZD | 02/21/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 02/21/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1170633 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller failed to recognize a purge cassette disc during setup of the device. a new console was used without any issues. no patient was involved. | [5 empty] | The investigation has been completed. the console was sent back for further investigation. visual inspection of the console in danvers revealed missing (broken) disk detect flag. the root cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
16527 | 1220648-2024-21328 | 20541565 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541565&pc=OZD | 02/16/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 02/17/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1382144 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller failed to recognize a purge cassette disc during setup of the device. a new console was used without any issues. no patient was involved. | [5 empty] | The purge assembly area was inspected, specifically the purge disc retainer. it was observed that the purge flag was missing/broken. the root cause of the "purge disc not detected" alarm is due to a missing purge flag. this report is being filed as part of a retrospective review of historical records.
16528 | 1220648-2024-21329 | 20541566 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541566&pc=OZD | 02/15/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 02/16/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1218086 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced damage to the console impacting the purge sensor. a new console was used without any issues. no patient was involved. | [5 empty] | The investigation has been completed. the console was sent back for further investigation. the purge disc retainer and flag were confirmed to be broken. the root cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
16529 | 1220648-2024-21330 | 20541567 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541567&pc=OZD | 02/16/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 02/16/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1218088 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller experienced damage to the purge disc/clip. a new console was used without any issues. no patient was involved. | [5 empty] | The investigation has been completed. the console was sent back for further investigation. the "purge disk disconnected" message, however, was an actual failure, as during visual inspection of the console the purge disk retainer was confirmed to have been broken off and purge disk detect flag was missing. the root cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
16530 | 1220648-2024-21331 | 20541573 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541573&pc=OZD | 01/27/2023 | 10/27/2024 |  | 10/27/2024 |  |  | 01/27/2023 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1145127 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported patient was placed on impella cp with smart assist for hemodynamic support. while on support the console alarmed controller error and placement signal was no longer present. the nurse switched to a backup console and placement signal re-appeared. no reported of patient injury or harm. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the aic issue was most likely a defective 4in1 board. this report is being filed as part of a retrospective review of historical records.
16531 | 1220648-2024-21332 | 20541581 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541581&pc=OZD | 01/22/2023 | 10/27/2024 |  | 10/27/2024 |  |  | 01/22/2023 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1230742 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the customer was placed on impella pump support due to cardiomyopathy. while on support, the aic exhibited a "purge disc not detected" alarm. the staff removed and reinserted the purge disk, changed the cassette, cleaned the sensor cleaned. however, this did not resolve the issue. the device was replaced with another aic and the alarms stopped. it was noted that there were no issues with the impella pump, purge pressure and flow were still in range, and the patient was stable. no report of patient harm or injury. | [5 empty] | The investigation into the purge disc/flag issues issue has been completed. the automated impella controller (aic) was returned for investigation. the device logs from the complaint date revealed that the console issued the purge disc not detected alarm several times when purge cassette was inserted. re-insertion of purge disc was also observed, yet the console still had problem detecting purge disc. visual inspection of the purge assembly area confirmed a hairline crack on blue retainer. the failure mode was due to a broken purge retainer. this report is being filed as part of a retrospective review of historical records.
16532 | 1220648-2024-21333 | 20541583 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541583&pc=OZD | 02/15/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 02/15/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1320564 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) experienced damage observed at the purge disc, which was resolved by changing the console. no patient harm reported. | [5 empty] | The investigation has been completed. the console (ic5705) was sent back for further investigation. visual inspection of the console in danvers confirmed broken purge disk retainer and missing purge disk detect flag. the root cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
16533 | 1220648-2024-21334 | 20541584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541584&pc=OZD | 05/24/2021 | 10/27/2024 |  | 10/27/2024 |  |  | 05/24/2021 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1426097 |  | 0042-0000-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  | Asian | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was placed onto impella cp support due to acute myocardial infarction/cardiogenic shock. while on support, the automated impella controller exhibited a battery level low that occurred once. no harm or injury to the patient was reported. | [5 empty] | The investigation for the reported issue has been completed. a review of the console logs from the reported day of event are consistent with complaint and showed single occurrence of alarm (battery level low) 9 seconds after aleternating current (ac) power was disconnected. when the ac cord was reconnected, the alarm resolved itself after 7 seconds. the battery level low alarm was triggered erroneously due to communication glitch with the power battery manager (pbm). the impella device was not received from the customer and therefore, an evaluation of the device was not possible. any other issue with the pbm was unable to be determined. this report is being filed as part of a retrospective review of historical records.
16534 | 1220648-2024-21335 | 20541586 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541586&pc=OZD | 02/10/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 02/10/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1213827 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced not detecting a purge cassette during preparation of the device for a case. a new console was used without any issues. no patient was involved. | [5 empty] | The investigation has been completed. the console (ic4026) was sent back for further investigation. the console was tested with a known-good pump and purge. there were no issues maintaining purge pressure through the entirety of the testing. the purge assembly was visually inspected for any damage to the stepper motor and purge pressure sensor/ retainer, and a hairline crack was found on the purge pressure retainer. this damage to the purge pressure retainer did not impact the ability to transduce pressure onto the purge pressure sensor and was determined not to be related to the failure mode as investigated. the purge fluid not detected window was not reproduced under normal testing conditions but likely the result of an open line. this report is being filed as part of a retrospective review of historical records.
16535 | 1220648-2024-21336 | 20541587 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541587&pc=OZD | 02/08/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 02/09/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1210714 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced a purge disc broken. a new console was used without any issues. no patient was involved. | [5 empty] | The investigation has been completed. the console (ic3980) was sent back for further investigation. visual inspection of the purge assembly area confirmed a broken blue retainer. the root cause of the aic issue was a defective component. this report is being filed as part of a retrospective review of historical records.
16536 | 1220648-2024-21337 | 20541588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541588&pc=OZD | 02/04/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 02/04/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1281592 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella ld for mechanical circulatory support. while on support, the automated impella controller (aic) experienced damage observed at the purge disc, which was resolved by changing the console. no patient harm reported. | [5 empty] | The investigation has been completed. the console (ic4970) was sent back for further investigation. the purge disc retainer was confirmed to be broken which likely effected the sensor¿s ability to detect the purge disc. the root cause of the issue was a defective component. this report is being filed as part of a retrospective review of historical records.
16537 | 1220648-2024-21338 | 20541589 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541589&pc=OZD | 02/04/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 02/04/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1210713 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced a purge disc not able to be placed and could not be detected. no patient was involved. | [5 empty] | The investigation has been completed. the console (ic3979) was sent back for further investigation. console was tested with purge cassette and disc however the issue did not reproduce. console was run with pump and purge overnight with no issues. the root cause of intermittent "purge disc not detected" alarms could not be determined due to the inability to reproduce. this report is being filed as part of a retrospective review of historical records.
16538 | 1220648-2024-21339 | 20541590 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541590&pc=OZD | 02/04/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 02/04/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1249754 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced a purge disc issue and could not be detected. a new console was used without any issues. no patient was involved. | [5 empty] | The investigation has been completed. the console ((b)(6)) was sent back for further investigation. visual inspection of the purge assembly area confirmed a broken blue retainer. the root cause of the aic issue was a defective component. this report is being filed as part of a retrospective review of historical records.
16539 | 1220648-2024-21340 | 20541591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541591&pc=OZD | 01/20/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 01/20/2022 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1077402 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced a purge disc issue. no clinical details regarding resolution or patient condition were provided. no patient was involved. | [5 empty] | The investigation has been completed. the console (ic1001) was sent back for further investigation. the console was tested thoroughly on an engineering lab bench. as received, the stepper motor spindle was freely moving (not sticking irregularly) which would have explained the clicking sound heard in the clinical setting. the root cause of the purge system failure was most likely due to a defective cassette (demo product). no problem was found with the console ic1001. this report is being filed as part of a retrospective review of historical records.
16540 | 1220648-2024-21341 | 20541608 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541608&pc=OZD | 01/23/2023 | 10/27/2024 |  | 10/27/2024 |  |  | 01/23/2023 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1427009 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | Failure to Read Input Signal | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the aic experienced a purge disc/flag issue. the sensor for the purge pressure disk was broken. it was reported that the procedure was successful and there was no harm or injury to the patient. | [5 empty] | The investigation into the purge disc/flag issues issue has been completed. the automated impella controller (aic) was returned for investigation. visual inspection of the purge assembly area confirmed a ripped/ missing blue retainer. this report is being filed as part of a retrospective review of historical records.
16541 | 1220648-2024-21342 | 20541612 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541612&pc=OZD | 01/20/2023 | 10/27/2024 |  | 10/27/2024 |  |  | 01/20/2023 | 10/27/2024 |  |  |  |  | 10/27/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1306676 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  |  | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a controller error yellow alarm and loss of opm data. it was reported that this aic was replaced with a loaner aic. there was no patient harm reported. | [5 empty] | The impella device was returned, and console logs from the reported day of event are consistent with complaint and show that after a day since pump start, the placement signal became not reliable, and corresponding alarm #1036 was presented. after another day of uninterrupted hemodynamic support the pump had been unplugged and re-connected in an apparent attempt to recover placement signal, however this time the console showed a controller error alarm indicating defective optical sensor lamp, and there was no placement signal. after two hours the console had been shut down, restarted, a new pump had been connected (409648), and placement signal was present. when the console was tested in danvers, the issue was not reproduced. optical bench operated within specification, and there were no interruptions or losses of placement signal during lab testing. we were able, however, to replicate identical symptoms to those observed in the field by using a known-defective pump with broken optical fiber. unfortunately, pump 409649 had been discarded before it could be analyzed and the defect confirmed. repair: perform functional check of the console. this report is being filed as part of a retrospective review of historical records.
16542 | 1220648-2024-21343 | 20541705 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20541705&pc=PYX | 07/31/2022 | 10/27/2024 |  | 10/27/2024 |  |  | 07/31/2022 | 10/27/2024 |  |  | 05/31/2024 |  | 10/27/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2023213705 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella rp device. the complainant reported a purge side arm leak originating from the reservoir. a side-arm bypass configuration resolved the issue. no patient harm and support proceeded as planned. | [5 empty] | The investigation for the reported issue has been completed. sidearm was visually inspected and no abnormalities were observed. a syringe was attached to the side arm and a leak was observed at the reservoir. the cause of the purge leak was sidearm reservoir damage. leakage was from the membrane. this report is being filed as part of a retrospective review of historical records.
16543 | 1220648-2024-21137 | 20540164 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540164&pc=OZD | 04/23/2023 | 10/25/2024 |  | 10/26/2024 |  |  | 04/23/2023 | 10/26/2024 |  |  | 08/31/2024 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023262835 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the device became malpositioned during use, resulting in high motor current and left ventricle tracings/aorta placement signal emitting abnormal readings. the pump was repositioned, and the issues resolved with no harm to the patient. patient was determined to be in multi organ failure and the patient was left on support until the family was able to arrive. once the family arrived, the decision was made to discontinue all therapies. the patient expired moments later. use of the device cannot conclusively be associated with the outcome. | [5 empty] | The investigation for the reported positioning issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. data confirmed the pump was not in proper position, then the issue was resolved and the pump was in good position and ran without issue to the rest of the case. based on clinical description & data analysis, the root cause of positioning issue was most likely due to use issue. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings, cautions and precautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ correct impella catheter position: ¿for optimal positioning of the impella catheter, the inlet area of the catheter should be 3.5 cm below the aortic valve annulus and well away from papillary muscle and subannular structures. the outlet area should be well above the aortic valve.¿ this report is being filed as part of a retrospective review of historical records.
16544 | 1220648-2024-21152 | 20540228 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540228&pc=OZD | 05/25/2023 | 10/25/2024 |  | 10/26/2024 |  |  | 05/25/2023 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1419816 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fracture | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) was found to have a broken purge clip. there was no reported patient involvement. | [5 empty] | The investigation for the reported purge disc/flag issues has been completed. the device logs were reviewed but are not relevant to the complaint. visual inspection of the purge assembly area confirmed a ripped/ missing blue retainer. the cause of the issue was broken/ missing purge disk retainer. this report is being filed as part of a retrospective review of historical records.
16545 | 1220648-2024-21156 | 20540349 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540349&pc=OZD | 03/27/2023 | 10/26/2024 |  | 10/26/2024 |  |  | 03/27/2023 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1486033 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Damaged Prior to Use | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) damaged and was unable to roll or stand. no clinical details regarding resolution. no patient was involved. | [5 empty] | The investigation has been completed. there are no logs necessary to analyze for review of a broken cart. the product did not need to be returned and a replacement was sent out to the customer. the root cause of the issue was not determined since product was not returned. this report is being filed as part of a retrospective review of historical records.
16546 | 1220648-2024-21157 | 20540351 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540351&pc=OZD | 04/05/2023 | 10/26/2024 |  | 10/26/2024 |  |  | 04/05/2023 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1542729 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a controller failure red alarm. no clinical details regarding patient condition or resolution were provided. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the cause of the issue was not determined since issue could not be reproduced. this report is being filed as part of a retrospective review of historical records.
16547 | 1220648-2024-21158 | 20540357 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540357&pc=OZD | 11/29/2022 | 10/26/2024 |  | 10/26/2024 |  |  | 11/29/2022 | 10/26/2024 |  |  | 07/31/2024 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023259260 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support the cannula kinked. the pump was removed and replaced with a new pump to continue with patient support. | [5 empty] | The investigation for the reported issue has been completed. data logs show there was sudden drop in flow and loss of motor current pulsatility accompanied by suction alarms and no change in p-level. the placement signal briefly went ventricular prior to the alarms. visual inspection of the returned pump revealed a significant cannula kink near the inflow cage. no other damage or abnormalities were found. after reviewing the clinical information, it was determined that the cause of the low pump flow was a kinked cannula stemming from the advancement of the repositioning sheath. this report is being filed as part of a retrospective review of historical records.
16548 | 1220648-2024-21159 | 20540359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540359&pc=OZD | 12/06/2022 | 10/26/2024 |  | 10/26/2024 |  |  | 12/06/2022 | 10/26/2024 |  |  | 07/31/2024 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023254337 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  |  | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the impella pump was inadvertently pulled back across the aortic valve. an attempt to reposition the device was unsuccessful as the pump was noted to have folded. due to the condition of the patient, the initial impella was removed, and a new device was placed. use of the device cannot conclusively be associated with the outcome. no patient harm related to the device was identified. | [5 empty] | The investigation for the reported issue has been completed. the field rt log was reviewed and pump in the ventricle alarm triggered during support confirming that positioning issues were occurring. the returned pump was visually inspected and a cannula kink was observed. no other abnormalities were observed. the cause of the issue was wireless re-access resulting in a cannula kink. the impella device is not indicated for wireless re-access. this report is being filed as part of a retrospective review of historical records.
16549 | 1220648-2024-21160 | 20540360 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540360&pc=OZD | 10/28/2022 | 10/26/2024 |  | 10/26/2024 |  |  | 10/28/2022 | 10/26/2024 |  |  | 07/31/2024 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023255554 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Malfunction | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was attempted to being implanted with an impella cp device for mechanical circulatory support. the complainant reported that the impella cp pump could not be fully advanced into the patient because of calcified aortic arch; therefore, the insertion was aborted. additional information was provided that noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation has been completed for the reported issue of delivery issue. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was due to patient anatomy, most likely due to patient condition. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
16550 | 1220648-2024-21161 | 20540362 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540362&pc=OZD | 12/15/2022 | 10/26/2024 |  | 10/26/2024 |  |  | 12/15/2022 | 10/26/2024 |  |  | 08/31/2024 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023263332 |  | 0048-0045 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Deformation Due to Compressive Stress | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported failure to successfully place the impella device after the 0.018" guidewire fell out of the ventricle during delivery attempt. additional information was provided that noted the patient expired. there was no reported relation of the impella with the patient's outcome. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. after reviewing the clinical information, it was determined that the cause of the delivery issues was most likely wireless insertion. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings: ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
16551 | 1220648-2024-21162 | 20540363 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540363&pc=OZD | 01/13/2023 | 10/26/2024 |  | 10/26/2024 |  |  | 01/13/2023 | 10/26/2024 |  |  | 01/31/2024 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023188109 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported a failure to deliver the device due to the patient's anatomy. the pump was removed and replaced with a smaller impella cp. additional information was provided that noted the patient expired. there was no reported relation of the impella with the patient's outcome. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. based on the clinical data root cause was patient condition due to the acute vascular angulation, the device was unable to pass. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
16552 | 1220648-2024-21163 | 20540366 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540366&pc=OZD | 02/05/2022 | 10/26/2024 |  | 10/26/2024 |  |  | 02/15/2022 | 10/26/2024 |  |  | 06/30/2023 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | IMPELLA CP |  |  | US |  | 2022100452 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced bleeding including hematoma on (b)(6) 2021. the patient was transfused with an unknown amount of blood products due to the blood loss. it was also reported that the patient experienced thrombocytopenia on (b)(6) 2021. there was no information provided on any treatments provided to the patient for the thrombocytopenia. it was reported that the patient remained on impella support and the patient's cardiac function recovered. the patient was successfully weaned and the impella was explanted. it was reported that the patient then experienced a hemorrhagic cerebrovascular accident after explant on (b)(6) 2021. it was stated that it was possibly related to the impella device. the patient then expired 3 days later. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
16553 | 1220648-2024-21164 | 20540367 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540367&pc=OZD | 06/17/2022 | 10/26/2024 |  | 10/26/2024 |  |  | 06/17/2022 | 10/26/2024 |  |  | 02/29/2024 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023195285 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, there was a temporary impella cp pump stop. the device was restarted but it had saturated flows. the survival outcome at explant noted the patient expired. use of the device cannot conclusively be associated with the outcome. | [5 empty] | The investigation for the reported temporary pump stop and saturated flow issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the product was returned and large amounts of biomaterial were found in the outflow and inside of the cannula. the root cause of the pump stop was determined to be ingested biomaterial. the pump stop occurred on day 11 of use which is beyond the design indication of 8 days per design trace (b)(4). as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high-risk cpi section: entering cardiac output ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
16554 | 1220648-2024-21165 | 20540368 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540368&pc=OZD | 06/16/2022 | 10/26/2024 |  | 10/26/2024 |  |  | 06/16/2022 | 10/26/2024 |  |  | 02/29/2024 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023197511 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was presented with cardiomyopathy and implanted with an impella cp device for mechanical circulatory support. after the catheter prep, prior to insertion there was a continued yellow ¿placement signal not reliable¿ alarm that would not clear. the decision was made to swap out to new a cp catheter. the second cp was prepped and inserted without incident. the procedure outcome noted that the patient expired. use of the device cannot conclusively be associated with the outcome. | [5 empty] | The investigation for the reported optical signal issue has been completed. the impella device was not received from the customer nor was clinical information provided sufficient to determine the root cause. therefore, the root cause of the reported issue could not be determined. this report is being filed as part of a retrospective review of historical records.
16555 | 1220648-2024-21166 | 20540370 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540370&pc=OZD | 12/29/2022 | 10/26/2024 |  | 10/26/2024 |  |  | 12/29/2022 | 10/26/2024 |  |  | 10/31/2024 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023279181 |  | 1000080 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Cardiac Arrest | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient had significant stenosis at the illiac bifurcation. the doctor was not able to advance introducer sheath all of the way into the patient due to the significant calcification and the impella would not pass either. additional information was provided that noted the patient was actively coding. cardiopulmonary resuscitation was attempted but not successful and the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. after reviewing the clinical information, it was determined that the cause of the delivery issues was patient condition related, specifically the patient's diseased vasculature. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
16556 | 1220648-2024-21167 | 20540371 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540371&pc=OZD | 01/17/2023 | 10/26/2024 |  | 10/26/2024 |  |  | 01/17/2023 | 10/26/2024 |  |  | 10/31/2024 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024285094 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in st elevation myocardial infarction was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the patient experienced bilateral limb ischemia. a femoral to femoral bypass was placed due to a lack of perfusion to the right leg after the impella was placed. the medical team decided to escalate the patient to an impella 5.5 which was inserted through the axillary artery. after the impella cp was removed from the right femoral, the patient still did not have bilateral pulses. it was reported the patient was later withdrawn from care on the impella 5.5 due to poor neurological function. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported limb ischemia has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. per the clinical description, the patient did not have bilateral pulses which did not change after femoral to femoral bypass or the removal of the impella cp. therefore, the root cause of limb ischemia was most likely due to patient condition. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
16557 | 1220648-2024-21168 | 20540373 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540373&pc=OZD | 12/27/2022 | 10/26/2024 |  | 10/26/2024 |  |  | 12/27/2022 | 10/26/2024 |  |  | 08/31/2024 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023273324 |  | 0048-0045 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  |  |  | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient needing an impella cp device for mechanical circulatory support. the complainant reported a failed delivery of the impella device due to calcification. balloon was attempted but was still unable to deliver the impella. additional information was provided that noted the patient expired. there was no reported relation of the impella with the patient's outcome. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the issue was stenotic artery patient condition related. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ this report is being filed as part of a retrospective review of historical records.
16558 | 1220648-2024-21169 | 20540375 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540375&pc=OZD | 12/22/2021 | 10/26/2024 |  | 10/26/2024 |  |  | 12/22/2021 | 10/26/2024 |  |  | 10/31/2023 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022142029 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Death | Yes | Yes | Fluid/Blood Leak | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the purge cassette leaked and then the purge filter leaked. a purge bypass was done, and there was no medical intervention, and there was no harm to the patient. however, it was later reported that the patient expired while on support. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The investigation for the reported purge leak issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the purge leak pump was sidearm filter damage. the cause of the purge leak cassette leak was not determined. impella cp with smartassist for use during cardiogenic shock and high risk cpi instructions for use for the related event are as follows: section: cleaning: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ this report is being filed as part of a retrospective review of historical records.
16559 | 1220648-2024-21170 | 20540376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540376&pc=OZD | 08/31/2022 | 10/26/2024 |  | 10/26/2024 |  |  | 08/31/2022 | 10/26/2024 |  |  | 04/30/2024 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023229966 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during implantation, there were delivery issues due to vascular anomalies of the right femoral artery. a new pump was successfully delivered through the left femoral artery. the patient expired after impella support due to other patient condition, not related to impella and extracorporeal membrane oxygenation support. use of the device cannot conclusively be associated with the outcome. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that a visual inspection of the pump did not reveal any abnormalities. the cause of the delivery issue was patient condition related. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this report is being filed as part of a retrospective review of historical records.
16560 | 1220648-2024-21171 | 20540378 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540378&pc=OZD | 12/05/2022 | 10/26/2024 |  | 10/26/2024 |  |  | 12/05/2022 | 10/26/2024 |  |  | 07/31/2024 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023248483 |  | 0048-0014 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced severe cerebral hemorrhage and died. no report of device malfunction and no additional information is available. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
16561 | 1220648-2024-21172 | 20540379 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540379&pc=OZD | 08/10/2022 | 10/26/2024 |  | 10/26/2024 |  |  | 08/10/2022 | 10/26/2024 |  |  | 04/30/2024 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023198429 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported while on impella cp support, the impella cp pump failed to purge appropriately. the impella cp pump, purge cassette and different aic were employed with success. the patient later expired on the second impella cp pump after the family requested care to be withdrawn. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | Data logs show purge system open alarms upon initial priming. since the product was not returned, the root cause of the priming issue could not be determined. this report is being filed as part of a retrospective review of historical records.
16562 | 1220648-2024-21173 | 20540380 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540380&pc=OZD | 03/16/2023 | 10/26/2024 |  | 10/26/2024 |  |  | 03/16/2023 | 10/26/2024 |  |  | 12/31/2024 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024304113 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that an "optical sensor system error" alarm occurred during the start of the case. the medical team attempted to troubleshoot the issue by power-cycling the console, however the issue did not resolve. the pump was then replaced with a new impella cp, which went through case start and functioned with no issues. there were no reported patient injuries due to the optical sensor alarm. | [5 empty] | The investigation for the reported optical signal issue has been completed. the cause of the optical signal issue was not determined since neither the device nor the data logs were returned. this report is being filed as part of a retrospective review of historical records.
16563 | 1220648-2024-21174 | 20540381 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540381&pc=OZD | 07/17/2022 | 10/26/2024 |  | 10/26/2024 |  |  | 07/17/2022 | 10/26/2024 |  |  | 02/29/2024 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023200343 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | [6 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was admitted to the icu and following admission went into pulseless electrical activity arrest, was given epinephrine, and was transported wo the cardiac catheterization lab where the impella was placed during cardiopulmonary resuscitation. it was reported after impella cp support was initiated, the impella cp experienced an abrupt pump stop following a motor spike and no flows were noted. the impella was removed and replaced with a new impella cp. the patient subsequently died two hours later while supported on the new impella. medical safety review of the event noted there is not enough information to exclude the initial impella cp device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported pump stop has been completed. the data logs confirm multiple high motor current pump stops along with spikes in motor current. the product was returned to the lab and ran without issue. the root cause of the pump stop was most likely ingested biomaterial that dislodged upon explant due to the evidence in the data logs and the failure mode being unable to reproduce in the lab. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ this report is being filed as part of a retrospective review of historical records.
16564 | 1220648-2024-21175 | 20540382 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540382&pc=OZD | 02/11/2023 | 10/26/2024 |  | 10/26/2024 |  |  | 02/11/2023 | 10/26/2024 |  |  | 10/31/2024 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024290764 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Embolism/Embolus | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the complainant reported that extracorporeal membrane oxygenation (echo) showed bubbles from the ventricle to the aorta every 55 seconds. it was seen that systemic vascular resistance (svr) was normal, mean arterial pressure was stable, and there was no ventricular septal defect (vsd), or structural issues identified. changed console and purge fluid and cassette with no change. patient was not a candidate for escalation and so the plan was made to continue with therapy. additional information noted the patient was made a do not resuscitate, care was withdrawn, and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported issue has been completed. the field lt log was reviewed, and a long purge cassette/bag change was observed. no air in purge alarms triggered throughout support. the returned pump and 2 purge cassettes were visually inspected and no abnormalities were observed. one purge cassette (gen2) was run through de-air without issue, connected to the returned pump, and no air in purge alarms triggered. high purge pressure eventually triggered due to biomaterial in the outflow. the next purge cassette was run through de-air without issue, connected to the returned pump, and no air in purge alarms triggered. high purge pressure again triggered due to biomaterial in the outflow. the cause of the issue was most likely patient condition related since returned device operated to specification. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." this report is being filed as part of a retrospective review of historical records.
16565 | 1220648-2024-21176 | 20540383 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540383&pc=OZD | 09/03/2021 | 10/26/2024 |  | 10/26/2024 |  |  | 09/04/2021 | 10/26/2024 |  |  | 12/31/2022 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2022042956 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Female |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, extensive cerebral hemorrhage occurred immediately after impella support. the survival outcome at explant noted the patient expired. per the physician, the death was not related to impella. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ this report is being filed as part of a retrospective review of historical records.
16566 | 1220648-2024-21177 | 20540391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540391&pc=OZD | 04/20/2023 | 10/26/2024 |  | 10/26/2024 |  |  | 04/20/2023 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1323490 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a battery failure red alarm. no clinical details regarding patient condition or resolution were provided. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the battery failure alarm was due to a defective battery. the root cause of the defective battery was due to single cell failure. this report is being filed as part of a retrospective review of historical records.
16567 | 1220648-2024-21178 | 20540392 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540392&pc=OZD | 04/24/2023 | 10/26/2024 |  | 10/26/2024 |  |  | 04/24/2023 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1645321 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced system self-check errors. no patient was involved. | [5 empty] | The investigation has been completed. the console (b)(6) was sent back for further investigation. the root cause of the aic issue was a loose component. this report is being filed as part of a retrospective review of historical records.
16568 | 1220648-2024-21179 | 20540393 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540393&pc=OZD | 09/06/2022 | 10/26/2024 |  | 10/26/2024 |  |  | 09/06/2022 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1431282 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported the aic pressure sensor was not reading. action taken to resolve the issue are unknown. there was no patient harm reported. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, there was special attention to the purge disc retainer and purge disc flag. no issues related to purge pressure disc detection or function could be found. no corrective action is recommended at this time because the failure was determined to have a risk as low as practical (alap). failures of this type will be monitored for trends. this report is being filed as part of a retrospective review of historical records.
16569 | 1220648-2024-21181 | 20540394 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540394&pc=OZD | 08/31/2022 | 10/26/2024 |  | 10/26/2024 |  |  | 08/31/2022 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1337720 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No |  |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. it was reported that the aic had a broken purge clip. the aic ifu was followed with a backup controller utilized. no patient harm was reported. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, the issue with broken purge pressure disc was visually confirmed. the purge retainer has a broken stand and cracks on it; therefore, the purge flag was not being held properly at blue disc retainer. the cause of the issue was a defective component. no corrective action is recommended at this time because the failure was determined to have a risk as low as practical (alap). failures of this type will be monitored for trends. this report is being filed as part of a retrospective review of historical records.
16570 | 1220648-2024-21180 | 20540396 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540396&pc=OZD | 08/24/2022 | 10/26/2024 |  | 10/26/2024 |  |  | 08/24/2022 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1401344 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported a purge disc connection issue during use of the device. actions taken to resolve the issue are unknown. no patient harm was reported. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken. the cause of the issue was a defective component. no corrective action is recommended at this time because the failure was determined to have a risk as low as practical (alap). failures of this type will be monitored for trends. this report is being filed as part of a retrospective review of historical records.
16571 | 1220648-2024-21183 | 20540397 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540397&pc=OZD | 08/26/2022 | 10/26/2024 |  | 10/26/2024 |  |  | 08/26/2022 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1547850 |  | 0042-4201 |  | Yes |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the aic had a purge disc malfunction/damage. the ifu was followed and a backup controller was utilized for patient support. no patient harm was reported. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, a broken purge retainer was confirmed and replaced. during functional check of the console it¿s been determined that full charge capacity of battery 1 was below specification. the battery set was replaced, and aic data re-analyzed. it¿s been determined that battery 1 failed due to cell imbalance (a known pattern failure). this malfunction, although already present while the console was in the field, did not manifest with any controller alarm. the cause of the issue was a defective component. no corrective action is recommended at this time because the failure was determined to have a risk as low as practical (alap). failures of this type will be monitored for trends. this report is being filed as part of a retrospective review of historical records.
16572 | 1220648-2024-21182 | 20540398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540398&pc=OZD | 08/26/2022 | 10/26/2024 |  | 10/26/2024 |  |  | 08/26/2022 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1239550 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. it was reported that there was a purge disc issue at patient transfer. the ifu was followed, and the backup aic was utilized. no patient harm was reported. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, the issue with properly detecting the purge pressure disc was reproduced, and purge flag was confirmed to be broken (missing at blue disc retainer). the cause of the issue was a defective component. no corrective action is recommended at this time because the failure was determined to have a risk as low as practical (alap). failures of this type will be monitored for trends. this report is being filed as part of a retrospective review of historical records.
16573 | 1220648-2024-21186 | 20540399 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540399&pc=OZD | 05/03/2023 | 10/26/2024 |  | 10/26/2024 |  |  | 05/03/2023 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1323599 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced system self-check errors. no clinical details regarding resolution or patient condition were provided. no patient was involved. | [5 empty] | The investigation has been completed. the console ((b)(6)) was sent back for further investigation. reported syscall error on boot-up was caused by loss of communication to pbm due to pbm corrosion near supercapacitors. this report is being filed as part of a retrospective review of historical records.
16574 | 1220648-2024-21185 | 20540400 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540400&pc=OZD | 12/22/2021 | 10/26/2024 |  | 10/26/2024 |  |  | 12/22/2021 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1106949 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the automated impella controller experienced a controller error yellow alarm. no clinical details regarding resolution or patient involvement/condition were provided. no patient was involved. | [5 empty] | The investigation has been completed. the console (ic1691) was sent back for further investigation. the root cause of the purge system failure alarm was most likely build-up of glucose on the stepper motor gasket. this report is being filed as part of a retrospective review of historical records.
16575 | 1220648-2024-21184 | 20540401 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540401&pc=OZD | 08/23/2022 | 10/26/2024 |  | 10/26/2024 |  |  | 08/23/2022 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1250419 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the complainant reported that the aic had a broken purge disc. actions taken are unknown. no patient harm was reported. | [5 empty] | The console (aic) was returned for evaluation. upon inspection of the console, the purge disc retainer was confirmed to be broken (broken blue disc retainer). the cause of the purge disc flag issues was due to a broken retainer. the failure mode will be monitored and trended. this report is being filed as part of a retrospective review of historical records.
16576 | 1220648-2024-21188 | 20540402 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540402&pc=OZD | 12/10/2021 | 10/26/2024 |  | 10/26/2024 |  |  | 12/10/2021 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1133147 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. during prep, the automated impella controller experienced a controller error yellow, which was resolved with replacement console. no patient harm reported. | [5 empty] | The investigation has been completed. the console (ic2383) was sent back for further investigation. the root cause of the controller error could not be determined due to the inability to reproduce but involved the 4-in-1, kbd, or connecting ribbon cable (decision was made to replace all three). this report is being filed as part of a retrospective review of historical records.
16577 | 1220648-2024-21187 | 20540403 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540403&pc=OZD | 12/23/2021 | 10/26/2024 |  | 10/26/2024 |  |  | 12/23/2021 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1084502 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the automated impella controller (aic) experienced a controller error yellow alarm, which was resolved by changing the console. no patient harm was reported. | [5 empty] | The investigation has been completed. the console ((b)(6)) was sent back for further investigation. the root cause of the controller error was most likely due to degradation of the kbd ribbon cable, which was due to fluid ingress. this report is being filed as part of a retrospective review of historical records.
16578 | 1220648-2024-21189 | 20540404 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540404&pc=OZD | 12/10/2021 | 10/26/2024 |  | 10/26/2024 |  |  | 12/10/2021 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1204377 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 26 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. while on support, the automated impella controller experienced a controller error yellow alarm, no clinical details regarding resolution. no patient harm reported. | [5 empty] | The investigation has been completed. the console (ic3882) was sent back for further investigation. this is a known pattern failure in which all signals received from optical bench (placement, snr, contrast, lamp, gain) are interpreted as -16384 values. this lasts between 2 to 10 minutes, after which all signals go back to normal. complaint cause is a known pattern failure characterized by temporary loss of placement signal. the root cause of the transient loss of opm-related signals could not be determined due to the inability to reproduce. this report is being filed as part of a retrospective review of historical records.
16579 | 1220648-2024-21191 | 20540405 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540405&pc=OZD | 12/01/2021 | 10/26/2024 |  | 10/26/2024 |  |  | 12/01/2021 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1145537 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the automated impella controller experienced a controller error yellow alarm, which was resolved through console replacement. no patient was involved. | [5 empty] | The investigation has been completed. the console (ic2759) was sent back for further investigation. console was investigated on engineering bench, but the controller failure alarm was reproduced during power cycling and overnight running with pump simulator. console was opened, and the ribbon cable from purge pressure sensor was found slightly bent. the cause of the controller error alarm was most likely due to the defective purge pressure sensor. this report is being filed as part of a retrospective review of historical records.
16580 | 1220648-2024-21193 | 20540406 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540406&pc=OZD | 11/21/2021 | 10/26/2024 |  | 10/26/2024 |  |  | 11/21/2021 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | IMPELLA CP |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1138942 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the automated impella controller experienced a controller error yellow alarm. no clinical details regarding resolution or patient involvement/condition were provided. no patient was involved. | [5 empty] | The investigation has been completed. the console ((b)(6)) was sent back for further investigation. console interior connections were inspected, with specific attention to the optical connectors. no issues were found. the root cause of the controller error was a result of optical placement signal loss, due to opm communication corruption. this report is being filed as part of a retrospective review of historical records.
16581 | 1220648-2024-21192 | 20540407 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540407&pc=OZD | 12/02/2021 | 10/26/2024 |  | 10/26/2024 |  |  | 12/02/2021 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1245288 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the automated impella controller experienced a controller error yellow alarm. no clinical details regarding resolution or patient involvement/condition were provided. no patient was involved. | [5 empty] | The investigation has been completed. the console ((b)(6)) was sent back for further investigation. after the console arrive in danvers, it was found that the power supply cable for optical bench was loosened. therefore, the supply voltage to optical bench was insufficient and led to the controller error alarm. the root cause of the controller error (yellow alarm) was the loose optical bench power supply cable. this report is being filed as part of a retrospective review of historical records.
16582 | 1220648-2024-21190 | 20540408 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540408&pc=OZD | 12/07/2021 | 10/26/2024 |  | 10/26/2024 |  |  | 12/07/2021 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1576134 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 24 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. while on support, the automated impella controller experienced a controller error yellow alarm, which was resolved through console replacement. no patient harm reported. | [5 empty] | The investigation has been completed. the console ( (b)(6) ) was sent back for further investigation. this known pattern failure, which is characterized by a temporary loss of optical data (placement, snr, contrast, lamp, gain) and triggering of a controller error alarm, has a low probability of reoccurrence. if needed, patient hemodynamics and impella position can be monitored with imaging. this report is being filed as part of a retrospective review of historical records.
16583 | 1220648-2024-21194 | 20540409 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540409&pc=OZD | 11/19/2021 | 10/26/2024 |  | 10/26/2024 |  |  | 11/19/2021 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1426097 |  | 0042-0000-JP |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. while on support, the automated impella controller experienced a controller error yellow alarm, which was resolved through console replacement. no patient harm reported. | [5 empty] | Console ((b)(6)) was sent back to japan fs for further investigation. this known pattern failure, which is characterized by a temporary loss of optical data (placement, snr, contrast, lamp, gain) and triggering of a controller error alarm, has a low probability of reoccurrence. if needed, patient hemodynamics and impella position can be monitored with imaging. the root cause of the transient loss of optical data was unable to be determined because the issue was unable to be reproduced. this report is being filed as part of a retrospective review of historical records.
16584 | 1220648-2024-21195 | 20540410 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540410&pc=OZD | 11/10/2021 | 10/26/2024 |  | 10/26/2024 |  |  | 11/10/2021 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1494422 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the automated impella controller experienced a controller error yellow alarm. no clinical details regarding resolution or patient involvement/condition were provided. no patient was involved. | [5 empty] | The investigation has been completed. the console (ic8190) was sent back for further investigation. the defective console was tested on an engineering bench, console was boot-up without syscall error, but the battery capacity showed 0% (batteries fully discharged). after console being opened, the voltage supply for impellatronic was measured 20v on power battery manager (pbm) even with console powered down. visual inspection further confirmed the pbm malfunction by a damaged capacitor r319 mounted on the bottom of pbm. the console failure was due to defective pbm. this report is being filed as part of a retrospective review of historical records.
16585 | 1220648-2024-21196 | 20540411 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540411&pc=OZD | 10/25/2021 | 10/26/2024 |  | 10/26/2024 |  |  | 10/25/2021 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1431282 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the automated impella controller experienced a controller error yellow alarm. no clinical details regarding resolution or patient involvement/condition were provided. no patient was involved. | [5 empty] | The investigation has been completed. the console (ic7368) was sent back for further investigation. the root cause of the intermittent opm communication data and the controller error is likely due to a faulty u21 component. this report is being filed as part of a retrospective review of historical records.
16586 | 1220648-2024-21197 | 20540412 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540412&pc=OZD | 10/18/2021 | 10/26/2024 |  | 10/26/2024 |  |  | 10/18/2021 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1347481 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. while on support, the automated impella controller experienced a controller error yellow alarm, which was resolved through console replacement. no patient harm reported. | [5 empty] | The investigation has been completed. the console (ic6106) was sent back for further investigation. the root cause of the controller error alarm is due to lack of communication due to an optical bench malfunction. this report is being filed as part of a retrospective review of historical records.
16587 | 1220648-2024-21198 | 20540413 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540413&pc=OZD | 09/20/2021 | 10/26/2024 |  | 10/26/2024 |  |  | 09/21/2021 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1431282 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the automated impella controller experienced a controller error yellow alarm in training. no clinical details regarding resolution or patient involvement/condition were provided. no patient was involved. | [5 empty] | The investigation has been completed. the console (ic7366) was sent back for further investigation. the controller failure alarm was reproduced during investigation, and confirmed by visual inspection of the dim lamp of optical bench once booted up the console. after lamp assembly was momentarily replaced with a known-good one, the controller failure alarm was resolved. the controller failure alarm was resulted from a defective lamp assembly. this report is being filed as part of a retrospective review of historical records.
16588 | 1220648-2024-21199 | 20540414 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540414&pc=OZD | 08/23/2021 | 10/26/2024 |  | 10/26/2024 |  |  | 08/23/2021 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1158455 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the automated impella controller experienced a controller error yellow alarm in training. no clinical details regarding resolution or patient involvement/condition were provided. no patient was involved. | [5 empty] | The cause of the controller alarm was communication error between 4-in-1 and kbd due to kbd malfunction. this report is being filed as part of a retrospective review of historical records.
16589 | 1220648-2024-21200 | 20540415 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540415&pc=OZD | 05/15/2023 | 10/26/2024 |  | 10/26/2024 |  |  | 05/15/2023 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1691800 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 91 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with high-risk percutaneous coronary intervention was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) experienced an optical signal issue that was discovered prior to impella therapy, which was resolved by changing the console. no patient harm reported. | [5 empty] | The investigation has been completed. the console (ic10175) was sent back for further investigation. the console was tested on an engineering lab bench. the failure mode was not reproduced while running an optical pump simulator. there was no issue with the console hardware. the root cause of the controller error was due to an aic software issue. this report is being filed as part of a retrospective review of historical records.
16590 | 1220648-2024-21201 | 20540416 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540416&pc=OZD | 08/18/2021 | 10/26/2024 |  | 10/26/2024 |  |  | 08/18/2021 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1112878 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the automated impella controller experienced a controller error yellow alarm. no clinical details regarding resolution or patient involvement/condition were provided. no patient was involved. | [5 empty] | The investigation has been completed. the console ((b)(6)) was sent back for further investigation. console booted with the controller failure seen in the field. the cause of the controller failure is corruption of the epm firmware on the impellatronic board. this report is being filed as part of a retrospective review of historical records.
16591 | 1220648-2024-21202 | 20540417 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540417&pc=OZD | 08/11/2021 | 10/26/2024 |  | 10/26/2024 |  |  | 08/11/2021 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1390619 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the automated impella controller experienced a controller error yellow alarm in training. no clinical details regarding resolution or patient involvement/condition were provided. no patient was involved. | [5 empty] | The investigation has been completed. the console (ic6772) was sent back for further investigation. during inspection of the console it's been observed that the optical connector of the "blue receptacle" was contaminated, and cleaning was ineffective. this is unrelated to optical bench malfunction, however the blue receptacle needs to be replaced. the cause of the controller error upon boot up was defective optical bench pca. this report is being filed as part of a retrospective review of historical records.
16592 | 1220648-2024-21203 | 20540418 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540418&pc=OZD | 07/15/2021 | 10/26/2024 |  | 10/26/2024 |  |  | 07/21/2021 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1107076 |  | 0042-0000-EU |  | No |  |  | Not Available |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the automated impella controller experienced a controller error yellow alarm, which was resolved through backup console. no patient was involved. | [5 empty] | Clinical staff reported that console ic1699 issued a yellow controller error as it was powered on. this complaint was not related to a clinical procedure, and there were no adverse events related to this product malfunction. after replacing the optical bench, the failure mode persisted. the cause of the controller error / communication issues with optical bench were most likely due to the 4-in-1. the exact root cause of the defective 4-in-1 is unknown. this report is being filed as part of a retrospective review of historical records.
16593 | 1220648-2024-21205 | 20540419 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540419&pc=OZD | 06/28/2021 | 10/26/2024 |  | 10/26/2024 |  |  | 06/28/2021 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1561819 |  | 0042-0040-JP |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. during support, the automated impella controller (aic) experienced a controller error yellow alarm that was resolved by replacing the console. there was no reported patient harm. | [5 empty] | The investigation for the reported controller error (yellow alarm) has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was patient anatomy, most likely due to patient condition. console logs from the reported day of event are consistent with complaint and show "controller error" alarms related to optical bench. rtlogs show that for about 12 minutes all signals from optical bench became invalid, with reported values of -16384. this is a known pattern failure "transient loss of optical data", caused by an unspecified communication corruption between optical bench and epc. currently there is no known root cause, and no evidence indicating that replacing any console component would resolve this issue. the cause of the abnormal placement signal was unable to be determined because the issue was unable to be reproduced. this report is being filed as part of a retrospective review of historical records.
16594 | 1220648-2024-21204 | 20540420 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540420&pc=OZD | 07/01/2021 | 10/26/2024 |  | 10/26/2024 |  |  | 07/01/2021 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  | 1188000 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the automated impella controller (aic) produced a controller error yellow alarm during maintenance. there was no patient involvement. | [5 empty] | The device was repaired in the field but not returned for evaluation. however, clinical details provided were sufficient to determine a root cause. console logs from the reported day of event are consistent with complaint and show alarm #1039 (defective optical sensor lamp) preceded by two "bad_lamp" errors reported by optical bench. replacing the lamp assembly did not resolve the issue, so the optical bench was replaced and issue finally resolved. the cause for optical bench malfunction was not investigated. the cause of the controller alarm was defective optical bench. the cause of optical bench defect was not determined. this report is being filed as part of a retrospective review of historical records.
16595 | 1220648-2024-21206 | 20540421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540421&pc=OZD | 06/24/2021 | 10/26/2024 |  | 10/26/2024 |  |  | 06/24/2021 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, EU |  |  | ABIOMED, INC. |  |  | US |  | 1127920 |  | 0042-0000-EU |  | Yes |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. the automated impella controller (aic) experienced a controller error yellow alarm. no patient involvement was reported. | [5 empty] | Note: d8/9 - device return date is unknown. the device was returned and evaluated by aachen field service, and the investigation for the reported controller error (yellow alarm) has been completed. the console was tested at the field service center. console logs from field service - when the issue was independently identified - are consistent with complaint and show alarm #1039 (defective optical sensor lamp) after boot-up, as well as "bad_lamp" errors reported by optical bench. during troubleshooting, the voltages provided to optical bench and to the lamp assembly were verified to be within specification, and all cables were mechanically stable. optical fiber connectors were clean. after optical bench was replaced, the issue was resolved. the cause of the controller alarm was optical bench malfunction. the cause of the malfunction was not determined. this report is being filed as part of a retrospective review of historical records.
16596 | 1220648-2024-21207 | 20540422 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540422&pc=OZD | 04/28/2023 | 10/26/2024 |  | 10/26/2024 |  |  | 04/28/2023 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1076603 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Connection Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with high-risk percutaneous coronary intervention was implanted with an impella cp for mechanical circulatory support. while on support, the automated impella controller (aic) experienced the usb to be malfunctioning, which was resolved by changing the console. no patient harm reported. | [5 empty] | The investigation has been completed. the console (ic1062) was sent back for further investigation. the root cause of the connectivity issues was likely due to a hospital network configuration issue. this report is being filed as part of a retrospective review of historical records.
16597 | 1220648-2024-21208 | 20540423 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540423&pc=OZD | 05/28/2023 | 10/26/2024 |  | 10/26/2024 |  |  | 05/28/2023 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER JP |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER JP |  |  | ABIOMED, INC. |  |  | US |  | 1379919 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 44 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced a controller error. it was reported the error appeared twice and resolved spontaneously in a few minutes. additionally, when the alarm appeared, the aortic and left ventricle position waveform were not displayed, there was no change in pump flow rate, motor waveform, or purge flow rate/pressure. the console was not replaced as there was no problem with the pump drive itself, and there was no change in the patient's hemodynamics. reportedly the team was observing the issue and would take action to replace the aic if needed. there was no patient harm reported, the aic was not replaced and the procedure was successfully completed. | [5 empty] | The investigation into the controller error/loss of opm data has been completed. the impella automated controller was returned for investigation. the data analysis confirmed the reported failure mode given there was a controller error alarms (opm comm crc error ¿ event set # 1045) triggered during the clinical procedure. real time (rt) logs confirmed that this was a pattern failure mode in which all signals received from optical bench (placement, snr, contrast, lamp, gain) are interpreted as -16384 values. this is a known pattern failure, which is characterized by a temporary loss of optical data (placement, signal to noise ratio, contrast, lamp, gain) and triggering of a controller error alarm, the event has a low probability of reoccurrence. if needed, patient hemodynamics and impella position can be monitored with imaging. the cause of the transient loss of optical data was not determined due to the inability to reproduce the event. this report is being filed as part of a retrospective review of historical records.
16598 | 1220648-2024-21209 | 20540441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540441&pc=OZD | 07/01/2023 | 10/26/2024 |  | 10/26/2024 |  |  | 07/01/2023 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1161391 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced an optical signal issue that was resolved by changing the console. no patient was involved. | [5 empty] | The investigation has been completed. the console (b)(6) was sent back for further investigation. the root cause of the issue was most likely that the optical connector was not fully connected causing the optical signal issues.this report is being filed as part of a retrospective review of historical records.
16599 | 1220648-2024-21210 | 20540442 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540442&pc=OZD | 07/05/2023 | 10/26/2024 |  | 10/26/2024 |  |  | 07/05/2023 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1504376 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility's biomedical engineer reported the automated impella controller (aic) experienced replaced due to the loss of ao signal. no patient was involved. | [5 empty] | The investigation has been completed. the console (ic8322) was sent back for further investigation. the root cause of the controller error was due to an aic software issue. this report is being filed as part of a retrospective review of historical records.
16600 | 1220648-2024-21211 | 20540454 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540454&pc=OZD | 05/01/2023 | 10/26/2024 |  | 10/26/2024 |  |  | 05/01/2023 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | AUTOMATED IMPELLA CONTROLLER |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1478138 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Fragmentation | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that while raising the bed, the lift team hospital member accidentally knocked over the automated impella controller and cart. this caused the cover over the fiber optic port to break off. no patient harm was reported due to the issue. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the damage of optical pump connector was due to use issue. this report is being filed as part of a retrospective review of historical records.
16601 | 1220648-2024-21212 | 20540460 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540460&pc=OZD | 04/25/2023 | 10/26/2024 |  | 10/26/2024 |  |  | 04/25/2023 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1180866 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The surgical consultant reported that the automated impella controller (aic) experienced a controller error during preparation. this aic was replaced with a backup aic and the procedure successfully completed. there was no patient harm reported. | [5 empty] | E1: name of the initial reporter is unknown. the investigation into the controller error/keyboard issue have been completed. the impella automated controller was returned for investigation. the data analysis of the logs indicated a controller error #24 on (b)(6) 2023 when the console was booted up. during investigation of the returned console, the error was able to be reproduced only once during testing, the clock to the qt1060 microcontroller was pulled low until unplugging and plugging back in the ribbon cable to the keyboard at which point the error resolved. the cause of the aic controller error issue was most likely a defective keyboard. this report is being filed as part of a retrospective review of historical records.
16602 | 1220648-2024-21213 | 20540462 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540462&pc=PYX | 10/01/2024 | 09/30/2025 |  | 10/26/2024 |  |  | 10/01/2024 | 10/26/2024 |  |  | 03/31/2025 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024394922 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Ventricular Fibrillation | Thrombosis/Thrombus | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | Followup | [5 empty] | The complainant had an impella rp flex placed to support a 73-year-old male patient admitted with aortic stenosis and right ventricle failure with postcardiotomy cardiogenic shock. the impella rp flex was placed but there were issues during the support. the chest had repeated washouts for bleeding and was given blood products. the impella rp flex had low flows. the patient's condition deteriorated with the flow reduction and the patient was escalated with medical therapy/intravenous drips maxed. the patient went into ventricular fibrillation and was shocked. care was withdrawn and the patient expired. | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured coagulopathy with a "definite" relationship to the impella device used: ¿impella device became clotted. patient taken to cath lab, impella was reposition and re-started¿. the patient recovered with sequelae. | The us complainant reported that the device exhibited an optical sensor issue. |  |  |  | The pump and data stick were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller instructions for use & clinical reference manual section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ | The investigation into low pump flow and vf/vt/af/at has been completed. the product was returned and there were no abnormalities seen with the device. the device was run in a pressure loop to see if the drift in the placement signal or low flows could be reproduced, but they remained stable over the course of a 3 day run. the data logs were observed and there was a decrease in cvp, motor current, and flows with an increase in placement signal. the patients condition was deteriorating and they were on crrt and losing blood volume. the root cause of the low pump flow is most likely due to the patients condition. as the necessary information was not provided, the root cause of the vt/vf was not determined. | B.5 additional information was received. h.6 codes were added for the additional information received. h.10 investigation results were added for the new information received. the investigation of thrombus has been completed. the product was returned and there were no abnormalities seen with the device. there were no motor current spikes observed or positioning issues. as the necessary information was not provided, the root cause of the thrombus was not determined b.1 product problem was inadvertently omitted from follow up 1 manufacturer device report number 1220648-2024-21213. b.2 required intervention was inadvertently omitted from manufacturer device report number 1220648-2024-21213. g.1 revised reporting contact email since the initial manufacturer device report number 1220648-2024-21213 was submitted in accordance with updated procedures. g.2 study was inadvertently omitted from manufacturer device report 1220648-2024-21213. | The investigation of low pump flow and vf/vt/af/at has been completed. the data logs were observed and on 10/1 there was a decrease in cvp, motor current, and flows with an increase in placement signal. there were no motor current spikes observed or positioning issues. the patients condition was deteriorating and they were on continuous renal replacement therapy (crrt) and losing blood volume. the root cause of the placement signal issue is most likely due to the patients condition. the root cause of the low pump flow is most likely due to the patients condition. as the necessary information was not provided, the root cause of the vt/vf was not determined. h10 the investigation results for low pump flow and vf/vt/af/at were inadvertently omitted on manufacturer device report number 1220648-2024-21213-1. | This report is being filed as part of a retrospective review of historical records. this mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the fm optical signal issue was changed to placement signal issue to address the abnormal fluidic or differential placement signal issue. the product was returned and there were no abnormalities seen with the device. the device was run in a pressure loop to see if the drift in the placement signal or low flows could be reproduced, but they remained stable over the course of a 3-day run. the data logs were observed and on 10/1 there was a decrease in cvp, motor current, and flows with an increase in placement signal. there were no motor current spikes observed or positioning issues. the patient's condition was deteriorating, and they were on crrt and losing blood volume. the root cause of the placement signal issue is most likely due to the patient's condition. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d6: added optical signal issue "2917" as part of a retrospective review of optical signal issues in accordance with fda recommendation.
16603 | 1220648-2024-21214 | 20540463 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540463&pc=OZD | 04/24/2023 | 10/26/2024 |  | 10/26/2024 |  |  | 04/24/2023 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1110486 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) experienced battery charging issues. no clinical details regarding resolution or patient involvement/condition were provided. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the console was tested with demo pump and pc for days, with ac power disconnected and reconnected occasionally. but the failure mode was not reproduced. based on the observation from data analysis, the battery level low alarm was most likely due to a momentary communication glitch between battery 1 and pbm. the root cause of the failure mode was unable to determine due to not reproduced. this report is being filed as part of a retrospective review of historical records.
16604 | 1220648-2024-21215 | 20540466 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540466&pc=OZD | 04/08/2023 | 10/26/2024 |  | 10/26/2024 |  |  | 04/14/2023 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | CONTROL,, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1489955 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Decrease in Pressure | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that when the automated impella controller (aic) was preparing to begin priming, the purge fluid did not flow and a controller failure alarm occurred, which were resolved by replacing internal components. no patient was involved. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the failure mode was due to contamination around purge motor assembly. this report is being filed as part of a retrospective review of historical records.
16605 | 1220648-2024-21216 | 20540471 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540471&pc=OZD | 04/14/2023 | 10/26/2024 |  | 10/26/2024 |  |  | 04/17/2023 | 10/26/2024 |  |  |  |  | 10/26/2024 |  | AUTOMATED IMPELLA CONTROLLER |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1645321 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 47 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Read Input Signal | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was placed onto impella cp support due cardiomyopathy. there was an automated impella controller (aic) failure alarm during impella priming. there was an indication that the purge liquid was not detected, which were resolved by replacing internal components. no patient harm was reported. | [5 empty] | The investigation for the reported issue has been completed. the impella device was received from the customer and an evaluation of the device was completed. the failure mode was reproduced. inspection of the console confirmed that pbm board was in normal operation in providing voltage to purge system. fs noted purge liquid at the purge disk, gasket of purge motor and its surrounding. once the contamination was cleaned, the problem was not reproduced the failure mode was due to contamination around purge motor assembly. this report is being filed as part of a retrospective review of historical records.
16606 | 1220648-2024-21217 | 20540477 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540477&pc=OZD | 01/21/2023 | 10/26/2024 |  | 10/26/2024 |  |  | 01/21/2023 | 10/26/2024 |  |  | 09/30/2024 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023276119 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, pump was having lower than expected flows. the previous echocardiogram verified proper positioning; however, due to the condition of the patient, the decision was made to replace the pump. the patient subsequently coded and expired prior to the replacement. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported low pump flow has been completed. the product was returned and the low pump flow was confirmed. cannula kink was observed. some biomaterials were built up around the impella too. and based on the clinical data of confirmed health conditions of the patient caused biomaterial build up. penumbra retrieved large clot. data log reviewed and many suction alarms appeared. root cause was most likely biomaterial, with a clear cannula kink observed. this report is being filed as part of a retrospective review of historical records.
16607 | 1220648-2024-21218 | 20540478 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540478&pc=OZD | 11/16/2022 | 10/26/2024 |  | 10/26/2024 |  |  | 11/16/2022 | 10/26/2024 |  |  | 07/31/2024 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023256297 |  | 0048-0045 |  | No |  |  | Yes |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Pumping Stopped | [6 empty] | Cardiac Arrest | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, there were high motor current and pump stop alarms. cardiopulmonary resuscitation remained in progress. the decision was made to replace the pump. additional information was provided that noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation for the reported pump stop has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the cause of the pump stop was most likely improper positioning of the device due to cardiopulmonary resuscitation during support. as noted in the impella ifu: pump stop: impella cp with smartassist for use during cardiogenic shock and high-risk cpi & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock. section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high-risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller. section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ this report is being filed as part of a retrospective review of historical records.
16608 | 1220648-2024-21219 | 20540479 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540479&pc=OZD | 03/24/2023 | 03/07/2025 |  | 10/26/2024 |  |  | 03/24/2023 | 10/26/2024 |  |  | 09/30/2024 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | PUMP 371 14F LT CMR SET |  |  | ABIOMED, INC. |  |  | US |  | 2023277660 |  | 0048-0045 |  | Yes |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that there was an immediate pump stop. eventually pump became ¿functional¿ but not with high reliability. the pump was explanted and replaced. during the explant there was vascular damage discovered at left femoral artery (lfa) closure site. ten units of packed red blood cells (prbc ) were given. additional information was provided that noted the patient expired. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation for the reported pump stop and vascular damage issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the returned pump was visually inspected and horizontal crease was observed on the cannula. under borescope inspection, a tangled fractured red lumen was observed in the cannula near the cannula crease. the device passed all post-sterile inspection checks. the cause of the issue was a red lumen removal issue. a crease in the cannula was observed that likely contributed to the red lumen fracture. the root cause of the vascular damage could not be determined as insufficient clinical details were provided. this report is being filed as part of a retrospective review of historical records.
16609 | 1220648-2024-21220 | 20540480 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540480&pc=OZD | 05/13/2023 | 10/01/2025 |  | 10/26/2024 |  |  | 05/13/2023 | 10/26/2024 |  |  | 12/31/2024 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2024322129 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. it was reported while on impella cp support, the patient experienced a large ooze and hematoma the impella access site. it was also reported that a femoral to femoral bypass was proactively performed to the patient for distal leg perfusion. due to the oozing at the impella access site, the patient was taken to the cardiac catheterization laboratory for a new impella cp and a larger sheath to resolve the bleeding. it was noted that care was withdrawn on the patient four days later. medical safety review of the event noted use of the device cannot conclusively be associated with the outcome (patient's demise). | The complainant reported that the staff received a placement signal unreliable alarm following insertion of the impella pump. access site bleeding and limb ischemia was also reported. lastly, hematoma was noted at impella insertion site after several rounds of cpr. dr decided to remove initial pump and replace using a 16fr dry lock for sheath. hematoma pressed out. a fem-fem bypass proactively performed for distal leg perfusion. the pump was removed and replaced. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. due to the limited clinical information and lack of available data/product the root cause of this investigation is not determined. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical signal issue as part of a retrospective review of historical records.
16610 | 1220648-2024-21221 | 20540481 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540481&pc=OZD | 06/15/2022 | 10/26/2024 |  | 10/26/2024 |  |  | 06/15/2022 | 10/26/2024 |  |  | 03/31/2024 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023191916 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  |  | Death | Yes | Yes | Positioning Failure | [6 empty] | Cardiac Arrest | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock needing an impella cp device for mechanical circulatory support. the complainant reported a purge cassette failing to prime during preparation. a new cassette was used. during use of the pump the pump had positioning issues that could not be resolved with normal troubleshooting attempts. a new impella cp was attempted to be inserted; however, the patient required cardiopulmonary resuscitation without success and the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records.
16611 | 1220648-2024-21222 | 20540482 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=20540482&pc=OZD | 06/26/2022 | 10/26/2024 |  | 10/26/2024 |  |  | 06/26/2022 | 10/26/2024 |  |  | 01/31/2024 |  | 10/26/2024 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP SMARTASSIST JP |  |  | ABIOMED, INC. |  |  | US |  | 2023180585 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Death | Yes | No | Patient-Device Incompatibility | [6 empty] | Ventricular Fibrillation | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock needing an impella cp device for mechanical circulatory support. it was reported that during implantation, the guidewire would not pass due to a block. the team tried to insert the guide wire by pulling out the guide tube, but could not insert it because the guide wire did not come out along the long axis marker of the discharge part. the implantation was aborted as the patient suffered from ventricular fibrillation. additional information was provided that noted the patient expired. medical safety review of the event noted there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this report is being filed as part of a retrospective review of historical records. as noted in the impella ifu: impella cp with smart assist system section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿