9638 | 1220648-2025-28184 | 21978589 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21978589&pc=PYX | 04/17/2025 | 06/06/2025 |  | 05/07/2025 |  |  | 04/17/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2026625261 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 83-year-old male patient was admitted in need of bipella support. the nstemi patient had biventricular failure. the cp was placed however the rp flex failed to deliver. the pump kinked and was aborted. the patient outcome was unknown. | [5 empty] | Reported pump was unable to pass femoral vein/ivc junction and kinking behind the motor housing. device was removed and another device was attempted. delivery issues - the root cause of the delivery issue was most likely patient condition related since the pump was unable to pass the femoral vein/ivc junction. product damage - the root cause of the product damage was most likely a catheter kink due to patient condition. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. the ifu states: rp flex with smart assist system with automated impella controller. section: warnings & cautions: warnings - ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves - ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings - ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
9639 | 1220648-2025-28185 | 21978795 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21978795&pc=OZD | 04/17/2025 | 05/29/2025 |  | 05/07/2025 |  |  | 04/17/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639375 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 75 year old on an impella cp which was implanted in the right femoral artery. the left ventricle was small with a short aorta. during the cardiac procedure, the impella was shallow, and cannula was kinked just distal to the outlet. kink caused suction, so impella was turned down to p2, flows 1.5l throughout the pci case. the impella was removed post case and there was no visible kink remaining on cannula. the patient survived the procedure. | [5 empty] | The investigation of low pump flow, product damage (cannula kink), and positioning issues has been completed. the pump was returned for investigation. cannula kink was reported on the returned product. no other physical damage was noted on the returned product. the pump was run as per the specifications in a bucket of water. the data log shows low pump flow with suction alarm from the start of the case. there were no positioning-related alarms or motor current spikes reported during the case. the cause of the product damage issue was a kinked cannula. the cause of the low pump flow issue was a kinked cannula. the cause of the positioning issue was most likely patient condition related to small heart. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
9640 | 1220648-2025-28186 | 21978889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21978889&pc=OZD | 04/17/2025 | 06/06/2025 |  | 05/07/2025 |  |  | 04/17/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026627481 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Device Handling Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the 61-year-old male patient was being prepped for an impella cp implant and during the preparation, the pump was prepped and purged. staff member accidently started pump while on field with "red cheater" still in cannula. motor stopped with red alarm impella stopped. team opened a new impella cp and proceeded with case. the impella introducer was implanted in the patient. the patient was eventually weaned and explanted. | [5 empty] | Impella cp was returned for evaluation. upon visual inspection of returned pump, the full red lumen was present within the impella cannula. insufficient data logs were returned for analysis at time of pump stop. clinical details noted that the pump was accidentally started outside of the body with the red lumen still in the cannula. as the pump failure occurred within the first 5 minutes of activation, the failure mode was change from "pump stop" to "pump did not start". the root cause of the pump did not start was due to red lumen issue as the pump was accidentally started outside of the body with the red lumen still inserted. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high-risk pci states the following: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
9641 | 1220648-2025-28187 | 21979131 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21979131&pc=OZD | 08/10/2023 | 06/11/2025 |  | 05/07/2025 |  |  | 08/29/2024 | 05/07/2025 |  |  | 06/30/2024 |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023247278 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient on impella cp support experienced hemolysis due to low platelets. the patient also experienced oozing at the groin site while activated clotting time (act) was above 200 and it is unknown if or what intervention was done. additionally, suction alarms and low flows occurred which were improved with one unit of blood and additional volume. furthermore, the impella cp was repositioned due to being too deep in the ventricle via echocardiogram. the site bleeding resolved and there were no signs of a hematoma. the patient was bridged to a left ventricular-assist device and was stable. | [5 empty] | The investigation of hemolysis, access site bleeding, positioning issues, and low pump flow has been completed. the impella device was not returned for evaluation. access site bleeding: the clinical states that during the time of reported oozing from the access site, the activated clotting time (act) was over 200. based on the clinical details, the root cause of the access site bleeding is most likely due to anticoagulation management due to higher act's during the bleeding. low pump flow: the clinical details state that the patient experienced suction and low flows that were improved with 1 unit of blood and additional volume. there were slightly lower flows seen in the data logs around that time and a trend in the placement signal. based on the clinical details and the data log analysis, the root cause of the low pump flow is most likely due to the patients condition. hemolysis: due to lack of clinical information and multiple events happening during this time, the true root cause of the hemolysis could not be confirmed. positioning issue: due to lack of clinical information and multiple events happening during this time, the true root cause of the positioning issues could not be confirmed. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9642 | 1220648-2025-28188 | 21979158 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21979158&pc=OZD | 04/17/2025 | 05/07/2025 |  | 05/07/2025 |  |  | 04/17/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026645874 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support and during support, the patient had bleeding noted surrounding arteriotomy site. two units of packed red blood cells were given to the patient. additionally, the patient had ventricular tachycardia. the patient was cardioverted. additionally, the patient had limb ischemia. the pump was explanted and post explant, the pulses were palpable. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9643 | 1220648-2025-28189 | 21979284 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21979284&pc=OZD | 03/16/2025 | 05/07/2025 |  | 05/07/2025 |  |  | 03/16/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026619794 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 66 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Respiratory Insufficiency | Delirium | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was documented via an impact study that a patient was implanted with an impella 5.5 and experienced respiratory insufficiency and delirium. both conditions were believed to have a possible relationship to the impella 5.5 procedure. the delirium was treated with haloperidol. the respiratory insufficiency was believed to have stemmed from a pleural effusion that developed following removal of the drainage. the pump was successfully weaned and no harm was reported as the patients outcome. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
9644 | 1220648-2025-28190 | 21979298 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21979298&pc=OZD | 02/10/2025 | 09/24/2025 |  | 05/07/2025 |  |  | 02/10/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025569310 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male | 109 KG | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | Thrombocytopenia | [10 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant also reported that when the patient was ambulating in the hall, a placement signal unreliable alarm was active. the alarm was disabled. | The complainant reported that after impella 5.5 insertion there was bleeding at the axillary cutdown site. the dressing was changed and one unit of packed red blood cells was transfused. the bleeding was resolved. the patient developed thrombocytopenia, and systemic anticoagulation was not ordered. additionally, the patient was septic. the team is unsure of the cause. the patient survived. |  |  |  |  | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal, access site bleeding, thrombocytopenia, and infection/sepsis. no product was returned. on 01may, the clinician indicated that the placement signal became unreliable when the patient was ambulating in the halls. the data logs show that the placement signal (ps) suddenly goes out of range, while the signal-to-noise ratio (snr) drops to 0, contrast rails, lamp goes to 100, and gain and light both decreases. additionally, the motor current has some small spikes in it after this, which can indicate stress on the motor current lines. based on the data logs, the root cause is most likely due to a damaged fiber line based on the clinical details and data log pattern. due to limited clinical details, the root cause of the access site bleeding cannot be determined. the cause of thrombocytopenia was not determined due to insufficient clinical details. an infection/sepsis can occur during or after impella support. the root cause of the infection cannot be determined. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include placement signal issue description. h6 updated to include placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-28190. d10 concomitant medical products and therapy dates updated. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock states the following: section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9645 | 1220648-2025-28191 | 21979505 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21979505&pc=OZD | 04/18/2025 | 06/11/2025 |  | 05/07/2025 |  |  | 04/18/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026616547 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Atrial Flutter | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella 5.5 pump to support a patient admitted in with cardiomyopathy. the pump was supporting while transplant conversations were ongoing. the pump had suction repeatedly noted and the patient had to be cardioverted due to atrial flutter. the plan is to let the patient rest on impella 5.5 support while waiting for a heart transplant. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | No product had been returned for analysis as of 28-may-2025. low pump flow: clinical mentioned resolution of suction after volume was administered. the data logs show placement signal trending down over the duration of support. there were suction alarms as mentioned in the clinical details. the logs do not show any motor current spikes outside p-level changes. the logs show 'impella position unknown' later on during support but not within suction occurrence. the root cause of the low pump flow was most likely patient condition related as evidenced by the trending placement signal and clinical details mentioning volume issues and resolution after volume given. vt/vf/af/at: cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease ¿timing of the arrythmia event in relation to impella support (during or post-implant) ¿electrocardiogram (ekg) recordings of the arrhythmia episodes ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities ¿response to any antiarrhythmic treatments or interventions during the event ¿any documented device-related issues or complications during impella support ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand ¿review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined.
9646 | 1220648-2025-28192 | 21979572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21979572&pc=OZD | 04/17/2025 | 06/10/2025 |  | 05/07/2025 |  |  | 04/17/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026640425 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was attempted to be implanted with an impella device for mechanical circulatory support. due to anatomy/bleeding/calcium in common femoral arteries and concerns for instability/bleeding, the decision was made by to abort the case. no patient harm was reported. | [5 empty] | Due to anatomy/bleeding/calcium in common femoral arteries case aborted. the root cause of the delivery issue was most likely patient condition related, due to anatomy/bleeding/calcium in common femoral arteries and concerns for instability/bleeding. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves - ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings - ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions - ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings - ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves - ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings - ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves - ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings - ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
9647 | 1220648-2025-28193 | 21979709 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21979709&pc=OZD | 04/18/2025 | 06/06/2025 |  | 05/07/2025 |  |  | 04/18/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026622724 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Malposition of Device | [6 empty] | Atrial Fibrillation | Tachycardia | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs), and approximately two days after start of the support, the patient had runs of supraventricular and ventricular tachycardia. the impella cp device was repositioned under echocardiogram and pulled back to 3.9 centimeters from 4.3 centimeters. the inlet was within mid-ventricular space. the next day, the patient experienced atrial fibrillation with rapid ventricular response, was converted with aminodarone bolus. the patient continued on impella mcs. following, some days later, plans were made and executed upon to remove the impella cp due to low platelet count (thrombocytopenia). the patient remained hospitalized for further post-op recovery and was noted to be in stable condition. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of vf/vt/af/at, positioning issues, and thrombocytopenia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
9648 | 1220648-2025-28194 | 21979797 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21979797&pc=OZD | 04/17/2025 | 06/12/2025 |  | 05/07/2025 |  |  | 04/17/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026650958 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that the transseptal ablation catheter hooked on impella pigtail and caused it to be pulled back into aorta. dr (b)(6) decided to remove impella. there was no patient harm. | [5 empty] | The pump was pulled out of position by the patient's transseptal ablation catheter as stated in the clinical. positioning issues: the cause of the positioning issue most likely was an interaction with another device as the transseptal ablation catheter hooked on the impella pigtail and caused it to be pulled back into aorta. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9649 | 1220648-2025-28195 | 21980220 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21980220&pc=OZD | 04/17/2025 | 05/07/2025 |  | 05/07/2025 |  |  | 04/17/2025 | 05/07/2025 |  |  |  |  | 05/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025592038 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. there was a noted loss of pulsatility during balloon inflations, but pulsality recovered. the impella was weaned and removed. impella sheath removed and groin closed with prostar. pt had color change of impella leg, but palpable distal pulse. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
9650 | 1220648-2025-28098 | 21966963 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21966963&pc=OZD | 04/14/2025 | 06/04/2025 |  | 05/06/2025 |  |  | 04/14/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025574314 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella 5.5 was implanted in a patient found unresponsive in their home and who had received cardiopulmonary resuscitation for two hours. upon insertion it was found that the impella guidewire was kinked. the guidewire was exchanged and the impella was backloaded in standard fashion. the impella was noted to be deep and was repositioned. the patient's family ultimately withdrew care and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of delivery issues and positioning issues has been completed. the impella device was not returned for evaluation. the failure mode (fm): product damage (cannula kink) changed to fm: delivery issue since the guidewire kink prior to insertion per salesforce reported information. delivery issue: the cause of the delivery issue cannot be determined since complaint device not returned for analysis and insufficient clinical data provided. positioning issues: the cause of the positioning issues cannot be determined since insufficient clinical data provided why position issue occurred. instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist system with smart assist system with automated impella controller: section: warnings & cautions: warnings, section: pre-support evaluation, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system with automated impella controller: section: warnings & cautions: warnings, section: pre-support evaluation, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
9651 | 1220648-2025-28099 | 21967035 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21967035&pc=OZD | 04/15/2025 | 06/06/2025 |  | 05/06/2025 |  |  | 04/15/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025599974 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that roughly two minutes into impella cp support of a patient, the left ventricle waveform decoupled and suction alarms appeared. despite multiple attempts to reposition the device, suction persisted. the scheduled percutaneous coronary intervention was successfully performed while being supported with the decreased flow rates. the decision was ultimately made to explant the pump when the flow rate remained low and the alarms persisted. there was no resulting patient harm. | [5 empty] | Low flow - reported initially starting the pump at 3.1l/min on auto and placed on p7 with no change in flow but decreased to 1-1.5l/min within a few minutes with suction alarms. repositioning attempted but did not resolve the issue. patient had tavr valve. pump was returned and damaged impeller blades were found. data logs show two motor current spike within the first few minutes of the pump running followed by suction. shortly after there is a step down in pump flow and pump remains in low for for the duration of the case. the root cause of the low flow was most likely fractured impeller blades since the blades were damaged on the returned pump. the fracture was characteristic of an interaction with another device but that was unable to be definitively determined since fluoroscopic imaging and/or additional clinical details were not provided. access site bleeding - reported bleeding through the sheath after inserting the pump and the sheath was exchanged for the long sheath. sheath was found and a scoreline split was found with evidence of buckling. the root cause of the access site bleeding was an introducer sheath split along the scoreline since a scoreline split was found on the returned introducer. cause of the sheath split was unknown. | The impella pump, introducer and data stick were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
9652 | 1220648-2025-28100 | 21967128 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21967128&pc=OZD | 04/14/2025 | 05/29/2025 |  | 05/06/2025 |  |  | 04/14/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026634119 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella cp was implanted in a patient admitted with multiple cardiac and systemic comorbidities. an access site bleed occurred that prompted intervention. the patient was transfused with three units of red blood cells and survived. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9653 | 1220648-2025-28101 | 21967158 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21967158&pc=OZD | 04/12/2025 | 05/06/2025 |  | 05/06/2025 |  |  | 04/12/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025586007 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 5.5 pump for a 62 year old female patient admitted in with cardiogenic shock and cardiomyopathy. the pump supported for just under 51 days and was then explanted for the heart transplant. the pump did have positional issues throughout the support. the pump was often assessed and repositioned. the pump access site also had treatment for infection. the site had grown staph and was medically managed. the patient continued on support. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smartassist for use during cardiogenic shock states the following: section: table 3.2 impella catheter components: "the infusion filter prevents bacterial contamination and air from entering the purge lumen." section: warnings & cautions: warnings: "to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers."
9654 | 1220648-2025-28102 | 21967193 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21967193&pc=OZD | 04/14/2025 | 05/06/2025 |  | 05/06/2025 |  |  | 04/14/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026630530 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella 5.5 and experienced torsades de pointes and required defibrillation. additionally, the pump was repositioned under echocardiogram as lactate dehydrogenase was elevated. two days later, the impella 5.5 was repositioned again at bedside due to alarms and the patient went into ventricular tachycardia. the patient was defibrillated, and the pump was observed to be too shallow. again, the pump was repositioned to normalize pressure and flow. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9655 | 1220648-2025-28103 | 21967291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21967291&pc=OZD | 04/14/2025 | 06/04/2025 |  | 05/06/2025 |  |  | 04/14/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025585479 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 76-year-old male patient developed hemolysis during impella cp support. the p-level was adjusted and repositioning was completed but the condition persisted. due to the ongoing hemolysis, the decision was made to explant the pump. following removal, a manta was deployed but was unsuccessful in closing the site. manual pressure was held and a balloon was placed until vascular consult could close the site. patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and hemolysis has been completed. the impella device was not returned for evaluation. the root cause of the hemolysis was most likely biomaterial since data logs show a motor current spike prior to reported hemolysis. the root cause of the access site bleeding was not determined since product was not returned and insufficient clinical details were provided. instructions for use: hemolysis. impella cp with smartassist system: section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ h10 instructions for use for hemolysis were inadvertently omitted on the initial manufacturer device report number 1220648-2025-28103.
9656 | 1220648-2025-28104 | 21967296 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21967296&pc=OZD | 02/11/2025 | 07/18/2025 |  | 05/06/2025 |  |  | 04/14/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025578566 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | Tachycardia | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported via abiomed's long-term outcome and quality indicator (loqi) impella registry that a patient endured recurring non-sustained ventricular tachycardia (vt). the vt was believed to have a possible relationship to the impella device. as support continued, the patient was non-compliant, and the decision was ultimately made to withdraw care. the patient was transferred to palliative care and passed away. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | Upon later notification via loqi, it was documented that on (b)(6) 2025, the patient endured sepsis believed to have a possible relationship to the impella device used. a blood culture came back positive for enterobacter cloacae. antibiotics were given to treat the condition |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | B5 updated with additional information received from the clinical site. h6 health effect - clinical code and health effect - impact code updated to reflect the additional adverse event that was reported.
9657 | 1220648-2025-28105 | 21967381 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21967381&pc=OZD | 04/14/2025 | 06/11/2025 |  | 05/06/2025 |  |  | 04/14/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025609631 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella cp. during the peel away sheath removal, the pump was pulled out of left ventricle. after several attempts, the doctor was able to recross into the left ventricle. the pump was sutured in a forward fashion. care was withdrawn and the patient expired. | [5 empty] | The cause of the positioning issues most likely was use related due to improper technique since during peel away sheath removal pump pulled out of lv and was able to recross it. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9658 | 1220648-2025-28106 | 21967570 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21967570&pc=OZD | 04/14/2025 | 06/05/2025 |  | 05/06/2025 |  |  | 04/14/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1301547 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Nurse | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the automated impella controller (aic) beeped, and the screen turned black with white text. after it "rebooted", staff reports "unexpected controller shutdown" alarm was present. the aic was replaced. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | Data analysis: console logs aligned with the complaint, showing an unexpected shutdown at approximately 11:08:28 on 04/14, followed by "unexpected shutdown" alarm after the system rebooted. there were no warning signs or precursor events leading up to the shutdown, and no indications of any software failures. after the unexpected shutdown, the console's software automatically restarted and pump was operating at the previous p-level. the aic ic(b)(6) was exchanged with ic(b)(6) within 10 mins of unexpected shutdown alarm was triggered. device analysis: the console was tested for 100 powercycles in danvers, but the issue could not be reproduced. no unexpected shutdowns or performance issues were observed. a thorough visual inspection of the internal components was conducted, but no physical problems were found. there was no interruption in the supply were detected between the pbm and the 4-in-1 assembly. root cause: the root cause of the intermittent, unexpected controller shutdown could not be determined, as there was no clear indication of the cause, and the issue could not be reproduced. corrective actions: before returning the console to the customer, the following actions are instructed to perform: 1. replace the tested 4-in-1 gen2 assy (5500-0036) as a precaution. 2. replace the compact flash (2025-0019) as a precaution. 3. replace the pbm (0042-1025/0042-1125) as a precaution. 4. perform a full functional check on the console (0042-7316).
9659 | 1220648-2025-28107 | 21967633 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21967633&pc=OZD | 04/14/2025 | 12/11/2025 |  | 05/06/2025 |  |  | 04/14/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636394 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was being prepped for an impella implant, and the impella cp pump malfunctioned after it was opened and connected to automated impella controller. pump was not implanted into the patient, however the introducer was used. a new pump was opened and used. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. ection: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | Additional information was received in the form of medical safety review which noted the following: an adult patient status post mitral valve replacement (mvr) experienced cardiac arrest and was taken to the cardiac catheterization lab (cvl) for revascularization. despite intervention, the patient continued to code, prompting initiation of veno-arterial extracorporeal membrane oxygenation (va-ecmo) for mechanical circulatory support. transesophageal echocardiography (tee) revealed severe left ventricular (lv) distension while on ecmo. the decision was made to place an impella cp device for lv unloading. impella cp pump 1 (malfunction): not implanted due to purge priming malfunction. it is unclear if an alternate purge cassette was attempted. impella cp pump 2: (not reportable) not implanted due to contamination upon opening by the healthcare provider (hcp). impella cp pump 3: (malfunction) (successfully implanted without issue. patient remained on impella cp support for approximately 2 days before escalation to impella 5.5 due to positioning concerns with the impella cp. impella 5.5 pump 4: (death) implanted for continued lv unloading. patient remained on impella 5.5 and ecmo with maximal medical therapy for approximately 1 day. subsequently, the family elected to withdraw care. the patient expired while supported on ecmo and impella 5.5. note: type of reportable event and investigation with coding remains unchanged as previously reported. note: this impella cp pump 1 is one of three device reports associated with the event story. the manufacturer report number for impella cp pump 3 (malfunction) and impella 5.5 pump 4 (death) associated complaints will be generated in the new complaint handling system and once these regulatory reports have been generated, they will capture this impella cp pump 1 device report number (in the related report number field).
9660 | 1220648-2025-28108 | 21967765 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21967765&pc=OZD | 04/15/2025 | 06/06/2025 |  | 05/06/2025 |  |  | 04/15/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026637015 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a decision for an impella cp to be implanted in a patient for repeat coronary angiogram. it was noted that they attempted to place impella cp but met resistance at the aorta-iliac bifurcation. the impella cp was removed and the sheath was exchanged for a 14 french long sheath. however, they were unable to advance the sheath past the same area of resistance. a 6 millimeter peripheral balloon angioplasty was then performed to that area. the pigtail and 0.035 were reinserted into the left ventricle and the wire exchanged for an 0.018 wire and the pigtail was removed. the impella cp was then advanced over the wire into left ventricle without difficulty. the wire was removed and the impella cp was started on auto. it was initially ramped up to 2.7 literms per minute and was placed on p7 but within a minute, suction alarms presented and flow decreased. despite repositioning, impella cp appeared to be kinked on fluoroscopy image just distal to impella motor/outflow. if pulled back, the kink would resolve but only if the impella cp was almost entirely in the aorta. if advanced, the kink was still present, as were suction alarms and decreased flow. there was concern that the impella cp may have kinked due to resistance on first attempt at insertion, so the impella cp was removed and a new impella cp was inserted without resistance. the patient was successfully weaned. | [5 empty] | Impella cp was returned for evaluation. low pump flow: upon visual inspection a cannula kink was present just distal of the motor housing. data logs were reviewed and pump flows were lower than set p-level. suction alarm triggered at time of low pump flows. clinical details noted that pump was kinked at the cannula when positioned in the lv and low flows were occurring. the root cause of the low pump flows was most likely due to a cannula kink due to reported resistance at aortic-iliac bifurcation and short ascending aorta. product damage (cannula kink): returned product showed a cannula kink just distal to the motor housing and outflow cage. clinical details noted that the pump was found to be kinked after pump was inserted and positioned. the root cause of the product damage was due to a cannula kink due to reported resistance at aortic-iliac bifurcation and short ascending aorta. second pump also resulted in a kink. positioning issues: clinical details noted that pump was kinked when in position across valve and in lv. when attempting to resolve the kink, the pump would be pulled back to almost entirely in aorta to remove the kink. clinical details noted that patient had a short ascending aorta and when a second pump was inserted, it was having the same issue when placed in the lv. the root cause of the positioning issues was most likely patient condition related due to reported resistance at aortic-iliac bifurcation and short ascending aorta. delivery issues: a failure mode of "delivery issues" was added to the investigation based on clinical details noting difficulty inserting pump. clinical details noted that resistance was met at aortic-iliac bifurcation and both long and short introducers were used in attempt to advance pump without success. an angioplasty of the area was performed and the pump was then able to successfully be inserted. the root cause of the delivery issues was most likely patient condition related due to reported resistance at aortic-iliac bifurcation and short ascending aorta. | The impella pump and data stick were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
9661 | 1220648-2025-28109 | 21967902 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21967902&pc=OZD | 04/14/2025 | 06/04/2025 |  | 05/06/2025 |  |  | 04/14/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026632717 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, there was oozing at the access site. anticoagulation was withheld. when systemic heparin was given, the ooze started once again. physician did a figure of 8 suture instead of forward tension. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: "assess access site for bleeding and hematoma." section: entering cardiac output: use of the repositioning sheath and peel-away introducer: "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site." section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was anticoagulation management due to high patient activated clotting time (act) values, anticoagulation was weaned down after procedure to achieve hemostasis as per the clinical description.
9662 | 1220648-2025-28110 | 21967973 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21967973&pc=OZD | 04/25/2025 | 09/22/2025 |  | 05/06/2025 |  |  | 04/25/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026630537 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that an impella 5.5 was supporting a patient when the impella stopped but they were able to restart. the next day, the impella 5.5 stopped and restarted again. no spikes or changes in motor current were noted prior to the pump stop. the pump was eventually explanted and the patient continued on extracorporeal membrane oxygenation support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock entering cardiac output (impella cp with smart assist) section: direct aortic insertion. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ impella 5.5 with smartassist for use during cardiogenic shock section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the pump was not returned for investigation. on the seventh day of use, data log shows optical fiber line damage trend on all 5s optical parameters, which persisted until the end of the case run. the exact location of the damage fiber line could not be confirmed without the returned product. the cause of the temporary pump stop issue was most likely an open electrical line, based on the data log. the exact location of the electrical open line could not be confirmed without the returned product. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: add optical signal issue per a retrospective review of optical signal issues in accordance with fda recommendation. | The investigation of temporary pump stop has been completed. the impella device was not returned for evaluation. the cause of the temporary pump stop issue was most likely an open electrical line, based on the data log. the exact location of the electrical open line could not be confirmed without the returned product.
9663 | 1220648-2025-28111 | 21968031 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21968031&pc=OZD | 04/14/2025 | 06/10/2025 |  | 05/06/2025 |  |  | 04/14/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026642656 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp was placed over impella 0.018 wire. pigtail was pointing toward the mitral valve. during repositioning, impella came out of the aortic valve and was unable to cross back. the impella implantation was unsuccessful. the patients outcome is unknown. | [5 empty] | Positioning issues: the cause of the positioning issues most likely was use related because the pump pulled out of the lv upon initial repositioning attempt. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9664 | 1220648-2025-28112 | 21968147 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21968147&pc=OZD | 04/14/2025 | 06/05/2025 |  | 05/06/2025 |  |  | 04/14/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026643587 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Contamination | Failure to Advance | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was being prepped for an impella cp implant and the pump was opened and was inadvertently contaminated. the introducer was used on the patient but the pump was not used. the patient's outcome is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues and packaging issues- sterility has been completed. the impella device was not returned for evaluation. the root cause of the packaging issues - sterility is most likely due to use as the clinical stated that the pump was inadvertently contaminated and they had the same issue with the first pump. the root cause of the delivery issue is most likely due to use as pump insertion was aborted as the user contaminated the device.
9665 | 1220648-2025-28061 | 21968171 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21968171&pc=OZD | 04/09/2025 | 06/12/2025 |  | 05/06/2025 |  |  | 04/11/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026640017 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that after implantation of an impella cp, no placement signal or motor current signal was shown. the pump was removed and replaced. the patient survived. | [5 empty] | The investigation of optical signal issue has been completed. the pump was returned for investigation. data logs were not provided, and they could not be retrieved from the remote server. the returned pump passed the laser continuity test for the optical sensor, and all the 5s parameters were within the specification limits. the pump was further tested in a bucket of water and under pressure testing, where no issues were found. the cause of the optical signal issue was not determined since sufficient clinical information was not provided including data logs. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
9666 | 1220648-2025-28113 | 21968264 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21968264&pc=OZD | 04/14/2025 | 05/06/2025 |  | 05/06/2025 |  |  | 04/14/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026623953 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 24 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was unable to advance through the patient's vasculature during attempted delivery. the implant was aborted and an impella cp pump was placed for patient support. the patient's outcome is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
9667 | 1220648-2025-28114 | 21968552 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21968552&pc=OZD | 04/15/2025 | 07/11/2025 |  | 05/06/2025 |  |  | 04/15/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026641272 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a male patient was on impella cp due to acute myocardial infarction. it was reported that during an ultrasound, mitral valve insufficiency was noted and the pump was found to be positioned towards the mitral valve. attempts to reposition the pump failed and the decision was made to remove the device. there was no patient harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. based on the clinical information provided, the root cause of the positioning issue was patient condition since patient had a diseased mitral valve which was causing mitral valve insufficiency.
9668 | 1220648-2025-28115 | 21968570 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21968570&pc=OZD | 04/14/2025 | 09/24/2025 |  | 05/06/2025 |  |  | 04/14/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026613569 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported at patient on impella 5.5 support and it was documented that the medical docotr had difficulty pulling the catheter back during attempted repositioning of the pump. substantial force was said to have been used in order to pull the device back one centimeter. support continued with the implanted pump despite the noted issue. there was no resulting patient harm. | The complainant reported that an implanted impella 5.5 pump exhibited an optical signal issue. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | There was no product returned. this mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. positioning issues: the clinical details state that the md was trying to reposition 5.5, but had difficulty pulling catheter back. he had to use substantial force to pull device back one centimeter, meeting a lot of resistance. when asked if he encountered difficulty previously, md denied. it is not confirmed why the pump was repositioned or why there was resistance met. due to the lack of clinical information provided, the root cause of the positioning issues cannot be determined. optical signal issue: it was reported on 4/14 that there were continuing impella position in aorta alarms with confirmed positioning. the data logs show that the impella position in aorta alarms were accurately triggered as the placement signal would increase while the motor current loses pulsatility. the ps, contrast and snr increase in a cyclic pattern, but the cause of this is unknown. due to lack of clinical information provided, the root cause of the optical signal issue cannot be determined. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical signal issues for "2917" as part of a retrospective review of optical signal issues in accordance with fda recommendation.
9669 | 1220648-2025-28116 | 21968637 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21968637&pc=OZD | 04/15/2025 | 06/06/2025 |  | 05/06/2025 |  |  | 04/15/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1215711 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an automated impella controller (aic) did not connect to the wi-fi network when it was used for a case. when the device became available, they troubleshot the issue. the behavior on the rlm during boot up and would blink the cloud light and the airplane light together non-stop and continued to blink both lights for over five minutes. it never blinked error code of 1-6 blinks on the cloud light. the console was removed from service. | [5 empty] | The impella console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - battery comm. failure (yellow alarm) has been completed. based on data and device analysis there was no issue found with a battery communication failure.
9670 | 1220648-2025-28117 | 21968663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21968663&pc=OZD | 03/30/2025 | 10/07/2025 |  | 05/06/2025 |  |  | 04/07/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026631325 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the impella 5.5 pocket wound opened. wound care was initially consulted and the wound was packed with iodoform. anticoagulants were stopped and the patient was taken for multiple washouts. support continued uninterrupted. | The complainant also reported that the placement signal stopped working. |  |  |  |  | D6: explant date unknown. the investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. (only the cassette was returned for investigation, not the pump). | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue: data logs were reviewed, however, the logs were only available from case start. through the available logs, and even prior to starting the pump, signal-to-noise ratio (snr) was below 2000. since data logs were not available during the period of the reported psnr alarms and product was not returned for investigation, the optical signal issues could not be determined. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include placement signal issue description. h6 updated to include placement signal coding. corrections that were made from the initial manufacturer device report number. d6b updated. d10 concomitant medical products and therapy dates updated.
9671 | 1220648-2025-28118 | 21968715 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21968715&pc=OZD | 04/15/2025 | 05/29/2025 |  | 05/06/2025 |  |  | 04/15/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026640268 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella cp was going to be implanted but the patient refused to lay flat due to back pain. the patient was moving constantly and refused to lay still. the decision was made to take the patient to the operating room for impella 5.5 placement. the impella cp insertion was aborted. the patients outcome is stable. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related. the patient refused to lay flat due to back pain. the patient was moving constantly and refused to lay still. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system: section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
9672 | 1220648-2025-28119 | 21968878 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21968878&pc=OZD | 04/28/2025 | 05/06/2025 |  | 05/06/2025 |  |  | 04/28/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility |  |  | 67 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Disability | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A literature piece titled, "a case of acute lower limb compartment syndrome caused by impella" discussed a patient who received an impella implant due to having ventricular fibrillation and needing a percutaneous coronary intervention. mechanical cardiac support was weaned off, however swelling and poor coloration of their left lower extremities appeared 12 hours after impella support. although angiography showed no occlusion, the artery including superficial femoral artery and below the knee was generally shrunken. unfortunately, limb salvage was failed despite relief incision. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
9673 | 1220648-2025-28120 | 21968925 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21968925&pc=OZD | 04/15/2025 | 06/04/2025 |  | 05/06/2025 |  |  | 04/15/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026614955 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Tachycardia | Ectopic Heartbeat | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A patient drove himself into urgent care with shortness of breath. he was transferred to the hospital where they intubated him. they did an electrocardiogram and saw a st-segment elevated myocardial infarction. the patient was treated with portneuf and was brought to the catheterization lab. it was found that the patient had totaled right coronary artery. they fixed that and pressures dropped. they ended up coding him and got return of spontaneous circulation. an impella cp was subsequently placed for mechanical circulatory support. upon positioning the pump, it was noted that the patient experienced more ectopy. the pump was subsequently left deeper to alleviate the condition. the following day, the patient was noted to be experiencing hemolysis. a unit of blood was given to help stabilize the patient. the patient's symptoms were noted to be clearing. five days later, the patient experienced ventricular tachycardia and was shocked four times. the patient expired later that day. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
9674 | 1220648-2025-28121 | 21969125 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21969125&pc=OZD | 04/15/2025 | 05/06/2025 |  | 05/06/2025 |  |  | 04/15/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025604751 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. the patient experienced a gastrointestinal bleed during impella 5.5 support. heparin was put on hold and over six units of blood were given to the patient. support was continued uninterrupted. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
9675 | 1220648-2025-28122 | 21969171 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21969171&pc=OZD | 05/01/2025 | 07/11/2025 |  | 05/06/2025 |  |  | 05/02/2025 | 05/06/2025 |  |  | 12/31/2025 |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5-JP |  |  | ABIOMED, INC. |  |  | US |  | 2025444069 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella 5.5 pump was unable to advance through the patient's vasculature during attempted delivery. the implant was aborted and an impella cp was placed for planned support. the patient's outcome was unknown. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, the impella was unable to pass through due to vascular anatomical factors.
9676 | 1220648-2025-28123 | 21969242 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21969242&pc=OZD | 05/01/2025 | 07/11/2025 |  | 05/06/2025 |  |  | 05/01/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026623724 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 46 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during escalation from an impella cp to an impella 5.5 pump, the team found that the native anatomy would not allow the impella 5.5 delivery. the team performed ballooning but, the delivery failed and was aborted. the patient outcome was unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿. | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be patient condition because pump was unable to pass the anastomosis and the shaft was bent during insertion as per clinical details and the insertion was aborted.
9677 | 1220648-2025-28124 | 21969358 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21969358&pc=OZD | 04/15/2025 | 05/30/2025 |  | 05/06/2025 |  |  | 04/15/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653732 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. there were noted suction alarms during the night which resolved with one liter of fluid. additionally, ventricular tachycardia was also present. the impella was repositioned, and the p-level was lowered. the patient is stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of low pump flow, vf/vt/af/at, and positioning issues has been completed. the impella device was not returned for evaluation. as per the clinical information provided and data log review, the root cause of the low pump flow issue was most likely patient condition related. to patients low blood volume which was given as troubleshooting to suction alarms. as the necessary information was not provided, the root cause of the vt/vf was not determined. the root cause of the positioning issue was not determined since product was not returned and insufficient clinical information was provided.
9678 | 1220648-2025-28125 | 21969616 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21969616&pc=OZD | 04/24/2025 | 07/10/2025 |  | 05/06/2025 |  |  | 04/30/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025513752 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a heavy bleed was observed between the 16fr sheath (sheath made by another company) and the blood vessel. the impella was removed to control the bleed. the patient remained supported via extracorporeal membrane oxygenation. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding is determined to be use issue since non-abiomed product was used which is not recommended per instructions for use. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-28125 as it is unknown if the initial reporter also sent the report to the food and drug administration.
9679 | 1220648-2025-28126 | 21969628 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21969628&pc=OZD | 04/15/2025 | 06/11/2025 |  | 05/06/2025 |  |  | 04/15/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026637014 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was attempted to be implanted with an impella cp device for mechanical circulatory support. it was reported that there was inability to successfully deliver the pump. the pump had a kink and low flows prompted removal. no patient harm was reported. | [5 empty] | The investigation of product damage (cannula kink) and low pump flow has been completed. the data was not returned for review. visual inspection found a kink on can about 15 mm from of cage and cannula bonding site. no other issues were found on the rest of the pump. low flow: the root cause of low flow was most likely kinked cannula. product damage (cannula kink): the root cause of kinked cannula was most likely patient condition related since the patient had aorta-iliac bifurcation. | The impella pump and data stick were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
9680 | 1220648-2025-28127 | 21969794 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21969794&pc=OZD | 04/15/2025 | 06/23/2025 |  | 05/06/2025 |  |  | 04/15/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025611101 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Male | 80 KG |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that access site bleeding was noted following impella cp implant. vascular surgery was performed at the site and the bleeding stopped. two units of blood were also given to the patient in response to the condition. the patient survived. | [5 empty] | The investigation of vascular damage - peripheral vessel has been completed. the pump was returned with no abnormalities found. the root cause of the vascular damage peripheral vessel was not determined due to insufficient clinical details. d9: device returned to manufacturer date added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9681 | 1220648-2025-28128 | 21969836 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21969836&pc=OZD | 04/14/2025 | 06/13/2025 |  | 05/06/2025 |  |  | 04/14/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025560113 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | Ventricular Fibrillation | Thrombosis/Thrombus | Retroperitoneal Hemorrhage | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella cp support and during the removal, the impella cp was removed from the peel away sheath and then the peel away sheath was removed. the peel away and impella cp were found to have clots and calcium on them. a wire was inserted into the peel away sheath, which was used to deploy per close sutures. the patient then went into ventricular fibrillation (vf) and was defibrillated. the patient was hypotensive, and their groin was blue and blown up like a balloon. the patient started hemorrhaging into the groin area. emergently, the right groin was opened and found a retroperitoneal bleed coming, likely from the iliac artery. the artery was exposed via cut down. the patient survived the procedure. | [5 empty] | The investigation of thrombus, vf/vt/af/at, and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of thrombus and vf/vt/af/at could not be determined. the root cause of the vascular damage - peripheral was most likely patient condition related as evidenced by clinical mention of calcified anatomy and patient having severe peripheral vascular disease. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
9682 | 1220648-2025-28129 | 21969999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21969999&pc=OZD | 05/01/2025 | 05/06/2025 |  | 05/06/2025 |  |  | 05/01/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 30 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Complainant reported the patient was on impella cp support and when peeling away the 14fr sheath, a large amount of bleeding occurred. a retention sheath was inserted but bleeding could not be stopped. the impella was removed and hemostasis was achieved by surgical sutures. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. h.4 device manufacture date is unknown. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9683 | 1220648-2025-28130 | 21970128 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21970128&pc=OZD | 04/16/2025 | 05/29/2025 |  | 05/06/2025 |  |  | 04/16/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1814768 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an upon arrival, the automated impella controller (aic) was found to have a cracked housing, bent handle, and open rlm cover. there was no patient harm. | [5 empty] | The investigation of automated impella controller (aic) - damage before therapy has been completed. there were no concerning alarm or evidence found in the logs to confirm the reported failure mode. this console was tested upon arriving. it was visually confirmed that the rear housing was broken, and the rlm was partially unmounted from the aic, and unable to mount it back appropriately. these two claims align with the reported failure mode. no issues were found with the aic handle. root cause: the root cause of the aic damage was most likely a use issue. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
9684 | 1220648-2025-28131 | 21970161 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21970161&pc=PYX | 04/07/2025 | 06/02/2025 |  | 05/06/2025 |  |  | 04/07/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC |  |  | US |  | 2025536967 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient on bipella support experienced hemolysis. bleeding and renal failure were also noted. the source of bleeding was unknown. the patient was scheduled for pump repositioning and renal therapy. the family decided to withdraw care, and the patient expired. per review by abiomed's medical safety officer, the impella rp flex device experienced positioning issues that were most likely associated with the patient outcome. | [5 empty] | A4 is unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral. vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis, positioning issues, renal failure, and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. hemolysis - reported hemolysis on (b)(6) 2025 with unknown details on troubleshooting or any resolution of the hemolysis. data logs show motor current spikes with speed deviations and step up in placement signal on (b)(6) 2025. the motor current (mc) remains slightly elevated and has fluctuations while at a constant p-level. the last mc spike occurs on (b)(6) 2025. the root cause of the hemolysis was most likely biomaterial based on mc spikes and speed deviations observed in the data logs prior to the clinical report of hemolysis. positioning issue - reported planning on repositioning on (b)(6) 2025 but unknown if repositioning occurred or any details on how/why. the root cause of the positioning issue was not determined since product was not returned and insufficient clinical information was provided. vascular damage - reported unspecified bleeding on (b)(6) 2025. the root cause of the vascular damage was not determined since product was not returned and insufficient clinical details provided. renal failure - reported the patient was anuric on (b)(6) 2025. the cause of the renal failure was not determined due to insufficient clinical details. b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28131. b2 selection of required intervention was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28131.
9685 | 1220648-2025-28132 | 21970171 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21970171&pc=OZD | 04/16/2025 | 05/06/2025 |  | 05/06/2025 |  |  | 04/16/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026616553 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. percutaneous coronary intervention was completed to the circumflex and left anterior descending arteries. thereafter, the peel-away sheath was removed and the repositioning sheath was inserted and a hematoma at the access site was noted. the physician assessed and determined that the impella cp device would need to be removed if the patient tolerated the weaning. weaning was performed and the patient was able to tolerate the wean. the impella cp device was explanted. percloses and 8 french angioseal was used. hemostasis was obtained. manual pressure and sterile dressing were applied, and the patient was sent to the unit in stable condition. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9686 | 1220648-2025-28133 | 21970353 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21970353&pc=OZD | 04/15/2025 | 06/06/2025 |  | 05/06/2025 |  |  | 04/15/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1356082 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella aic had a controller error on the screen. the aic was exchanged. no patient was involved in the event. | [5 empty] | Data analysis: console logs from the reported day of the event are consistent with the complaint. on the complaint date immediately after the initial console boot-up, there was a ¿controller error (defective optical sensor lamp) - alarm,¿ event #1039. this indicates that the reported failure occurred. the console was then shut down and restarted multiple times, but event #1039 happened each time. further investigation into the console logs revealed a similar situation with a previously reported case involving pump sn: (b)(6). that case displayed psnr alarm with multiple "algs suspended opm signal invalid" error messages as well as the controller error alarm. additionally, the lt log and rt log showed fluctuating contrast values, with the signal-to-noise ratio (snr) dropping to zero. see lt_log.png and rt_log.png. device analysis: this console was tested after arriving in fs. unable to reproduce the reported failure by power cycling the console. voltage from the pbm at connector (j3) was 4.9v and the optical bench lamp connector (p4) was 4.8v. obtained the analog output from the optical bench without any noise or distortion with pump and without pump connected as well as 5s parameter testing showed the light output passing the required specification and the lamp was visibly outputting light. see fringe_wop.jpg, fringe_wp.jpg and 5sparameter.jpg. root cause: the root cause for the controller error (defective optical sensor lamp) was most likely due to the defective optical bench lamp assembly. corrective actions: pia-2024-0000252 was opened in response to an upward complaint trend in aic-caused optical signal issues. failure of the optical bench¿s bulb was identified as the primary cause of the optical signal issues. capa-014132 was opened to identify a potential enhancement opportunity related to proactively changing the tungsten bulb in aic's optical bench. before returning the console to the customer, the following actions are instructed to perform: 1. replace the optical bench (0042-1012) due to defective. 2. perform section 23 of the preventative maintenance procedure (0042-7309). 3. perform a full functional check on the console and optical system (0042-7316). | The device was received for evaluation, upon completion of the investigation a supplemental report will be submitted.
9687 | 1220648-2025-28134 | 21970357 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21970357&pc=OZD | 04/16/2025 | 06/04/2025 |  | 05/06/2025 |  |  | 04/16/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026613612 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the insertion of an impella 5.5 in a patient with postcardiotomy cardiogenic shock was unsuccessful. the impella 5.5 was advanced through the graft into axillary artery. however, it was unable to advance through innominate artery. after multiple attempts, the decision was made to abort the impella 5.5. an attempt was made with an impella cp, which was successfully placed. there was no harm to the patient as a result of the event, and the patient was reported to be in stable condition. | [5 empty] | The pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist: section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the pump was returned and no damages or abnormalities were found. the root cause of the delivery issue was most likely patient condition related since vessel stenosis was reported.
9688 | 1220648-2025-28135 | 21970568 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21970568&pc=OZD | 04/16/2025 | 05/22/2025 |  | 05/06/2025 |  |  | 04/16/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025590046 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support, and the following day, the patient lost pulse in the right foot distal to the impella cp. an antegrade sheath was placed for bypass perfusion from left to right. doppler signals were present post sheath bypass placement. however, the decision was made to explant and replace the device preferably with an impella 5.5. there was no harm to the patient as a result of the event, and the patient was noted to be in stable condition. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
9689 | 1220648-2025-28136 | 21970797 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21970797&pc=OZD | 03/27/2025 | 05/06/2025 |  | 05/06/2025 |  |  | 04/15/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2026618537 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female | 51 KG | Non Hispanic | White | Injury | Yes | No | Increase in Pressure | [6 empty] | Thrombosis/Thrombus | Limb Pain | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient with acute systolic heart failure, ventricular free wall rupture contained, and postinfarction ventricular septal defect nstemi was implanted with an impella 5.5. during support there were concerns about a thrombus in the apex of the left ventricle. ptt levels were being kept lower due to the suspected blood clot in the apex. a purge pressure increase alarm occurred. the impella was explanted and debris was found in the pump. over the course of treatment, the patient had pain in right arm where they went for a thrombectomy. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
9690 | 1220648-2025-28137 | 21970822 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21970822&pc=OZD | 04/16/2025 | 06/06/2025 |  | 05/06/2025 |  |  | 04/16/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026640406 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Cardiac Arrest | Hemodynamic instability | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp for mechanical circulatory support (mcs) and experienced suboptimal positioning and pump flow. the patient expired the same day. an echocardiogram was completed and showed ejection fraction between 15-20% with moderate mitral regurgitation and severe tricuspid regurgitation. left and right heart catheterizations were ordered and completed and showed multivessel disease involving all major vessels including the left main. end-diastolic pressure was greater than 30. at completion of the procedure, the patient was stable. once the patient was moved to a bed, the patient arrested. the patient was, therefore, moved back to catheterization lab table and coded. code involved cardiopulmonary resuscitation, intubation, emergency medications, and defibrillation. once return of spontaneous circulation was achieved, the physician opted to place an impella cp for mcs. the impella cp device was successfully completed, and the patient was transferred to the cardiovascular intensive care unit on support. for continued evaluation. the following day the patient remained intubated and was noted to be following the nurses' commands. nephrology was consulted for continuous renal replacement (crrt) therapy. the next morning the patient's groin site began to bleed and there were waveform changes. manual pressure was applied to the groin site. regarding the waveform changes, an echocardiogram for placement assessment was completed and confirmed the impella cp device was "very" shallow. repositioning at bedside was completed. the device was initially advanced to approximately 4.6 centimeters and then pulled back to 3.3 centimeters. it was noted that the pigtail of the impella appeared to be being pulled by ventricle cord or papillary bed. the physician was unable to reposition without risk of the device being pulled through the aortic valve. the device was only able to achieve a performance level of p-5 without suction and getting approximately 2.8 liters of flow. the patient continued to deteriorate. there were several discussions made regarding the initiation of crrt and escalation of impella therapy. it was noted that the impella position and flow were to be suboptimal. however due to the family's decision, no measures were taken to rectify. the family made the decision to place the patient on do not resuscitate status. the patient would expire shortly thereafter. | [5 empty] | There is not enough evidence to exclude the impella device suboptimal position and flow as an associated factor to the overall outcome. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | Access site bleeding - minor: clinical states that mild bleeding was observed at the access site after the linen of the patients bed was changed with was easily controlled by manual pressure. pump was not returned for investigation. therefore, based of the clinical information provided, the root cause of the access site bleeding issue was most likely patient movement since the bleeding started right after the linen change. positioning issue: clinical states that after the bed linen change, the patient had access site ooze and the echo found pump to be advanced further into the ventricle initially and pulled back again due to pigtail appearing to be pulled by ventricle cord or pap bed. pump was not returned for investigation. therefore, as per the clinical information provided, the root cause of the positioning issue was most likely patient movement related to the movement of patient during the linen change. saturated flow: fm changed from low pump flow to saturated flow per review of the data logs. clinical states that the pump could only achieve p5 without any suction alarms after the positioning issue. pump was not returned for investigation. data logs confirm suction alarms when p-level is turned above p5. placement signal is seen trending downward. on (b)(6) 2025, after 12pm the pump was auto adjusted to p5 with saturated flows. similar saturated flows at p4 and p3 as well till pump explant. no motor current spike or purge issues. patient had positioning issues earlier on (b)(6) 2025 and pump was repositioned with risk involved but the outcome of repositioning is unknown. therefore it is not conclusive the true reason for suction alarms. based on the clinical information provided and data logs review, the root cause of the saturated flow was not determined since pump was not returned and clinical data and data log review were inconclusive.
9691 | 1220648-2025-28138 | 21971039 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21971039&pc=OZD | 04/16/2025 | 05/29/2025 |  | 05/06/2025 |  |  | 04/16/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025610807 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the insertion of an impella cp in a patient undergoing high-risk percutaneous coronary intervention was unsuccessful. the impella cp access with ultrasound and iliac looked appropriate although known to have significant calcification. dilated up without difficulty except a little resistance at the bifurcation but passed. the decision was made to use a v18 for added support. however, after multiple attempts the impella cp was unable to pass/advance and the implant was aborted. there was no harm to the patient as a result of the event. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, unable to pass through the vasculature due to calcification. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system, section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
9692 | 1220648-2025-28139 | 21971217 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21971217&pc=OZD | 04/15/2025 | 06/10/2025 |  | 05/06/2025 |  |  | 04/15/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025582843 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male | 78 KG | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hematoma | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 had a hematoma at the graft anastomosis. leakage was found and the surgical site was repaired. additionally, the patient had non-sustained ventricular tachycardia (vt). the patient survived. | [5 empty] | Access site bleeding: hematoma on surgical insertion site. the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical details provided. vt/vf: patient had some non-sustained vt this am. "cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. ¿timing of the arrythmia event in relation to impella support (during or post-implant) ¿electrocardiogram (ekg) recordings of the arrhythmia episodes. ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. ¿response to any antiarrhythmic treatments or interventions during the event. ¿any documented device-related issues or complications during impella support. ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. ¿review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined." | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9693 | 1220648-2025-28140 | 21971632 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21971632&pc=OZD | 04/15/2025 | 05/19/2025 |  | 05/06/2025 |  |  | 04/15/2025 | 05/06/2025 |  |  |  |  | 05/06/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639660 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female | 59 KG | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a left heart catheterization was performed on a patient and showed severe disease in the left circumflex artery and left main coronary artery/left anterior descending artery. the patient was hypotensive and was intubated. implantation of an impella cp was begun. the impella was advanced without issue however, oozing was noted at access site. a purse string suture was placed and lidocaine with bicarbonate was injected. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of access site bleeding was most likely use issue since the oozing was resolved by placing purse string suture. g1 reporting contact facility name was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28140.
9694 | 1220648-2025-28056 | 21959189 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21959189&pc=OZD | 04/23/2025 | 07/11/2025 |  | 05/05/2025 |  |  | 04/30/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026638290 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 89 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the purge pressure increased immediately after the impella cp was started on a patient. troubleshooting was done but was unsuccessful. the impella cp was exchanged. the patient was eventually weaned and explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of high purge pressure has been completed. the data logs show that the purge pressure was stable for about 3 hour and 20 minutes, then there was a rise in purge pressure that took about 6 minutes to trigger purge pressure high alarms, with no motor current spikes before or any changes in pump flows. the purge pressure remained high for about 30 minutes and did not resolve with cassette/bag changes. the pump was then explanted. the pump was returned and there was biomaterial seen around the impeller and in the purge gap. the pump was soaked in warm water with tergazyme and the biomaterial still remained. the pump was connected to a lab console and purged with a lab cassette. there were no kinks seen in the catheter. while connected, high purge pressure was reproduced. a window was cut in the catheter to expose the purge tubing and no blood was seen in the line. the catheter was cut and purge fluid expelled from the non motor side portion of the catheter and high purge pressure resolved indicating there was a blockage isolated to the motor. the root cause of the high purge pressure is most likely due to biomaterial based on data log and returned product analysis. d9 device returned to manufacturer date added.
9695 | 1220648-2025-28057 | 21959260 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21959260&pc=OZD | 02/19/2025 | 05/05/2025 |  | 05/05/2025 |  |  | 02/19/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026613586 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Detachment of Device or Device Component | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 was attempted to be implanted in a patient but difficulty with delivery was experienced. the pump was removed and damage to the catheter was found. the pump was discarded and a new pump was used. | [5 empty] | The impella pump was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist: section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
9696 | 1220648-2025-28059 | 21959333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21959333&pc=OZD | 04/07/2025 | 05/15/2025 |  | 05/05/2025 |  |  | 04/07/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025589741 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombocytopenia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella 5.5 and was positive for heparin induced thrombocytopenia. three days later, some suction alarms were noted and addressed. care was withdrawn two after the alarms and the patient expired. | [5 empty] | The investigation of low pump flow and thrombocytopenia has been completed. the cause of the low pump flow issue was not determined since product was not returned and sufficient clinical information was not provided. the cause of the thrombocytopenia issue was not determined due to insufficient clinical details. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9697 | 1220648-2025-28058 | 21959335 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21959335&pc=OZD | 04/08/2025 | 05/05/2025 |  | 05/05/2025 |  |  | 04/08/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026637151 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | Ventricular Fibrillation | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient (unknown race) in cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support (mcs) and while placing the impella, pre percutaneous coronary intervention, the patient went into ventricular fibrillation and was shocked one time. pci via thrombectomy to the left anterior descending artery (lad) and angioplasty with one stent to the mid lad was completed and the patient was sent to the unit on impella mcs. the following day after start of support the patient was showing signs of sepsis. position was noted as "good" under echocardiogram. surgery was consulted and will consider escalation if still requiring support in three days. it was noted the following day the patient remained in distributive septic and cardiogenic shock. the patient was noted to be improving however, and the patient was slowly weaned from the impella cp device, and all drips were off. the impella cp device was eventually explanted with a surviving outcome. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. however, product history review was completed, and the pump passed all post sterile inspection checks. the investigation of the complaint was completed and noted the following: ventricular fibrillation, cardiac arrythmia, can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: -patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease -timing of the arrythmia event in relation to impella support (during or post-implant) -electrocardiogram (ekg) recordings of the arrhythmia episodes -evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements -presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities -response to any antiarrhythmic treatments or interventions during the event -any documented device-related issues or complications during impella support -evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand -review of any concomitant medications that may influence cardiac electrophysiology -with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. therefore, to determine the true root cause of the ventricular fibrillation, the above-noted clinical information would need to be assessed in conjunction with evidence from the returned device. as the necessary clinical information was not provided, and the device was not returned, the root cause of the ventricular fibrillation was not determined. with regards to infection/sepsis, this can occur during or after impella support. when diagnosing the cause of an infection in a patient, the following clinical evidence would need to be considered holistically: -patient symptoms and medical history obtained by the treating clinician -physical examination findings -results of special investigations : laboratory test results & imaging studies -microbiological culture and sensitivity results -response to previous treatments and therapies -any relevant epidemiological information or exposure history -concomitant devices in use -preexisting patient condition such as underlying infection, remote septic sources such as indwelling lines/catheters or other bioprosthetic material in situ. -care of access site because this evidence has not been provided, the root cause of the infection, as well, cannot be determined. instructions for use for the related event are as follows: cp with smartassist system section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
9698 | 1220648-2025-28060 | 21959434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21959434&pc=OZD | 04/11/2025 | 05/05/2025 |  | 05/05/2025 |  |  | 04/11/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025527110 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 84 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that following explant of an impella cp the patient's blood pressure suddenly dropped and cardiopulmonary resuscitation was performed for over an hour. the suspected cause for the drop in blood pressure was a bleed in the iliac artery. the bleed could not be treated and the patient passed away. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
9699 | 1220648-2025-27977 | 21959473 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21959473&pc=OZD | 04/11/2025 | 05/05/2025 |  | 05/05/2025 |  |  | 04/11/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1095635 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automated impella controller (aic) would not recognize the purge disc being inserted. new tubing attempting with same result. staff requested the aic be sent back for investigation. there was no patient harm. | [5 empty] | The investigation has been completed. the impella pump was returned for investigation. imc logs from the complaint date revealed that the console issued purge disc not detected alarm. the console was examined after arriving in field service. a broken purge disc flag was identified. the purge disc flag was observed to be broken and was the root cause of the reported purge disc not detected alarm issue.
9700 | 1220648-2025-28062 | 21959605 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21959605&pc=OZD | 04/12/2025 | 05/05/2025 |  | 05/05/2025 |  |  | 04/12/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639781 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Gastrointestinal Hemorrhage | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the 62-year-old male patient on impella cp support developed gi bleed and gi was consulted. the plan was to scope the patient for source of bleeding. two units of blood were transfused. it was further reported that suction events were noted overnight and impella was pulled back with echo guidance in the morning. there was no-good long-term options for the patient as substance abuse and noncompliance issues complicate advanced therapy strategy. the patient¿s family made decision to stop support, and donate organs. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9701 | 1220648-2025-28063 | 21959659 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21959659&pc=OZD | 04/09/2025 | 09/23/2025 |  | 05/05/2025 |  |  | 04/12/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026629249 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the 49 year old male patient was on impella 5.5 support and during support, the pump was found pointed towards the mitral valve. the pump was repositioned. additionally, there was placement signal drift. the pump was reading 10-20pts lower than aline map. support continued. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. optical signal issue: data logs were reviewed lt log at time of reported issue showed stable optical parameters throughout log with no hard failures observed. no reoccurring positioning or optical alarms noted on returned logs. clinical details noted that pump position was confirmed via echocardiography but clinical team noted that aorta reading was 10 to 20 points off from patient's arterial line reading. the root cause of the optical signal issue could not be definitively determined as no hard failures of the optical sensor was observed during the case and no product was returned for evaluation. positioning issues: clinical details noted that pump was found to be out of position and towards mitral valve. pump was repositioned successfully. no details were provided on when and how pump came out of position. the root cause of the positioning issues could not be definitively determined as limited clinical details were provided. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-28063. b7 other relevant history, including preexisting medical conditions updated. d6b explant date updated. d10 concomitant medical products and therapy dates updated.
9702 | 1220648-2025-28064 | 21959767 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21959767&pc=OZD | 04/11/2025 | 05/05/2025 |  | 05/05/2025 |  |  | 04/11/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026618435 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that after implantation of an impella cp, a hematoma developed. heparin was withheld and manual pressure and femstop were applied. the hematoma resolved and the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9703 | 1220648-2025-28065 | 21959827 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21959827&pc=OZD | 04/12/2025 | 05/05/2025 |  | 05/05/2025 |  |  | 04/12/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653455 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the 67-year-old female patient was on impella 5.5 support and during support, the patient was having runs of ventricular tachycardia (vt). echo was performed and the pump appeared shallow. the doctor did not want to reposition at that time. the patient bridged to transplant. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
9704 | 1220648-2025-28066 | 21959884 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21959884&pc=OZD | 04/11/2025 | 05/23/2025 |  | 05/05/2025 |  |  | 04/11/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026617533 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 19 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient experienced multiple complications during impella cp support. the patient initially experienced a trachea bleed prompting a halt in the heparin drip. the following day, the pump stopped unexpectedly. the pump was restarted, but evidence of saturated flow ensued. the pump was explanted as a result of the failure. | [5 empty] | The investigation of vascular damage - peripheral vessel, pump stop, and saturated flow has been completed. the root cause of the vascular damage - peripheral vessel cannot be determined due to limited clinical details provided. saturated flow and pump stop. the data logs show that the pump was stable up until the 24th day where around 9 there was a spike in mc and pump stop. following there was an increase in pp, decrease in purge flow, and increase in pump flow. the motor current and flows remained elevated. around 10:30 there were spikes in motor current again and a pump stop. it is able to turn on two more times before it completely stops. due to lack of product returned, it cannot be confirmed the true root cause of this pump stop. the root cause of the saturated flows cannot be determined. g2 report source foreign was erroneously selected on the initial manufacturer device report number 1220648-2025-28066. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: using the automated impella controller with the impella catheter. ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock. section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
9705 | 1220648-2025-28067 | 21959999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21959999&pc=OZD | 04/13/2025 | 06/04/2025 |  | 05/05/2025 |  |  | 04/13/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577226 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that upon initiating impella cp support on a 76-year-old male patient, the positioning of the pump looked too deep. repositioning was attempted, but suction alarms were encountered and the catheter appeared to be kinked. the pump was subsequently replaced as a result of the noted issue. the patient was explanted in the procedural area. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of low pump flow, positioning issues, and product damage (cannula kink) has been completed. the impella device was not returned for evaluation. low pump flow: clinical states that the pump was getting low flows from case start. pump looked deep and was attempted reposition. patient at tortuous vessels. there was a catheter kink observed and pump was replaced. pump was not returned for investigation. data logs confirm low pump flow at case start along with suction alarms. flows of 3l/min observed at p9 and pump constantly attempting to auto-correct p-level base on the low volume. "impella in ventricle" and "impella in aorta" alarms observed followed by 0l/min pump flow at p9. therefore, as per data log review and clinical information provided, the root cause of the low pump flow issue was not determined since product was not returned for review. product damage (cannula kink): clinical states that log sheath was used to dilate patient due to tortuosity. pump was repositioned and reduced flows with suction alarms were observed. the catheter appeared to be kinked upon further inspection and pump was replaced. pump was not returned for investigation. as per the clinical information provided, the root cause of the product damage issue was a catheter kink most likely due to patients tortuous vessels. positioning issues: clinical states that after initial implant, the pump looked deep and was repositioned. pump was not returned for investigation. data logs confirm "impella in ventricle" and "impella in aorta" alarms. the pump was repositioned to get better flow but unsuccessful and a kink was observed on the catheter. pump was not returned for investigation. as per the clinical information provided, the root cause of the positioning issue was use related to improper technique off initial pump implant as the pump appeared to deep at initial placement.
9706 | 1220648-2025-28068 | 21960051 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21960051&pc=OZD | 04/13/2025 | 06/05/2025 |  | 05/05/2025 |  |  | 04/13/2025 | 05/05/2025 |  |  | 09/30/2025 |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024395704 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella 5.5 was scheduled for implant in a patient with heart failure. implantation could not be completed due to patient anatomy. the physician elected to implant an impella cp instead. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist: section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be patient condition because the pump was unable to pass past innominate artery and the delivery was aborted.
9707 | 1220648-2025-28069 | 21960132 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21960132&pc=OZD | 04/13/2025 | 06/05/2025 |  | 05/05/2025 |  |  | 04/13/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025607758 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient was on impella cp support and during support, the impella crossed over into aorta and was unable to push back across aortic valve. the impella was removed. the patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of positioning issues has been completed. the product was not returned for review. based on the clinical details the root cause of positioning issue was most likely use related since the pump was accidentally pulled back into aorta.
9708 | 1220648-2025-28070 | 21960143 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21960143&pc=OZD | 04/13/2025 | 05/05/2025 |  | 05/05/2025 |  |  | 04/13/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025610365 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that an impella cp was implanted in a patient who had a coronary artery bypass graft. the patient developed a thrombus in the left leg which was resolved by thrombectomy and fasciotomy. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
9709 | 1220648-2025-28071 | 21960299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21960299&pc=OZD | 04/13/2025 | 05/05/2025 |  | 05/05/2025 |  |  | 04/13/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026630033 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a 48 year old male patient on impella 5.5 support as an escalation of therapy. it was noted that they were monitoring neuro status and treating cardiogenic shock and sepsis. additionally the patient had some bleeding from mouth/nose so they decreased bleeding time. the patient also had confirmed occult blood in stool. the patient received 2 units blood transfusion and new labs were sent to reassess blood count while continuing to "monitor" platelets. the patient was not following commands and was taken for ct scan to evaluate for head bleed. team reports managing gastrointestinal bleed and oozing from all lines. it was noted that there were no new findings on head ct, upper endoscopy nonspecific, remains off heparin pending neuro surgery and gi recommendations. the patient remained on support. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
9710 | 1220648-2025-28072 | 21960396 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21960396&pc=OZD | 04/13/2025 | 06/11/2025 |  | 05/05/2025 |  |  | 04/13/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025584147 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp was being explanted when bleeding occurred after applying the closure device. hemostasis was achieved by placing a cover stent. the patient survived. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the access site bleeding issue was most likely use related, bleeding occurred after applying the closure device. hemostasis was achieved by placing a cover stent. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9711 | 1220648-2025-28073 | 21960529 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21960529&pc=OZD | 04/11/2025 | 06/03/2025 |  | 05/05/2025 |  |  | 04/11/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026632951 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Prime | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The patient presented with complaints of chest pain approximately 1 hour prior to arrival that was described as heaviness to his chest. he then stated he had pain down his left arm and up his left side of his neck. patient denied any previous coronary artery disease. upon arrival to the er, ekg was obtained and revealed possible anteroseptal infarct. there was concern about thoracic aortic aneurysm that was being monitored in wagoner so he was sent for a ct which came back negative. patient was transferred to the cath lab as st-segment elevation myocardial infarction and percutaneous coronary intervention (pci) was performed. during the pci, the physician performed thrombectomy, ballooned and stented the left anterior descending and the left circumflex artery. once stents were in, arteries kept closing off and there was no flow to the arteries. patient become hemodynamically unstable so the decision was made to place impella cp. impella cp access was gained and the device was then advanced over the wire. they then entered the body with the device but it was after the device was inserted that staff realized the cp had not been primed. they had not spiked the purge solution and the decision was made to pull the device from the body. staff was under the impression that once the device has entered the body, it could not be re-inserted so they opted to open a new impella cp. the patient expired on support of the replacement pump. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of priming issue has been completed. no product was returned. data log review shows that the pump was connected and the console prompted case start. logs also show an elevated placement signal before priming is started, which matches the clinical details that they did not prime the pump before inserting inside the patient. the root cause of the priming issue was most likely use-related since clinical details note that they did not prime or spike the purge solution prior to inserting the pump into the patient.
9712 | 1220648-2025-28074 | 21960572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21960572&pc=OZD | 04/13/2025 | 05/29/2025 |  | 05/05/2025 |  |  | 04/13/2025 | 05/05/2025 |  |  | 09/30/2025 |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024395704 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was attempted to be implanted with an impella 5.5 but was unsuccessful due to patient anatomy. the implant was aborted and the patient was stable. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, unable to successfully implant due to patient anatomy. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist: section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
9713 | 1220648-2025-28075 | 21960740 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21960740&pc=OZD | 04/13/2025 | 05/15/2025 |  | 05/05/2025 |  |  | 04/13/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026644294 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella cp and lost pedal pulses on the left side and exhibited deep vein thrombosis (dvt) in the left greater saphenous vein. the impella was explanted at the bedside. the patient survived. | [5 empty] | The investigation of limb ischemia and thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
9714 | 21960762 | 21960762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21960762&pc=OZD | 04/14/2025 | 04/29/2025 | 04/29/2025 | 05/05/2025 | 04/29/2025 | 05/05/2025 |  | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | 0042-0010-US |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Not Available | [6 empty] | User Facility report | Yes | Other Health Care Professional | [6 empty] | 76 YR |  | Male | 88 KG |  |  | Injury | Yes | Yes | No Audible Alarm | Therapeutic or Diagnostic Output Failure | [5 empty] | Low Blood Pressure/ Hypotension | Respiratory Failure | Gastrointestinal Hemorrhage | [10 empty] | Life Threatening | Hospitalization | Required Intervention | [22 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | The patient's (pt's) impella cp device alarmed suddenly with "impella stopped". recommendation of impella is a complete pump exchange. attending notified and decision to place on va ecmo [venoarterial extracorporeal membrane oxygenation] due to impella pump failure. vasopressors and inotropes started as pt's map [mean arterial pressure] dropped to the 30's. perfusion contacted, and vascular surgery contacted for emergent operating room following cannulation. impella device failure led to the pt's intubation, ecmo cannulation, and emergent or for removal of the impella pump. the patient required immediate ecmo placement with va ecmo, this has delayed his ability to get a lvad [left ventricular assist device] which patient was in the middle of being worked up for, while on va ecmo has now developed a gi [gastrointestinal] bleed (secondary to being on ecmo) and will likely have to be terminally extubated. no indication that it was anything other than a sudden failure of the device, no warning was found, no changes in the readings to indicate that this could happen soon. impella representatives are aware and also stated that it is a mechanical failure.
9715 | 1220648-2025-28076 | 21961014 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21961014&pc=OZD | 04/12/2025 | 05/15/2025 |  | 05/05/2025 |  |  | 04/12/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026633593 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  |  | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella cp to treat a stemi. the staff reported that they started the impella outside of the body. no ao/lv placement signals were present however, they had motor current, flows, and the purge was running. it was assumed they had damaged the optical sensor and the only way to remedy would be to remove the catheter and replace it with a new one. the pump was exchanged for a new impella cp pump. the patient survived. | [5 empty] | The investigation of optical signal issue has been completed. clinical mention pump was started before inserted into patient. the returned product was inspected and there were no physical abnormalities. the 5s were within specification and the pump passed the light continuity test. the pump was run in the lab and it functioned as expected with 5s in normal range. the data logs from the field show a pattern identified as a large air gap. the signal not reliable (snr) drops to zero, contrast and light decrease, gain and lamp increase. the placement signal not reliable alarm was confirmed and placement signal went out of range. the air gap was likely as a result of a user error of not pushing the patient cable all the way in. the cause of the optical signal failure was an air gap from between the automated impella controller (aic) and the patient cable plug as evidenced by the log pattern and failure to reproduce the issue on the returned product. g1 reporting contact facility name was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28076. | The device has been received for evaluation, upon completion of the investigation a supplemental report will be submitted.
9716 | 1220648-2025-28077 | 21961045 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21961045&pc=OZD | 04/13/2025 | 09/23/2025 |  | 05/05/2025 |  |  | 04/13/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026624683 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient was on impella 5.5 for support after removal of intra-aortic balloon pump (iabp). it was noted that they had a successful insertion of the impella 5.5 and removal of iabp. it was reported that throughout support, placement signal waveforms were decoupled and suction events were experienced. the physician repositioned the pump. however, waveforms were still uncoupled and suction alarms were present at p levels higher than p-4. the patient was given a fluid bolus. an echocardiogram was ordered to verify the placement. additionally, it was reported that repositioning attempt failed. the nurse also reported intermittent suction events and inquired about negative left ventricle diastolic of -60. the patient continues on support. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. all optical 5s parameters are stable without variations. placement signal (ps) spiking and variable during the case from 04/11 to 4/30. no issue with optical 5s parameters. the root cause of the lv waveform (optical signal) issue was likely patient condition related based on log and clinical review. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. corrections that were made from the initial manufacturer device report number 1220648-2025-28077. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The root cause of the positioning issue was not determined since product was not returned and due to insufficient clinical information. the root cause of the low pump flow issue was most likely patient condition related based on log and clinical review.
9717 | 1220648-2025-28078 | 21961226 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21961226&pc=OZD | 04/06/2025 | 05/27/2025 |  | 05/05/2025 |  |  | 04/06/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025600069 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump support ongoing for the patient in with cardiac history/coronary artery bypass graft( cabg) and new acute myocardial infarction. during support, bleeding was noted from the chest tubes and required heparin to be put on hold and blood products were to be given. low pump flows were also noted. several blood products and various sources of volume were administered. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of low pump flow and access site bleeding has been completed. the impella device was not returned for evaluation. data log was reviewed and there were no motor current spikes outside of p-level changes observed. no position issue alarmed. suction occurred on 05-apr at p-6 with no flow issue observed which match with clinical state: occasional suction alarms as patient moves and coughs, resolves without intervention. at last couple hours purge pressure low alarmed. no drop in flow observed. no sufficient details provided regarding access site bleeding; pump was removed at bedside and manual pressure held for hemostasis. low pump flow: per post procedure narrative: cp inserted without issues. several blood products and various sources of volume administered. flows starting to get better on impella. the cause of the low pump flow was most likely patient condition related since the issue resolved after volume was administered. the cause of the access site bleeding cannot be determined since complaint device not returned for analysis and insufficient clinical data provided. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9718 | 1220648-2025-28079 | 21961397 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21961397&pc=OZD | 04/06/2025 | 05/21/2025 |  | 05/05/2025 |  |  | 04/06/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025493706 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic |  | Death | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Tachycardia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump support was being initiated for a patient admitted in with acute myocardial infarction/cardiogenic shock. the team was unable to advance the pump and met much resistance. native anatomy precluded placement of the pump and the patient went into ventricular tachycardia arrest. catheterization lab team attempted resuscitation efforts but were unable to obtain return of spontaneous circulation. the patient expired afterwards. abiomed's medical safety officer reviewed the death and has stated there is not enough evidence to exclude the impella device used as an associated factor to the patient's outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. abiomed's medical safety officer reviewed the death, and has stated there is not enough evidence to exclude the impella device used as an associated factor to the patients outcome. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since the impella met resistance at aortic arch which made it unable to advance. b2 date of death was erroneously entered on the initial manufacturer device report number 1220648-2025-28079.
9719 | 1220648-2025-28080 | 21961401 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21961401&pc=OZD | 04/13/2025 | 05/05/2025 |  | 05/05/2025 |  |  | 04/13/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026645749 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced access site oozing that stopped with quick clot and dressing change. the patient outcome was noted to be stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9720 | 1220648-2025-28081 | 21961648 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21961648&pc=OZD | 04/13/2025 | 06/02/2025 |  | 05/05/2025 |  |  | 04/13/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026627479 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  |  |  | Injury | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | Stroke/CVA | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the 71-year-old female patient on impella 5.5 support had occlusion of distal intradural vertebral artery. no real deficit except for slow recall on some words. it was managed by anticoagulation therapy and closely monitored. additionally, the impella was repositioned as the patient was having some non-sustained ventricular tachycardia (vt). the team switched to dopamine. the patient also had some bigeminy and atrial fibrillation while on impella support. the patient was successfully weaned and explanted. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | The investigation of stroke/cva, positioning issues, and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
9721 | 1220648-2025-28082 | 21961753 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21961753&pc=OZD | 04/14/2025 | 05/28/2025 |  | 05/05/2025 |  |  | 04/14/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639490 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  | MW 3902560000-2025-8070 | Yes | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during impella cp support, it was reported that the pump stopped. multiple attempts were made to restart. however, they were unsuccessful. the patient was taken to the operating room for removal and the patients outcome is unknown. | [5 empty] | Included reference to medwatch (b)(4) in section h(10). | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | Medwatch report (b)(4) received following field were updated. e4 revised to reflect yes. h9 revised to reflect medwatch report number. medwatch report attached. on the initial manufacturer device report number 1220648-2025-28082. h10 narrative erroneously stated the device was not received, the device had been received at the time of the initial report.
9722 | 1220648-2025-28083 | 21961826 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21961826&pc=OZD | 04/14/2025 | 05/22/2025 |  | 05/05/2025 |  |  | 04/14/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1146949 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Fell | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the automated impella controller (aic) dropped to the floor by accident. this issue occurred during use. no patient harm was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | D6: unknown implant/explant dates the investigation of automated impella controller (aic) - damage during therapy has been completed. the cause of the issue was use related i.e. the crack in the housing was caused due to the console falling to the floor.
9723 | 1220648-2025-28084 | 21962015 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21962015&pc=OZD | 04/14/2025 | 05/05/2025 |  | 05/05/2025 |  |  | 04/14/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025611097 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Increase in Pressure | Pumping Stopped |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the pump experienced continuous suction. positioning was verified and volume was added, but steady flows could not be achieved. high purge pressure alarms then appeared and the pump subsequently stopped. the pump was explanted as a result of the reported issues. there was no patient harm. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
9724 | 1220648-2025-28085 | 21962220 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21962220&pc=OZD | 04/14/2025 | 06/13/2025 |  | 05/05/2025 |  |  | 04/14/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1323487 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during aic maintenance a controller error was displayed. troubleshooting did not resolve the issue. there was no patient involvement. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. data analysis: on the complaint date, april 14, 2025, during the service, the console displayed ¿controller error (defective optical sensor lamp) ¿ alarm¿ 1039. this confirms the reported failure mode. the log recorded abnormal optical parameters, particularly the light: 0.46v and the snr: 0.7. after a manual console reboot, event 1039 persisted. note: the pump (sn: (b)(6)) was the last commercial pump used with ic5762 from (b)(6) 2025, to (b)(6) 2025, and it did not have a controller error. later, the console was turned on and off without pump usage on (b)(6) 2025, and it also did not display a controller error. device analysis: this console was tested after arriving at the engineering lab and consistently reproduced the ¿controller error.¿ there was no analog output from the optical bench, both with and without the optical test cavity connected. the 5s parameters were abnormal with the optical test cavity. root cause: given that the persistent controller error occurred during the service due to a broken optical bench fiber, and there was no controller error at the client site, the root cause of the controller error was most likely technician error.
9725 | 1220648-2025-28086 | 21962223 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21962223&pc=OZD | 04/14/2025 | 06/04/2025 |  | 05/05/2025 |  |  | 04/14/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026624514 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult to Remove | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following a successful impella cp implant, some resistance was met when attempting to remove the introducer. the sheath needed to be split while still in the arteriotomy due to the patient's size. the repositioning sheath was subsequently placed. however, oozing was noted at the site. manual pressure was applied and forward tension sutures were reinforced to help manage the condition. vascular consult arrived at bedside and the bleeding subsided. support was continued with the implanted pump. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of access site bleeding and difficulty inserting/removing introducer has been completed. the impella device was not returned for evaluation. access site bleeding: clinical details noted that peel-away sheath need to be split while still in arteriotomy during removal. bleeding occurred and manual pressure was applied with forward sutures and angle matching and resolved once transferred to unit. the root cause of the access site bleeding - major was most likely a use related issue as the peel-away sheath was split while still in vasculature. difficulty inserting/removing introducer: clinical details noted that peel-away sheath had to be split while still in arteriotomy due to patient's anatomy. the root cause of the difficulty removing introducer was most likely patient condition related due to the patient's size per clinical details.
9726 | 1220648-2025-28087 | 21962241 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21962241&pc=OZD | 04/23/2025 | 07/16/2025 |  | 05/05/2025 |  |  | 04/23/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that an impella pump could not be advanced through the patient's vasculature during attempted delivery. following insertion, the pump catheter met strong resistance at the artificial blood vessel. because the patient was in an open chest state, the surgeon assisted the artificial blood vessel at the resistance part from the outside, but the pump catheter could not be advanced. the implant was subsequently aborted. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist for use during cardiogenic shock and high risk pci ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ a1- a.5 are unknown. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. e.1 name prefix/title, first name and last name are unknown. h.4 device manufacture date is unknown. | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since impella pump could not be advanced through the patient's vasculature following insertion as the pump catheter met strong resistance at the artificial blood vessel. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-28087 as it is unknown if the initial reporter also sent the report to the food and drug administration.
9727 | 1220648-2025-28088 | 21962265 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21962265&pc=OZD | 04/23/2025 | 05/05/2025 |  | 05/05/2025 |  |  | 04/28/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026631777 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that the patient on impella cp developed bleeding for the impella access site. hemolysis also appeared after the impella insertion. the impella cp was removed and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. impella cp with smart assist system: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9728 | 1220648-2025-28089 | 21962284 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21962284&pc=OZD | 04/14/2025 | 06/13/2025 |  | 05/05/2025 |  |  | 04/14/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1548488 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during preventive maintenance when the automated impella controller was powered "on", "system self check failure" was displayed. there was no patient involvement. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - system self-check error has been completed. based on the device and data analysis the cause of the issue was a defective component.
9729 | 1220648-2025-28090 | 21962391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21962391&pc=OZD | 04/14/2025 | 05/05/2025 |  | 05/05/2025 |  |  | 04/14/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1768506 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an automated impella controller (aic) had a battery failure. the aic was exchanged and there was no patient harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9730 | MW5169972 | 21962471 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21962471&pc=OZD | 04/28/2025 | 05/01/2025 |  | 05/05/2025 |  |  |  | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US |  |  |  | 1000100 |  | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Voluntary report | Yes | Other Health Care Professional | [6 empty] | 62 YR |  | Male | 85 KG | Non Hispanic | White | Injury | Yes | Yes | Material Integrity Problem | Therapeutic or Diagnostic Output Failure | Intermittent Energy Output |  |  |  |  | Failure of Implant | [12 empty] | Hospitalization | Life Threatening | Required Intervention | [26 empty] | Initial submission | [9 empty] | Patient experienced intermittent device failure of lmpella 5.5 assumed to be associated with power wire damage within the driveline. patient was emergently taken to operating room the next day for pump exchange.
9731 | 1220648-2025-28091 | 21962530 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21962530&pc=OZD | 04/14/2025 | 05/22/2025 |  | 05/05/2025 |  |  | 04/14/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1096323 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an console (aic) displayed an unexpected controller shutdown message. the device was removed as a result of the failure. there was no patient harm. | [5 empty] | The impella device was received for evaluation. upon completion of the investigation, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - shutdown or restart issue has been completed. the root cause of the unexpected shutdown could not be determined due to the inability to reproduce the issue.
9732 | 1220648-2025-28092 | 21962532 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21962532&pc=OZD | 04/14/2025 | 06/05/2025 |  | 05/05/2025 |  |  | 04/14/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026623808 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 90 YR |  | Female |  |  | White | Injury | Yes | Yes | Pumping Stopped | Infusion or Flow Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and as soon as the impella was started it went to auto mode and then immediately to p0 and a ¿impella stopped, motor current high¿ alarm was present. there was also an ¿ impella stopped, retrograde flow¿ alarm after. attempts were made to restart the pump and the alarms remained present. the pump was removed and flushed in heparinized saline. the doctor stated that they saw something come out in the basin. the pump was reinserted and the impella was restarted. the pump was flowing however, motor current was high, flows in max and min matched, left ventricle waveform was severely distorted and pump appeared saturated. the doctor stated at the end of the case all waveforms appeared normal. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella pumpe and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: impella cp with smart assist catheter insertion, section: axillary insertion of the impella cp with smart assist catheter, section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: using the automated impella controller with the impella catheter: ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock, section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The pump was returned for investigation. no physical damage was noted on the returned product. biomaterial was observed in the purge gap after soaking the pump in warm water. the pump was run in a bucket of water with saturated pump flow. purge flow was observed to be within the normal specification range during testing. the data log shows several pump stops during the start of the case, but the pump was able to restart during the second insertion with saturated pump flow. the doctor also confirmed that the pump stopped after initial placement but was able to restart after being flushed with heparinized saline. the doctor also found something coming out into the basin during the flushing of the pump. the pump was utilized for one hour and then explanted. the cause of the temporary pump stop issue was most likely biomaterial, based on data log review and returned product investigation. the cause of the saturated pump flow issue was most likely biomaterial, based on data log review and returned product investigation. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions. description of the location and characteristics of the thrombus (e.g., size, shape, consistency). relevant laboratory test results, such as coagulation profiles and platelet counts. imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). medications or treatments that the patient was receiving at the time of thrombus formation. surgical or procedural details if applicable (e.g, cardiac catheterization) any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. the pump was returned for investigation. no physical damage was noted on the returned product. biomaterial was observed in the purge gap after soaking the pump in warm water. the pump was run in a bucket of water with saturated pump flow. purge flow was observed to be within the normal specification range during testing. the data log shows several pump stops during the start of the case, but the pump was able to restart during the second insertion with saturated pump flow. the doctor also confirmed that the pump stopped after initial placement but was able to restart after being flushed with heparinized saline. the doctor also found something coming out into the basin during the flushing of the pump. the pump was utilized for one hour and then explanted. the cause of the temporary pump stop issue was most likely biomaterial, based on data log review and returned product investigation. the cause of the saturated pump flow issue was most likely biomaterial, based on data log review and returned product investigation. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions. description of the location and characteristics of the thrombus (e.g., size, shape, consistency). relevant laboratory test results, such as coagulation profiles and platelet counts. imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). medications or treatments that the patient was receiving at the time of thrombus formation. surgical or procedural details if applicable (e.g, cardiac catheterization). any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined.
9733 | 1220648-2025-28093 | 21962615 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21962615&pc=OZD | 04/14/2025 | 05/21/2025 |  | 05/05/2025 |  |  | 04/14/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025504133 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was on impella cp support and while on suport a small hematoma was noted at the access site and manual pressure was held. it was also reported that the p-level dropped from p-9 due to suspected hemolysis. the hematoma became severe and the decision was made to remove the impella cp and place the impella 5.5. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.7¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis and access site bleeding has been completed. the root cause of the access site bleeding - major was not determined as no product was returned and insufficient information related to failure was provided. clinical noted patient had hemolysis with red urine (on 4/14 at 10:50pm) during period of low volume. p-level dropped and volume administered for suspected hemolysis. next patient update (4/15 at 8:02am) revealed patient urine yellow with no further reports of hemolysis during support. the data logs do not show any motor current spikes, positioning issues or suction alarms. the root cause of the hemolysis was most likely patient condition related as evidenced by hemolysis occurring during low volume and resolving after volume administered with no abnormalities in the logs.
9734 | 1220648-2025-28094 | 21962848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21962848&pc=OZD | 04/14/2025 | 05/05/2025 |  | 05/05/2025 |  |  | 04/14/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026634496 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Malposition of Device | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that an impella cp was implanted in a patient with cardiogenic shock. during support, an echocardiogram was performed to confirm placement and it revealed that the pump was stuck under the mitral apparatus. attempts were made to reposition and torque the cannula multiple times. however, the pigtail consistently got stuck. the physician decided to leave the pump in place. the echocardiogram also revealed an encapsulated clot and heparin was started. the patient was successfully weaned and the pump was explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
9735 | 1220648-2025-28095 | 21963002 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21963002&pc=OZD | 04/14/2025 | 05/27/2025 |  | 05/05/2025 |  |  | 04/14/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025548290 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that while attempting to implant an impella cp and dilate the access point from an 8 french to a 10 french sheath, it was noted that the wire used was not stable and the dilator moved. the sheath could not be readvanced and resistance was felt when the device was pushed. based on the results of an angiogram, the physician believed the sheath went extra-luminal. the implant was subsequently aborted and the site was closed. the patient's outcome is unknown. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since the placement of pump was aborted due to patient's anatomy. d6a and d6b were erroneously entered on the initial manufacturer device report number 1220648-2025-28095. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
9736 | 1220648-2025-28096 | 21963003 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21963003&pc=OZD | 04/13/2025 | 05/29/2025 |  | 05/05/2025 |  |  | 04/13/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025568673 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support it was reported that the impella alarmed ¿placement signal low¿ and ¿impella in ventricle¿ multiple times. echo was performed, impella was repositioned with echo guidance. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation of positioning issues has been completed. the data logs from the field were not recovered. the pump and purge cassette were returned for investigation. investigation showed no abnormalities and nominal 5s parameters. the pump was soaked in warm water and tergezyme and run for 10 minutes at nominal parameters. the cause of the positioning issue was most likely patient movement due to the bedside nurse report of patient bending their leg 'several' times before placement signal low alarms occurred. after the physician repositioning, issue was not reported to have occurred form the remainder of the case. | The impella was received for evaluation. upon completion of the investigation, a supplemental mdr will be filed.
9737 | 1220648-2025-28097 | 21963426 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21963426&pc=OZD | 04/13/2025 | 05/05/2025 |  | 05/05/2025 |  |  | 04/14/2025 | 05/05/2025 |  |  |  |  | 05/05/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636378 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the 56 year old male patient on impella cp support, right lower extremity began to exhibit early signs of limb ischemia, so the care team decided to wean and explant impella. after cp was removed fasciotomy was performed. patient remained on va ecmo. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9738 | 1220648-2025-28019 | 21948702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21948702&pc=OZD | 04/08/2025 | 05/29/2025 |  | 05/02/2025 |  |  | 04/08/2025 | 05/02/2025 |  |  |  |  | 05/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026642585 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaint reported that a patient was implanted with an impella cp and experienced suction alarms due to low pulsatility. albumin was given and the impella was set at p4 to correct the alarm. it was further reported that the impella was pulled back under echocardiogram and was corrected and set. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of low pump flow and positioning issues has been completed. the impella device was not returned for evaluation. low pump flow: data log was reviewed and few suction alarms were seen on (b)(6) 2025. p-level was reduced to p-4 and suction alarms resolved. low flows were seen during the whole case with slight declining ps. rt log showing suction # 14 alarm(low pulsatility) on (b)(6) 2025. suction alarm triggers on and dpress. is lost during this time(reads zero) and when suction alarm is off, dpress. is again reading values. per clinical patient was given some fluid bolus and then p-level was reduced to p-4 to resolve suction alarms. the root cause of the low pump flow issue was most likely patient condition related due to extracorporeal membrane oxygenation (ecmo) used during support and based on log, clinical review. positioning issues: the root cause of the positioning issue was most likely use related since impella was pulled by doctor during support. a5 race and ethnicity was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28019.
9739 | 1220648-2025-28020 | 21948809 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21948809&pc=OZD | 04/11/2025 | 09/29/2025 |  | 05/02/2025 |  |  | 04/11/2025 | 05/02/2025 |  |  |  |  | 05/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026631943 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | No | Other | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Material Split, Cut or Torn | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Sepsis | Cardiogenic Shock | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient was implanted with an impella 5.5 to treat heart failure/cardiogenic shock and experienced developed a hematoma. the patient became hypotensive requiring pressors and was taken back to the operating room to stop the bleeeding. the hematoma was evacuated, but active bleed from anastomosis was not found. the wound was stabilized and vac was placed. four units of packed red blood cells and one unit of platelets were transfused. pressor support was weaned down with plans to turn off vasactive medications. the patient survived. | Additional information received that on 27-may-2025, the impella 5.5 was explanted due to anti-contamination sleeve tear and replaced with new impella 5.5. additionally, the patient continues to have septic picture with elevated white blood cells (wbc), fever. decision made to change out impella 5.5 for new device. patient continues to demonstrate mixed picture of cardiogenic shock and septic shock. |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the following fields were updated based on the additional information received regarding product malfunction (sleeve tear) and septic shock. should the device or any new information be received, a supplemental manufacturer device report will be filed. b1 revised to reflect both adverse event and product problem. b5 revised to reflect the additional information. d6b explant date added. h6 health effect - clinical code, health effect - impact code, medical device problem code, and component codes were added to reflect the additional information received. | The investigation of access site bleeding, infection/sepsis, and product damage (sterile sleeve damage) has been completed. pump not returned. access site bleeding: patient developed a hematoma within an hour post op. the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical details provided. infection/sepsis: based on the clinical review, it was found that the patient had developed an infection during impella support. the root cause of the infection/sepsis was unable to be determined due to insufficient clinical details. product damage (sterile sleeve damage): the pump was explanted and replaced with another pump because of sterile sleeve damage. the root cause of the product damage (sterile sleeve damage) was not determined as no product was returned for investigation.
9740 | 1220648-2025-28021 | 21948889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21948889&pc=OZD | 04/11/2025 | 05/02/2025 |  | 05/02/2025 |  |  | 04/11/2025 | 05/02/2025 |  |  |  |  | 05/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025578948 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 37 YR |  | Female |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [11 empty] | Death | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured bleeding with a "possible" relationship to the impella device used: ¿patient arrived at facility intubated and was found to be in biventricular failure. worsening septic shock and acidosis prompted initiation of v-a ecmo and impella placement. patient coded on table before initiation of ecmo or impella. patient was given 2 units of rbc¿s on (b)(6) 2025 post impella and ecmo placement as indicated by her unstable angina, cardiogenic shock, or a h&h below the 24 and 8 thresholds. the patient¿s hemoglobin dropped to 6.9 g/dl but improved to 9.0 g/dl after the administration of rbc¿s not requiring any additional blood products at that time.¿ the patient recovered with sequelae. two days post impella cp implant and extracorporeal membrane oxygenation (ecmo) placement indicated active bleeding and 750 cubic centimeter¿s of blood loss were reported. the patient's hemoglobin dropped to 6.5 g/dl, probably due to blood loss. after administration of red blood cell¿s, the patient's hemoglobin was found to be 10.2 g/dl and her blood volume was stable enough to transport the patient to an outlying facility. loqi reported renal failure. the relationship is noted as "possibly" related. loqi states: patient developed very severe metabolic acidosis and with the creatinine rising, the decision was made to place the patient on crrt. with low potassium the patient started on crrt with phosphate containing solutions. potassium, co2, and magnesium were all maintaining well on crrt. additionally, the patient prognosis was poor, and care was withdrawn. the patient expired. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9741 | 1220648-2025-28022 | 21948951 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21948951&pc=OZD | 04/11/2025 | 05/29/2025 |  | 05/02/2025 |  |  | 04/11/2025 | 05/02/2025 |  |  |  |  | 05/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025582399 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Detachment of Device or Device Component | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Ectopic Heartbeat | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella 5.5 and experienced ventricular tachycardia. the impella 5.5 was repositioned but the patient still had occasional ectopy. additionally, the sterile sleeve was pulled back when the impella 5.5 was repositioned and left back. there was a tear in the sterile locking sleeve. the site was covered with occlusive sterile dressing. it was further reported that on day seven of impella support, bleeding developed from the axillary cutdown pocket and impella site. heparin was withheld. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of access site bleeding, vf/vt/af/at, and product damage (sleeve damage) has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
9742 | 1220648-2025-28023 | 21949000 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21949000&pc=OZD | 04/11/2025 | 05/02/2025 |  | 05/02/2025 |  |  | 04/11/2025 | 05/02/2025 |  |  |  |  | 05/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026644893 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient implanted with an impella cp exhibited access site bleeding. the pump was explanted and the patient survived. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9743 | 1220648-2025-28024 | 21949073 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21949073&pc=OZD | 04/11/2025 | 07/10/2025 |  | 05/02/2025 |  |  | 04/11/2025 | 05/02/2025 |  |  |  |  | 05/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026624820 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that despite a successful implantation, the impella cp pump was unable to advance through the patient's vasculature during attempted delivery. the implant was aborted due to the noted issue. the patient survived. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system: section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issue has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be patient condition because the pump was unable to pass through vasculature due to vascular stenosis and the insufficient diameter of the vessel within the groin.
9744 | 1220648-2025-28025 | 21949204 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21949204&pc=OZD | 05/15/2023 | 06/11/2025 |  | 05/02/2025 |  |  | 04/11/2025 | 05/02/2025 |  |  | 11/30/2024 |  | 05/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024315166 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 81 YR |  | Female |  | Non Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Tamponade | Pericardial Effusion | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported via abiomed¿s long-term outcome and quality indicator (loqi) impella registry that an impella cp pump was implanted to support a patient admitted in with acute myocardial infarction/cardiogenic shock. the patient endured cardiac tamponade with a "possible" relationship to the impella device used. after impella placement, the patient later had a pericardiocentesis due to small-moderate pericardial effusion with cardiac tamponade. the patient was transferred to the intensive care unit after the procedure. after five hours of support, care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of ventricle perforation has been completed. based on insufficient clinical details, product not being returned, and unclear evidence from data logs whether or not the position in ventricle alarms occurred at the time of the complication, the root cause of the ventricle perforation could not be determined. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-28025.
9745 | 1220648-2025-28026 | 21949260 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21949260&pc=OZD | 04/12/2025 | 06/12/2025 |  | 05/02/2025 |  |  | 04/12/2025 | 05/02/2025 |  |  |  |  | 05/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026627478 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility a patient was on an impella cp for support during cardiogenic shock. it was reported that due to the impella cp purge reservoir leaking, the impella was removed. the patient remained stable with extracorporeal membrane oxygenation. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: cleaning: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." | The investigation of purge leak ¿ pump has been completed. neither data logs nor product were available for investigation. though product wasn't returned for investigation, it was reported that the pump was leaking from the purge sidearm. therefore, the root cause of the purge leak was determined to most likely be a damaged sidearm.
9746 | 1220648-2025-28027 | 21949309 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21949309&pc=PYX | 04/12/2025 | 10/03/2025 |  | 05/02/2025 |  |  | 04/12/2025 | 05/02/2025 |  |  |  |  | 05/02/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025599042 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Cardiovascular Dysfunction/Insufficiency | [11 empty] | Required Intervention | Death |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The user facility reported a patient with an impella rp flex required cardiac pacing. during support, the patient began having increased suction alarms and the central venous pressure (cvp) was much lower than it had been the previous day. blood loss with line placement and continuous renal replacement initiation could have played a role. given the patients hemodynamics, they elected to change impella flow to p6. however, the family decided to withdraw care and the patient expired. that there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | During impella rp flex support, the patient experienced a substantial bleed during a line placement for dialysis. the patient's systemic heparin was discontinued and the patient was transfused with a unit of packed red blood cells. the following day, there were reported suction alarms and the p-level was lowered from 8 to 6 and the patient was transfused with an additional unit of packed red blood cells and 2 units of fresh frozen plasma. | The complainant also reported that there were alarms for placement signal not reliable. the pump was confirmed to be in good position, and the motor control was still pulsatile. there was no disruption of patient support due to the placement signal issue. |  |  |  | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of low pump flow has been completed. based on the investigation, low pump flow was removed as a failure mode because the logs do not show any abnormal change in motor current and low pump flow occurrence for matching p-level outside when flow calculation was disabled as a result of the placement signal issue. insufficient product was returned for evaluation. the returned pump was severed at the pump plug so only the cannula with catheter was inspected. biomaterial was observed on the impeller. | B5 revised to include additional information that was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28027. h6 revised to reflect the additional information that was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28027. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for placement signal issue. the data logs confirm placement signal not reliable alarm and confirm suspected sensor drift. placement signal was drifting down while pump flows drifted up before placement signal went out of range and flow calculation was lost. the pattern observed has been characterized as a sensor drift. the root cause of the placement signal issue was most likely due to sensor drift as evidenced by the pattern observed in the logs. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include placement signal issue description. h6 updated to include placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-28027. b7 other relevant history, including preexisting medical conditions updated. d4 expiration date and unique identifier (udi) # corrected. d10 concomitant medical products and therapy dates updated. h4 device manufacture date corrected..
9747 | 1220648-2025-28028 | 21949374 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21949374&pc=OZD | 04/04/2025 | 05/21/2025 |  | 05/02/2025 |  |  | 04/04/2025 | 05/02/2025 |  |  |  |  | 05/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026629274 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient experienced persistent bleeding and suction issues during impella 5.5 support. roughly two days into support, a graft bleed was noted, prompting repeated washouts and explorations to manage the condition. heparin was eventually paused in response to the ongoing bleeds. they patient also later became symptomatic from a possible gastrointestinal bleed. transfusions were performed throughout support in response to the bleeds. coinciding with the bleeds, the pump experienced persistent suction events. it was noted that flows higher than p-4 would result in immediate suction or ectopy. repositioning was attempted, but was ultimately unsuccessful and the device remained in suboptimal positioning for support. the patient passed away after 27 days of support. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: warnings & cautions: cautions: "physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta." caution: "physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance." section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position." "to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality." "to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance." | The investigation of vascular damage - peripheral vessel, positioning issues, and low pump flow has been completed. the impella device was not returned for evaluation. data logs were reviewed and persistent suction alarms were present throughout case. clinical details note that patient was having bleeding issues and required multiple blood products throughout support. pump was unable to flow higher than p-4 without suction alarms. the root cause of the low pump flow was most likely patient condition related as patient required multiple suction alarms throughout support and patient needed many units of blood and required volume resuscitation. the root cause of the vascular damage - peripheral vessel could not be definitively determined as limited clinical details were provided of bleeding on support. the root cause of the positioning issues could not be definitively determined as limited clinical details were provided and no positioning issues were observed in data logs. b2 required intervention selection was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28028.
9748 | 1220648-2025-28029 | 21949377 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21949377&pc=OZD | 04/12/2025 | 05/30/2025 |  | 05/02/2025 |  |  | 04/12/2025 | 05/02/2025 |  |  |  |  | 05/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026645746 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | VENOARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient on impella cp support and venoarterial extracorporeal membrane oxygenation (va ecmo) had cut down incisions bilaterally with jp drains in place. is it unknown if the bleeding occurred from around ecmo sites or impella site, or both. the patient required volume resuscitation with unknown units of blood products. a second venous drainage cannula was also placed for the va ecmo circuit. bleeding appeared to have been resolved with the cut downs and extra cannula. it was further reported that on day four of impella support an echocardiogram was done and the impella cp was repositioned in the left ventricle. the patient was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and positioning issues has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
9749 | 1220648-2025-28030 | 21949461 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21949461&pc=OZD | 04/12/2025 | 05/13/2025 |  | 05/02/2025 |  |  | 04/12/2025 | 05/02/2025 |  |  |  |  | 05/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026620924 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella cp support and during support, the patient had a hematoma and bleeding from access site despite manual pressure. the doctor removed the impella despite recommendations. the patient survived the procedure. | [5 empty] | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. the device passes all post sterile inspection checks. no device was received for evaluation. the root cause of access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9750 | 1220648-2025-28031 | 21949484 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21949484&pc=OZD | 04/11/2025 | 05/28/2025 |  | 05/02/2025 |  |  | 04/11/2025 | 05/02/2025 |  |  |  |  | 05/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026613394 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was noted that during access the patient vomited and moved extremities causing bleeding at the impella sheath site. the physician was in the process of removing sheath when the patient vomited and moved. the patient was bleeding a decent amount. the patient as aspirated and vagaled down dropping the patient's pressure and heart rate and required emergent intubation. manual pressure was held and forward tension suture was placed. two units of blood were administered. bleeding was controlled. the patient weaned from impella and was noted to be recovering well. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the patient vomited and moved extremities causing bleeding at the impella sheath site. the root cause of the access site bleeding issue was most likely patient movement. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9751 | 1220648-2025-28032 | 21949662 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21949662&pc=OZD | 04/11/2025 | 05/21/2025 |  | 05/02/2025 |  |  | 04/11/2025 | 05/02/2025 |  |  |  |  | 05/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025609918 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was noted that the patient experienced bleeding and leg ischemia. this procedure had challenges due to perclosing not working. 3 perclose were eventually successfully deployed. the pump was explanted in the procedural area. surgical intervention was done the next day by vascular: surgical endarterectomy for a left fem artery dissection. after the procedure was finished the flow resumed in the leg. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation of limb ischemia and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. h6 health effect - clinical code 3160, health effect - impact code 4624 and component code 3038 were inadvertently omitted on the initial manufacturer device report number 1220648-2025-28032.
9752 | 1220648-2025-28033 | 21949839 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21949839&pc=OZD | 04/11/2025 | 05/20/2025 |  | 05/02/2025 |  |  | 04/11/2025 | 05/02/2025 |  |  |  |  | 05/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026645882 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient developed a hematoma at their groin site during impella cp support. the heparin drip was put on hold to manage the condition. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system: section: general patient care considerations: "assess access site for bleeding and hematoma." section: entering cardiac output: use of the repositioning sheath and peel-away introducer: "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site." section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of access site bleeding was most likely patient movement since hematoma appeared at groin after transport and after the patient moved onto her side when transferred from stretcher. h6 health effect - clinical code 1884 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28033.
9753 | 1220648-2025-28034 | 21949889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21949889&pc=OZD | 04/12/2025 | 09/29/2025 |  | 05/02/2025 |  |  | 04/12/2025 | 05/02/2025 |  |  |  |  | 05/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026637027 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Hispanic |  | Death | Yes | Yes | Defective Device | Device Sensing Problem | [5 empty] | Tachycardia | [12 empty] | Required Intervention | Death |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient was on impella cp support for cardiogenic shock. during the placement with a 14 french sheath, the sheath kinked. the team replaced the sheath and placed an impella cp but during the first hour of support, the patient suffered ventricular tachycardia and had optical signal issues. the patient and family made the decision to not escalate care and the impella cp and lucas device were explanted and the patient expired. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of introducer damage ¿ mechanical and vf/vt/af/at has been completed. the impella device was not returned for evaluation. during implant procedure, attempted to exchange sheath for long impella sheath in rfa however impella sheath kinked and unable to pass a wire. short sheath then used successfully. the pump and introducer were not returned. the cause of the issue was a introducer sheath kink. the cause of the sheath kink was unknown. as the necessary information was not provided, the root cause of the vt/vf was not determined | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. additional clinical information supplied by the representative indicated the patients actual blood pressure may have been responsible for the placement signal (ps) low alarms. the pump was not returned. data log review showed placement signal low alarms with no 5s parameter abnormalities related to sensor failure. the cause of the optical signal issue was most likely patient condition related due to the patients condition being unstable and deteriorating with non-nominal blood pressure. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-28034. d10 concomitant medical products and therapy dates updated.
9754 | 1220648-2025-28035 | 21949975 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21949975&pc=OZD | 04/11/2025 | 05/02/2025 |  | 05/02/2025 |  |  | 04/11/2025 | 05/02/2025 |  |  | 10/31/2025 |  | 05/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2024418849 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Unspecified Infection | Thrombocytopenia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient implanted with an impella 5.5. the patient exhibited a hematoma with a "probable" relationship to the impella device used and had a hematoma evacuation. additionally, the patient exhibited an insertion site infection with a "probable" relationship to the impella device used. the patient had impella washout with debridement of skin, tissue, muscle and fascia. cultures of the insertion site came back positive for proteus and pseudomonas. wound vac was placed, and the patient was treated with antibiotics. it was further reported by loqi that the patient endured thrombocytopenia with a "possible" relationship to the impella device used. the patient received four units of platelets. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 for use during cardiogenic shock, section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella 5.5 with smartassist for use during cardiogenic shock states the following: section: table 3.2 impella catheter components, ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings, ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
9755 | 1220648-2025-28036 | 21950019 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21950019&pc=OZD | 04/12/2025 | 05/29/2025 |  | 05/02/2025 |  |  | 04/12/2025 | 05/02/2025 |  |  |  |  | 05/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026629218 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  |  | Black Or African American | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Hematoma | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was on day four of impella 5.5 support and had suction alarms with ventricular tachycardia (vt). on day five of impella 5.5 support, the patient was taken to operating room for extracorporeal membrane oxygenation decannulation. the decannulation was successful. as the team was about to get the patient off the table, the patient was rolled to be clean. as the patient was rolled back, the impella 5.5 got pulled out of the axillary graft. the patient began to exsanguinate out the axillary graft. at this time, the patient coded and required two rounds of cardiopulmonary resuscitation. the decision was made to emergently recannulate via the groins and start rapid blood product administration. the impella 5.5 was reinserted into the left ventricle which got pulled back but not across the aortic valve. transesophageal echocardiogram confirmed the device was in proper position in the ventricle. hematoma was noted. systemic anticoagulation was held. the next day, low flow was noted on impella 5.5 and suction alarms were intermittent on p4. pocus showed the left ventricle was notably empty. crystalloid bolus was given and flow was improved. the impella was pulled back one centimeter. the patient was successfully weaned. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of positioning issues, access site bleeding, low pump flow, and vf/vt/af/at has been completed. the impella device was not returned for evaluation. data log reviewed: no motor current spike outside of p-level changes observed. many suctions occurred during the support. many impella unknown position alarms seen. no placement signal issue observed. drop in flow observed on 14- apr as reported then back to range and flow improved. access site bleeding - major: the cause of the access site bleeding - major most likely was patient movement since 5.5 impella got pulled out of the axillary graft while the patient rolled back for cleaning. positioning issues: the cause of the positioning issues most likely was patient movement since 5.5 impella got pulled out of the axillary graft while the patient rolled back for cleaning. low pump flow: the cause of the low pump flow most likely was patient condition related since crystalloid bolus given and flow improved. vt/vf: as the necessary information was not provided, the root cause of the vt/vf was not determined. g4 pma/510(k)number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28036.
9756 | 1220648-2025-28037 | 21950057 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21950057&pc=OZD | 04/04/2025 | 05/02/2025 |  | 05/02/2025 |  |  | 04/04/2025 | 05/02/2025 |  |  |  |  | 05/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639368 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a physician elected to place impella cp for cardiogenic shock. it was noted that as impella was being opened the patient began to brady down, atropine was given, however the patient did not respond. the cardiopulmonary resuscitation (cpr) was started. return of spontaneous circulation (rosc) was achieved but in ventricular tachycardia (vt). the patient was defibrillated once; cpr was stopped temporarily for pulse check when impella was on the wire at the peel-away sheath. vt again - encouraged team to hold defib at this time, impella loaded over wire in left ventricle (lv) and on in auto. impella flows ramping up, patient was defibrillated once and rosc was achieved with accelerated junctional rhythm. the impella was locked down with prolapsed wire to verify position as it was originally deep in the lv with cpr. position optimized after groin dressing. additionally, it was noted that following suturing of the blue suture hub, a final echocardiogram was being completed to optimize positioning as some arrhythmia was noted on the monitor. the patient vf arrested requiring a single defibrillation. following the defibrillation, there was a shadowing noted on the aortic leaflet on transesophageal echocardiography. lv appeared free of clot and 2 anesthesiologists stated there was no evidence of clot on induction or throughout any of the procedure. 2nd anesthesiologist had shown me the mitral valve prior with moderate mitral regurgitation (mr) noted and an odd appearance around the mitral leaflet that he was originally concerned the mitral valve had a tear in the leaflet. however, following the echocardiogram he and the ic / ctx stated that was likely due to the low flow state, high pa pressures and mr. no clot was observed elected to proceed with the implant. following the implant - dr's both stated they are not sure if this is a clot-like substance or not. they elected to resume heparin at a fixed rate in cardiovascular intensive care unit and will follow on echocardiogram daily. ultimately, the patient was escalated to a 5.5. the patient's demise is associated with the impella 5.5 device accordingly under complaint file (b)(4). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. abiomed's medical safety officer reviewed the patient's outcome and it was determined that the patient demise would be reported under the impella 5.5 (serial number (b)(6)). a manufacture device report will be sent to reflect the 5.5 and the death. as noted in the impella instructions for use: impella cp with smart assist system: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
9757 | 1220648-2025-28038 | 21950158 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21950158&pc=OZD | 06/04/2024 | 05/02/2025 |  | 05/02/2025 |  |  | 04/11/2025 | 05/02/2025 |  |  | 12/31/2025 |  | 05/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025455180 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 85 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported via abiomed¿s long-term outcome and quality indicator (loqi) impella registry that a patient implanted with an impella cp endured gi bleed with a "possible" relationship to the impella cp device used. "patient presented with gi bleed and acute blood loss anemia for which she received blood transfusion. patient had ct enterography which was negative for any acute bleed. patient also had an esophagogastroduodenoscopy and small bowel capsule which were negative as well. patient received three units of red blood cells." the device was explanted and the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
9758 | 1220648-2025-28039 | 21950332 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21950332&pc=OZD | 04/04/2025 | 05/02/2025 |  | 05/02/2025 |  |  | 04/04/2025 | 05/02/2025 |  |  |  |  | 05/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026615990 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Positioning Failure | [6 empty] | Ventricular Fibrillation | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella cp and was being escalted to an impella 5.5 and during the implant, the pump was being repositioned and the patient went into ventricular fibrillation (vf). the patient was defibrillated. additionally, the patient had recurrent persistent ventricular tachycardia. impella 5.5 was repositioned twice. the impella was touching what appears to be the septum per team. the pump was repositioned again facing towards mitral valve. care was withdrawn and the patient expired. the patient received suboptimal hemodynamic support and contributing to arrhythmic events. there is not enough evidence to exclude impella 5.5 as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. abiomed's medical safety officer reviewed the patients outcome and there is not enough evidence to exclude impella 5.5 as an associated factor in the patient's outcome. the impella cp for the same patient is reported on manufacture report number 1220648-2025-28037.
9759 | 1220648-2025-28040 | 21950340 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21950340&pc=OZD | 04/13/2025 | 05/29/2025 |  | 05/02/2025 |  |  | 04/13/2025 | 05/02/2025 |  |  |  |  | 05/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026622084 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Pumping Stopped | [5 empty] | Arrhythmia | Bradycardia | Fever | Low Blood Pressure/ Hypotension | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced bradycardia and was administered dopamine. echocardiogram verified placement was deep. the pump was pulled back to around 3.6 centimeters in the left ventricle. additionally, the patient went into atrial fibrillation and became hypotensive. amiodarone was started and albumin was administered which improved pressure. furthermore, the patient experienced a temperature of 103 degrees fahrenheit so antibiotics was started. it was noted that the plan was to remove the pump. however, during the weaning process the pump stopped. they were able to restart the pump and continued to wean and remove the device. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: entering cardiac output ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation of temporary pump stop, positioning issues, vf/vt/af/at, and infections/sepsis has been completed. the pump was returned for investigation. no physical damage was observed on the returned product. further investigation revealed a large purge gap, and some biomaterial was observed on the impeller. the pump was able to restart during the testing in a bucket of water. the pump was in for 15 minutes with a few pump stops ¿ motor current high alarm ¿ and observed saturated flow. the data log shows the pump ran for 9 days, and a pump stop was observed at the end of the case with a motor current spike and an impella stop-motor current high alarm. the pump was able to restart with observed saturated pump flow. the cause of the temporary pump stop was most likely bearing wear due to pump use beyond design life of 8 days. the cause of the positioning issue was most likely use related since it occurred after initial placement and securing the pump. as the necessary information was not provided, the root cause of vt/vf and infection/sepsis was not determined.
9760 | 1220648-2025-28041 | 21950593 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21950593&pc=OZD | 09/27/2024 | 05/02/2025 |  | 05/02/2025 |  |  | 04/11/2025 | 05/02/2025 |  |  |  |  | 05/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025510631 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported via abiomed¿s long-term outcome and quality indicator (loqi) impella registry that a patient implanted with an impella 5.5 exhibited thrombocytopenia with a "possible " relationship to the impella device used. multiple units of blood products were transfused. the patient recovered with no sequelae. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9761 | 1220648-2025-28042 | 21950793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21950793&pc=OZD | 04/13/2025 | 10/09/2025 |  | 05/02/2025 |  |  | 04/13/2025 | 05/02/2025 |  |  |  |  | 05/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025607750 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. during support, positioning issues were noted as the device was pulled from 6.8 to 4.9 centimeters from aortic valve annulus. it was further reported that the patient was successfully bridged to an left ventricular assist device. | The complainant also reported that a placement signal not reliable alarm occurred and the placement signal became absent. |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated to include placement signal coding. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated. | Optical signal issue. about 419 hours into support, there was a drift seen in the ps, from about 85mmhg to about 20mmhg over the span of 38 minutes, and then it slowly drifts down over the course of 7 hours until the ps spikes to 3000 and it becomes unreliable. based on the data logs, the root cause of the optical signal issue is most likely due to sensor wear within design life. refer to the design traceability matrix (b)(4). positioning issues: the clinical stated that the impella was repositioned from 6.8 to 4.9 from the aortic valve annulus. due to ack of clinical details provided, the root cause of the positioning issues cannot be determined.
9762 | 1220648-2025-28043 | 21950800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21950800&pc=OZD | 04/04/2025 | 05/02/2025 |  | 05/02/2025 |  |  | 04/04/2025 | 05/02/2025 |  |  |  |  | 05/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026642038 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was being implanted with an impella cp device for mechanical circulatory support and encountered a placement signal issue. the sensor read -18/-18/-18 while the device was being primed and in the plastic packaging. it was confirmed the device had not been implanted or ever taken out of the plastic packaging. the decision was made to replace with a new impella cp device which was successfully implanted via the right axillary artery. | [5 empty] | Further information noted the patient would eventually expire one day later and, however, from complications unrelated to the impella cp devices. support was withdrawn related to ischemic bowel injury, rising pressors requirements, and loss of pulses to bilateral lower extremities. medical safety review noted placement signal issue occurred during priming. new impella cp was implanted without incident with minimal delay. there is no indication or report of hemodynamic instability. patient expired from complications that are dissociated from the impella devices used. the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9763 | 1220648-2025-28044 | 21951244 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21951244&pc=OZD | 04/14/2025 | 06/05/2025 |  | 05/02/2025 |  |  | 04/14/2025 | 05/02/2025 |  |  |  |  | 05/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026644051 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the impella cp was placed and percutaneous transluminal coronary angioplasty (ptca) was done. the physician also performed a right axillary angiogram in case of impella 5.5 escalation and then the patient had ventricular fibrillation arrest and required defibrillation. impella peel-away sheath was removed, and final positioning was confirmed. suction alarms occurred during suturing. echocardiogram and transesophageal echocardiogram showed the impella under the papillary. the physician was unable to reposition and decided to leave as is as suction was no longer occurring. the impella cp was observed to be shallow and the physician was able to reposition the impella cp. additionally, the patient developed bleeding from the impella access site which has not completely resolved. re-suturing, tranexamic acid was administered, fem stop was placed, manual pressure was held, heparin was paused, and thromboelastography was run in the lab trying to resolve the bleeding issue. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding, positioning issues, and vf/vt/af/at has been completed. the impella device was not returned for evaluation. access site bleeding: the pump was not returned. therefore, as per the clinical information provided, the root cause of the access site bleeding issue was most likely patient condition related since patient was bleeding from multiple sites and multiple troubleshooting attempts were unsuccessful. the root cause of the positioning issue and vt/vf was not determined due to lack of clinical information and no product returned.
9764 | 1220648-2025-28045 | 21951441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21951441&pc=OZD | 04/14/2025 | 06/13/2025 |  | 05/02/2025 |  |  | 04/14/2025 | 05/02/2025 |  |  |  |  | 05/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639713 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during support, the patient had a large hematoma at the access site. pressure was held, forward tension sutures were redone and a fem stop was applied. support was continued. additionally, the patient had hemolysis. the position was verified. the staff considered escalating the pump. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of access site bleeding and hemolysis has been completed. the product was not returned for review. based on clinical description and data analysis, the root cause of hemolysis was most likely patient condition related since the clinical team confirmed the high afterload of the patient's ventricles caused hemolysis and resolved by reducing afterload. the root cause of access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review.
9765 | 1220648-2025-28046 | 21951537 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21951537&pc=OZD | 04/11/2025 | 06/04/2025 |  | 05/02/2025 |  |  | 04/11/2025 | 05/02/2025 |  |  |  |  | 05/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025536169 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Positioning Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp was anuric and had hemolyzed labs. the impella was entrapped in chordae and the pump was eventually repositioned. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system: section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter, ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle, ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis, ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions, including catheter position, pre-existing medical conditions, and small left ventricular volumes, may also play a role in patient susceptibility to hemolysis.¿ | The investigation of hemolysis and positioning issues has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28046. g1 reporting contact facility name was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28046. h6 medical device problem code 3009 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28046.
9766 | 1220648-2025-28047 | 21951703 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21951703&pc=OZD | 04/11/2025 | 05/02/2025 |  | 05/02/2025 |  |  | 04/11/2025 | 05/02/2025 |  |  |  |  | 05/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025601350 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male | 76 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 pump exhibited a gi bleed. heparin was withheld. the patient remained on support for 11 days until he was successfully weaned and the pump was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during. section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves, ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings, section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves, ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
9767 | 1220648-2025-28048 | 21952096 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21952096&pc=OZD | 04/12/2025 | 05/27/2025 |  | 05/02/2025 |  |  | 04/12/2025 | 05/02/2025 |  |  |  |  | 05/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026647343 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | Gastrointestinal Hemorrhage | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a 65 year old male patient collapsed and was taken to hospital with. taken to ccl emergently, where lhc revealed severe vessel disease. thrombectomy catheter was used to try and aspirate thrombus without success. a stent was placed in the distal left main into the left anterior descending artery (lad) causing the left circumflex artery (lcx) to go down. a wire was able to pass through the struts into the lcx and vessel was successfully ballooned and stented. timi 3 flow achieved. two perclose were placed. impella cp was inserted using best practices. during support, it was reported that a echo was done and showed impella device was deep. the pump was repositioned to 4.1cm to av. nephrology and g.i. consulted due to scope indicating g.i. bleed that was cauterized successfully. additionally, the patient had runs of ventricular tachycardia (vt) that required cardioversion in addition of received five unit units blood in past 24 hours. the patient was successfully weaned and explanted. | [5 empty] | The investigation of positioning issues, vf/vt/af/at, and vascular damage - peripheral vessel has been completed. the pump was not returned for investigation. data log review was not required for this case run. the cause of the positioning issue was most likely use related since pump was found deep on second day of implant. the cause of the vascular damage issue was not determined since product was not returned and sufficient clinical information was not provided. the relation between gi bleed and impella was unknown. as the necessary information was not provided, the root cause of the vt/vf was not determined. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9768 | 1220648-2025-28049 | 21952278 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21952278&pc=OZD | 04/12/2025 | 06/05/2025 |  | 05/02/2025 |  |  | 04/12/2025 | 05/02/2025 |  |  |  |  | 05/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025579695 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Pumping Stopped | [5 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Gastrointestinal Hemorrhage | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported while inserting the impella 5.5 into a 77 year old male, that there was resistance upon passing the impella 5.5 over the .018 wire. the 5.5 would not take the bend down the ascending aorta from the subclavian. several wires were used and the impella was successfully implanted. the patient went into ventricular tachycardia (vt) that required shocking. the patient was stabilized once the wire was removed, and the pump started. the doctor was also concerned with the amount of blood lost during case due to removing and inserting the pump so many times. 2 units of blood were given. there was also oozing blood from below the graft site noted even when graft was clamped. once the patient was closed up, no hematoma or bleeding noted. additionally, the patient kept having occasional short runs of vt while on impella support. on day three of 5.5 support multiple suction alarms occurred, which were resolved. it was further reported that the patient continued to have gi bleeding. heparin was on hold. the patient have received 4 units of blood in total as well as 1l of albumin. upper gi scope showed a bleeding ulcer that was cauterized. it was further reported that the patient was just getting back to bed when the impella abruptly stopped. several attempts were made to restart with no success. the impella was explanted. | [5 empty] | The investigation of access site bleeding, vf/vt/af/at, vascular damage - peripheral vessel, pump stop, and low pump flow has been completed. low pump flow: clinical states that the pump was given multiple units of blood on (b)(6) 2025 due to gi bleed on previous midnight. 4 units of blood were given. suction alarms along with vt runs were observed as well. pump was returned cut from the catheter so the product investigation cannot be performed for low flow. data logs were investigated and low flows along with suction alarms were confirmed on (b)(6) 2025 midnight along with suction alarms. placement signal was seen slightly trending down. therefore as per the clinical information provided and data log investigation, the root cause of the low pump flow issue was most likely patient condition related. access site bleeding: clinical states that the access site was oozing after pump implant. blood was lost during multiple attempts of removing and inserting the pump so 2 units of prcb was delivered. oozing was also observed below the graft site even when clamped which stopped once patient was closed up. graft and introducer were not returned for investigation. based on the clinical information provided, the root cause of the access site bleeding issue was most likely use related to improper technique of inserting the pump which led to multiple attempts and incorrect clamping the graft since it was leaking even after clamping the graft to the vessel. vt/ vf: as the necessary information was not provided, the root cause of the vt/vf was not determined." vascular damage - peripheral vessel: clinical states that the patient had ongoing gi bleed on (b)(6) 2025 and 4 units of blood were given as a result. no other information related to the bleed has been given. therefore, the root cause of the vascular damage issue was not determined since insufficient clinical information was provided regarding the reason of the gi bleed. pump stop: clinical states that there was "impella stop" alarm observed. multiple restart attempts failed. pump was returned cut from the 75cm mark on catheter. there was biomaterial in the outflow cage and the resistance was 0 from the cut area to the pump. there was a hard kink near the motor housing of the pump. data logs show "alarm 115" with no motor current spikes or purge pressure issue. therefore, based on the clinical review and the product and data logs investigation, the root cause of the pump stop issue was most likely an open electrical line due to excessive force. d9 device returned to manufacturer date added. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for theimpella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (unites states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
9769 | 1220648-2025-27963 | 21952355 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21952355&pc=PYX | 04/10/2025 | 09/24/2025 |  | 05/02/2025 |  |  | 04/10/2025 | 05/02/2025 |  |  |  |  | 05/02/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025564027 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female | 91 KG | Hispanic |  | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella rp flex to treat right sided failure and cardiogenic shock. due to bleeding at the groin the team had to place additional suture. the patient survived. | Additionally, the complainant reported that the pump was observed to have lost the placement signal. the pump was not replaced and continued to support the patient. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ impella rp with smart assist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. no product was returned. access site bleeding: clinical stated that there was ooze noted so an additional suture was placed and the ooze resolved. the root cause of the access site bleeding is most likely due to use as an additional suture was placed and it stopped the oozing. placement signal issue: the clinical stated there was a loss of all metrics except motor current. no alarms were triggered at this time. flows were 0.0/0.0. data logs show that about 7 hours into support the placement signal began to drift up while flows drifted down. the drift lasted for about one hour and 45 minutes before going completely out and triggering alarm 1051. no issues were found with the motor current. the root cause of the placement signal issue is most likely due to dp sensor drift based on the clinical details and data log review. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include placement signal issue description. h6 medical device problem code updated to include placement signal coding. b1, product problem b7, other relevant history, including preexisting medical conditions, updated.
9770 | 1220648-2025-28050 | 21952546 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21952546&pc=OZD | 04/04/2025 | 05/02/2025 |  | 05/02/2025 |  |  | 04/04/2025 | 05/02/2025 |  |  |  |  | 05/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026622867 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Disability | Death | Required Intervention | Life Threatening |  |  | ECMO. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was initially implanted with an impella cp device for mechanical circulatory support (mcs) and the next day was escalated to the impella 5.5 device due to limb ischemia complication related to the impella cp. the impella 5.5 device was implanted with some complications and days following the patient would experience a stroke and subsequently expire. the impella 5.5 (pump 2) was placed and during the tying on of the axillary graft, the patient went into atrial fibrillation with rapid ventricular response and was successfully cardioverted after three shocks. the surgeon made the decision to go on extracorporeal membrane oxygenation (ecmo) first and then insert the impella 5.5. access in left femoral vein to cannulate for ecmo was made. second graft on axillary used for arterial cannulation of ecmo and ecmo was started. the impella cp device support level was lowered to p-2 and pulled back into aorta. the plan for vascular closure of right femoral artery (rfa) was at end of the case. the patient was noted to have a significant amount of bleeding from axillary access site and was given volume and blood products throughout. the physician placed a 23ff axillary sheath in graft with graft locks x2. left ventricle was accessed with al1 catheter and crossed with 0.035 straight wire. the guidewire was exchanged for 0.018 abiomed wire and impella 5.5 backloaded. the impella 5.5 device was started at p-4 with flows of approximately 2 lpm. ecmo flows were 3.4 lpm. the impella 5.5 was repositioned using fluoroscopy and transesophageal cardiogram guidance due to flipping up. the impella measured 5.6cm, and the physician was good with the position. the graft was trimmed down and blue suture hub placed and secured. the right axillary site was closed. large dressing applied over graft that was on the outside of the body. graft was secured and included impella and ecmo catheters to chest. vascular then removed the impella cp from the rfa and site was surgically closed. at end of procedure, the patient was transferred to the unit on ecmo and impella mcs. approximately three days later, however, the patient experienced a stroke and expired the next day. care was withdrawn. | [5 empty] | Medical safety review of the event noted the impella cp is a serious injury given the limb ischemia. the patient's demise is associated with the impella 5.5 device. with impella 5.5 related complications, there is not enough evidence to exclude impella as an associated factor in the patient's outcome. this is one of two devices associated with the event. refer to initial medical device report for impella cp serial number: (b)(6) with date of event 04-apr-2025 and patient identifier ending in (b)(6). the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
9771 | 1220648-2025-28051 | 21952805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21952805&pc=OZD | 04/11/2025 | 05/05/2025 |  | 05/02/2025 |  |  | 04/11/2025 | 05/02/2025 |  |  |  |  | 05/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1349165 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) would not recognize the purge cassette when it was inserted. the aic was replaced. no adverse impact to patient management was reported. | [5 empty] | Patient information a.1, a.2, a.3, a.4 and a.5 are unknown. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. the logs from the complaint date revealed that the console issued purge disc not detected alarm several times. the console was examined, and a broken purge disc flag was identified. the purge disc flag was observed to be broken and was the root cause of the reported purge disc not detected alarm issue. h6 type of investigation, investigation findings, investigation conclusions, and investigation conclusions codes were erroneously entered on the initial manufacturer device report number, as the investigation was complete at the time of the initial manufacturer device report number 1220648-2025-28051. h10 investigation results was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28051.
9772 | 1220648-2025-28052 | 21952983 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21952983&pc=OZD | 04/11/2025 | 06/06/2025 |  | 05/02/2025 |  |  | 04/11/2025 | 05/02/2025 |  |  |  |  | 05/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1756407 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the automated impella controller (aic) was connecting on and off to impella connect multiple times a day. once the patient was off support, they were able to troubleshoot and get the aic to connect successfully to impella connect multiple times. there was no harm to the patient reported. | [5 empty] | Patient information a1, a.2, a.3, a.4 and a.5 are unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - battery comm. failure (yellow alarm) has been completed. based on data and device analysis there was no issue found with a battery communication failure. d9 device returned to manufacturer date added.
9773 | 1220648-2025-28053 | 21953296 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21953296&pc=OZD | 04/04/2025 | 05/23/2025 |  | 05/02/2025 |  |  | 04/04/2025 | 05/02/2025 |  |  |  |  | 05/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026616502 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and the next day after the start of the support, a purge pressure high alarm triggered. the plan was to start tissue plasminogen activator (tpa) through the purge system. an echocardiogram was completed and confirmed optimal impella position. however, it was further noted that there was an inability to doppler pedal pulses on the impella side. the right leg was cold. the patient was planned for escalation to an impella 5.5 which was successfully completed. the patient would expire four days later however while on impella 5.5 device. | [5 empty] | Medical safety review of the event noted the impella cp is a serious injury given the limb ischemia. the patient's demise is associated with the impella 5.5 device. with impella 5.5 related complications, there is not enough evidence to exclude impella as an associated factor in the patient's outcome. refer to manufacturing report number 1220648-2025-28050 for the report on the impella 5.5 device. the impella cp device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The investigation of high purge pressure and limb ischemia has been completed. high purge pressure: upon visual inspection of returned product, blood was present in purge gap. pump was returned damaged and partially cut on catheter. high purge pressure reproduction testing was unable to be completed as purge line was cut at catheter damage. data logs were reviewed and lt logs showed a sharp rise in purge pressure with "high purge pressure" alarms being triggered. real time (rt) log showed a short purge rise of approximately 5 minutes with multiple motor current spikes and speed deviations right before purge rise. pump flow became almost fully saturated after motor current spikes and had clipping flows for remainder of rt log. clinical details noted purge pressure alarms were resolved after tissue plasminogen activator (tpa) use. the root cause of the high purge pressure was most likely due to biomaterial ingestion based on returned data log analysis. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
9774 | 1220648-2025-28054 | 21953560 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21953560&pc=OZD | 04/05/2025 | 05/02/2025 |  | 05/02/2025 |  |  | 04/05/2025 | 05/02/2025 |  |  |  |  | 05/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026622291 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had placed a impella cp for the support of a patient admitted in with acute myocardial infarction and cardiogenic shock. the day after implant the team observed left ventricle thrombus and 2 days later a thrombus in the atrium. the pump remained on despite the thrombus. after 6 days of support the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: impella cp with smart assist catheter insertion, section: axillary insertion of the impella cp with smart assist catheter, section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
9775 | 1220648-2025-28055 | 21953812 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21953812&pc=OZD | 04/10/2025 | 05/02/2025 |  | 05/02/2025 |  |  | 04/10/2025 | 05/02/2025 |  |  |  |  | 05/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026636275 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient developed heparin induced thrombocytopenia during impella 5.5 support. the purge was switched to bival and bicarb in response to the condition. support continued uninterrupted with the implanted pump. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9776 | 1220648-2025-27988 | 21941350 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21941350&pc=OZD | 04/27/2025 | 07/11/2025 |  | 05/01/2025 |  |  | 04/27/2025 | 05/01/2025 |  |  |  |  | 05/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026633534 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient had an attempted impella cp implant. due to severe vascular tortuosity, the implant was aborted. the patient's outcome is unknown. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ a2, a3 are unknown. | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be patient condition because the pump was unable to pass through patient anatomy.
9777 | 1220648-2025-27989 | 21941462 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21941462&pc=OZD | 04/03/2025 | 06/16/2025 |  | 05/01/2025 |  |  | 04/03/2025 | 05/01/2025 |  |  |  |  | 05/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025609326 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a purge system rupture occurred during impella cp support. it was noted that fluid was leaking from between the red plug and purge sidearm. the staff was advised to perform a purge filter bypass and the leak site was secured to ensure a tight seal. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. d.6b explant date is unknown as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of purge leak ¿ pump has been completed. the impella device was not returned for evaluation. clinical details noted that a leak was observed between the red plug and purge sidearm of the pump. the leak was temporarily secured but it was not noted if the leak fully resolved. the root cause of the purge leak - pump was most likely due to a damaged sidearm based on clinical details, however, the exact location of the leak could not be definitively determined as no product was returned for evaluation.
9778 | 1220648-2025-27990 | 21941598 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21941598&pc=OZD | 04/06/2025 | 06/10/2025 |  | 05/01/2025 |  |  | 04/06/2025 | 05/01/2025 |  |  |  |  | 05/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026614954 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Detachment of Device or Device Component | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that while a patient was on the impella 5.5 support the anticontamination sleeve broke during repositioning. there was no patient harm. there was a discussion of a 5.5 exchange however the patient remains on the original 5.5 support. | [5 empty] | The investigation of positioning issues and product damage (sleeve damage) has been completed. the impella device was not returned for evaluation. the cause of the product damage sleeve damage was most likely use related as it occurred during repositioning of the pump. the cause of the positioning issue was not determined due to lack of clinical details and no product was returned. d6b explant date added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Access site bleeding minor / pain at insertion site: hematoma at 5.5 pocket. slightly painful to patient. the cause of the access site bleeding minor was not determined due to lack of clinical details. the cause of the pain at insertion site was not determined due to insufficient clinical details. positioning issues / sterile sleeve damage: the pump was repositioned while on support during which the sterile sleeve reported to have ripped. the cause of the positioning issue was not determined due to lack of clinical details and no product was returned. the cause of the product damage sterile sleeve damage was most likely use related as it occurred during repositioning of the pump.
9779 | 1220648-2025-27991 | 21941718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21941718&pc=OZD | 04/06/2025 | 06/03/2025 |  | 05/01/2025 |  |  | 04/06/2025 | 05/01/2025 |  |  |  |  | 05/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026613454 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the positioning of an implanted impella cp pump appeared to be pointing posterior with the pigtail under the papillary muscle. attempts to reposition the pump were unsuccessful and the decision was made to explant the device. it was also noted that the patient had hemolyzed labs at this time and up until explant. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis and positioning issues has been completed. the impella device was not returned for evaluation. positioning issue: clinical states that the initial placement of the pump appeared to be pointing posterior with pigtail under the papillary muscle and repositioning was unsuccessful. pump was not returned. as per the clinical information provided, the pump's initial placement was incorrect and therefore, the root cause of the positioning issue was most likely use related to improper technique due to incorrect initial positioning. hemolysis: clinical states that after implant the echo showed pump in a bad position where the pigtail was under the pap muscle. pump was not returned and data logs were not returned as well. therefore, as per the clinical information provided, the root cause of the hemolysis issue was most likely improper positioning of the device where the pump was facing the posterior with pigtail under the pap muscle due to improper technique.
9780 | 1220648-2025-27992 | 21941904 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21941904&pc=OZD | 04/08/2025 | 06/18/2025 |  | 05/01/2025 |  |  | 04/08/2025 | 05/01/2025 |  |  |  |  | 05/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026618048 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the insertion of impella 5.5. device to a patient in acute myocardial infarction/cardiogenic shock was unsuccessful due to the vascular anatomy. during placement, several maneuvers for advancing were performed. the impella shaft looped during this and the decision was made for an impella cp device implant, which was completed via the femoral artery as via the right axillary artery was also not possible. the physician noted a "good procedural outcome. there was no report or indication of harm to the patient. | [5 empty] | Reported unable to place impella 5.5 due to torturous vascular anatomy. multiple maneuvers were attempted and impella shaft looped. a catheter kink was noted. impella 5.5 delivery was aborted and cp delivery was attempted but was also unsuccessful. patient was obese. impella 5.5 was returned and a catheter kink was found around 21cm mark. delivery issue - the root cause of the delivery issue was most likely patient condition related since the patient was obese with tortuous vascular anatomy. product damage - the root cause of the product damage was most likely a catheter kink due to tortuous vascular anatomy leading to the challenging delivery. d9 device returned to manufacturer date added g1 manufacturing site name, address, country, and fax number were erroneously entered on the initial manufacturer device report number 1220648-2025-27992. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
9781 | 1220648-2025-27993 | 21941963 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21941963&pc=OZD | 04/08/2025 | 05/01/2025 |  | 05/01/2025 |  |  | 04/08/2025 | 05/01/2025 |  |  |  |  | 05/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026629014 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [11 empty] | Required Intervention | Death | Disability | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support and during support, low flows were encountered intermittently. repositioning of the pump worked with varied results and blood volume was given to help stabilize the flow rate. it was additionally noted that biomaterial was observed at the time of the flow issues. as support continued, the patient became agitated, and an evaluation confirmed that the patient had suffered a stroke. support continued with the implanted pump for five more days, at which point the patient's plan of care was changed to comfort measures only and the patient expired. | [5 empty] | Given the details of the event, there is not enough evidence to exclude the impella device used as an associated factor. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smart assist system. section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
9782 | 1220648-2025-27994 | 21942115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21942115&pc=OZD | 04/08/2025 | 05/29/2025 |  | 05/01/2025 |  |  | 04/08/2025 | 05/01/2025 |  |  |  |  | 05/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025574133 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Defective Device | Malposition of Device | [5 empty] | Arrhythmia | Hemolysis | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs). however, during insertion using the 14x13 peel-away sheath, half of the orange collar of the sheath fell off, and there was oozing from around the diaphragm on the half missing the orange part. impella was already through the sheath at this time. therefore, the physician proceeded with placement and sheath was exchanged. the bleeding resolved when 14x13 peel-away sheath was removed. the patient was sent to the unit on support. discussion was made regarding transferring of the patient to be evaluated by heart failure for ventricular assist device/orthotopic heart transplantation. the next day after start of the impella mcs, the patient was having hemolysis. the impella cp device was deeper than expected, and the pigtail was interacting with the papillary muscle causing ectopic rhythms. multiple attempts to reposition were made. the impella cp pump lost its position, coming of ventricle after each repositioning. after multiple repositioning attempts, the physician decided to not reposition any further. support continued. two days later, it was noted the patient was listed for heart and kidney transplant, currently at that time on continuous renal replacement therapy. of further note was the impella position unknown alarms were occurring due to low pulsatility. support was noted to have continued, and approximately two days later, the patient would expire. the patient changed to do not resuscitate and withdrew care. even though the patient withdrew the impella support, with the impella-related hemolysis, arrhythmic events and device malpositionings there is not enough evidence to exclude the impella device used as an associated factor. | [5 empty] | The impella device was returned, and an evaluation/analysis is underway. upon completion of the analysis, a supplemental report will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of introducer damage ¿ mechanical, hemolysis, positioning issues, and vf/vt/af/at has been completed. the introducer was returned and inspected. visual inspection revealed insufficient adhesive was used on the orange hub which allowed the hub to become detached during use. the cause of the introducer damage - mechanical was a damaged valve due to the returned orange introducer hub showing insufficient adhesive to remain attached during use. hemolysis: the cause of the hemolysis was not determined since pump s/n (b)(6) was not returned and insufficient clinical information provided. positioning issues: during implantation procedure, it was evident impella was behind the papillary muscle and md repositioned. pump s/n (b)(6) was not returned. data logs were not recovered. the cause of the positioning issues was most likely use related due to the pump being inserted too deep into the pap muscles during implantation procedure. as the necessary information was not provided, the root cause of the vt/vf was not determined. h6 component code 3038 was erroneously entered on the initial manufacturer device report number 1220648-2025-27994.
9783 | 1220648-2025-27995 | 21942391 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21942391&pc=OZD | 04/08/2025 | 05/01/2025 |  | 05/01/2025 |  |  | 04/08/2025 | 05/01/2025 |  |  |  |  | 05/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026625146 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Malposition of Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock implanted with an impella cp device for mechanical circulatory support experienced access site bleeding, device malpositioning, and would later expire. the patient presented with st-segment elevation myocardial infarction with suprapubic pain and burning, purple fingers and toes hypotensive. it was noted four days prior the patient had a transcarotid artery revascularization. the decision was made to implant an impella cp device which was completed. temporary pacer was placed left femoral vein. the left heart catheterization revealed multivessel disease and surgical consult initiated. the swan-ganz catheter replaced the temporary pacer. the peel-away sheath was removed and the repositioning sheath advanced accordingly. the patient was transferred to the unit on support. once on the unit, purse sutures were placed for track oozing. additionally, attempts to reposition was made but unsuccessful under echocardiogram guidance. the pump was seen less than optimal position at the mitral. however, the pump was eventually repositioned. the patient continued on support but would subsequently expire a little over seven hours of impella mcs. the patient made the decision for no further treatment and care was withdrawn. | [5 empty] | Although the patient decided not to continue with care, there is not enough evidence to exclude the impella device (related complications) an associated factor to the overall outcome. the impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
9784 | 1220648-2025-27996 | 21942486 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21942486&pc=OZD | 04/08/2025 | 06/02/2025 |  | 05/01/2025 |  |  | 04/08/2025 | 05/01/2025 |  |  |  |  | 05/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026641431 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Anemia | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received from protect iv study regarding a patient (unknown race and ethnicity) implanted with an impella cp device for mechanical circulatory support. the study reported the patient experienced normocytic anemia, hemolysis. the patient was administered blood products, and the outcome was noted as resolved. the study reported the adverse event is probably related to the impella study, impella procedure and possibly related to the percutaneous coronary intervention. after approximately 20 hours of support, the patient was successfully weaned and the impella cp device was explanted with a survived outcome. | [5 empty] | The impella device was not received from the customer; therefore, an evaluation/analysis of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions- including catheter position, pre-existing medical conditions, and small left ventricular volumes- may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. clinical states that anemia was reported with possible relationship with impella. no signs of hemolysis next day. pump was not returned. data logs show no relevant abnormalities. no motor current spike, no suction alarms or positioning issues observed. therefore, the root cause of the hemolysis issue was not determined since the pump was not returned for investigation and insufficient clinical information was provided.
9785 | 1220648-2025-27997 | 21942844 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21942844&pc=OZD | 04/09/2025 | 05/01/2025 |  | 05/01/2025 |  |  | 04/09/2025 | 05/01/2025 |  |  |  |  | 05/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025572068 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male | 66 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that while an impella cp was implanted the patient had episodes of ventricular tachycardia (vt) with defibrillation. the pump was repositioned and support continued. the patient expired on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9786 | MW5169721 | 21943612 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21943612&pc=OZD | 05/09/2024 | 04/27/2025 |  | 05/01/2025 |  |  |  | 05/01/2025 |  |  |  |  | 05/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD | [5 empty] | ABIOMED, INC. | [8 empty] | No | [6 empty] | Health Professional |  |  | Voluntary report | Not Available | Unknown | [6 empty] | 81 YR |  | Female | 68 KG | Non Hispanic | White | Death | Yes | Yes | Therapeutic or Diagnostic Output Failure | [6 empty] | Failure to Anastomose | Hemorrhage/Blood Loss/Bleeding | Cardiogenic Shock | Thrombocytopenia | [9 empty] | Hospitalization | Required Intervention | Death |  |  |  | CYCLOBENZPRINE 10MG 1 TAB ORAL DAILY. | HYDRALAZINE 25MG ORAL ONCE DAILY. | HYDROCODONE-ACETAMINOPHEN 10MG-325MG. | IMPELLA CP. | IPRATROPIUM500/MCG/2.5ML. | LEVOTHYROXINE 25 MCG TAB 1 ORAL DAILY. | LOSARTAN 100MG ORAL TAB DAILY. | METFORMIN 500MG ORAL TABLET 1 BID. | MONTELUKAST 10MG TAB 1 ORAL QPM. | NEB RTQ6H. | NEXIUM 40 MG ORAL DELAYED RELEASE CAPSULE 1 CAP DAILY. | PRAVACHOL 80MG 1 TAB DAILY. | SOLN-INH0.5MG 2.5ML. | ZOLPIDEM 5MG TAB ORAL AT BEDTIME. | [9 empty] | Initial submission | [9 empty] | On (b)(6) 2024 dr. (b)(6) planned to place the impella 5.5 through the subclavian to bridge until cabg/mv repair on (b)(6) 2024. during the in initial procedure the impella 5.5 "would not fit" per dr. (b)(6) so the decided to place the impella cp and do surgery the following morning (b)(6) 2025. at 3:32 (b)(6) 2024 it was discovered that "the impella cp had basically quit working by the time they got ms. (b)(6) to the operating room" per dr. (b)(6). also, there was difficult exposure due to the leak around the impella device in the ascending aorta. the impella was removed and replaced with an iabp (intra-aortic balloon pump). ms. (b)(6) continued to have bleeding which led dr. (b)(6) to return her to the operating room. there was epicardial bleeding on the anterior surface of the heart and there was also bleeding from the proximal anastomosis vein graft to the aorta. bleeding was controlled and patient was sent back to cicu. patient expired on (b)(6) 2025 at 2:12 am from cardiogenic shock. thrombocytopenia most likely related to impella.
9787 | 1220648-2025-27999 | 21943968 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21943968&pc=OZD | 04/26/2025 | 07/10/2025 |  | 05/01/2025 |  |  | 04/26/2025 | 05/01/2025 |  |  |  |  | 05/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026622027 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp inserted for venting purposes, but it could not pass the aortic valve, so it was discontinued. blood volume analysis, two wire, 0.035' wire, etc. were also tried but could not be passed. the patient outcome is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since in spite of multiple attempts made, surgeon was unable to pass through innominate artery. d6a and d6b were erroneously entered on the initial manufacturer device report number.
9788 | 1220648-2025-27998 | 21943982 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21943982&pc=OZD | 04/09/2025 | 06/02/2025 |  | 05/01/2025 |  |  | 04/09/2025 | 05/01/2025 |  |  |  |  | 05/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1399060 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male | 98 KG | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an automated impella controller (aic) alarmed twice with the yellow alarm. support was successful and uninterrupted, and the aic was not replaced. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Data analysis: console logs from the reported day of event are consistent with complaint and show an controller error alarm (opm comm crc error ¿ event set # 1045) triggered during the clinical procedure. see the attached imc_ic6897.pdf. real time (rt) logs confirmed that all signals received from optical bench (placement, snr, contrast, lamp, gain) were negative values due to the crc error. see the attached image rt_log.png. device analysis: the console was sent back to field service for further investigation and was tested after arriving in fs. voltage from the pbm at connector (j3) was 4.9v and the voltage on the optical bench lamp connector (p4) was 4.8v. visible light was detected from the optical pump connector in the console, and "5s" parameter testing confirmed that the optical system was within specification. the controller error was due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed. root cause: the root cause of the ¿controller error (opm comm crc invalid) [optical pump connected] - alarm issued (i.e. aic - loss of opm data) and controller error optical bench ambient pressure out of range) [optical pump connected - alarm issued" was due to a firmware issue (optical bench fw anomaly). corrective actions: before returning the console to the customer, the following actions are instructed to perform: perform section 13 ¿ 15 from preventive maintenance procedure (0042-7309). perform a full functional check on the console and optical system (0042-7316).
9789 | 1220648-2025-28000 | 21944042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21944042&pc=OZD | 04/09/2025 | 05/13/2025 |  | 05/01/2025 |  |  | 04/09/2025 | 05/01/2025 |  |  |  |  | 05/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The publication entitled: "impella 5.5 in left ventricular noncompaction syndrome as bridge to heart transplant" reports a cerebrovascular accident of unclear etiology, inclusive of either bypass, peri-aortic clamp calcification, transient hypoxia, and/or impella explant. no patient information or device information was provided. | [5 empty] | The investigation of stroke/cva has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. a.1, a.2, a.3, a.4 and a.5 are unknown. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. h.4 device manufacture date is unknown. impella 5.5 with smart assist for use during cardiogenic shock: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ publication citation: sharma, s., ruiz, j., garg, p., leoni, j., patel, p., nativi, j., lyle, m., & goswami, r. (2024). impella 5.5 in left ventricular noncompaction syndrome as bridge to heart transplant. jhlt open, 4, 100051. https://doi.org/10.1016/j.jhlto.2023.100051.
9790 | 1220648-2025-28001 | 21944049 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21944049&pc=OZD | 04/23/2025 | 05/01/2025 |  | 05/01/2025 |  |  | 04/23/2025 | 05/01/2025 |  |  |  |  | 05/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | It was reported that information from a literature case was reviewed. a patient reliant on hemodialysis due to end-stage kidney disease, was admitted owing to deteriorating heart failure caused by triple-vessel coronary artery disease and low-flow low-gradient aortic stenosis. his cardiac index was 1.2 ml/min/m2 despite continuous inotropes administration. he had chronic limb-threatening ischemia, which was not assumed to disturb percutaneous cardiac intervention. the heart team decided to prefer percutaneous cardiac intervention (pci) to surgical intervention considering his advanced frailty. they performed pci under impella cp® support following balloon aortic valvuloplasty, resulting in complete revascularization. on the next day, just after removing impella cp® and closing the puncture site of the right femoral artery using perclose prostyle¿ closure device, the coldness and pain in his right lower extremity appeared. contrast-enhanced ct demonstrated the thrombus occlusion of his right femoral artery. emergent thrombectomy using fogarty catheters was performed, leading to complete restoration of arterial flow within 3 hours of symptom onset. surgical findings confirmed impella removal-related thrombosis as the etiology. the patient was transferred to a rehabilitation institute. the literature discussion is as follows "our case highlighted an instance of acute lower extremity ischemia subsequent to the removal of impella cp®, a rare yet potentially critical complication demanding specialized attention and management." the patient was noted to be in stable condition. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
9791 | 1220648-2025-28002 | 21944222 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21944222&pc=OZD | 04/28/2025 | 07/14/2025 |  | 05/01/2025 |  |  | 04/30/2025 | 05/01/2025 |  |  | 01/31/2026 |  | 05/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5-JP |  |  | ABIOMED, INC. |  |  | US |  | 2025459937 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 5.5 pump being placed to support a 54-year-old female patient admitted in with suspected myocarditis. the team had delivery issues that prevented the pump to be delivered. the pump kinked in the process of delivery. the pump was exchanged out and a new 5.5 pump was placed. the patient outcome was unknown. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. | The investigation of delivery issues and product damage (cannula kink) has been completed. the impella device was not returned for evaluation. delivery issue: pump unable to pass through vasculature and had a kink. patient had narrow vessels and suspected myocarditis. the root cause of the delivery issue was most likely patient condition related as the pump was unable to pass through due to the patients narrow vasculature and myocarditis. product damage (cannula kink): the root cause of the product damage cannula kink issue was most likely patient condition related to the patient's narrow vasculature and myocarditis. instructions for use: impella 5.5 with smartassist with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use during impella with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-28002. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-28002 as it is unknown if the initial reporter also sent the report to the food and drug administration. h10 instructions for use were incomplete on the initial manufacturer device report number 1220648-2025-28002.
9792 | 1220648-2025-28003 | 21944388 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21944388&pc=OZD | 04/10/2025 | 05/21/2025 |  | 05/01/2025 |  |  | 04/10/2025 | 05/01/2025 |  |  |  |  | 05/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026637987 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient had an attempted impella 5.5 implant and the pump had difficulty advancing impella through axillary artery due to anatomy and steep turn. decision made to place lunderquist buddy wire to aid in insertion. impella still unable to be advanced through artery. decision made to abort 5.5 insertion and proceed with impella cp insertion. there was no noted patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since physician had difficulty advancing impella through axillary artery due to anatomy and steep turn.
9793 | 1220648-2025-28004 | 21944640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21944640&pc=OZD | 03/06/2025 | 05/01/2025 |  | 05/01/2025 |  |  | 04/09/2025 | 05/01/2025 |  |  |  |  | 05/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025592336 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 71 YR |  | Female | 69 KG | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | Cardiogenic Shock | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported via a notification that was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured cardiogenic shock with a relatedness to the impella listed as unknown (undetermined). it was further reported that the patient was exhibiting arrhythmia and hypotension and impella was implanted. aicd was placed, but ventricular tachycardia/ventricular fibrillation was persistent, and the family decided to withdraw care while still on impella support. the patient expired. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9794 | 1220648-2025-28005 | 21944694 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21944694&pc=OZD | 04/10/2025 | 05/28/2025 |  | 05/01/2025 |  |  | 04/10/2025 | 05/01/2025 |  |  | 01/31/2026 |  | 05/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025458137 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported position in aorta alarm kept happening with impella cp. the physician did not think that the cp was in aorta after performing echocardiogram. the position could have possibly been a bit shallow due to the patient¿s small stature and small left ventricle. the patient remained stable and care was not altered from issue. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. based on the clinical details the cause of the positioning issues most likely was patient condition related due to small left ventricle.
9795 | 1220648-2025-28006 | 21944939 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21944939&pc=OZD | 04/10/2025 | 05/01/2025 |  | 05/01/2025 |  |  | 04/10/2025 | 05/01/2025 |  |  |  |  | 05/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026644213 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient on impella cp experienced a retroperitoneal bleed. the cp was removed, the patient received three units of red blood cells and the bleed was repaired. the patient was escalated to a 5.5. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9796 | 1220648-2025-28007 | 21945044 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21945044&pc=OZD | 04/11/2025 | 05/01/2025 |  | 05/01/2025 |  |  | 04/11/2025 | 05/01/2025 |  |  |  |  | 05/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026653704 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the sterile sleeve ripped during attempted repositioning at bedside. tegarderm was placed over the sleeve and support continued uninterrupted. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9797 | 1220648-2025-28008 | 21945106 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21945106&pc=OZD | 03/28/2025 | 05/01/2025 |  | 05/01/2025 |  |  | 04/10/2025 | 05/01/2025 |  |  |  |  | 05/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026619076 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Low Blood Pressure/ Hypotension | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on impella cp support and during support, the patient developed a hematoma at the access site and was hypotensive requiring blood products. the patient remained stable and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
9798 | 1220648-2025-28009 | 21945127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21945127&pc=OZD | 04/11/2025 | 10/17/2025 |  | 05/01/2025 |  |  | 04/11/2025 | 05/01/2025 |  |  |  |  | 05/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025610391 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Other | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support it was reported that they were having suction alarms occasionally. they checked position and were satisfied. additionally, a ultrasound completed of right arm, showed a small non occlusive deep vein thrombosis. the patient continued on support without intervention. | A complainant reported that the decision for impella 5.5 was made when a patient started to decompensate. while on support, a placement signal low alarm occurred. additionally, the surgeon reported that the pump was angled toward the mitral valve, however the pump was flowing well with no hemolysis. the surgeon decided to leave the pump in for support. the patient was successfully weaned from the pump and the device explanted. |  |  |  |  | The investigation of thrombus, positioning issues, and low pump flow has been completed. the impella device was not returned for evaluation. data log analysis did not show major positioning issues during the case. the root cause of the suction alarms was most likely improper positioning since clinical details stated the suction alarms were resolved by repositioning and no other abnormalities in the data logs. the root cause of thrombus and the positioning issues was not determined since insufficient clinical details were provided. d6b explant date was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28009. h6 health effect - impact code 4648 was erroneously entered on the initial manufacturer device report number 1220648-2025-28009. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. optical signal issue: placement signal not reliable (psnr) alarms were reported on (b)(6) 2025 with mention of placement signal low alarms as well. data log review shows period of low snr throughout the first two weeks of the case with variable gain and low contrast. an abrupt rise in snr is seen on (b)(6) 2025 with a decrease in gain and increase in contrast afterwards. after these changes, there were no abnormalities with the optical parameters. just before this abrupt rise in snr, the ac power was disconnected and then reconnected shortly after the rise. the root cause of the optical signal issue was most likely an air gap from between the aic and the patient cable plug since data logs showed characteristics of an air gap and the issue was eventually resolved. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: optical signal "2917" issue added for fda reporting. | D10 concomitant medical products and therapy dates updated.
9799 | 1220648-2025-28010 | 21945154 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21945154&pc=OZD | 04/10/2025 | 05/05/2025 |  | 05/01/2025 |  |  | 04/10/2025 | 05/01/2025 |  |  |  |  | 05/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025609615 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with an impella cp for support due to cardiogenic shock. during support it was reported that pulses were absent in impella leg. they transitioned to axillary cp. distal pulse improved after groin cp removal. additionally, a thrombectomy was performed to right lower extremity. the patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia and thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
9800 | 1220648-2025-28011 | 21945263 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21945263&pc=OZD | 04/07/2025 | 05/28/2025 |  | 05/01/2025 |  |  | 04/07/2025 | 05/01/2025 |  |  |  |  | 05/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026634600 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | Ventricular Fibrillation | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient on impella cp experienced a gastrointestinal bleed and anticoagulant was paused. it was further reported multiple episodes of ventricular tachycardia (vt)/ ventricular fibrillation (vf) requiring cardiopulmonary resuscitation and electrical cardioversion. an echocardiogram performed which showed the cp was deep and had to be pulled back into a good position. the patient continued to have persistent episodes of vt/vf. the patient was taken back to cath lab for left heart catheterization. the patient was stable and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of vf/vt/af/at, positioning issues, and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. the root cause of the vascular damage issue and positioning issue was not determined since product was not returned and insufficient clinical information was provided. positioning issue: clinical states that after the patient had vt runs, the patient had cardiopulmonary resuscitation (cpr) performed which led to pump pushed deep and was pulled back into a good position. pump was not returned for investigation. therefore, as per the clinical information provided, the root cause of the positioning issue was use related to the cpr performed to the patient. h6 health effect - impact code 4647 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28011.
9801 | 1220648-2025-28012 | 21945356 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21945356&pc=PYX | 04/11/2025 | 05/28/2025 |  | 05/01/2025 |  |  | 04/11/2025 | 05/01/2025 |  |  | 08/31/2025 |  | 05/01/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024404271 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was attempted to be implanted with an impella device for mechanical circulatory support. the device would not advance past the right ventricle. multiple attempts were made advance pump unsuccessfully. a buddy wire was added and was also unsuccessful. decision was made to abort and attempt rp flex. rp flex went into correct position without issues. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella rp smartassist system with automated impella controller. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related.
9802 | 1220648-2025-28013 | 21945369 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21945369&pc=OZD | 04/10/2025 | 06/19/2025 |  | 05/01/2025 |  |  | 04/11/2025 | 05/01/2025 |  |  | 12/31/2025 |  | 05/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025441854 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Defective Device | Material Split, Cut or Torn | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during the impella cp procedure, after positioning the guidewire in the ventricle the wire couldn¿t be removed. the wire stopped in the peel away sheath and they both were removed together. it was noted the tip of the introducer was split. it was also reported that the patient expired while on impella cp support. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Delivery issues: the fm "product damage (guidewire damage - peripheral vessel)" was updated to the fm "delivery issues" since there were no full/partial detachments observed on the returned guidewire. clinical details state that the 0.018" wire could not be removed after delivering the pump, stopping in the peel-away sheath. the guidewire was returned and minor bending was observed near the tip as well as some unraveling at the beginning of the coiled section on the wire. the root cause of the delivery issues was not determined since insufficient clinical details were provided. introducer damage - mechanical: clinical details state that the 0.018" guidewire could not be removed through the sheath and sheath tip damage was observed upon removal of the sheath. it was noted that there were no vessel conditions and it's unknown if excessive force was used. the 14frx13cm sheath was returned and there was damage found on the tip as well as a slight kink at the base of the sheath near the hub. the root cause of the introducer damage was most likely sheath tip split but the cause of the damage is unknown.
9803 | 1220648-2025-28014 | 21945542 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21945542&pc=OZD | 04/11/2025 | 06/19/2025 |  | 05/01/2025 |  |  | 04/11/2025 | 05/01/2025 |  |  |  |  | 05/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026627454 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Female |  |  |  | Injury | Yes | Yes | Break | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was further reported that the patient on impella cp support was turned in bed and the connection between the purge disc and filter broke. | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced limb ischemia. the plan is to place a fem to fem bypass. |  |  |  |  | The investigation of limb ischemia and puge leak pump has been completed. purge leak - reported the connection between purge disc and filter broke when turning the patient in bed. impella cp and purge cassette were returned. pump was cut at the catheter around 60cm and purge side arm was broken between the purge disc and filter. the root cause of the purge leak was damage to the reservoir to filter joint since it was broken when turning the patient. limb ischemia: as all the necessary information was not provided, the root cause of the limb ischemia was not determined. instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." d9 device returned to manufacturer date added b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28014. b5 revised with additional information that was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28014. h6 medical device problem code 1069 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28014. h10 instructions for use was inadvertently omitted on the initial manufacturer device report number 1220648-2025-28014. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
9804 | 1220648-2025-28015 | 21945553 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21945553&pc=OZD | 04/11/2025 | 05/01/2025 |  | 05/01/2025 |  |  | 04/11/2025 | 05/01/2025 |  |  |  |  | 05/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026641911 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that the impella cp was explanted due to the patient developing limb ischemia. the patient is stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
9805 | 1220648-2025-28016 | 21945593 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21945593&pc=OZD | 04/11/2025 | 06/30/2025 |  | 05/01/2025 |  |  | 04/11/2025 | 05/01/2025 |  |  |  |  | 05/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1379222 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The automated impella controller intermittently displayed ¿ac power disconnected¿ although it was plugged into the wall. multiple outlets were tried. the device was removed due to failure mode. there was no patient harm reported. | [5 empty] | The investigation of automated impella controller (aic) - a/c power issues has been completed. the console logs from the reported day of event are consistent with complaint and show several recurring instances of ac power disconnected alarm (109), being set and cleared. console was inspected and a defect of ac power cord was discovered: strain relief in the plug was not engaged, and wires were loose and frayed. the cause of the issue was a defective cable. h6 component code 925 was erroneously entered on the initial manufacturer device report number 1220648-2025-28016. new code entered | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
9806 | 1220648-2025-28017 | 21945808 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21945808&pc=OZD | 04/11/2025 | 06/03/2025 |  | 05/01/2025 |  |  | 04/11/2025 | 05/01/2025 |  |  |  |  | 05/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1108091 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the placement signal on an automated impella controller (aic) suddenly disappeared during patient support. pump flows remained stable, but a yellow notification shortly followed indicating a controller error. the aic was swapped and the placement signal returned on the new aic. there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - controller error (yellow alarm) and optical signal issue has been completed. the console logs from the reported day of the event are consistent with the complaint. the case began on (b)(6) 2025, when the pump was connected and then disconnected within a few seconds. within 30 minutes of the pump starting, the impella position alarm and adjust lv signal alarm were displayed. additionally, controller error alarm (opm comm stopped ¿ event set 1043) triggered during the clinical procedure. the logs also reported process sampling data from optical bench: sent stop acquisition cmd ack, indicating ossiopsens receive thread gets a data sampling packet with an unexpected length. the logs confirmed that all signals received from optical bench (placement, signal not reliable (snr), contrast, lamp, gain) are lastly reported as large positive values and then no more samples were recorded. device analysis: the console was sent back for further investigation. the reported complaint could not be reproduced. visible light was detected from the optical pump connector in the console, and analog output from the array of the optical bench (fringe pattern) was seen without any distortion and the "5s" parameter using optical cavity found that the values were within the specified values as per the preventative maintenance procedure. root cause: the root cause of the loss of placement signal and the controller error alarm (opm comm stopped) [optical pump connected] - event 1043 was most likely due to the aic software issue. this was entered into jama with defect/ bug
9807 | 1220648-2025-28018 | 21945824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21945824&pc=OZD | 04/11/2025 | 05/01/2025 |  | 05/01/2025 |  |  | 04/11/2025 | 05/01/2025 |  |  |  |  | 05/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026633147 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, the motor current increased and the pump stopped. the pump could not be restarted. the pump was explanted and replaced. there was no patient harm reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. e.1 name prefix/title, first name and last name are unknown. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk cpi & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
9808 | 1220648-2025-27958 | 21933760 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21933760&pc=OZD | 04/03/2025 | 05/22/2025 |  | 04/30/2025 |  |  | 04/03/2025 | 04/30/2025 |  |  |  |  | 04/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026642186 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombocytopenia | Hemodynamic instability | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs). the patient presented with lethargy and extreme fatigue and discomfort. radial access was gained and the right coronary artery was 100% proximally occluded. the following day after start of impella mcs, it was noted that the patient continued to have limitations to impella flow despite adequate position and preload. the next day it was noted that the patient's hemodynamics looked "worse" suggesting right ventricular dysfunction. approximately four days later, the patient experienced hemolysis and was followed by thrombocytopenia during the impella mcs. the patient was given three units of platelets and one unit of red blood cells in response to the conditions. support was continued with the implanted pump. however, the patient continued to become worse. the physician discussed starting continued renal replacement therapy through the extracorporeal membrane oxygenation (emco) circuit. approximately six days later, day 11 of support, extremities were mottled, oozing yellowish blood from previous femoral ecmo sites. the patient was noted to be in irreversible organ failure and subsequently expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis and thrombocytopenia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
9809 | 1220648-2025-27959 | 21933778 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21933778&pc=OZD | 04/08/2025 | 05/16/2025 |  | 04/30/2025 |  |  | 04/08/2025 | 04/30/2025 |  |  |  |  | 04/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025610837 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was noted the groin looking "great" and no oozing departing from the catheterization lab. however, upon arrival to the unit the patient had oozing from all line sites. activated clotting times were high. heparin was stopped to control the bleeding. the patient was also given two units of packed red blood cells. the following day the bleeding was noted to have subsided. the patient was noted to be stable following the event. support continued for an additional two days. the patient was successfully weaned, and the device was explanted with a survived outcome. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system, section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events, ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the groin site noted to be bleeding post arrival to icu. the root cause of access site bleeding was most likely patient condition since the groin started to bleed after the patient (b)(6) arrived to icu.
9810 | 1220648-2025-27960 | 21933926 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21933926&pc=OZD | 01/23/2025 | 05/28/2025 |  | 04/30/2025 |  |  | 04/09/2025 | 04/30/2025 |  |  |  |  | 04/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025534767 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported through abiomed¿s long-term outcome and quality indicator (loqi) impella registry that a patient implanted with an impella cp experienced ventricular tachycardia (vt) with a "possible" relationship to the impella device used. amiodarone was given and the patient recovered with no sequelae. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of low pump flow and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact facility name was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27960. h6 medical device problem code 1248 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27960.
9811 | 1220648-2025-27961 | 21934059 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21934059&pc=OZD | 04/01/2025 | 04/30/2025 |  | 04/30/2025 |  |  | 04/09/2025 | 04/30/2025 |  |  |  |  | 04/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025551010 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported via abiomed¿s long-term outcome and quality indicator (loqi) impella registry that a patient implanted with an impella cp exhibited a hematoma at the left groin impella access site with a " probable" relationship to the impella device used. a femstop was placed on the left femoral artery to manage hemostasis. post procedure pulses became thready and the patient was positive for persistent sciatic artery ischemia. the patient received a thrombin injection, pulses became normal, and the patient was cleared for discharge. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9812 | 1220648-2025-27962 | 21934170 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21934170&pc=OZD | 04/09/2025 | 05/01/2025 |  | 04/30/2025 |  |  | 04/09/2025 | 04/30/2025 |  |  |  |  | 04/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1094547 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an automated impella controller (aic) would not recognize multiple purge discs. the issue was identified during an in-service training and no patient was involved in the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. the logs from the complaint date revealed that the console issued purge disc not detected alarm several times. the console was examined, and a broken purge disc flag was identified. the purge disc flag was observed to be broken and was the root cause of the reported purge disc not detected alarm issue g1 reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27962. h6 type of investigation code 4114 was erroneously entered on the initial manufacturer device report number 1220648-2025-27962. h10 investigation results were inadvertently omitted on the initial manufacturer device report number 1220648-2025-27962.
9813 | 1220648-2025-27964 | 21934433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21934433&pc=OZD | 04/10/2025 | 05/28/2025 |  | 04/30/2025 |  |  | 04/10/2025 | 04/30/2025 |  |  |  |  | 04/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026643653 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male | 75 KG | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that while implanted with an impella cp the patient experienced a hematoma that required surgical intervention. the vessel had to be repaired and the pump was explanted. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation of vascular damage - peripheral vessel has been completed. the pump was returned for investigation. catheter was returned cut at the sterile sleeve. no abnormalities were found upon investigating the blue butterfly. hematoma that required the vessel had to be repaired. the root cause of the vascular damage was unable to be determined as insufficient clinical details were provided and no abnormalities were noted with returned product related to the failure mode upon investigation. g1 reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27964. h10 narrative "device not received" was erroneously stated on the initial manufacturer device report number 1220648-2025-27964.
9814 | 1220648-2025-27965 | 21934732 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21934732&pc=OZD | 04/10/2025 | 04/30/2025 |  | 04/30/2025 |  |  | 04/10/2025 | 04/30/2025 |  |  |  |  | 04/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026622585 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient was placed on impella 5.5 for hemodynamic support. during support it was reported that the echo showed a apical lv thrombus. when the patient went to the or for further evaluation last week, tee showed there was no thrombus. patient was successfully weaned and explanted yesterday. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
9815 | 1220648-2025-27966 | 21935110 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21935110&pc=OZD | 04/10/2025 | 07/02/2025 |  | 04/30/2025 |  |  | 04/10/2025 | 04/30/2025 |  |  |  |  | 04/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026645922 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Perforation of Vessels | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 71-year-old female patient on an impella cp for support during a high-risk procedure. it was reported that during repositioning, the impella was inadvertently pulled back into the aorta. pressor requirements were increased as a result of the issue. attempts to readvance the pump were unsuccessful as a result of the patient's aortic stenosis. it was then noted that the impella site was hemorrhaging. the patient experienced mixed cardiogenic/hemorrhagic shocked. manual pressure was held and over five units of blood were given to try to manage the condition. hemostasis was not achieved. the patient was made dnr by the family. vascular surgery was made aware of the arterial injury at the impella site, but the patient ultimately passed away due to the hemorrhagic event. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | No product was returned for evaluation. access site bleeding - major: clinical details noted that the surgical guidewire present in the femoral artery next to impella site and when repositioned, bleeding at impella site was observed. it was noted that angle matching and forward suturing were applied when impella was inserted. multiple units of blood were given and manual pressure was in attempt to achieve hemostasis without success. as the clinical details noted that excessive bleeding but no vascular damage occurred, the failure mode was changed from "vascular damage - peripheral vessel" to "access site bleeding - major". the root cause of the access sure bleeding could not be definitively determined as pump was angle matched and forward sutured, blood was given, and manual pressure was held in attempt to resolve bleeding. positioning issues: clinical details noted that when guidewire was being repositioned at impella access site, "impella in aorta" alarms were triggered and echo showed pump in aorta with pigtail across av and in aorta. the root cause of the positioning issues was most likely an interaction with other devices as the pump came out of position when repositioning of a guidewire at the access site based on clinical details. | No product was returned for evaluation. access site bleeding - major: clinical details noted that the surgical guidewire present in the femoral artery next to impella site and when repositioned, bleeding at impella site was observed. it was noted that angle matching and forward suturing were applied when impella was inserted. multiple units of blood were given and manual pressure was in attempt to achieve hemostasis without success. as the clinical details noted that excessive bleeding but no vascular damage occurred, the failure mode was changed from "vascular damage - peripheral vessel" to "access site bleeding - major". the root cause of the access sure bleeding could not be definitively determined as pump was angle matched and forward sutured, blood was given, and manual pressure was held in attempt to resolve bleeding. positioning issues: clinical details noted that when guidewire was being repositioned at impella access site, "impella in aorta" alarms were triggered and echo showed pump in aorta with pigtail across av and in aorta. the root cause of the positioning issues was most likely an interaction with other devices as the pump came out of position when repositioning of a guidewire at the access site based on clinical details.
9816 | 1220648-2025-27967 | 21935374 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21935374&pc=OZD | 04/09/2025 | 04/30/2025 |  | 04/30/2025 |  |  | 04/10/2025 | 04/30/2025 |  |  | 01/31/2026 |  | 04/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025463322 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 79-year-old male with an impella cp due to acute myocardial infarction. it was reported that during the impella cp procedure one perclose attempted and failed due to calcium in the vessel. the impella started successfully to support the patient. the next day the patient had a large hematoma at impella site. anticoagulation was held. it was also reported that the patient started to cough up clotted blood, which turned into frank blood. patient emergently intubated and two units of blood were administered. bleeding continued to worsen in et tube, patient was bronched three more times to identify bleeding, but no bleeding was found. patient continued to deteriorate, and decision was made for comfort care. patient passed. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
9817 | 1220648-2025-27968 | 21935604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21935604&pc=OZD | 04/10/2025 | 04/30/2025 |  | 04/30/2025 |  |  | 04/10/2025 | 04/30/2025 |  |  |  |  | 04/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026644892 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the 77-year-old male patient after impella cp placement developed hematoma from the impella access site. manual pressure was held and hematoma resolved. in icu the patient developed hematoma again. the patient was taken to the cath lab and impella cp was exchanged. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9818 | 1220648-2025-27969 | 21936018 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21936018&pc=OZD | 04/10/2025 | 06/03/2025 |  | 04/30/2025 |  |  | 04/10/2025 | 04/30/2025 |  |  |  |  | 04/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1285929 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a console failed to boot up. this aic issue was discovered outside of patient use. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - system self-check error has been completed. the console logs showed a communication issue with the power battery manager (pbm) during the initial bootup. due to the communication issue, the aic console rebooted automatically (as designed) to attempt another power on. after rebooting, the console displayed the "system self check failed" screen, and the console was then forcefully shut down. device analysis: the reported failure mode was reproduced during testing. upon bootup, the console rebooted automatically and showed the "system self check failure" screen. visual inspection confirmed the pbm contamination which has impacted a neighboring communication channel. root cause: the root cause of the ¿system self check failure¿ was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps.
9819 | 1220648-2025-27970 | 21936468 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21936468&pc=OZD | 03/09/2025 | 05/19/2025 |  | 04/30/2025 |  |  | 03/09/2025 | 04/30/2025 |  |  |  |  | 04/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025547728 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male | 99 KG | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Detachment of Device or Device Component | [5 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella 5.5 experienced ventricular tachycardia (vt) and ectopy throughout support. the position of the pump was verified and the p level was decreased. additionally, there were suction events that occurred throughout support. furthermore, the anticontamination sleeve was disconnected from distal white lock. the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of vf/vt/af/at, product damage (sleeve damage), and low pump flow has been completed. the impella device was not returned for evaluation. the clinical states that the patient had multiple suction alarms throughout the case and patient unable to go over p-5 and p-4 on certain days. patient had low volume on (b)(6) 2025. pump was not returned for investigation. data logs were investigated and no motor current spike was observed. patient had suction alarms on multiple days and at the attempts of p-level increase. patient was on the lowest end of flow range at p5 for most of the case. the patient most likely couldn't tolerate higher p-levels. based on the data logs and clinical information, the root cause of the low pump flow issue was most likely patient condition related. as the necessary information was not provided, the root cause of the vt/vf was not determined. the root cause of the product damage- sterile sleeve issue was not determined since product was not returned for investigation. d6b explant date added g1 reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27970.
9820 | 1220648-2025-27971 | 21936589 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21936589&pc=OZD | 04/10/2025 | 06/04/2025 |  | 04/30/2025 |  |  | 04/10/2025 | 04/30/2025 |  |  |  |  | 04/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1713948 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an issue during prep when the team was attempting to start new case with the console. it was reported that the gear would not rotate when it came time to insert the purge cassette. therefore, the team was unable to insert the cassette. there was no patient harm. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. e1-initial reporter unknown. | The investigation of automated impella controller (aic) - purge issue ¿ other has been completed. based on the data and device analysis there was no issue found with the console. the device functioned/operated to specification.
9821 | 1220648-2025-27972 | 21936774 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21936774&pc=OZD | 04/10/2025 | 04/30/2025 |  | 04/30/2025 |  |  | 04/10/2025 | 04/30/2025 |  |  |  |  | 04/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026630613 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient was on bipella support and when the impella rp was removed the impella 5.5 needed to be repositioned. on day 10 of 5.5 support the impella was facing toward the mitral valve and measuring 6.8cm. decision made to reposition and there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9822 | 1220648-2025-27973 | 21937027 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21937027&pc=OZD | 04/10/2025 | 04/30/2025 |  | 04/30/2025 |  |  | 04/10/2025 | 04/30/2025 |  |  |  |  | 04/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026612984 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella 5.5 support and during support, the pump was towards the posterior wall. the pump was then repositioned. after repositioning, there was bleeding into sterile sleeve coming from around the white catheter at the blue button lock. intervention is unknown for the bleeding. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9823 | 1220648-2025-27974 | 21937262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21937262&pc=OZD | 04/10/2025 | 05/21/2025 |  | 04/30/2025 |  |  | 04/10/2025 | 04/30/2025 |  |  |  |  | 04/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026629253 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Restricted Flow rate | Material Separation | [5 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that when the patient was in the operating room for impella 5.5 placement blood products and platelets were required for impella suction events. the next day, impella was repositioned. it was pulled back 1 cm to the 42 cm marker. impella flow increased after repositioning. on day two of 5.5 support the patient received one unit packed red blood cells for ventricular tachycardia and suction alarms. on day 16 of support, it was reported that tegarderm was applied to the distal end of sterile sleeve where it was separated from total blood volume. the patient remains on support and outcome is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of vf/vt/af/at, low pump flow, and product damage (sleeve damage) has been completed. as the necessary information was not provided, the root cause of the vt/vf was not determined. as per clinical information provided, the root cause of the low pump flow issue was most likely improper positioning of the device as observed under transesophageal echocardiogram (tee) which was resolved with repositioning. clinical states on (b)(6) 2025, the sleeve was torn off and tegaderm was use to wrap it. pump was not returned for investigation. therefore the root cause of the product damage (sleeve) was not determined since the product was not returned for investigation. b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27974. d6a and d6b were erroneously entered on the initial manufacturer device report number 1220648-2025-27974. h6 medical device problem code 1562 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27974.
9824 | 1220648-2025-27975 | 21937658 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21937658&pc=OZD | 04/10/2025 | 06/24/2025 |  | 04/30/2025 |  |  | 04/11/2025 | 04/30/2025 |  |  |  |  | 04/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1445594 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller (aic) failed. this caused the screen to turn off and reboot on its own, displaying only letters, but not completing the reboot process. the aic was switched without any issues. there was no patient harm. | [5 empty] | The investigation of automated impella controller (aic) - shutdown or restart issue has been completed. data analysis: unable to download the logs from the console, as it couldn¿t boot a hard drive or enter service mode. this causes the inability to confirm the unexpected reboot. device analysis: this console was tested and the reported failure mode was not reproduced. however, a different symptom, the inability to bootup was observed which likely has a related cause to the reported reboot. the root cause of the reported aic reboot and inability to boot up (observed during testing) was due to corrupted compact flash cards. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-27975. d6a and d6b should have been left blank on the initial manufacturer device report number 1220648-2025-27975. h8 usage of device was erroneously selected on the initial manufacturer device report number 1220648-2025-27975. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
9825 | 1220648-2025-27976 | 21937889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21937889&pc=OZD | 04/11/2025 | 04/30/2025 |  | 04/30/2025 |  |  | 04/11/2025 | 04/30/2025 |  |  |  |  | 04/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026620757 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Detachment of Device or Device Component | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support and during the insertion, the catheter sleeve got bunched up and partially underneath the toughy valve causing difficulty advancing. physician thought it was an issue with the sleeve. it was advised to attempt to readjust the sleeve to allow for ease of movement, but physician opted to exchange for new pump. after removal, able to adjust the catheter sleeve with ease. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high-risk pci. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
9826 | 1220648-2025-27978 | 21938452 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21938452&pc=OZD | 04/17/2025 | 07/11/2025 |  | 04/30/2025 |  |  | 04/28/2025 | 04/30/2025 |  |  |  |  | 04/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5-JP |  |  | ABIOMED, INC. |  |  | US |  | 2025574182 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | Yes | Positioning Failure | [6 empty] | Insufficient Information | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that a patient had an impella 5.5 pump placed for support during a high-risk procedure. the pump was supporting for six days when accidentally the pump was pulled out of position. the pump was not re-delivered across the valve to the left ventricle so the pump was explanted and replaced. further details were received stating that the patient expired due to a cause unrelated to the pump. | It was further reported that the patient expired on support. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. clinical states that the pump was pulled out into ventricle when the patient was changing position for an x-ray. pump was explanted and replaced with another pump. as per the clinical information provided, the root cause of the positioning issue was most likely use related due to improper technique. b1 product problem was inadvertently omitted on manufacturer device report number 1220648-2025-27978-1. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-27978. | B1-adverse event b2-selected death and added date of death. b5-additional details provided states that the patient expired while on support. h6-clinical code 4580, impact code 1802. h10-medical safety officer (mso) review. the impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ it was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and although the cause of death is provided and noted as multiple organ failure (mof), it cannot be said that the loss/lack of hemodynamic support had nothing to do with the final outcome. there is not enough evidence to excluded the impella as an associated factor to the outcome. h1-updated to death.
9827 | 1220648-2025-27979 | 21938631 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21938631&pc=OZD | 03/12/2025 | 04/30/2025 |  | 04/30/2025 |  |  | 03/12/2025 | 04/30/2025 |  |  |  |  | 04/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025592178 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Material Split, Cut or Torn | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was implanted with an impella 5.5 as a bridge to left ventricular assist device or transplant. during support, the patient was getting some suctioning events at p-8. the team lowered to p-5 without suction alarms. later, it was noted that they were unable to increase the p-level to above p-4 due to suction alarms. the nurse expressed that she thought that there may have been a small leak around the purge reservoir. they changed the purge cassette and monitored it, which there had been no leaks. additionally, there was a small tear noticed on the anti-contamination sleeve of impella. they addressed according to their hospital protocol. the patient received an orthotopic heart transplantation. the impella 5.5 was successfully weaned and explanted for bridge to transplant. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9828 | 1220648-2025-27980 | 21938788 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21938788&pc=OZD | 03/24/2025 | 05/08/2025 |  | 04/30/2025 |  |  | 03/24/2025 | 04/30/2025 |  |  |  |  | 04/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025610823 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 pump placed to support a cardiogenic shock patient. support was ongoing with the pump position "in aorta" alarms occurring. the pump position was assessed repeatedly as it was close to the valve and there were signs of hemolysis. the labs were hemolyzing and lactate dehydrogenase and plasma-free hemoglobin were drawn. the patient remains on support and the outcome is unknown. | [5 empty] | The impella device was not received from the customer as it is still supporting the patient. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of the positioning issues and hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
9829 | 1220648-2025-27981 | 21938863 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21938863&pc=OZD | 04/05/2025 | 05/19/2025 |  | 04/30/2025 |  |  | 04/05/2025 | 04/30/2025 |  |  |  |  | 04/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025611353 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support and during support, the patient had dysrhythmia. the p level was decreased and the pump was attempted to be repositioned. the repositioning failed. support was continued. weaning was successful and the device was explanted. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of vf/vt/af/at and positioning issues has been completed. the cause of the vf/vt/af/at and positioning issues was not determined since product was not returned and lack of clinical information.
9830 | 1220648-2025-27982 | 21938999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21938999&pc=OZD | 04/06/2025 | 05/19/2025 |  | 04/30/2025 |  |  | 04/06/2025 | 04/30/2025 |  |  |  |  | 04/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026635260 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 26 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support and during support, there were suction events. the suction events occurred when the patient was coughing or when the patient awakens. additionally, the pump was repositioned during support. flows improved after repositioning. the patient remains on support. | [5 empty] | The impella device was not received from the customer as it is still supporting the patient. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | D6: explant date unk. the investigation of low pump flow and positioning issues has been completed. low flows: pump not returned. data logs show suction alarms and positioning alarms. no motor current (mc) spikes outside of p-level changes. the cause of the low pump flow issue was most likely improper positioning of the device as clinical stated flows improved after repositioning of the pump. positioning: the cause of the positioning issue was not determined due to lack of clinical information.
9831 | 1220648-2025-27983 | 21939108 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21939108&pc=OZD | 04/07/2025 | 07/02/2025 |  | 04/30/2025 |  |  | 04/07/2025 | 04/30/2025 |  |  |  |  | 04/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026613209 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Tachycardia | Ectopic Heartbeat | Renal Impairment | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient on impella 5.5 had recurrent ventricular tachycardia and ventricular ectopic beats with change of position and manipulation. it was noted of renewed tachycardia and cardioversion. no other information was provided. the procedural outcome was the patient survived the surgery. | Additional information received from the clinical site that the patient experienced aki / kidney failure, that is possibly related to the study impella device. the patient was started on dialysis. |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | B5 additional information received from the clinical site. h6 revised to reflect the additional failure mode. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9832 | 1220648-2025-27984 | 21939463 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21939463&pc=OZD | 04/08/2025 | 05/23/2025 |  | 04/30/2025 |  |  | 04/08/2025 | 04/30/2025 |  |  |  |  | 04/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026644323 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the insertion of impella cp device to a patient undergoing a high-risk percutaneous coronary intervention was unsuccessful. the patient anatomy made it very hard to pass the impella cp through the patient's descending aorta. the impella cp eventually would not advance further than the aortic arch and the decision was made to abort the impella cp implant procedure. there was no report or indication of adverse effects to the patient. | [5 empty] | The investigation of delivery issues has been completed. based on the clinical details the root cause of delivery issue is determined to be patient condition because pump was unable to pass through the tortuous patient anatomy. introducer damage was added failure mode: the root cause of the introducer damage was determined to be patient condition because sheath was kinked during delivery due to patient anatomy. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
9833 | 1220648-2025-27985 | 21939535 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21939535&pc=OZD | 04/08/2025 | 05/30/2025 |  | 04/30/2025 |  |  | 04/08/2025 | 04/30/2025 |  |  |  |  | 04/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026623688 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and shortly after implant, the patient developed hemolysis. the physician did not want to reposition, and the pump was explanted in response to the condition. it is unknown if another impella device was used. however, the patient was reported to be in stable condition following the event. | [5 empty] | The pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. the data was not returned for review. the pump was returned for analysis. visual inspection found no issues on the pump. pump passed hemolysis test. conclusion: the root cause of hemolysis was most likely due to pump was in improper position in addition to low central venous pressure (cvp) patient condition per clinical report and no issue was found on the pump under product analysis.
9834 | 1220648-2025-27986 | 21939563 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21939563&pc=OZD | 04/08/2025 | 04/30/2025 |  | 04/30/2025 |  |  | 04/08/2025 | 04/30/2025 |  |  |  |  | 04/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025591371 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and developed a hematoma at the groin site within a couple of hours of the impella cp support. blood products were given to the patient to treat the condition. it was noted that the patient was escalated to impella 5.5 device mcs. however, the escalation was independent of the reported hematoma. upon removal, the site was closed with sutures to achieve hemostasis. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9835 | 1220648-2025-27987 | 21939658 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21939658&pc=OZD | 04/08/2025 | 05/28/2025 |  | 04/30/2025 |  |  | 04/08/2025 | 04/30/2025 |  |  |  |  | 04/30/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025599003 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Material Puncture/Hole | Tear, Rip or Hole in Device Packaging | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and approximately eight hours into the support bleeding from the red plug was noted. consequently, the impella cp device was explanted and replaced for continued mcs. | [5 empty] | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of product damage (hole in catheter) has been completed. the impella cp was returned for evaluation. upon visual inspection, blood was present in the red handle. a hole in the catheter was present at the 90cm mark. clinical details noted that blood was coming from the red plug after patient was brought back to unit after insertion. the root cause of the product damage (hole in catheter) was most likely a use related issue. h6 medical device problem code 2385 model was erroneously entered on the initial manufacturer device report number 1220648-2025-27987.
9836 | 1220648-2025-27913 | 21924358 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21924358&pc=OZD | 04/11/2025 | 04/29/2025 |  | 04/29/2025 |  |  | 04/11/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026630049 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Ectopic Heartbeat | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. positioning issues were noted. the impella 5.5 was initially pointed towards the mitral valve. the device was able to be turned. echocardiogram found that the device was about 4.5-5 centimeters from the aortic valve. positioning was deeper which caused the patient to experience ectopy. later, there was an impella in aorta alarm. the pump was in the aorta just past the valve. the decision was made to explant the device and the patients outcome was stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9837 | 1220648-2025-27914 | 21924385 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21924385&pc=OZD | 04/10/2025 | 05/05/2025 |  | 04/29/2025 |  |  | 04/10/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026634325 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was on impella cp support and developed a small hematoma, but it was growing. pressure was held and one unit of blood was administered. a decision was made to remove the impella cp due to the hematoma. the patient was stable. the impella cp supported for 10 hours. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9838 | 1220648-2025-27915 | 21924423 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21924423&pc=OZD | 04/05/2025 | 06/11/2025 |  | 04/29/2025 |  |  | 04/05/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025611099 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support, the pump became dislocated into the aorta. the doctor's plan is to explant the pump when activated clotting time is in the normal range. the patient remains on support. | [5 empty] | H6: explant date is unknown the investigation of positioning issues has been completed. clinical states that the pump was pulled out but the reason is unknown. an "impella in aorta" alarms was observed and due to difficult repositioning, the pump was explanted. pump was not returned. therefore, the root cause of the positioning issue was not determined since product was not returned and clinical information was inconclusive since the reason for the repositioning is unknown. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9839 | 1220648-2025-27916 | 21924475 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21924475&pc=OZD | 01/26/2025 | 05/16/2025 |  | 04/29/2025 |  |  | 04/04/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025579839 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 71 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Pseudoaneurysm | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed's long-term outcome and quality indicator (loqi) impella registry that the patient endured left lower extremity ischemia. relationship to pump is "definitely" related. loqi states: presented to er w/ c/o l le pain that started 9 days ago; numbness, discoloration & purple skin color over l distal toes that began 4 days ago. ct- occluded l sfa prior to bifurcation. heparin drip started. morphine prn, tylenol prn, oxydone prn for pain control. 05feb2025- thrombectomy of cfa & sfa. continue monitoring foot- appears to continue demarcating, possible amputation. demarcating worsening to (6) 2025- l bka performed. ordered: dressing changes, maintain knee immobilizer * pain management prn. | A notification was received via abiomed's long-term outcome and quality indicator (loqi) impella registry about a patient who was admitted in for high-risk percutaneous coronary intervention and was placed on the impella cp for support in (b)(6) 2025. the loqi database reported in (b)(6) 2025 a left groin pseudoaneurysm. the patient was treated with heparin, thrombin injection, and vascular surgery was consulted. this was noted 6 days post pump explant. the procedural outcome was the patient survived the surgery. |  |  |  |  | Additional information received from abiomed's long-term outcome and quality indicator (loqi) impella registry and the following fields have been updated based on information received from the clinical site. b2 additional outcome added b5 updated with the additional information h6 health effect - clinical code, health effect - impact code and component code updated. ¿vascular damage peripheral vessel: pseudoaneurysm after explant. the patient was treated with heparin, thrombin injection, and vascular surgery consulted.¿ the root cause of vascular damage peripheral vessel cannot be determined since the product was not returned and insufficient clinical details were provided. ¿limb ischemia: limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition ¿ any concomitant medication ¿ vascular size or variations thereof ¿ detailed description of the symptoms experienced by the patient ¿ physical examination findings, including pulse assessment and visual examination of the affected limb ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension ¿ patient's response to previous treatments or interventions for vascular conditions with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. instructions for use impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9840 | 1220648-2025-27917 | 21924500 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21924500&pc=OZD | 04/11/2025 | 05/28/2025 |  | 04/29/2025 |  |  | 04/11/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026626846 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was attempted to be implanted with an impella 5.5 for mechanical circulatory support. the pump was unable to advance past the innominate during attempted delivery. the implant was aborted and an impella cp was placed for planned support. there was no patient harm. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, could not get past innominate on patient anatomy. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
9841 | 1220648-2025-27918 | 21924551 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21924551&pc=OZD | 04/10/2025 | 06/03/2025 |  | 04/29/2025 |  |  | 04/10/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1438552 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported that an automated impella controller (aic) turned on with an active controller failure alarm outside of patient support. they switched to a backup aic. the patient was never on support when the alarm occurred. | [5 empty] | H.6 health effect - impact code 2645 was inadvertently omitted from the initial manufacturer device report number 1220648-2025-27918. | The investigation of automated impella controller (aic) - controller failure (red alarm) has been completed. data analysis: on the complaint date of april 10, 2025, immediately after the console booted up, there was a ¿controller error (defective optical sensor lamp) ¿ alarm." the optical parameters were abnormal. the user manually shut down and rebooted the console twice. each time after the reboot, the console displayed event. this confirms the reported failure mode. device analysis: this console was tested and the controller error event was reproduced. it was visually confirmed that the lamp was intermittently turning off by itself. when the lamp was off, 4.9 v was measured at the power battery manager (pbm) j3 connector, an average of 4.8 v was measured at the p4 connector of the optical bench, and the lamp resistance measured 5.1 ohms. no issues were found with the analog output from the optical bench. root cause: the root cause of the controller error was a lamp malfunction. there was no controller failure alarm recorded in the case associated with the reported complaint. the controller error was misreported as a controller failure. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
9842 | 1220648-2025-27919 | 21924889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21924889&pc=OZD | 04/05/2025 | 05/14/2025 |  | 04/29/2025 |  |  | 04/05/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026613584 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella 5.5 pump for a surgical patient who was sent to the intensive care unit with an open chest. the patient had been bleeding some from chest tubes but also oozing at line insertion sites. the patient received three units of fresh frozen plasma and two units of packed red blood cells overnight. bleeding was not felt to be related to th impella. bicarbonate was in purge solution. systemic heparin had not been started yet but will continue to monitor bleeding and start when approved by the doctor. weaning was successful and the device was explanted. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of access site bleeding is determined to be patient condition because of the bleeding from multiple location, from chest tube and the access sites. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9843 | 1220648-2025-27920 | 21924965 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21924965&pc=OZD | 04/04/2025 | 05/28/2025 |  | 04/29/2025 |  |  | 04/04/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026615135 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  |  |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support, the patient had a lack of flow to their left lower extremity. the patient did not have pulses prior to impella support. however, it is unknown if the impella contributed to the ischemia. additionally, the access site was resutured due to oozing. furthermore, there were constant suction alarms. the staff discussed disabling the suction alarm as the patient has no limited pulsatility and administration of fluid has not helped. later, weaning was successful and the device was explanted. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of access site bleeding, low pump flow, and limb ischemia has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the access site bleeding - major was use related due to site re-sutured and issue resolved. the cause of the low pump flow issue cannot be determined since the complaint device not returned for analysis and lack of clinical details. no data logs available to review. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9844 | 1220648-2025-27921 | 21925031 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21925031&pc=OZD | 04/11/2025 | 05/22/2025 |  | 04/29/2025 |  |  | 04/11/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025612000 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 for mechanical circulatory support. the patient experienced hemorrhagic stroke during impella 5.5 support. heparin was halted in response to the stroke. support was continued with the implanted pump. the patient was noted to be stable. | [5 empty] | The impella device was not received from the customer as it is still supporting the patient. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock, section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | D6: explant date unknown. the investigation of stroke/cva has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
9845 | 1220648-2025-27922 | 21925056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21925056&pc=OZD | 04/04/2025 | 06/17/2025 |  | 04/29/2025 |  |  | 04/04/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1093078 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the automated impella controller was turned on and a controller error message was received. this complaint is not related to a clinical procedure therefore no patient involvement. | [5 empty] | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. based on the data and device analysis there was no issue found during the case. the device functioned/operated to specification. | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. the console logs from the reported day of event are not consistent with complaint, and there were no boot-up failures or controller error alarms within (or near) specified time frame. console was tested in danvers, and was operating within specification. visual inspection did not reveal any anomalies. console was power-cycled several times, and no alarms upon boot-up occurred. there was no issue found during the case. the device functioned/operated to specification. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
9846 | 1220648-2025-27923 | 21925150 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21925150&pc=OZD | 04/08/2025 | 04/29/2025 |  | 04/29/2025 |  |  | 04/08/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1456795 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that an automated impella controller (aic) had a system self check failure error. this issue occurred during preparation for impella support. no patient or pump information was provided. | [5 empty] | The investigation for the reported system self check failure has been completed. data analysis of the console logs confirm the customer's complaint and the failure mode was reproduced during testing of the customer's console. the root cause of the system self check failure was determined to be printed circuit board module corrosion due to leakage of the electrolyte from the battery failure alarm super caps.
9847 | 1220648-2025-27924 | 21925185 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21925185&pc=OZD | 04/10/2025 | 05/29/2025 |  | 04/29/2025 |  |  | 04/10/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026632639 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was on impella 5.5 support and had some sustained ventricular tachycardia overnight and was shocked twice and reintubated. per the registered nurse, they repositioned the impella overnight, pulled back about a centimeter, and did an echocardiogram to confirm the position. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of positioning issues and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9848 | 1220648-2025-27925 | 21925303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21925303&pc=OZD | 04/08/2025 | 04/29/2025 |  | 04/29/2025 |  |  | 04/08/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025593568 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient implanted with an impella 5.5 exhibited a splenic infarct. the impella was exchanged for a second impella 5.5 pump. later, the patient successfully received a heart transplant and was reported to be in stable condition. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves, ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
9849 | 1220648-2025-27926 | 21925398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21925398&pc=OZD | 04/09/2025 | 09/23/2025 |  | 04/29/2025 |  |  | 04/09/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026615201 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Fever | [12 empty] | Other | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant also reported that during support some position unknown alarms occurred at the time of walking or coughing, however the alarms resolved immediately. | It was reported that while implanted with an impella 5.5, the patient developed a fever of 101.7 degrees fahrenheit. all culture results were negative. the patient survived the event. |  |  |  |  | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the root cause of the optical signal issue was not determined since product was not returned, data logs were not recovered, and there were insufficient clinical details. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial. manufacturer device report number 1220648-2025-27926. d10 concomitant medical products and therapy dates updated. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of infection/sepsis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
9850 | 1220648-2025-27927 | 21925418 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21925418&pc=OZD | 04/08/2025 | 04/29/2025 |  | 04/29/2025 |  |  | 04/08/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026636135 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported while in the intensive care unit, the impella 5.5 was pulled back under echocardiogram at bedside with stable waveform. it was further reported that heparin was on hold due to splenic and right chest wall hematoma. the patient¿s chest was washout. the patient's outcome is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9851 | 1220648-2025-27928 | 21925519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21925519&pc=OZD | 04/04/2025 | 05/15/2025 |  | 04/29/2025 |  |  | 04/04/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026643188 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was further reported that care was withdrawn and the patient expired seven minutes after explant. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the patient experienced leg ischemia. there was no flow past the impella sheath. a femorofemoral bypass was performed. the patient's pulse recovered. the physician considered escalating the patient to an impella 5.5 or axillary impella placement. |  |  |  |  | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. limb ischemia: "limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition. any concomitant medication. vascular size or variations thereof. detailed description of the symptoms experienced by the patient. physical examination findings, including pulse assessment and visual examination of the affected limb. diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. laboratory test results, including blood tests for markers of inflammation or clotting disorders. any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined." b2 revised to reflect the patient's outcome of death. b5 revised to reflect the patient outcome of death. d6b explant date added. h1 revised to death. h6 health effect - impact code 1802 added to reflect patient outcome. | The impella device was not received from the customer as it is still supporting the patient. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9852 | 1220648-2025-27929 | 21925668 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21925668&pc=OZD | 04/09/2025 | 07/31/2025 |  | 04/29/2025 |  |  | 04/09/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1554466 |  | 0042-0010-EU |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 72 YR |  | Unknown |  |  |  | Death | Yes | Yes | Application Program Problem | [6 empty] | Hemodynamic instability | [12 empty] | Death | [5 empty] | IMPELLA CP SN (B)(6). | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during use on a patient the automated impella controller failure "verification failed, exchange the console" occurred. because the hospital did not have a backup console, the aic was switched off and could not switched on again. the patient was in a very critical condition with cardiac arrest and expired. no patient information was provided. | [5 empty] | The investigation of automated impella controller (aic) - controller failure (red alarm) has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. a2 patient age updated. b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27929. | The automated impella controller device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9853 | 1220648-2025-27930 | 21925685 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21925685&pc=OZD | 04/11/2025 | 05/13/2025 |  | 04/29/2025 |  |  | 04/11/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1366755 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The automated impella controller (aic) would not recognize the purge disc during set up. the aic was replaced. no patient harm was reported. | [5 empty] | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. the logs from the complaint date revealed that the console issued purge disc not detected alarm several times. the console was examined, and a broken purge disc flag was identified. the purge disc flag was observed to be broken and was the root cause of the reported purge disc not detected alarm issue. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
9854 | 1220648-2025-27932 | 21925889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21925889&pc=PYX | 04/09/2025 | 06/03/2025 |  | 04/29/2025 |  |  | 04/09/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | RP FLEX |  |  | ABIOMED, INC |  |  | US |  | 2025560601 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male | 63 KG | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [5 empty] | ECMO, RVAD. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient prepped for implantation of an impella rp flex exhibited a plasma free hgb value of 140 mg/dl and the patient experienced ventricular tachycardia. the patient was shocked, the rp flex implantation was completed, and a permanent right ventricular assist device was placed. the patient survived. | [5 empty] | The investigation of hemolysis has been completed. the pump was returned cut at the catheter. pump was borescoped and no biomaterial was found. no other damages or abnormalities were found. data logs show no motor current spikes outside of p-level changes. the root cause of the hemolysis was not determined since the pump was returned cut and insufficient clinical details provided. d9 device returned to manufacturer date added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9855 | 1220648-2025-27931 | 21925908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21925908&pc=OZD | 04/10/2025 | 04/29/2025 |  | 04/29/2025 |  |  | 04/10/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1454057 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | [9 empty] | The complainant reported that an automated impella controller experienced a battery failure outside of patient support. there was no patient involvement. | [5 empty] | The investigation has been completed. the console was returned for investigation. the console logs were consistent with what was reported in the complaint. a single instance of battery failure (transport connection blown/missing) alarm was observed in the logs. the console was tested. the reported battery failure was unable to be reproduced while testing. results of running battest utility on telnet revealed that there were no evident anomalies with the consoles batteries. the reported battery failure was determined to be use related. it is likely that the console was booted up while the console was connected to alternating current power and the battery switch was disengaged since the issue occurred upon first boot of the console after preventative maintenance by field service.
9856 | 1220648-2025-27933 | 21925942 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21925942&pc=OZD | 04/09/2025 | 05/27/2025 |  | 04/29/2025 |  |  | 04/09/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026644891 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient implanted with an impella cp experienced oozing from both their venous and arterial sites during support. manual pressure was applied and a suture was placed to manage the condition. the patient survived. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the patient had some oozing from both venous and arterial site. the root cause of the access site bleeding issue was most likely patient condition related. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9857 | 1220648-2025-27934 | 21926102 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21926102&pc=OZD | 04/04/2025 | 05/20/2025 |  | 04/29/2025 |  |  | 04/04/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026639443 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during support it was reported that there was slight oozing from impella site. they were able to redress and re-angle match. however, it required the physician to come and re-suture the blue hub. pressure was applied to help with clotting as activated clotting times remained elevated. re-angle was matched and oozing improved. the impella was successfully weaned and explanted. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of access site bleeding is determined to be use issue because the bleeding was resolved by re-suturing impella and matching the angle. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) , ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9858 | 1220648-2025-27935 | 21926146 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21926146&pc=OZD | 04/09/2025 | 05/28/2025 |  | 04/29/2025 |  |  | 04/09/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026622741 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Priming Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during setup, the impella cp pump would not properly prime and was giving multiple impella stopped alarms. a new impella cp device was subsequently placed for patient support. | [5 empty] | The investigation of priming issue and pump did not start has been completed. pump was returned for investigation. data log was not provided for the case run. no physical abnormalities were reported on returned product. pump was primed without issue and ran as per the specification. the cause of the priming issue was not determined since no issue was noted on returned product and sufficient clinical information was provided. the cause of the pump did not start issue was not determined since no issue was noted on returned product and sufficient clinical information was provided. h6 medical device problem code 4040 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27935. h10 narrative erroneously stated, "device was not returned" on the initial manufacturer device report number 1220648-2025-27935. | The impella cp was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock, section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
9859 | 1220648-2025-27936 | 21926311 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21926311&pc=OZD | 04/04/2025 | 06/19/2025 |  | 04/29/2025 |  |  | 04/04/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1577697 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) had placement signal issues and a brief period when the aic disconnected from the impella connect. this is one of two related reports for the patient associated with this event. | [5 empty] | The investigation of optical signal issue has been completed. based on the data and device analysis there was no issue found with the console, during the case. the device functioned/operated to specification. the patient expired while on support and is captured on manufacturer device report number 1220648-2025-27937 which is against the pump. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
9860 | 1220648-2025-27937 | 21926480 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21926480&pc=OZD | 04/04/2025 | 05/28/2025 |  | 04/29/2025 |  |  | 04/04/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026629167 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Hispanic |  | Death | Yes | Yes | Positioning Failure | Restricted Flow rate | Device Sensing Problem |  |  |  |  | Arrhythmia | Ventricular Fibrillation | Ectopic Heartbeat | [10 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was brought to the catheterization lab today for a planned percutaneous coronary intervention of the saphenous vein graft and selective saphenous vein graft to the right coronary artery. the left subclavian was shockwaved to make the left radial approach feasible. saphenous vein graft to obtuse marginal was completed. during intervention, the patient had ventricular fibrillation arrest requiring intubation and advanced cardiac life support. the impella cp was placed as a bailout. the complainant reported that impella cp had ongoing suctions while in use. the impella was not able to go above p1 without suction. the impella position and preload status were reportedly good. the impella flows appeared to be normal. the physician verified the position and suction alarms were turned off. the patient was also supported with extracorporeal membrane oxygenation. it was further reported that on day five of impella cp support, the patient was taken to the lab for repositioning under fluoroscopy. the impella only moved forward at 2.7 centimeters the physician could not move it any more forward and opted to keep placement as is. the next day, the impella cp was still in poor position and was unable to establish full flows. the patient had runs of torsade and ventricular fibrillation. the impella was removed in an attempt to relieve ectopy. after the impella was removed, the patient expired. | [5 empty] | The investigation of positioning issues, low pump flow, vf/vt/af/at, and optical signal issue has been completed. the product was not returned for investigation. since the clinical reports that attempts to reposition the pump didn't resolve the issue was confirmed by data logs, but weaning extracorporeal membrane oxygenation (ecmo) flows aligned with pump flow improvement on data logs, the root cause of the low pump flow was determined to most likely be patient condition. additionally, there were no abnormalities noted in data logs around the time of the reported complication, so the root cause of the positioning issues could not be determined. as the necessary information was not provided, the root cause of the vt/vf was not determined. since data log review showed the motor current to quite low during both phases of the cardiac cycle even though the placement signal (ps) was pulsatile and within specification, the root cause of the optical signal issue was determined to most likely be a suction condition. the suction condition was most likely due to the patient's condition, per the investigation into low pump flow. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9861 | 1220648-2025-27938 | 21926487 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21926487&pc=OZD | 04/09/2025 | 05/29/2025 |  | 04/29/2025 |  |  | 04/09/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026637073 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Failure to Advance | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  | No Clinical Signs, Symptoms or Conditions | [18 empty] | TAVR, BAV. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that implantation of an impella cp was prepared for a patient who presented for primary percutaneous coronary intervention as part of the transcatheter aortic valve replacement. the patient was known to have a heavily calcified aorta. following bicuspid aortic valve repair, the impella cp was implanted in normal fashion. it was then found the patient had an extremely short aortic root that caused the impella cp to kink on entry to the left ventricle. the 0.018 guidewire was removed and the impella was started, however the pump failed to achieve flows higher than 1.5l/min with suction in auto. the impella cp was then successfully explanted from the patient without complication. the patient survived. | [5 empty] | The investigation of delivery issues, low pump flow, and product damage (cannula kink) has been completed. reported patient having a short aortic root and a heavily calcified aorta and observing a kink distal to the outflow cage of the cannula. two cannula kinks were found distal to the outflow cage. delivery issues - the root cause of the delivery issue was most likely patient condition related since the patient had a short aortic root and a heavily calcified aorta. low flow - reported being unable to run pump past 1.5l/min with suction alarms present in auto. data logs show pump running at p-2 on auto with flow around 1.3l/min and suction alarms. the root cause of the low flow was most likely a kinked cannula since a cannula kink was reported and found on the returned product. cannula kink - the root cause of the product damage was a kinked cannula most likely due to patient condition since the patient had a short aortic root and a heavily calcified aorta. h6 medical device problem codes 1248, 2682, and 2889 were inadvertently omitted on the initial manufacturer device report number 1220648-2025-27938. | Upon completion of the investigation a supplemental report will be submitted. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk pci states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
9862 | MW5169616 | 21926896 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21926896&pc=OZD | 04/19/2025 | 04/22/2025 |  | 04/29/2025 |  |  |  | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | REF 0048-0003 |  |  | ABIOMED, INC. |  |  |  |  | 1726123121576121 |  |  |  | Yes | [5 empty] | No | Health Professional |  |  | Voluntary report | Yes | Risk Manager | [6 empty] | 87 YR |  | Male |  |  |  | Malfunction | No | Yes | Material Invagination | Material Separation | Deformation Due to Compressive Stress | Physical Resistance/Sticking |  |  |  | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | [9 empty] | Provider obtained rfa (radiofrequency ablation) access using ultrasound and verifying access location with rfa angiogram. the 14 french impella sheath was inserted following several rfa dilitations. he then advanced the impella cp catheter but was met with resistance. attempts to advance the pump led to prolapsing of the shaft due to kinking of the impella sheath from severe tortuosity of the patient's anatomy. the impella sheath and catheter were then removed as a unit. after the device was removed it was observed that the tip of the impella catheter had broken off. nothing was left behind in the patient's body. no harm to the patient.
9863 | 1220648-2025-27939 | 21926952 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21926952&pc=OZD | 04/09/2025 | 06/04/2025 |  | 04/29/2025 |  |  | 04/09/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026641435 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female | 75 KG |  | White | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | Deformation Due to Compressive Stress |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that following implantation of an impella cp the pump persistently migrated to the papillary muscle despite multiple attempts to reposition. the pump began to experience suction issues and it was noted that the catheter had kinked. a large hematoma was also observed at the access site, for which femstop was placed and blood was given. the decision was ultimately made to remove the pump and place a new device in the opposite femoral artery for ongoing support. the patient survived. | [5 empty] | The investigation of low pump flow, access site bleeding, positioning issues, and product damage (cannula kink) has been completed. the data logs confirmed low pump flow when pump started and remained to the rest of the case that triggered auto suction response and suction alarms. product was not returned for review. based on clinical description and data review, the root cause of low flow was most likely due to kinked cannula. the root cause of cannula kink was most likely patient condition related since the patient had a small ventricle. the root cause of access site bleeding was most likely patient movement related during transportation to another hospital. the root cause of positioning issue was most likely patient condition related since the patient had a small ventricle resulted in difficult position of the pump. h6 medical device problem code 2889 and component code 3038 were inadvertently omitted on the initial manufacturer device report number 1220648-2025-27939. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
9864 | 1220648-2025-27940 | 21926959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21926959&pc=OZD | 04/03/2025 | 05/29/2025 |  | 04/29/2025 |  |  | 04/04/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025579592 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Fluid/Blood Leak | [6 empty] | Ischemia | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reportedly arrived onsite with the patient already in the intensive care unit (icu). intraoperatively, a mitral valve repair was added in addition to coronary artery bypass graft times four. the patient decompensated coming off bypass and an intraaortic balloon pump (iabp) was inserted. iabp did not provide any hemodynamic support, so the physician was asked to assist in inserting an impella cp in place of the iabp. the patient was in the icu on five times inotropes/pressors, skin mottled and was actively coding. during preparation, the automated impella controller (aic) alarmed that the purge system was open when it was not, after the white power cord was plugged in, the aic kept showing directions to connect the grey cable arrow to arrow. a controller failure alarm appeared. the aic was replaced and the same issue occurred. the purge was de-aired and the aic and pump were operational. the clinical representative did not know if this was an aic issue or a pump issue. the patient outcome at explant noted that the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation of limb ischemia and purge leak ¿ pump has been completed. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. there was no purge leak-pump issue found during the case since since purge leak issue was resolved after successful priming with second purge caseate. the device functioned/operated to specification during the field run and returned product confirms no issue on purge leak-pump. | The pump and purge cassette was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: cleaning, ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
9865 | 1220648-2025-27941 | 21927178 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21927178&pc=OZD | 04/03/2025 | 11/20/2025 |  | 04/29/2025 |  |  | 04/04/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1175558 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Read Input Signal | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP | [18 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | During preparation, the automated impella controller (aic) alarmed that the purge system was open when it was not, after white power cord was plugged in. the aic kept showing directions to connect the grey cable arrow to arrow. a controller failure alarm appeared. the aic was replaced and the same issue occurred. the purge was de-aired and the aic and pump were operational. the clinical representative did not know if this was an aic issue or a pump issue. | It was reported that the patient expired while on support. the medical safety officer reviewed and determined "product malfunction. the event occurring during preparation with a brief delay in start of the support. another automated impella controller (aic) was available and used. there was no interruption in support and no change, compromise in therapy once started." the patient outcome of death in captured in manufacturer device report 1220648-2025-27940. |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation was completed: data analysis: review of the data logs did confirm skipped steps on boot (steps.png) and the presence of red "purge system open" alarms (purge.png) on the complaint date. no "controller failure" alarms were found when reviewing the data logs. aic - purge issue - other. the cause of the issue was use related. aic - controller failure (red alarm). no problem found. the cause of the issue was use related. d4: corrected the catalog number. h3: added information regarding the investigation status. h6: added codes per the results of the investigation. h10: added the results of the investigation. | B5 mso statement was inadvertently omitted on manufacturer device report 1220648-2025-27941. it was reported that the patient expired while on support. the medical safety officer reviewed and determined "product malfunction. the event occurring during preparation with a brief delay in start of the support. another automated impella controller (aic) was available and used. there was no interruption in support and no change, compromise in therapy once started." the patient outcome of death in captured in manufacturer device report 1220648-2025-27940.
9866 | 1220648-2025-27942 | 21927281 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21927281&pc=OZD | 04/09/2025 | 04/29/2025 |  | 04/29/2025 |  |  | 04/09/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026623444 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that an impella cp was implanted to support a percutaneous coronary intervention (pci) procedure. after the pci, the patient was successfully weaned from the impella. two perclose were deployed successfully but distal flow was absent after the second perclose. a balloon was used to reestablish flow. the patient is stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9867 | 1220648-2025-27943 | 21927418 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21927418&pc=OZD | 04/09/2025 | 06/03/2025 |  | 04/29/2025 |  |  | 04/09/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026632968 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female | 72 KG | Non Hispanic | White | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella cp was implanted in a patient with acute myocardial infarction/cardiogenic shock. when the patient's foley catheter was adjusted the pump came out of appropriate position within the left ventricle. the left ventricle is small and as such the position was still shallow, but the pump was not replaced. the patient remained on support. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. the data logs confirmed pump was in improper position corresponded with clinical description that triggered alarms impella position in aorta. the issue then was resolved and the pump ran without issue to the rest of the case. no other issues were found on the logs. the root cause of positioning issue was most likely patient condition related since the patient had a small ventricular size. d6b explant date added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9868 | 1220648-2025-27944 | 21927450 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21927450&pc=OZD | 04/09/2025 | 04/29/2025 |  | 04/29/2025 |  |  | 04/10/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026638875 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that while a patient was implanted with an impella cp, the patient exhibited bleeding at the access site. the vascular surgeon placed a tobacco pouch suture and treated the groin with a tabotamp. three units of blood products were given to the patient. care was withdrawn and the patient expired. the attending physician stated he does not view the impella therapy as a cause or in connection with the patient's death. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9869 | 1220648-2025-27945 | 21927593 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21927593&pc=OZD | 04/17/2025 | 04/29/2025 |  | 04/29/2025 |  |  | 04/17/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported a screw fell out from the aic during preparation. after the impella was removed, the hospital staff noticed that the screw had fallen off. on (b)(6) 2023 abiomed was made aware that the screw that fell out was a part of the notch of the lid that opens when installing the purge housing. the affected part originally had 4 screws, but 1 fell out during the aic preparation, 2 were loose, and 1 was already missing (no knowledge of when or how it was missing). since the engineer have never touched this part during past repairs, they suspect that it might have been loose since manufacturing. no known patient harm or injury. | [5 empty] | The investigation for the console (aic) purge issue has been completed. no data relevant to the failure mode was found in the logs. the console was returned and upon evaluation there were two screws missing and two loose. the cause of the aic issue was two loose screws in the door latch.
9870 | 1220648-2025-27946 | 21927673 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21927673&pc=OZD | 04/10/2025 | 06/12/2025 |  | 04/29/2025 |  |  | 04/10/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026623637 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient had an impella cp pump placed for support during a high risk cardiac procedure. after implant, there was no appropriate flow due to an observation made of the kinked pump. the kink was located in front of the motor. the pump was explanted and replaced without any harm or injury. | [5 empty] | The investigation of product damage (cannula kink) and low pump flow has been completed. data was not returned for review. visual inspection found a kink on cannula about 15 mm from of cage and cannula bonding site. no other issues were found on the rest of the pump. the root cause of low flow was most likely kinked cannula as per clinical information and observed on the returned product. the root cause of kinked cannula was unable to be determined as limited clinical details were provided for investigation. d9 device returned to manufacturer date added. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9871 | 1220648-2025-27947 | 21927695 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21927695&pc=OZD | 04/22/2025 | 04/29/2025 |  | 04/29/2025 |  |  | 04/22/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026633630 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient developed hemolysis during impella cp support. the patient was escalated to impella 5.5 support in response to the noted condition. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9872 | 1220648-2025-27949 | 21927847 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21927847&pc=OZD | 04/09/2025 | 06/12/2025 |  | 04/29/2025 |  |  | 04/09/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1218792 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during use of the console that the aic alerted "controller error". the aic was switched to a alternate aic. | [5 empty] | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. data analysis: on the complaint date, (b)(6) 2025, during support, the console displayed the ¿controller error (loss of comm (pbm1)) ¿ alarm,¿ event 1023. this confirms the reported failure mode. device analysis: this console was tested and corrosion on the power batter manager (pbm) board was visually confirmed. root cause: the root cause of the controller error was determined to be pbm corrosion resulting from the leakage of electrolyte from the battery failure alarm supercapacitors. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
9873 | 1220648-2025-27948 | 21927868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21927868&pc=OZD | 04/11/2025 | 06/03/2025 |  | 04/29/2025 |  |  | 04/11/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1127739 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | Key or Button Unresponsive/not Working | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the automated impella controller (aic) screen went black with no response to buttons. after shutting the device off and back on, the console was showing a normal screen. this issue was discovered during use. no patient harm was reported. | [5 empty] | The investigation of automated impella controller (aic - display issue has been completed. there were no log entries that could confirm the reported black screen. the user requested a shutdown at 09:47:55, but the shutdown request was not confirmed, which could indicate issues with the soft key/ rotary push button knob interface. finally, at 10:12:55, the console auto-shutdown occurred due to the lack of user interface interaction, as per its design. the console was returned for investigation and was tested. there were no kbd communication errors reproduced even after performing 100 power cycles. the console was run with known-good pump and purge cassette for couple of hours, and no issues were observed. additionally, the internal circuitry and cables of the console were examined, and everything appeared to be in good condition. root cause: the cause of the kbd communication errors was not determined due to the inability to reproduce. the reported aic display issue was not confirmed via data logs and was not reproduced during testing. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
9874 | 1220648-2025-27950 | 21928128 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21928128&pc=OZD | 04/09/2025 | 04/29/2025 |  | 04/29/2025 |  |  | 04/09/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026632521 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the patient was on impella 5.5 support and during support, the patient had a hematoma at the access site. the hematoma was drained and a jp drain was placed. the patient remained on support. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock states the following: section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9875 | 1220648-2025-27951 | 21928441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21928441&pc=OZD | 04/11/2025 | 04/29/2025 |  | 04/29/2025 |  |  | 04/11/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026643193 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. during support, it was reported that the impella cp was running at p7 and had a sudden drop in flows and motor current. the left ventricle went negative on systolic/diastolic. they were unable to resolve. the pump was down to p2 and was still suctioning with average flows of 0.5 liters. echocardiogram confirmed proper positioning. a liter of fluid was administered. the right ventricle was fully functioning. the decision was made to explant the impella cp. the patient condition was noted to be stable. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
9876 | 1220648-2025-27952 | 21928703 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21928703&pc=OZD | 04/09/2025 | 06/03/2025 |  | 04/29/2025 |  |  | 04/09/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563876 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that impella cp was placed emergently on a 44-year-old male patient. hemolysis was present on labs. bedside echo showed depth at 5.5 cm. impella outlet on aortic valve. physician notified of hemolysis and repositioned at bedside. it was further reported that the family decided to withdraw care, and the patient expired. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of hemolysis and positioning issues has been completed. the impella device was not returned for evaluation. hemolysis: clinical states that the pump was placed with outlet on the aortic valve causing labs to hemolyze. pump was repositioned and hemolysis resolved. pump was not returned for investigation. data logs show "impella in ventricle" alarms and placement signal was spread and no other abnormalities were observed. therefore, the root cause of the hemolysis issue was improper positioning of the device due to improper placement technique. positioning issue: clinical states that the pump was implanted and was found to facing the aortic valve which was then repositioned. data logs show initial "impella in ventricle" alarms at 15:22 on (b)(6) 2025 with ventricular waveform which resolve shortly after. "impella in ventricle" alarms return again at about 17:00 and are resolved once again by 18:00. pump was not returned. the root cause of the positioning issue was not determined since the pump was found to be deep upon arrival to the icu and unknown what caused it to go deep. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the ifu states: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿
9877 | 1220648-2025-27953 | 21928817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21928817&pc=OZD | 04/18/2025 | 04/29/2025 |  | 04/29/2025 |  |  | 04/18/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026617598 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient with a history of ventricular fibrillation (vf) and was on impella cp support. bleeding from the mediastinum occurred possibly due to cardiopulmonary resuscitation, and thoracotomy was performed to stop bleeding. hemodynamics remained unstable thereafter and vf recurred. as a result, the patient's life was difficult to save and he died. the patient's cause of death was fatal arrhythmia. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
9878 | 1220648-2025-27955 | 21929183 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21929183&pc=OZD | 04/23/2025 | 04/29/2025 |  | 04/29/2025 |  |  | 04/23/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant in japan reported information was obtained from a literature that a patient was on impella cp support. while on support, the impella cp came out of position during the percutaneous coronary intervention and before admitted back to the intensive care unit. the impella was successfully pushed back to just above the aortic valve (av) gently but could not be advanced beyond the av because the av was closed due to myocardial stunning. finally, the transradial bav balloon proceeded at the av position and the impella slid into the left ventricle simultaneously with balloon deflation. the patient's outcome is unknown. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. a.1, a.2, a.4 and a.5 are unknown. d.4 serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. e.1 name prefix/title, first name and last name are unknown. h.4 device manufacture date is unknown.
9879 | 1220648-2025-27956 | 21929439 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21929439&pc=OZD | 04/09/2025 | 06/24/2025 |  | 04/29/2025 |  |  | 04/09/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2025570056 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient had an attempted impella 5.5 implant and the pump was not able to be inserted due to patient anatomy. the implant was aborted. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock states the following: section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, impella 5.5 implant and the pump was not able to be inserted due to patient anatomy.
9880 | 1220648-2025-27957 | 21929999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21929999&pc=OZD | 04/09/2025 | 12/03/2025 |  | 04/29/2025 |  |  | 04/09/2025 | 04/29/2025 |  |  |  |  | 04/29/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1786118 |  | 0042-0040-EU |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an console (aic) suddenly went blank during impella support. the perfusionist pushed the impella cable and power cord firmly towards the aic housing, but now a console defective and impella defective alarm appeared. impella restart was initially not possible. the aic was swapped and the issues resolved. there was no patient harm. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
9881 | 1220648-2025-27869 | 21916970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21916970&pc=OZD | 04/03/2025 | 04/28/2025 |  | 04/28/2025 |  |  | 04/03/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026625244 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and "some" oozing was noted at the site which was noted to have resolved. femstop was placed proximal to the impella insertion site. anticoagulants were on hold due to bleeding at the site. three days following, oozing at the site was noted again, and it was captured "now resolved. no systemic anticoagulation." the status of the patient as a result of the event was indicated as unknown. however, the latest update noted another three days later, the patient was successfully weaned, and the impella cp device was explanted with a survived outcome. | [5 empty] | The impella pump was diposed after explant from the patient by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9882 | 1220648-2025-27870 | 21917023 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21917023&pc=OZD | 04/03/2025 | 04/28/2025 |  | 04/28/2025 |  |  | 04/03/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025603458 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Hematoma | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and experienced an access site hematoma and pump malposition. the patient would subsequently expire. a patient with atrial fibrillation with rapid ventricular response on admission had a heparin drip on. one day prior to admission, the patient underwent a stress test which revealed anterior ischemia. overnight the patient went into ventricular tachycardia arrest with 10-minute down time. the patient was intubated and started on levophed and vasopressin. the following day, left heart catheterization revealed no need for intervention but ejection fraction was less than 10% on high dose pressors. a decision was made to place a pulmonary artery catheter and impella cp device for rest in the unit. the impella cp was successfully implanted and following removal of the arterial line, a hematoma was present at access site. systemic heparin was discontinued, new forward tension sutures were placed, and fem stop was applied to address the hematoma. following, impella position in aorta alarm triggered. the team did not want to give volume or turn impella pump flow down with suction alarms. an echocardiogram was completed and the impella was measuring 5.7 centimeters below the aortic valve. the impella cp pump was repositioned at bedside with echocardiogram guidance and thereafter measured 3.6 centimeters below the aortic valve. the patient was noted to be still hypotensive with tachycardia. electrocardiogram was obtained for possible cardioversion and decrease in milrinone was noted. approximately two days later, the patient would expire. care was withdrawn. | [5 empty] | The impella device was discarded by the customer and therefore, ¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9883 | 1220648-2025-27871 | 21917055 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21917055&pc=OZD | 04/06/2025 | 04/28/2025 |  | 04/28/2025 |  |  | 04/06/2025 | 04/28/2025 |  |  | 12/31/2025 |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025442919 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient developed limb ischemia distal to impella insertion site. no information about intervention was noted. however, care was withdrawn and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
9884 | 1220648-2025-27872 | 21917095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21917095&pc=OZD | 04/06/2025 | 05/27/2025 |  | 04/28/2025 |  |  | 04/06/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026619120 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump inserted to support a patient admitted in with acute myocardial infarction/cardiogenic shock. the pump was supporting and signs of hemolysis were observed. hemolysis improved overnight and urine color turned yellow. weaning was successful. the device was explanted and the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9885 | 1220648-2025-27873 | 21917153 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21917153&pc=OZD | 04/06/2025 | 04/28/2025 |  | 04/28/2025 |  |  | 04/06/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026622039 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient developed some oozing at their access site following impella cp catheter manipulation. manual pressure was initially applied with two units of packed red blood cells given. the patient was hemodynamically unstable so the decision was made to explant the catheter. the vessel was found to be damaged and the bleed was surgically managed following explant. the procedural outcome was the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9886 | 1220648-2025-27874 | 21917238 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21917238&pc=OZD | 04/07/2025 | 09/17/2025 |  | 04/28/2025 |  |  | 04/07/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1814730 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 | [18 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was further reported the patient expired while on support. the medical safety officer reviewed and determined "automated impella controller (aic) -to-aic exchange brief interruption. however, the patient expired from complications not associated with the aic." the patient death is captured in related manufacturer device report 1220648-2025-28059. | It was reported that an automated impella controller (aic) would not detect the purge disc. perfusion attempted to insert the purge disc. the purge disc was not detecting a "click". the purge disc was confirmed by representative onsite to have been inserted properly. the aic would still not detect the purge disc. the aic was switched out and the new aic worked perfectly fine. there was no adverse impact to the patient reported. |  |  |  |  | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. the logs from the complaint date revealed that the console issued purge disc not detected alarm several times. the console was examined, and a broken purge disc flag was identified. the purge disc flag was observed to be broken and was the root cause of the reported purge disc not detected alarm issue. d9 date device returned to manufacturer added. | B5 mso statement was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27874. it was further reported the patient expired while on support. the medical safety officer (mso) reviewed and determined "automated impella controller (aic) -to-aic exchange brief interruption. however, the patient expired from complications not associated with the aic." the patient death is captured in related manufacturer device report 1220648-2025-28059. | The automated impella controller was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9887 | 1220648-2025-27875 | 21917336 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21917336&pc=OZD | 04/07/2025 | 05/19/2025 |  | 04/28/2025 |  |  | 04/07/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026625357 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 90 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Difficult to Insert | Defective Device |  |  |  |  | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during the initial implant attempt of an impella cp pump, the introducer kinked and subsequently prevented advancing the pump into position. a cook sheath was placed which allowed for successful implant. following implant, it was noted that the sheath was occluding blood flow to the leg. a distal perfusion catheter was placed to treat the condition. after one day of support, the pump generated suction events following repositioning. escalation to an impella 5.5 pump was discussed, but due to the patient's age and lack of durable options, the decision was made to move the patient to comfort care and the patient later passed away. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of difficulty inserting/removing pump through introducer, introducer damage ¿ mechanical, limb ischemia, and low pump flow has been completed. the impella device was not returned for evaluation. due to lack of data logs or product returned, the root cause of the low pump flow cannot be determined. the root cause of the introducer damage - mechanical is most likely due to a sheath kink due to the patients tortuous vessels. the root cause of the difficulty inserting/removing pump through introducer is most likely due to the sheath kink. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention: section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
9888 | 1220648-2025-27876 | 21917434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21917434&pc=OZD | 04/07/2025 | 05/22/2025 |  | 04/28/2025 |  |  | 04/07/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025608840 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella cp and developed a hematoma at the access site. one unit of blood products were given. support was withdrawn and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of access site bleeding is determined to be anticoagulation management because the activated clotting time (act) values (act 228) were noted be higher the therapeutic range at the time of the bleeding and the heparin was kept on hold to mange the bleed.
9889 | 1220648-2025-27877 | 21917448 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21917448&pc=OZD | 04/04/2025 | 05/30/2025 |  | 04/28/2025 |  |  | 04/04/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025582252 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient undergoing an ablation was implanted with an impella cp device for mechanical circulatory support and a hematoma developed after leaving the catheterization lab. heparin drip was running and act was greater than 200. mean arterial pressure at time the hematoma was also over 110. manual pressure held and team elected to place femstop. it is unknown if the heparin was withheld. of note there were positive distal pulses during this time. the following day, labs hemolyzed. echocardiogram noted impella cp device was in "good" position. however, left and right ventricles were noted to be "severely" dilated with a significantly decreased left ventricle ejection fraction between 20-25%. minimal, if any, urine output was noted even with lasix challenge. a pa catheter was placed, and the patient was started on continuous renal replacement therapy to see if hemolysis would improve. the patient was reported to be in stable condition following the event. following, the patient was transferred to a higher level of care facility and upon transfer it was noted the impella cp device was immediately removed. the patient was noted to have survived the explant. | [5 empty] | The investigation of hemolysis and access site bleeding has been completed. the impella device was not returned for evaluation. the data logs showed pump was started and run without issue on p-9 and p-5. about 9 hours into the case, low flow occurred on p-4 but no alarm was triggered. this event remained to the rest of the case. placement signal observed to be trending down throughout case. pump positioning reported to be good on echo and no motor current spikes observed in the logs. note: add failure mode low flow since low flow occurred when pump was run on p-4. product was not returned for review. based on clinical description and data review, the root cause of hemolysis was most likely patient condition since the patient had severely dilated left and right ventricles and right ventricle dysfunction. the root cause of access site bleeding was most likely anticoagulation management related since activated clotting time was greater than 200 when bleeding happened. the root cause of low flow was most likely patient condition since the patient had severely dilated left and right ventricles and right ventricle dysfunction. g1 reporting contact email was erroneously entered on the initial manufacturer device report number 1220648-2025-27877. h6 health effect - clinical code 1884 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27877. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9890 | 1220648-2025-27878 | 21917492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21917492&pc=OZD | 04/07/2025 | 05/28/2025 |  | 04/28/2025 |  |  | 04/07/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026634601 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that while implanting an impella cp, a patient developed hemolysis. the device was repositioned and the doctor lowered the p level. continuous renal replacement therapy was started and support was continued. the patient survived. | [5 empty] | The investigation of hemolysis has been completed. data review: insufficient data from 4/9/25 - 4/10/25 was returned for review and revealed no issue for that period of time. product was returned in damaged condition (catheter was cut off) for review. visual inspection found no issue on the pump. hemolysis test was unable to be conducted on the damaged pump. conclusion: the root cause of hemolysis was most likely patient condition related since hemolysis was resolved post continuous renal replacement therapy (crrt) and no issue was found on the pump under product analysis. d9 device returned to manufacturer date added. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9891 | 1220648-2025-27879 | 21917507 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21917507&pc=OZD | 04/04/2025 | 04/28/2025 |  | 04/28/2025 |  |  | 04/04/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025605357 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 96 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Defective Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient undergoing high-risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support (mcs). during the impella mcs, the anticontamination sleeve was difficult to pull down to the red impella plug. the anticontamination sleeve was noted to be stuck to the catheter. the support was noted to be uninterrupted, and there was no harm to the patient as a result of this event. pci was performed. shockwave to left anterior descending artery (lad), intravascular ultrasound before and after the patient received one drug-eluting stent to the proximal diagonal 1 artery and three drug-eluting stents to the distal, mid, proximal lad. after the pci. the impella cp device was weaned and explanted. hemostasis was achieved. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9892 | 1220648-2025-27880 | 21917614 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21917614&pc=OZD | 04/07/2025 | 05/28/2025 |  | 04/28/2025 |  |  | 04/07/2025 | 04/28/2025 |  |  | 08/31/2025 |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024380592 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Renal Failure | [10 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient presented with constrictive pericarditis, right coronary artery occlusion, and a large ventricular septal defect (vsd). the vessel perforated while attempting intervention. the vessel was coiled and the patient was stabilized and worked up for vsd closure. the patient was implanted with bipella (impella 5.5/impella rp). during support, the patient experienced bleeding, renal failure, anuresis, and hemolysis. care was withdrawn and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation of hemolysis, vascular damage - peripheral vessel, and renal failure has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the medical safety officer reviewed and determined. the patient's demise is associated with the impella rp flex device accordingly under manufacturer device report number. 1220648-2025-28131 the following corrections were made to manufacturer device report number 1220648-2025-27880, based on the medical safety officer review. b2 death and date of death should have been left blank, included other serious event. h1 type of report revised to serious injury. h6 health effect - impact code 1802 revised to 4614. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9893 | 1220648-2025-27881 | 21917661 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21917661&pc=OZD | 04/04/2025 | 04/28/2025 |  | 04/28/2025 |  |  | 04/04/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026619192 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. an angiogram was completed and showed left anterior descending (lad) artery and right coronary artery (rca) stenosis. percutaneous transluminal coronary angioplasty was completed to the lad. the patient's pressures started to drop down into the 30s. the physician cannulated for extracorporeal membrane oxygenation (ecmo) during cardiopulmonary resuscitation (cpr). total down time at this time was approximately one and half hours. during cpr, the impella cp device was brought down to p-2 performance level. ecmo was started and increased to 5lpm. the impella was kept at p-2. a stent was then placed to lad and two stents were placed to rca. revascularization was complete. during peel away sheath removal and repositioning sheath placement, the patient had increased bleeding after repositioning sheath was in placed. a large hematoma was noted, and consequently, the impella cp device was explanted. treatment post impella explant is unknown. access was obtained to right radial artery. picture was taken of left groin and showed bleeding. the decision was made to place a stent to control bleeding. the status of the patient after the event was noted as unknown. however, the patient was noted to have survived the explant. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
9894 | 1220648-2025-27882 | 21917691 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21917691&pc=OZD | 04/08/2025 | 05/15/2025 |  | 04/28/2025 |  |  | 04/08/2025 | 04/28/2025 |  |  | 06/30/2025 |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024376449 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella cp pump was implanted in a patient with acute myocardial infraction/cardiogenic shock. the access site had a wound vac placed as access was obtained through a lymph node. the patient survived and the impella cp was replaced with an impella 5.5. | [5 empty] | The investigation of vascular damage peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. h5 was erroneously selected on the initial manufacturer device report number 1220648-2025-27882. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9895 | 1220648-2025-27883 | 21917764 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21917764&pc=OZD | 04/09/2025 | 06/09/2025 |  | 04/28/2025 |  |  | 04/09/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025610999 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella cp was implanted for the support of a cardiogenic chock/acute myocardial infarction patient who exhibited high purge pressure. the hospital team assessed the pump and observed no kinking in the purge line. the team made the decision to remove and replace the pump. the patient survived. | [5 empty] | The pump was returned for investigation. data logs were not provided for this case run. no physical damage was observed on the returned pump. the pump was run according to specifications in a bucket of water, with observed high purge pressure. the catheter was cut near the motor housing while priming, and purge fluid was observed from the cut catheter, which confirms no blockage in the purge line. close examination of the purge gap revealed biomaterial inside it. the cause of the high purge pressure issue was most likely biomaterial, based on returned product information. the cause of the access site bleeding issue was not determined without sufficient clinical information. thrombus failure mode was added as an investigated failure after returned product investigation. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions. description of the location and characteristics of the thrombus (e.g., size, shape, consistency). relevant laboratory test results, such as coagulation profiles and platelet counts. imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). medications or treatments that the patient was receiving at the time of thrombus formation. surgical or procedural details if applicable (e.g, cardiac catheterization). any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9896 | 1220648-2025-27884 | 21917777 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21917777&pc=OZD | 04/04/2025 | 04/28/2025 |  | 04/28/2025 |  |  | 04/04/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025569064 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Difficult or Delayed Positioning | [6 empty] | Perforation | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the impella 5.5 device was implanted for mechanical circulatory support of a surgical patient. during coronary artery bypass graft and valve surgery, the impella cp device was placed with some difficulty, and a lv perforation was observed requiring intervention. the impella 5.5 device was placed via 10mm graft directly to the right anterior lateral aorta and tunneled out the right second intercostal space. impella guide wire and kmp catheter passed through aortortomy under direct visualization. the decision was made to close aortortomy, fill the heart and implant impella cp device on 2l bypass. positioning the impella device was difficult due to the acute angle of the tunnel and angle of the new aortic valve. after initiating support and prior to closing the chest the team noticed increased bleeding in the chest. blood products were given followed by 30 minutes of observation. the decision was made to put the patient back on bypass and explore. the impella cp device was removed, and the physicians were able to safely assess the posterior side of the heart. there was a small leak on the posterior wall, likely caused by guidewire or repositioning of the impella cp device. the pump had been explanted, and the patient came off bypass without impella support. the patient survived the perforation and 1.5 hours of support. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
9897 | 1220648-2025-27885 | 21917901 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21917901&pc=OZD | 04/06/2025 | 04/28/2025 |  | 04/28/2025 |  |  | 04/06/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026640360 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient on impella cp support had right lower extremity distal limb ischemia. the impella was removed and immediately flow was restored. a new cp was implanted in the left side. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
9898 | 1220648-2025-27886 | 21918006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21918006&pc=OZD | 04/06/2025 | 04/28/2025 |  | 04/28/2025 |  |  | 04/06/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026638855 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that after impella cp placement an echocardiogram at bedside showed impella needing repositioning. the physician tried to reposition the cp, but the patient had ectopy to the point of amio bolus and drip given. the decision was made to leave it shallow. later weaning was successful, the device was explanted, and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed.
9899 | 1220648-2025-27887 | 21918058 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21918058&pc=OZD | 04/09/2025 | 05/15/2025 |  | 04/28/2025 |  |  | 04/09/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025582401 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Low Blood Pressure/ Hypotension | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient with an impella 5.5 due to cardiomyopathy. the impella 5.5 was placed successfully. during support it was noted that the patient had several episodes of ventricular tachycardia (vt). the pump was repositioned, which resolved the issue. the next day they reported again that the patient had more runs of vt, so impella was repositioned. no issues since repositioning. in addition, the patient had a hypotension episode that was treated with volume. the patient remained on support. | [5 empty] | The investigation of vf/vt/af/at and positioning issues has been completed. the cause of the positioning issue was not determined since product was not returned and sufficient clinical information was not provided. as the necessary information was not provided, the root cause of the vt/vf was not determined. | The impella device is still being used with the patient. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9900 | 1220648-2025-27888 | 21918100 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21918100&pc=OZD | 04/07/2025 | 04/28/2025 |  | 04/28/2025 |  |  | 04/07/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026632538 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a thrombus was detected in the left ventricle of a patient implanted with an impella 5.5. the thrombus was noted to be mobile at the time of discovery. no intervention was taken in response to the thrombus and support continued uninterrupted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9901 | 1220648-2025-27889 | 21918120 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21918120&pc=OZD | 04/09/2025 | 05/29/2025 |  | 04/28/2025 |  |  | 04/09/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026634666 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that seven days into support with a patient with an impella cp pump, the pump migrated into the aorta following a scheduled right ventricle biopsy and coronary angiography. the pump was readvanced. however, continuous suction was subsequently encountered. despite attempts to reposition the device, flows did not return to normal and the decision was ultimately made to replace the pump. no patient harm was reported. | [5 empty] | The investigation of positioning issues and low pump flow has been completed. the impella device was not returned for evaluation. positioning issue: clinical states that the inlet of the pump migrated into aorta after a medical procedure. pump and data logs were not returned for investigation. therefore, as per the clinical information provided, the root cause of the positioning issue was most likely use related since the pump migrated during right ventricle (rv) biopsy and coronary angiography procedure before the patient was off the table. low pump flow: clinical states that the pump was pulled into the aorta and low flows were observed. the pump was repositioned but no change in flows and suction alarms were persistent. pump was not returned. data logs were not returned as well. the root cause of the low pump flow issue was not determined since product and data logs were not returned and clinical mentions that repositioning did not resolve the issue. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9902 | 1220648-2025-27890 | 21918234 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21918234&pc=OZD | 04/09/2025 | 05/29/2025 |  | 04/28/2025 |  |  | 04/09/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | OZD |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025602765 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  |  |  | Injury | Yes | No | Pumping Stopped | Failure to Sense | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient developed hemolysis during impella rp flex support. two units of blood were given to the patient and the pump was pulled back from its deep positioning. the condition cleared up following the intervention. support continued uninterrupted and patient outcome is unknown. | [5 empty] | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. the data logs were received for investigation and revealed additional failure modes. hemolysis: clinical details noted that patient was started of continuous renal replacement therapy (crrt) due to some hemolysis occurring overnight on first day of support. data logs were reviewed and real time (rt) log around time of reported issue shows a motor current spike with speed deviation and momentarily drop in pump flows. additionally, a step up in placement signal after spike and increase in purge pressure. the root cause of the hemolysis was most likely due to biomaterial ingestion based on returned data log analysis. temporary pump stop: data logs were reviewed and lt logs show multiple "impella stopped - motor current high" alarms and a failure mode of "temporary pump stop" was added to the investigation. data logs showed about 7 hrs into support a pump stop with a motor current spike with speed deviation immediately before pump stop and flows dropping to 0. purge pressures rose slightly after pump stop occurred but no purge alarms were triggered. pump was attempted to be restarted twice before a successful restart and pump continued support with elevated motor current (mc) and placement signal (ps). no clinical details on temporary pump stop was provided. the root cause of the temporary pump stop was most likely biomaterial based on mc spike and elevated ps upon data log review. placement signal issue: upon review of returned data logs, it is apparent that the dp sensor is the cause of the issue and the failure mode has been changed from "optical signal issue" to "placement signal issue". about 22 hrs into support ps begins to drift. data logs were reviewed and lt log showed a placement signal not reliable (psnr) alarm with dp signal out of range noted about 28 hrs into support. pump motor current remained stable at time of issue with placement signal drifting down then suddenly spiking to 3000 and remaining out of range. pump flows drift upwards then drop to 0 when placement signal goes out of range. cvp is stable at time of observed sensor drift and only goes out of range when placement signal is out of range. clinical details noted that pump had a "placement signal not reliable" alarm and a loss of flow calculations and pump was removed. the root cause of the placement signal issue was most likely due to dp sensor drift based on returned data log analysis. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9903 | 1220648-2025-27891 | 21918323 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21918323&pc=OZD | 04/07/2025 | 05/21/2025 |  | 04/28/2025 |  |  | 04/07/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025547596 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient with extensive cardiac history, new onset chest pain, and pulmonary edema was implanted with an impella cp. after 76 hours of support was decided to escalate to an impella 5.5. multiple troubleshooting measures were attempted to place the 5.5 but the delivery failed. both the 5.5 and the cp were explanted. the patient survived and no harm was noted. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the angulation of the artery just past the anastomosis prevented passage of the impella. the cause of the delivery issue was most likely patient condition related, the angulation of the artery just past the anastomosis prevented passage of the impella after many failed attempts and trouble shooting techniques. instructions for use: impella 5.5 with smartassist with automated impella controller & rp flex with smart assist system with automated impella controller, section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella 5.5 with smartassist for use during cardiogenic shock with automated impella controller. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use during impella with automated impella controller, section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ h10 instructions for use were incomplete on the initial manufacturer device report number 1220648-2025-27891. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock, section: warnings & cautions: warnings, ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.
9904 | 1220648-2025-27892 | 21918384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21918384&pc=OZD | 04/09/2025 | 10/24/2025 |  | 04/28/2025 |  |  | 04/09/2025 | 04/28/2025 | [6 empty] | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026617577 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation | [9 empty] | Malfunction | No | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The user facility reported that purge system blocked alarm occurred approximately 24 hours prior to placement signal unreliable alarm and loss of aorta and left ventricle waveforms. the alarms were silence, as the patient hemodynamics were stable. the nurse was going to monitor flows, purge pressure/flow, and motor current metrics for pump failure. the next day, the patient was taken to the operating room for an impella 5.5 device replacement due to optical sensor failure. the patient survived the procedure. | [5 empty] | The investigation of optical signal issue and high purge pressure has been completed. the pump was not returned for investigation. data logs were reviewed and the placement signal not reliable (psnr) alarm was confirmed on 04/09. at this time, snr drops to zero, lamp maxes out at 100 percent, light decreases, gain decreases, and contrast oscillates between +/-3,000. the real time (rt) logs were closely examined during the time of the psnr event, and the oscillating contrast trend was seen. the pump was remained in use during psnr. additionally, the pump was used on the same console to verify the temporary oscillating contrast trend. upon review of data logs, it is apparent that the console is the potential cause of the issue. high pressure purge: the data log shows a 3-minute stepwise increase in purge pressure followed by a high purge pressure alarm. the purge flow dropped stepwise to 0 during hpp. additionally, the purge flow tick indicates a stepwise trend. this trend was resolved in a stepwise manner. purge parameter matches the known trend of a sidearm kink. the cause of the high purge pressure issue was most likely kinked purge line at sidearm, based on data log review. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | G3 ¿date received by manufacturer¿ corrected. | D6b "if explanted, give date" corrected.
9905 | 1220648-2025-27893 | 21918447 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21918447&pc=OZD | 04/04/2025 | 04/28/2025 |  | 04/28/2025 |  |  | 04/04/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026614239 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that after impella 5.5 placement and following transfer to intensive care unit, echocardiogram (echo) showed pump position deep in left ventricle. the team repositioned and withdrew 2 cm at bedside under echo guidance. eleven days of 5.5 support impella was repositioned again due to posterior orientation of inflow. additionally, there was an acute increase in jp drainage overnight from impella site, total of 750cc. dressing remained intact and 1 point drop in hemoglobin. one unit of red blood cells was given, protamine administered and systemic heparin was held for 24 hours. the patient remains on support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9906 | 1220648-2025-27894 | 21918563 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21918563&pc=OZD | 04/09/2025 | 09/11/2025 |  | 04/28/2025 |  |  | 04/09/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1380525 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that at the end of therapy during removal, the automated impella controller screen froze and no action was possible at the screen. there was no patient harm reported. | [5 empty] | The investigation of automated impella controller (aic) - display issue has been completed. the console and data logs were received for investigation and there was no issue found during the case. the device functioned/operated to specification. aic - kbd/rotary knob issues was added based on the device data and analysis. there were no errors related to purge or keyboard operation, and console could be controlled entire time. the root cause was unable to be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9907 | 1220648-2025-27895 | 21918652 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21918652&pc=OZD | 04/05/2025 | 04/28/2025 |  | 04/28/2025 |  |  | 04/05/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026624512 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support and during support, the patient had bleeding at the access site. angle matching was performed, a femstop was placed, three units of packed red blood cells were given and heparin was withheld. additionally, the patient had limb ischemia. staff was unable to doppler pulses in the bilateral lower extremities and both legs were cold and mottled. support continued. the patient expired on support. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9908 | 1220648-2025-27896 | 21918706 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21918706&pc=OZD | 04/09/2025 | 05/28/2025 |  | 04/28/2025 |  |  | 04/09/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026634378 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that while implanted with an impella cp pump, a patient had oozing around the repo sheath. manual pressure was held and heparin was withheld. impella support was continued. additionally, the patient had hemolysis and the patient's urine was pink. the patient survived and the pump was successfully weaned. | [5 empty] | The investigation of access site bleeding and hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9909 | 1220648-2025-27897 | 21919242 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21919242&pc=OZD | 04/09/2025 | 05/29/2025 |  | 04/28/2025 |  |  | 04/09/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026640211 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Difficult to Setup or Prepare | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that as an impella cp was being prepped for implantation it was noted that the tuohy-borst was locked in place and unable to be loosened despite loosening all the way. a new pump was opened and used. no adverse impact to the patient was reported. | [5 empty] | The investigation of positioning issues has been completed. the pump was returned and the sterile sleeve was found to be stuck in the touhy-borst near the red handle side of the pump (100 cm mark). the touhy-borst was completely stuck and unable to move. pump was unused with red lumen still in place. therefore as per the clinical information and product review, the root cause of the positioning issue was the sleeve stuck in the touhy-borst due to improper technique making it unable to move along the catheter. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
9910 | 1220648-2025-27898 | 21919354 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21919354&pc=OZD | 04/09/2025 | 11/21/2025 |  | 04/28/2025 |  |  | 04/09/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1813537 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | During routine preventive maintenance, the automated impella controller (aic) was found to have purge sensor issues. the purge pressure sensor readings were outside the limits at various input levels. the purge sensor was replaced to resolve the issue. there was no patient involvement reported. | [5 empty] | The investigation has been completed. data analysis: console logs were not relevant for this complaint, however no other issues were seen in the imc logs. device analysis: aic ic11568 was not returned to pme. the issue occurred during pm, purge pressure sensor readings were inaccurate. the purge assembly was replaced to resolve the issue. root cause: the root cause of the "wrong purge pressure sensor reading" was likely a defective purge assembly. h3: added information regarding the investigation status. h6: added codes per the results of the investigation. h10: added the results of the investigation. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9911 | 1220648-2025-27899 | 21919446 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21919446&pc=OZD | 04/09/2025 | 05/22/2025 |  | 04/28/2025 |  |  | 04/09/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1603745 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the automated impella controller (aic) was being utilized along with an impella5.5 pump and exhibited a battery charging issue. the aic remained on for support. no patient harm or injury was noted. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The console logs were consistent with what was reported in the complaint. the console was tested and the reported battery level low issue was unable to be reproduced while testing on an engineering lab bench. no evident battery issues were observed after running the batttest utility on telnet the likely cause of the reported issue was a momentary communication glitch between the power battery manager (pbm) and the batteries.
9912 | 1220648-2025-27900 | 21919460 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21919460&pc=OZD | 04/07/2025 | 06/02/2025 |  | 04/28/2025 |  |  | 04/07/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025484229 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male | 97 KG | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Embolism/Embolus | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella 5.5 was never able to get optimal flows while implanted in the patient. the volume was adequate and positioning was good. the impella was removed and an intra-aortic balloon pump was placed. the patient survived. | [5 empty] | The investigation of low pump flow has been completed. the returned 5.5 pump visually inspected. no broken blade or cannula kink observed. bit of biomaterial observed under impeller blades. then impella run in water on p-9 for more than 10 minutes to try and reproduce reported low flow issue, no low pump flow observed. flow started at 5.1l/min and raised up to 5.6l/min and it was in flow range rate. biomaterial observed under impeller blades after impella run. data log reviewed: many motor current (mc) spikes during support and many suction alarms observed during the case. impella position unknown alarms seen. clinical stated: impella was not able to get max flows on pump despite proper positioning and volume status. low pump flow: the cause of the low pump flow issue was most likely biomaterial ingestion as observed on the returned product and mc spikes in the data logs. failure mode thrombus added as a result of biomaterial observation. as the necessary information was not provided, the root cause of the thrombus was not determined. based on the investigation findings the following fields were updated. b1 revised to include both adverse event and product problem. b2 updated. h1 revised from malfunction to serious injury. h6 health effect - clinical code revised to include 4438. instructions for use (thrombus). impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The impella device has been received for evaluation, upon investigation closure a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock, section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿
9913 | 1220648-2025-27901 | 21919732 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21919732&pc=OZD | 03/29/2025 | 11/07/2025 |  | 04/28/2025 |  |  | 03/29/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025603954 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Malposition of Device | [6 empty] | Adult Respiratory Distress Syndrome | Hemorrhage/Blood Loss/Bleeding | Hematuria | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support (mcs). post implant, same day, an echocardiogram was completed and the impella measured 3.8 centimeters to the aortic valve annulus and mid ventricular cavity. the following morning the impella cp device was retracted with guidance under echocardiogram. two days later, the plan for turndown of extracorporeal membrane oxygenation (ecmo) and potential upgrade to an impella 5.5 device was noted. hematuria was present. however, action taken, if any, is unknown. ecmo and impella support continued. the next day drips in use were noted as amiodarone, heparin (14 u/kg/hr.) and other (unspecified). thereafter two days later it was noted the patient experienced respiratory distress and ecmo cannula site bleeding (of note patient heparinized) that required multiple blood products as reported under the abiomed's long-term outcome and quality indicator impella registry and noted as a definite relationship to the impella device used. heparin was withheld the exact time is unknown. however, the ecmo placement was removed and the impella cp device was weaned and explanted later this day following with the implant of an impella 5.5 device, which was successfully completed for escalation in therapy. | [5 empty] | The investigation of reported failure modes has been completed. the impella device was not returned for evaluation. positioning issues: the root cause of the positioning issues was not determined since insufficient clinical details were provided. hemolysis: hthe root cause of the hemolysis was not determined since insufficient clinical details were provided and no product was returned. respiratory distress (vt/vf fm): it was reported that the patient endured respiratory distress. the cause of the respiratory distress was unable to be determined due to insufficient clinical details. vascular damage - peripheral vessel: the root cause of the vascular damage was not determined since insufficient clinical details were provided. bleeding reported from the non-impella side access site. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9914 | 1220648-2025-27902 | 21919923 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21919923&pc=OZD | 04/10/2025 | 04/28/2025 |  | 04/28/2025 |  |  | 04/16/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2025584205 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 91 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hypovolemic Shock | Multiple Organ Dysfunction Syndrome | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that an impella patient's artery was damaged during central venous insertion. a blood transfusion was performed, but the blood loss was so great that circulation could not be maintained. the patient subsequently passed away from multiple organ failure due to hemorrhagic shock. due to anticoagulation in the patient's system from impella support, there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
9915 | 1220648-2025-27903 | 21920031 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21920031&pc=OZD | 04/09/2025 | 05/22/2025 |  | 04/28/2025 |  |  | 04/09/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1354144 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA CP | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an automated impella controller (aic) failed to recognize the impella cp pump. despite connecting the pump, the aic continued to request the grey connector to be attached. the nurse restarted the cycle twice before the aic finally recognized the pump correctly. there was no patient harm reported. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of pump not detected has been completed. the console was received for investigation. data analysis: the case associated with the reported complaint was initiated on mar 20, 2025. at step 3 of the case start wizard, since the pump connection was not recognized for 20 seconds, the epm was reset automatically. even after the reset the pump was not recognized, and the console displayed ¿connect grey connectors screen¿. the user re-initiated the case start wizard; at step 3, the pump was not recognized again. this time, the user canceled the case start wizard after 17 seconds. (note: by design, at step 3 of the case start wizard, if the pump is not recognized for 20 seconds, the epm will be reset automatically at the 20th second.) on the third attempt of re-initiating the case start wizard, the pump was recognized after the epm reset. support was then provided for 1 hour and 33 minutes. on april 9, 2025, the console encountered the same issue with recognizing the pump connection initially. on the second attempt of re-initiating the case start wizard, the pump was recognized after the epm reset. support was then given for 1 hour and 7 minutes. this confirms the reported failure mode. device analysis: this console was tested after arriving at fs. the reported failure mode was reproduced 1 out of 10 attempts. the console was power cycled 10 times. each time, via the ¿case start¿ wizard, upon connecting the pump after the epm reset, the pump connection was recognized for the initial 9 times. however, on the 10th attempt, the pump was not recognized at step 3 of the ¿case start¿ wizard. root cause: the pump detection issue was due to a rarely occurring epm crystal issue. the impellatronic pca does not have permanent damage, but this intermittent pump detection failure occurs infrequently and already has a software mitigation in place.
9916 | 1220648-2025-27904 | 21920138 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21920138&pc=OZD | 04/05/2025 | 06/03/2025 |  | 04/28/2025 |  |  | 04/05/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026613594 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered saturated flows. the impella 5.5 pump was inserted via the right axillary approach due to the patient's short aortic root. wire access to the left ventricle was lost on the first attempt and the pump was removed. the pump was cleared of any remaining blood and the doctor quickly re-accessed the left ventricle. the pump was then successfully placed into the left ventricle and turned on. however, shortly, after initiating support, the console began exhibiting signs of flow saturation and the team decided to abort impella device placement. there was no patient harm. | [5 empty] | The investigation of saturated flow and delivery issues has been completed. the impella device was not returned for evaluation. saturated flow: shortly, after initiating support, the console began exhibiting signs of flow saturation. the pump was not returned. data log review showed flows at 4.5l/min at p-4, higher than the expected 2.0-2.5l/min. this occurred after a motor current (mc) spike which led to not only the saturated flow, but a rise in purge pressure and a decrease in purge flow. the cause of the saturated flow was most likely biomaterial due to a mc spike found that was followed by 4.5l/min flows at p-4, purge pressure trending up, and purge flows trending down. delivery issues: the cause of the delivery issues was most likely patient condition related due to the difficulty accessing the left ventricle (lv) due to the patients short aortic root. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
9917 | 1220648-2025-27905 | 21920403 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21920403&pc=OZD | 04/02/2025 | 04/28/2025 |  | 04/28/2025 |  |  | 04/03/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025601382 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Adult Respiratory Distress Syndrome | Arrhythmia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A notification was received via abiomed's long-term outcome and quality indicator impella registry regarding a patient who experienced respiratory distress with a definite relationship to the impella 5.5 device (and the previously implanted impella cp device) used. this occurred the same day the impella 5.5 device was implanted. action taken, if any, is unknown. the patient recovered from the respiratory distress with no sequelae. however, it was further noted regarding this patient's implant experience that the complainant also reported this patient with cardiomyopathy was initially implanted with an impella cp device mechanical circulatory support (mcs) before being implanted with an impella 5.5 device for escalation in therapy. the following day after start of the impella 5.5 device mcs, there were occasional suction alarms and overnight the patient had increased ectopy with central venous pressure (cvp) of 8 mmhg. echocardiogram of the impella 5.5 device was pending position assessment. the following day, suction alarms were still present, and the impella 5.5 pump was repositioned. however, the pump still appeared deep on transesophageal echocardiogram. a repeat echocardiogram was planned, and additionally noted heparin (withheld when on the impella cp) was planned for restart once platelets were greater than 50k. the day after, the patient was still with intermittent suction. performance level was maintained on p-4, which appears, until two days later when the patient was successfully weaned, and the impella 5.5 device was explanted with a survived outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. this is one of two devices related to the patient's implant experience. refer to manufacturing report number: 1220648-2025-27901 for the medical device report on the impella cp device.
9918 | 1220648-2025-27906 | 21920604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21920604&pc=OZD | 04/06/2025 | 04/28/2025 |  | 04/28/2025 |  |  | 04/06/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026630308 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient on impella cp experienced a leaking purge line just below the spike for the dextrose bag. purge pressure and flow were inconspicuous, but glucose spread everywhere. it was noted that since the patient is stable, the physicians decided to remove the pump as soon as possible, they are just waiting for the activated clotting time to decrease. | [5 empty] | The impella device was not received from the customer as it is still in use. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9919 | 1220648-2025-27907 | 21920681 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21920681&pc=OZD | 10/03/2023 | 04/28/2025 |  | 04/28/2025 |  |  | 04/07/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1155606 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that there was a placement signal not reliable alarm. the pump positioning was confirmed. this was not reportable however during investigation and review of the log files numerous occurrences of controller error were seen therefore now reportable. there was no patient harm. | [5 empty] | The investigation has been completed. the impella pump was returned for investigation. per data analysis the imc logs from the console confirm that numerous occurrences the controller error (defective optical sensor lamp) - alarm even as earliest. device analysis; the failure was reproduced in boot up in the engineering bench. voltage supply to from the pbm to the optical bench was measured to be within spec. there was no light observed from the optical lamp despite having voltage supply to it. swapping out the bad lamp with a good one resolved the issue. the root cause for the controller error was a defective optical bench lamp.
9920 | 1220648-2025-27908 | 21920805 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21920805&pc=OZD | 04/06/2025 | 05/27/2025 |  | 04/28/2025 |  |  | 04/06/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636214 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella cp support and during support, the impella was caught in the papillary muscle and the doctor struggled to free it. the impella cp kept folding in on itself, so the doctor explanted the impella cp and placed a new one. the patient survived the procedure. | [5 empty] | The investigation of positioning issues has been completed. no product was returned. the root cause of the positioning issue was most likely use related since cardiopulmonary resuscitation (cpr) was performed. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9921 | 1220648-2025-27909 | 21921031 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21921031&pc=OZD | 04/07/2025 | 05/15/2025 |  | 04/28/2025 |  |  | 04/07/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1098596 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | During preventive maintenance a crack in the rear housing was noted. there was no patient involvement. | [5 empty] | The investigation of automated impella controller (aic) - damage before therapy has been completed. the logs were reviewed but no issues were observed in either of the logs. the console was tested and functioned/operated to specification. no problem found with this aic. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
9922 | 1220648-2025-27910 | 21921063 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21921063&pc=OZD | 04/06/2025 | 05/30/2025 |  | 04/28/2025 |  |  | 04/06/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2026613250 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient had an impella 5.5 pump for support during cardiogenic shock. the impella 5.5 pump was supporting when the pump had "purge system blocked" alarms. the pump was troubleshot. no kinks were observed, the personal computer had already been replaced, and the team discussed tissue plasminogen activator being added to the purge. the pump remains on despite the high purge pressure. | [5 empty] | The investigation of high purge pressure has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer as it is still in use. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
9923 | 1220648-2025-27911 | 21921460 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21921460&pc=OZD | 04/07/2025 | 04/28/2025 |  | 04/28/2025 |  |  | 04/07/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1080533 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella support when the automated impella controller (aic) an ¿unexpected controller shutdown¿ alarm. this after the screen turned completely off then showed startup mode before switching back to placement screen with the above alarm. the aic was replaced and there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9924 | 1220648-2025-27912 | 21921916 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21921916&pc=OZD | 04/08/2025 | 06/03/2025 |  | 04/28/2025 |  |  | 04/08/2025 | 04/28/2025 |  |  |  |  | 04/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1161464 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the automated impella controller (aic) would not progress multiple pumps past 1.6 l/min. no patient or pump information was provided. | [5 empty] | The investigation of automated impella controller (aic) - pump stop has been completed. the pump stop was not related to any issues with the aic. no problems were found within the aic. d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2025-27912. | The device has been received for evaluation, upon completion of the investigation a supplemental mdr will be filed.
9925 | 1220648-2025-27830 | 21906310 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21906310&pc=OZD | 02/24/2025 | 10/21/2025 |  | 04/25/2025 |  |  | 04/02/2025 | 04/25/2025 |  |  |  |  | 04/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1464780 |  | 0042-0010-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported the 77-year-old male patient was on impella cp support and during support, the catheters crossed in the left ventricle and when the field medical catheter was ablating, the impella would drop to p0 and then the impella was manually brought up to the previous p level. this occurred several times throughout the case. the aic team requested that the aic be investigated for the issue. | [5 empty] | Not applicable because console 9542 was not used with the pump (sn: (b)(6)) and is not associated with the reported failure mode. the correct console 7888 associated with manufacture device report 1220648-2025-27830 the root cause of the reported pump stop was not determined as console 7888 or data logs were returned for the investigation. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-27830. d4 serial number, lot number, and udi number were erroneously entered on the initial manufacturer device report number 1220648-2025-27830. d9 was erroneously entered on the initial manufacturer device report number 1220648-2025-27830. as the correct console was not received for investigation h6 type of investigation code 10 was erroneously entered on the initial manufacturer device report number 1220648-2025-27830. h10 narrative statement erroneously stated the device was received on the initial manufacturer device report number 1220648-2025-27830. | This report is being filed as part of a retrospective review of historical records. data analysis (ic9542): not applicable because the ic9542 was not used with the pump (sn: (b)(6) and is not associated with the reported fm. data analysis (ic7888): on the complaint date, february 24, 2025, the case associated with the reported complaint was initiated with pump sn: (b)(6). two minutes after the pump was started, the console displayed ¿placement signal not reliable (optical signal-to-noise ratio) ¿ alarm,¿ event #130, which was automatically resolved in 47 minutes. the rt log shows that the snr was below 2000 throughout the logs, and the ps spiked to 3000 (7_rt_log_1st_psnr.png). after 48 minutes, the log recorded ¿pmd ok staterec1937 0 2000 257 0 0 0 0x593 0x593 cpld 3,¿ indicating a pmd reset caused by a watchdog trigger. immediately following the reset, the console displayed the first ¿impella stopped (mechanical/electrical issue) ¿ alarm,¿ event #115. the rt log confirms that the motor current dropped to zero (10_rt_log_1st_pump_stop_z.png). support was provided for an additional 2 hours and 36 minutes. during this period, multiple 'impella stopped' events occurred due to pmd resets (imc_logs_ic7888.pdf). this further confirms the reported failure mode. the lt log shows several instances where the motor current and mean flow dropped to zero (1_lt_log.png). every time, the user manually restarted the pump in the previous p-level. note: imc/rt log entries show no signs of high motor current or current phase mismatch. device analysis (ic7888): this console was tested after arriving at fs. while running the test pump at p9, measured 20v at the impellatronic connector (j001.2) as intended. this confirms that there is no pbm malfunction. upon running the optical test pump for 16 hours at p-level p9, no pump stops occurred, confirming no impellatronic board malfunction. the analog output from the optical bench shows no issues and matches the known good optical bench analog output. the optical parameters are within specifications according to the pm procedure (0042-7309). the front panel optical connector has minor debris, which is not related to the reported failure mode. during testing, the console displayed a ¿purge piston blocked¿ alarm. upon visual inspection, minor debris was confirmed on the purge motor gasket (purge_motor_gasket_ic7888.png). this is unrelated to the reported failure mode. corrective action: failures of this type will be monitored for trends. pia-0000551 was opened in response to the potential pfa interaction with impella. root cause (ic7888): the root cause of the pump stop was not determined due to the inability to reproduce the pmd reset, or pump stop. it is most likely related to its potential interaction with the pfa catheter, which could have induced a pmd reset and caused the loss of motor current or speed. no pump stops or pmd resets were observed on ic7888 after running the optical test pump and cassette for 16 hours at p-level p9 in the engineering lab. the root cause for the psnr was not determined due to the inability to reproduce it. ad hoc task (ic7888): before returning this console back to the customer, fs was instructed to perform the following, clean the front panel optical connector. if unable to remove the debris, replace the front panel optical connector (0042-2750). clean the debris behind the purge cassette holder and on the purge motor gasket. perform sections 10, 11, 12, 13, 14, 15, 16 and 23 of the pm procedure (0042-7309) perform a full functional check (0042-7316). | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. this report is for the same patient as in mdr mfr report #: 1220648-2025-27793.
9926 | 1220648-2025-27831 | 21907923 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21907923&pc=OZD | 03/31/2025 | 05/02/2025 |  | 04/25/2025 |  |  | 03/31/2025 | 04/25/2025 |  |  |  |  | 04/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1663431 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) was making a loud cranking noise during patient support. the aic was swapped in response to the issue. there was no patient harm. | [5 empty] | The investigation of automated impella controller (aic) - damage during therapy has been completed. the root cause of the cranking noise was not determined due to the inability to reproduce it consistently during device analysis. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
9927 | 1220648-2025-27832 | 21907984 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21907984&pc=OZD | 03/31/2025 | 04/25/2025 |  | 04/25/2025 |  |  | 03/31/2025 | 04/25/2025 |  |  |  |  | 04/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025594021 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Ischemia | Thrombocytopenia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the patient encountered limb ischemia, thrombus, thrombocytopenia and positioning issues. echocardiogram showed that the pump was positioned 6 centimeters inside the left ventricle which then required repositioning for it to 3.8 centimeters. furthermore, the patient was found to have critical limb ischemia of the right left extremity (impella cp indwelling in the right femoral artery). the decision was then made to escalate to impella 5.5 and remove the impella cp. however during pump explant, there was significant dense thrombus throughout the iliac extending into the common femoral artery and there was also a thrombus seen mid femoral. the vascular surgeon suspected that the thrombus was directly related to prolonged indwelling time. they saw no direct injury to the vessel from insertion or catheter, only thrombus. extensive complex thrombectomy was performed with deployment of multiple stents and flow was restored distally. the patient returned to the intensive care unit in stable condition. the patient was also found to have a positive panel for heparin-induced thrombocytopenia and so the patient was transitioned to bivalirudin. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: impella cp with smart assist catheter insertion, section: axillary insertion of the impella cp with smart assist catheter, section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
9928 | 1220648-2025-27833 | 21908071 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21908071&pc=OZD | 03/31/2025 | 07/21/2025 |  | 04/25/2025 |  |  | 03/31/2025 | 04/25/2025 |  |  |  |  | 04/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1805755 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller (aic) had an unexpected controller shutdown alarm message appear. the console was plugged in and fully charged according to the screen display. no patient harm was reported. | [5 empty] | The investigation of automated impella controller (aic) - shutdown or restart issue has been completed. based on the device and data analysis the root cause of the intermittent, unexpected controller shutdown was not determined due to the inability to reproduce. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
9929 | 1220648-2025-27834 | 21908130 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21908130&pc=OZD | 04/05/2025 | 06/03/2025 |  | 04/25/2025 |  |  | 04/05/2025 | 04/25/2025 |  |  | 04/30/2024 |  | 04/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025496184 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient encountered limb ischemia. the patient's right leg was occluded due to the extracorporeal membrane oxygenation (ecmo) and impella cp. for this reason, a femorofemoral bypass was placed. additionally, the flows of ecmo and impella cp were adjusted. the patient remained stable and the leg became well perfusioned. the patients outcome is stable. | [5 empty] | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9930 | 1220648-2025-27835 | 21908255 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21908255&pc=OZD | 03/29/2025 | 05/19/2025 |  | 04/25/2025 |  |  | 03/29/2025 | 04/25/2025 |  |  |  |  | 04/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025569328 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the insertion of an impella 5.5 device from an impella cp device for escalation in therapy for a patient who was admitted in with cardiogenic shock from an acute myocardial infarction was unsuccessful. the plan was to explant the impella cp from the left femoral artery (lfa) and escalate to impella 5.5 via right axillary artery. the physician stated the axillary pocket was deep and a more ¿direct¿ approach to the artery was used. the patient had an aortic valve replacement. the impella 5.5 (pump 1) was primed and setup without incidence. the physician was successfully able to cross the aortic valve with the 0.18 guidewire with the impella cp in place. the impella 5.5 (pump 1) was successfully loaded on 0.18 wire and entered the sheath without difficulty. however, when the physician attempted to advance the impella through the patients vascular, the impella would not drop down. despite multiple attempts to reposition the device, rotate etc. as well as guidewire change to a v18, the impella 5.5 (pump 1) was unable to advance and would ¿buckle¿ in the patient. upon removal a notable kink was observed on the catheter proximal to the motor housing. a new impella 5.5 device (pump 2) was opened, primed and setup without a problem. the impella cp was removed from the lfa and the new impella 5.5 (pump 2) was implanted through the right axillary graft and was unable to advance as well. consequently, the physician aborted implanting an impella device. the physician noted possible direct approach with second device being discussed. the patient was maintained on extracorporeal membrane oxygenation support and was switched to veno-arterial-venous. the patient was reported to be in stable condition following the event. | [5 empty] | The investigation of delivery issues and product damage (cannula kink) has been completed. the impella 5.5 pump was returned and a catheter kink near the motor housing was observed. delivery issue - the root cause of the delivery issue was most likely patient condition related since the patient had an aortic valve replacement (avr) and both the complaint pump and replacement pump were not able to pass through patient vasculature. product damage - the root cause of the product damage was most likely a catheter kink due to patient condition since the patient had an avr and both the complaint pump and replacement pump were not able to pass through patient vasculature. | The pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. this report represents pump 1. another manufacturer device report will represent pump 2. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
9931 | 1220648-2025-27836 | 21908382 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21908382&pc=OZD | 04/05/2025 | 05/14/2025 |  | 04/25/2025 |  |  | 04/05/2025 | 04/25/2025 |  |  |  |  | 04/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026623510 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was placed on impella cp for hemodynamic support. it was noted that the impella cp was backloaded and started on auto. however, flows were only 2.2. there was concerned that the right ventricle function/volume status was causing low flows. calcium and volume were given, flows increased to 2.8. intermittent suction alarms were reported. neurology was consulted due to possible intracranial hemorrhage. heparin was paused. there was also poor flow distal to impella. ultimately, the patient expired on support. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of low pump flow and limb ischemia has been completed. the impella device was not returned for evaluation. there were no data logs or product returned. based on the clinical description, the root cause of the low pump flow is most likely due to the patient's condition. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9932 | 1220648-2025-27837 | 21908492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21908492&pc=OZD | 04/05/2025 | 04/25/2025 |  | 04/25/2025 |  |  | 04/05/2025 | 04/25/2025 |  |  |  |  | 04/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026613427 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemolysis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient had an impella cp pump placed for support during a high-risk cardiac procedure. the pump was supporting when there were concerns of hemolysis and bigeminy. the pump had supported for just over 48 hours when the harms prompted premature explant. the patient survived the procedure. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9933 | 1220648-2025-27838 | 21908593 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21908593&pc=OZD | 04/05/2025 | 04/25/2025 |  | 04/25/2025 |  |  | 04/05/2025 | 04/25/2025 |  |  |  |  | 04/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026617432 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient on impella cp support developed ischemia in their lower extremity and hemolysis. the patient's family opted to withdraw care and the patient passed away. there is not enough evidence to disassociate the patient's passing with impella support. | [5 empty] | The impella device is not available from the customer. therefore, the investigation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9934 | 1220648-2025-27839 | 21908778 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21908778&pc=OZD | 04/22/2025 | 04/25/2025 |  | 04/25/2025 |  |  | 04/22/2025 | 04/25/2025 |  |  |  |  | 04/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2025611235 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 36 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that upon explant of an impella cp on a 36-year-old male patient, a clot was noted on the impella cp catheter. no effects on support were noted leading up to the explant. the patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
9935 | 1220648-2025-27840 | 21908921 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21908921&pc=OZD | 04/01/2025 | 04/25/2025 |  | 04/25/2025 |  |  | 04/01/2025 | 04/25/2025 |  |  |  |  | 04/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026616484 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella cp pump placed to support a patient with cardiac history. the impella cp pump was placed femorally despite knowledge that the leg was edematous, necrotic and infected. the team attempted to place distal bilateral infusion to the leg with antegrade sheaths. care escalation to the impella 5.5 pump and extracorporeal membrane oxygenation (ecpella) were discussed, along with limb amputation possibility, but rather the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9936 | 1220648-2025-27841 | 21908973 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21908973&pc=OZD | 04/14/2025 | 05/19/2025 |  | 04/25/2025 |  |  | 04/14/2025 | 04/25/2025 |  |  |  |  | 04/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1756146 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | Battery Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that when replacing the purge system on the impella cp, the console (aic) displayed a message, the screen went black, and impella temporarily stopped. the aic started up again and then started supporting the patient. the aic was exchanged. there was no harm to the patient. | [5 empty] | The investigation of automated impella controller (aic) - battery comm. failure (yellow alarm) and aic - pump stop has been completed. the console logs were partially consistent with what was reported in the complaint. an unexpected reboot was observed in the imc logs which coincided with the user¿s report of the screen going black, the impella stopping, the aic restarting, and the impella resuming operation again. the ¿as reported¿ code associated with the battery comm. failure could not be confirmed through data analysis. no evidence any battery communication failures were observed in the imc logs. device analysis: console was tested and the unexpected reboot was unable to be reproduced during testing. the console operated as intended. conclusion: the root cause of the unexpected reboot and the reported pump stop was unable to be determined due to the inability to reproduce. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop." section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
9937 | 1220648-2025-27842 | 21909098 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21909098&pc=OZD | 03/10/2025 | 04/25/2025 |  | 04/25/2025 |  |  | 04/01/2025 | 04/25/2025 |  |  |  |  | 04/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025584035 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during implantation, the patient experienced a massive hemorrhage due to femoral artery dissection. blood products were administered. the procedure was unsuccessful. the patients condition was noted as critical. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings & cautions: cautions. ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
9938 | 1220648-2025-27843 | 21909188 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21909188&pc=OZD | 04/17/2025 | 05/07/2025 |  | 04/25/2025 |  |  | 04/18/2025 | 04/25/2025 |  |  |  |  | 04/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1448687 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that once when the automated impella controller (aic) was turned on, controller error alarm occurred. the incident occurred during impella support. the aic was replaced and the support continued. no health damage to the patient. | [5 empty] | H5 was erroneously entered on the initial manufacturer device report number 1220648-2025-27843. h8 was erroneously entered on the initial manufacturer device report number 1220648-2025-27843. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
9939 | 1220648-2025-27844 | 21909438 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21909438&pc=OZD | 04/04/2025 | 04/25/2025 |  | 04/25/2025 |  |  | 04/04/2025 | 04/25/2025 |  |  |  |  | 04/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026629099 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient on impella cp support was transferred to another facility. upon arrival it was noted that the cp was placed via a 16f sheath dry seal (non-abiomed product) which it was left inside the patient. the patient needed impella support but couldn¿t be left in place through the sheath and had to be removed by way of cutdown in the operating room. the right femoral artery was surgically repaired and a intra-aortic balloon pump was placed. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
9940 | 1220648-2025-27845 | 21909476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21909476&pc=OZD | 04/01/2025 | 05/15/2025 |  | 04/25/2025 |  |  | 04/01/2025 | 04/25/2025 |  |  |  |  | 04/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025603497 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  |  | White | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Tachycardia | Embolism/Embolus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. the impella 5.5 was primed and inserted. the position was rechecked, unclear visualization of inlet, still appeared mitral facing. upon attempt to re-orient, the impella 5.5 was pulled back into the ascending aorta with appropriate alarms. the impella level was down to p2. the chest was re-opened and the impella 5.5 advanced back across theaortic valve, briefly at p0. impella flow was restarted at p4, lower than expected flows, "impella in aorta" alarm, but position was confirmed on echocardiogram. initially appropriately oriented, then again toward mitral after chest closure. minimally pulsatile motor current. the medical doctor opted to keep the impella at p-2. the patient was hemodynamics stable. final transesophageal echocardiogram position at 4.3 centimeters was not apically directed. the physician was aware of the position and motor current concerns and opted for no interventions. it was also noted that the swan was out of position and unable to reposition. attempted flowtrack to trend cardiac output. the physician was aware of minimally pulsatile motor current, lower than expected flows for current p-level and false "impella in aorta" alarms. no immediate interventions was planned. additionally, the patient experienced an episode of torsades ventricular tachycardia requiring direct current cardioversion four times, and the addition of lidocaine and amiodarone. the patient was reintubated due to code and lethargy. during the event, the impella generated a false ¿impella in aorta¿ alarm with loss of motor current pulsatility but no loss in flow. impella reportedly was pulled back 0.5 centimeters after the event. | [5 empty] | The investigation of vf/vt/af/at and positioning issues has been completed. product analysis: pump was returned in damage condition as the catheter was cut off. visual inspection found biomaterial inside the pump housing. no other issues were found on the rest of the pump. the failure mode (fm)'s thrombus was added since the customer reported thrombus. the fm low pump flow was added as the customer claimed there were lower flows. positioning issues: data logs confirmed pump was in aortic area on (b)(6) 2025 corresponded with the time impella inadvertently pulled back into ascending aorta that triggered alarm impella position in aorta & suction alarms. there are two positioning issues occurring here. the root cause of the positioning issue for the pump facing the mitral apparatus cannot be determined due to limited clinical information. the root cause of the positioning issue for the pump being inadvertently pulled back into the ascending aorta is most likely due to use as it was accidentally pulled back. low pump flow: data logs show on the 21st there was "impella position unknown" and "suction" alarms that resolved, while placement signal (ps) and motor current (mc) were flat and not very pulsatile. flows were varying. the returned pump was cut so it was unable to be run. based on the clinical details and data log analysis, the root cause of the low pump flow is most likely due to improper positioning. the cause of the vt storm and thrombus was unable to be determined due to insufficient clinical details. instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
9941 | 1220648-2025-27846 | 21909549 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21909549&pc=OZD | 03/25/2025 | 04/25/2025 |  | 04/25/2025 |  |  | 03/26/2025 | 04/25/2025 |  |  |  |  | 04/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025582942 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | IMPELLA RP FLEX. | [17 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was escalated to bipella (impella 5.5 and impella rp flex) from an impella cp and during support, the patient had a gastrointestinal bleed. five units of packed red blood cells were given. heparin was not used in the purge. however, the representative stated the relationship to the impella is unknown. care was withdrawn and the patient expired. the patient's demise is associated with the impella cp device and captured under the manufacture report number 1220648-2025-27522-1. | [5 empty] | The investigation was completed. the patient had a gastrointestinal bleed and five units of packed red blood cells were given. heparin was not used in the purge. however, the representative stated the relationship to the impella is unknown. the root cause of vascular damage peripheral vessel gastrointestinal bleed cannot be determined since the product was not returned and insufficient clinical details were provided. abiomed's medical safety officer review the patient's outcome and it was determined that the patient demise would be reported under the impella cp as there is not enough evidence to exclude the impella cp as an associated factor in the patient's outcome. the death is captured in the impella cp manufacture report number 1220648-2025-27522-1. another manufacture device report will capture the impella rp (serial number (b)(6)) event.
9942 | 1220648-2025-27847 | 21909612 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21909612&pc=OZD | 04/04/2025 | 05/23/2025 |  | 04/25/2025 |  |  | 04/04/2025 | 04/25/2025 |  |  |  |  | 04/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025576476 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complaint reported that a patient presented to ed in cardiac arrest. the patient continued to de-compensate, so decision for impella cp was made. it was reported that the pump was started but never achieved flows >1lpm. after multiple repositioning attempts, the device was removed and a new cp was used. there was no patient harm. | [5 empty] | The investigation of low pump flow has been completed. the impella device was not returned for evaluation. the data log was reviewed. no motor current (mc) spike observed. no position alarms seen. pump was on p-9 and having suction alarms and auto turning down the speed with no placement signal issues. no indication observed to confirm the cause of the low flow issue. possible clot seen stated in clinical account but not confirming cause of the issue, complaint device not returned for analyzing. the cause of the low pump flow issue cannot be determined since complaint device not returned for analysis and insufficient clinical data provided. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed.
9943 | 1220648-2025-27848 | 21909737 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21909737&pc=OZD | 04/02/2025 | 04/25/2025 |  | 04/25/2025 |  |  | 04/02/2025 | 04/25/2025 |  |  |  |  | 04/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | GM |  | 2025554972 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | Infusion or Flow Problem | [5 empty] | Hemodynamic instability | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and following the implant the impella cp pump experienced low purge flow coinciding with a high motor current. the device was replaced. however, the patient expired the following day. the patient was at gym swimming when the patient went into ventricular fibrillation arrest and coded. the patient aspirated a lot of pool water. the patient was brought into the hospital and was worked on throughout the night. the left main artery was stented and the decision made to place impella cp device. the patient was sent to the unit in critical condition on impella mcs. shortly after the implant of the impella cp, there was low purge flow coinciding with a high motor current. the pump flow was also noted to be 4.2 liters per minute for the maximum, minimum, and average values. the staff was advised that the pump could stop based on the noted values. the decision was made to replace the pump after just under three hours of support. the patient was cannulated for veno-venous extracorporeal membrane oxygenation (vv ecmo) and was sent back to the unit on support remaining in critical condition. care was continued. vv ecmo was switched to venoarterial ecmo. the patient was very critical, with very low pressures. the physician was aware and was using the impella cp device as a vent. the following day, as the patient continued to decompensate and the decision was made to withdraw care and the patient expired. | [5 empty] | Medical safety review of the event noted pump flow was compromised. the patient may not have received appropriate hemodynamic support. the pump was replaced and the patient expired approximately a day and a half later. the impact to outcome is unknown, and there is not enough information to exclude the initial (pump 1) impella device used as an associated factor. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
9944 | 1220648-2025-27849 | 21909776 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21909776&pc=OZD | 04/04/2025 | 05/16/2025 |  | 04/25/2025 |  |  | 04/04/2025 | 04/25/2025 |  |  |  |  | 04/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026644043 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient's groin site developed a worsening hematoma following impella cp implant. manual pressure was applied, but the bleeding persisted..blood products were given to counteract the bleed, but the decision was ultimately made to explant the pump. femostop was placed at the site to help manage the bleed. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of access site bleeding was most likely use issue since non-abiomed product was used which is not recommended per instructions for use. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9945 | 1220648-2025-27850 | 21909853 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21909853&pc=PYX | 03/26/2025 | 09/19/2025 |  | 04/25/2025 |  |  | 03/26/2025 | 04/25/2025 |  |  |  |  | 04/25/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025502673 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was escalated to bipella (impella 5.5 and rp flex) from an impella cp and during support, the patient had a gastrointestinal bleed. five units of packed red blood cells were given. heparin was not used in the purge. however, the representative stated the relationship to the impella is unknown. care was withdrawn and the patient expired. the patient's demise is associated with the impella cp device and captured under the manufacture report number 1220648-2025-27522-1. | The complainant reported that placement signal not reliable alarm occurred with loss of placement signal while the patient was on impella rp support. the placement monitoring alarms was disabled. medical safety reviewed the event and noted the patient was quite ill. the patient had significant hemolysis while using the impella cp device (first implanted impella) and required transfusion. the massive bleed occurred days later when on bipella (impella 5.5 and rp flex) which could have occurred for multiple reasons including significant hemolysis. impella rr flex is a serious injury. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. abiomed's medical safety officer review the patient's outcome and it was determined that the patient demise would be reported under the impella cp as there is not enough evidence to exclude the impella cp as an associated factor in the patient's outcome. the death is captured in the impella cp manufacture report number 1220648-2025-27522-1. the impella 5.5 was captured under manufacture device report (1220648-2025-27846). | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. no product was returned for evaluation. optical signal issue: data logs were reviewed and rt log at time of issue shows psnr alarm #1055. cvp goes out of range and raw electrical placement signal remains stable. optical 5s parameters are consistent with a fiber torque. no lt log was returned for analysis. clinical details noted that a psnr alarm was triggered on the rp flex pump, however, flow calculations were still intact after alarm. the root cause of the optical signal issue was most likely due to a damaged fiber as a result of a torqued fiber, however, the exact location and cause of the break could not be definitively determined as no product was returned for evaluation. vascular damage - peripheral vessel: clinical details noted a gi bleed was present while on rp flex support. patient received multiple units of prbcs. the root cause of the vascular damage -peripheral vessel cannot be definitively determined as the rp flex was inserted via the right axillary access site and limited clinical details were provided. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
9946 | 1220648-2025-27851 | 21909983 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21909983&pc=OZD | 03/27/2025 | 04/25/2025 |  | 04/25/2025 |  |  | 03/27/2025 | 04/25/2025 |  |  |  |  | 04/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026627467 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Malposition of Device | [6 empty] | Hematoma | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and experienced a hematoma and positioning issues. the patient would subsequently expire after 11 days of total mcs. the patient presented with a st-segment elevation myocardial infarction one week prior and an intra-aortic balloon pump (iabp) was placed. the patient was also on pressers but according to the team had mixed cardiogenic, septic shock. the sepsis improved. however, the patient remained on dobutamine, levophed and the iabp. the decision was made to place an impella 5.5 device as a bridge. the impella 5.5 was placed, and the patient was sent to the unit on impella mcs. approximately six days after start of the support, a hematoma developed and was managed with manual pressure and heparin was withheld. the following day positioning difficulty was encountered. there were impella position unknown alarms with reduced pulsatility. the position was verified 6.3 centimeters but the left ventricle was distended. the patient had severe mitral valve regurgitation. the patient was reintubated yesterday. pulmonary edema was worsening. goals of care leaned towards palliative care instead of left ventricular assist device as the patient was no longer a candidate for any advanced support. approximately four days later, the patient expired. care was withdrawn. | [5 empty] | Medical safety reviewed the event and noted the patient was in critical condition and already in septic shock mixed with cardiogenic shock. there was minimal bleeding (small hematoma with manual pressure). however, the main issue was the positioning. the physician was unable to get the pump in the correct position. there is not enough evidence to exclude impella as an associated factor. however, to note the demise outcome is very unlikely related to the impella device used. the patient was not a candidate for any advanced support for survival. the investigation for the access site bleeding has been completed. the impella device was not received from the customer. the root cause of the access site bleeding issue was not determined since the product was not returned and insufficient clinical details were provided.
9947 | 1220648-2025-27852 | 21910115 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21910115&pc=OZD | 03/29/2025 | 05/20/2025 |  | 04/25/2025 |  |  | 03/29/2025 | 04/25/2025 |  |  |  |  | 04/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025574127 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the insertion of an impella 5.5 device from an impella cp device for escalation in therapy for a patient who was admitted in with cardiogenic shock from acute myocardial infarction was unsuccessful. the plan was to explant the impella cp from the left femoral artery (lfa) and escalate to an impella 5.5 via right axillary artery. the physician stated the axillary pocket was deep and a more ¿direct¿ approach to the artery was used. the patient had an aortic valve replacement. the impella 5.5 (pump 1) was primed and setup without incidence. the physician was successfully able to cross the aortic valve with the 0.18 guidewire with the impella cp in place. the impella 5.5 (pump 1) was successfully loaded on 0.18 wire and entered the sheath without difficulty. however, when the physician attempted to advance the impella through the patients vascular, the impella would not drop down. despite multiple attempts to reposition the device, rotate etc. as well as guidewire change to a v18, the impella 5.5 (pump 1) was unable to advance and would ¿buckle¿ in the patient. upon removal a notable kink was observed on the catheter proximal to the motor housing. a new impella 5.5 device (pump 2) was opened, primed and setup without a problem. the impella cp was removed from the lfa and the new impella 5.5 (pump 2) was implanted through the right axillary graft and was unable to advance as well. consequently, the physician aborted implanting an impella device. the physician noted possible direct approach with second device being discussed. the patient was maintained on extracorporeal membrane oxygenation support and was switched to veno-arterial-venous. the patient was reported to be in stable condition following the event. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since the impella would not cross patient's vasculature. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. this is one of two devices associated with the event. refer to manufacturing report number 1220648-2025-27835 for the impella 5.5 device (pump 1) report. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
9948 | 1220648-2025-27853 | 21910384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21910384&pc=OZD | 04/16/2025 | 07/18/2025 |  | 04/25/2025 |  |  | 04/21/2025 | 04/25/2025 |  |  |  |  | 04/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1409579 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller (aic) had a self-check error when the console was turned on. after forcibly shutting down the device and turning it back on, the self-check was completed and the screen moved to the assistance preparation start screen. the aic was replaced and there was no patient harm. | [5 empty] | The investigation of automated impella controller (aic) - system self-check error has been completed. data analysis: the case related to the reported complaint was initiated on (b)(6) 2025. during the boot-up process, multiple error messages were observed, signaling communication issues within the aic. it was observed that all subprocesses within the aic were failed. following the error, the aic was forcefully shutdown and after the reboot, no further problems were seen. device analysis: the console was sent back to field service for further investigation and was tested. after performing manual multiple power cycles, but the issue could not be reproduced and no communication issues were observed. a thorough visual inspection of the internal components was conducted, but no physical problems were found. there was no interruption in the supply were detected between the power battery manager (pbm) and the 4-in-1 assembly. root cause: the root cause of the "system self check error" could not be determined due to the inability to reproduce | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
9949 | 1220648-2025-27855 | 21910780 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21910780&pc=OZD | 04/04/2025 | 05/27/2025 |  | 04/25/2025 |  |  | 04/04/2025 | 04/25/2025 |  |  |  |  | 04/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026635584 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, the pump came out of ventricle as evident by echocardiogram and impella in aorta alarm occurred. the pump was then explanted. no patient harm was reported. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. clinical states that after patient transfer to another hospital, echocardiogram showed pump pulled out with "impella in aorta" alarms on the console. pump was explanted since patient was stable. therefore, the root cause of the positioning issue was patient movement at the time of patient transfer to another hospital. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9950 | 1220648-2025-27856 | 21910828 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21910828&pc=OZD | 04/03/2025 | 06/02/2025 |  | 04/25/2025 |  |  | 04/03/2025 | 04/25/2025 |  |  |  |  | 04/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1663431 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Application Program Problem | Unintended Application Program Shut Down | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support and on day four of support, the automated impella controller (aic) failure red alarm occurred and required replacement of the aic. reportedly the patient tolerated the unplanned aic-to-aic exchange and was in stable condition following the event. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - controller failure (red alarm) has been completed. the console logs showed the qnx process io-graphics crashed which caused the aic software to reboot (systemidle:process io-graphics 339995 died abnormally). this reboot was accompanied by an unexpected shutdown alarm, and the pump flow decreased to zero before it was resumed at the previous p-level pump flow. therapy continued before the pump was unplugged, and the console was then safely powered down. device analysis: the failure mode was not reproduced during investigation on engineering lab bench. the io-graphics was force shutdown using a telnet command and the console responded the same as in the field. root cause: the cause of the aic issue (unexpected shutdown alarm) was the 3rd party hardware/software that runs io-graphics. no issues were identified within the aic console related to the reported controller failure alarm.
9951 | 1220648-2025-27857 | 21910994 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21910994&pc=OZD | 03/05/2025 | 06/16/2025 |  | 04/25/2025 |  |  | 04/03/2025 | 04/25/2025 |  |  | 01/31/2026 |  | 04/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025509325 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Adverse Event Without Identified Device or Use Problem | [5 empty] | Pulmonary Edema | [12 empty] | Required Intervention | [5 empty] | CENTRIMAG. | EXTRACORPOREAL MEMBRANE OXYGENATION. | [17 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella 5.5 and extracorporeal membrane oxygenation (ecmo) support ongoing for a patient being bridged to transplant. while on impella support the doctors reduced the ecmo until 3 liters per minute and weaned. after two days the patient got worst and they put an ecmo again. the lungs were totally white and the decided to implant a centrimag in the left ventricle and the patient lungs got better. the pump was explanted as they had bridged to transplant and the procedural outcome was the patient survived surgery. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of low pump flow has been completed. clinical states that the patient was on ecmo and impella and doctors had removed and re-inserted ecmo when the flows were thought to be causing lung issues. patient had lung edema and pump flows were thought to be wrong. pump was returned for investigation. the pump cannula and impellers were clean of biomaterial. the purge gap of the pump had biomaterial. no catheter kink observed. data logs were not returned for investigation. without the data logs, it is not possible to confirm if the biomaterial was ingested or was pulled in from sheath during explant. therefore, based on the clinical information and product evaluation, the root cause of the low pump flow issue was not determined since data logs were not returned to confirm biomaterial ingestion. d9 device returned to manufacturer date added h6 health effect - impact code 4641 and medical device problem code 1248 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27857.
9952 | 1220648-2025-27858 | 21910995 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21910995&pc=OZD | 04/03/2025 | 05/12/2025 |  | 04/25/2025 |  |  |  | 04/25/2025 |  |  |  |  | 04/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026630017 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Followup | [15 empty] | The investigation of low pump flow and thrombus has been completed. low pump flow: the returned pump was inspected and there was no biomaterial or physical abnormalities such as kinks or broken impeller blades. the pump was run in the lab and low pump flows did not reproduce. the clinical details mention loss of flows as protek duo was turned on. the data logs show an ingestion pattern soon after case start with a motor current spike, speed deviation and shift in placement signal. there were suction alarms and 'impella position in aorta' alarms and a slow decrease in pump flows. the root cause of the low pump flow was most likely biomaterial ingestion as evidenced by log pattern and returned product testing. as the necessary information was not provided, the root cause of the thrombus was not determined.
9953 | 1220648-2025-27859 | 21911143 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21911143&pc=OZD | 04/03/2025 | 04/25/2025 |  | 04/25/2025 |  |  | 04/03/2025 | 04/25/2025 |  |  |  |  | 04/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026624108 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 43 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a thrombus was discovered in the patient's left ventricle on the lower part of the impella during support. it was noted that the patient had been off heparin intermittently due to scheduled procedures. no additional effects were noted resulting from the thrombus and the decision was made to continue support. a follow up transesophageal echocardiogram was performed and the thrombus could no longer be found. although not verifiable, it was believed that the earlier observation may have been an artifact and not a thrombus. roughly five days later, the patient suffered a cerebral hemorrhage and passed away. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
9954 | 1220648-2025-27860 | 21911177 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21911177&pc=OZD | 04/04/2025 | 05/21/2025 |  | 04/25/2025 |  |  | 04/04/2025 | 04/25/2025 |  |  |  |  | 04/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026625142 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Defective Device | Infusion or Flow Problem |  |  |  |  | Embolism/Embolus | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump placed to support a patient admitted in with acute myocardial infarction/cardiogenic shock. it was reported that upon pump insertion, suction alarms persistent even after repositioning and turning down to p2. low purge flow was also noted. the pump was removed, inspected and attempted to reinsert after no obvious issues were noted. waveforms were not appropriate after restarting. a kink in the peel away sheath made it difficult for the medical doctor to advance the short 14 french peel away from the sheath. they elected to replace the pump and the patient was noted to be in stable condition. | [5 empty] | The pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | D6: explant date unknown. the investigation of low pump flow, saturated flow, and introducer damage ¿ mechanical has been completed. the pump was returned and inspected. there was dried blood on and around impeller, blocking purge gap, which cleared upon soaking revealing biomaterial. there were no other visual abnormalities. during reproduction attempt, there were motor current high pump stops before the pump was eventually able to run. the pump was run in the lab after soaking and there were saturated flows for a while flowing at 4.3l/min until likely some biomaterial was ejected and flows were at just around 4.0l/min. low pump flows did not reproduce. the pump was torn down to check for biomaterial and no biomaterial was observed in the motor, confirming biomaterial was contained to the impeller. no damage or scratches to the magnet was observed. the failure mode thrombus was added based on the findings during the investigation. the root cause of the introducer damage - mechanical was not determined as no product or images were returned for evaluation and limited information related to failure was provided. the root cause of the low pump flow was most likely biomaterial as evidenced by biomaterial on returned product and no low pump flow occurrence during lab run. the root cause of the saturated flow was determined to be biomaterial as evidenced by the biomaterial on the returned product and the confirmation from the tear down that there were no damage or scratches to the magnet as the necessary information was not provided, the root cause of the thrombus was not determined instructions for use (thrombus): impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." based on the investigation findings of thrombus, the following fields were updated from the initial manufacturer device report number 1220648-2025-27860. b1 revised to reflect both adverse event and product problem. b2 updated. h1 revised to serious injury. h6 health effect - clinical code 4582 was revised to embolus 4438. h10 instructions for use added.
9955 | 1220648-2025-27861 | 21911315 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21911315&pc=OZD | 04/04/2025 | 04/25/2025 |  | 04/25/2025 |  |  | 04/04/2025 | 04/25/2025 |  |  |  |  | 04/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025486446 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that a clot was noted during impella cp support. the clot was said to be keeping the mitral open, but was not interacting with the impella pump. the pump was turned back up and support continued. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smartassist catheter insertion section: axillary insertion of the impella cp with smartassist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
9956 | 1220648-2025-27863 | 21911735 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21911735&pc=OZD | 04/05/2025 | 04/25/2025 |  | 04/25/2025 |  |  | 04/05/2025 | 04/25/2025 |  |  |  |  | 04/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026637062 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was escalated to a new impella cp and while on the second pump it was reported that the patient's leg was bleeding from an above-knee amputation procedure. the amputation procedure was reported on the first pump as that is when limb ischemia took place however bleeding was noted while on the second pump. additionally, it was noted that the impella was repositioned and now measuring 50 cm. the impella was successfully weaned and explanted. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. there are two mdrs for the same patient however this report is for pump two (579562) and the other is for the first pump (555585). the first report is for the same patient mfr report #1220648-2025-27827. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9957 | 1220648-2025-27862 | 21911831 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21911831&pc=PYX | 04/04/2025 | 05/27/2025 |  | 04/25/2025 |  |  | 04/04/2025 | 04/25/2025 |  |  |  |  | 04/25/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026625112 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after impella rp flex was placed. the patient was restless and bleeding was observed at the impella access site. manual compression and hemostatic suture were done. it was further reported that hemolysis was observed as labs were hemolyzed. the impella rp flex was confirmed to be in good position. the impella rp flex p-level was dropped and the team will be monitoring the hemolysis. | [5 empty] | The investigation of hemolysis and access site bleeding has been completed. the data logs showed pump ran without issue during support. pump was returned for analysis. visual inspection found no issues on the pump. pump was conducted hemolysis test and passed the test. the root cause of access site bleeding was most likely anticoagulation management related since activated clotting time (act) was high when the bleeding occurred and no issue was found on the pump under product analysis. the root cause of hemolysis was most likely patient condition related since no issue was found on the pump under product analysis and the patient had renal insufficiency and on continuous renal replacement therapy (crrt) before placing impella. pump position was good. | The pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound .¿
9958 | 1220648-2025-27864 | 21912104 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21912104&pc=OZD | 04/04/2025 | 05/29/2025 |  | 04/25/2025 |  |  | 04/04/2025 | 04/25/2025 |  |  |  |  | 04/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1468360 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  |  | White | Malfunction | No | Yes | Application Program Problem | Battery Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella 5.5 pump for a patient awaiting either a ventricular assist device or a transplant. the pump was supporting when the team had a controller issue in which the battery failed. the console was swapped. there was no patient harm. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - battery charging/other issues has been completed. data analysis: 5.5 pump 551855 was connected to (b)(6) on 4/2. the console performed as expected until 9:17:05 when there were power battery manager (pbm)1 and pbm2 communication controller errors, the sau_purge process was power cycled, there were sau_eep communication difficulties, and the epm had a communication controller error and was power cycled. the real time (rt) log at the moment of the communication losses confirms that purge flow drops to 0 while the purge is reset. what the user saw as 0% battery power, was actually an inability to communicate the battery power but the battery was functioning as expected during this time as confirmed in the lt power logs. device analysis: (b)(6) was returned for investigation and tested thoroughly on and off ac power but the failure mode did not reproduce. root cause: there was no issue with battery charging on (b)(6) during this case. the cause of the widespread communication loss on (b)(6) resulting in loss of communication controller errors was not determined as the failure mode was not reproduced but could have been caused by loss of power (pbm) or communication failures.
9959 | 1220648-2025-27865 | 21912227 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21912227&pc=OZD | 04/05/2025 | 04/25/2025 |  | 04/25/2025 |  |  | 04/05/2025 | 04/25/2025 |  |  |  |  | 04/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025608757 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 37 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombocytopenia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the 37-year-old male patient on impella cp support had onset of heparin-induced thrombocytopenia (hit). the patient was switched from heparin to argatroban. the next day the patient¿s platelets were 65. the patient continued to require several units of blood. it was further reported that the patient developed some mild hemolysis, so p-level was reduced. the patient was eventually weaned and explanted. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9960 | 1220648-2025-27866 | 21912321 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21912321&pc=OZD | 04/05/2025 | 05/27/2025 |  | 04/25/2025 |  |  | 04/05/2025 | 04/25/2025 |  |  |  |  | 04/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025582378 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Hemolysis | Gastrointestinal Hemorrhage | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that there was concern for hemolysis for a patient on impella 5.5 support during a high risk cardiac procedure. after the p-level was increased, the hemolysis was resolved. the patient received one unit of blood and one unit of platelets. additionally, the patient kept having atrial fibrillation during impella 5.5 support. the physician stated that the patient had a history of atrial fibrillation. it was further reported that the patient developed a gastrointestinal bleed on day 15 of impella support. four units of blood and one unit of platelets, and one unit fresh frozen plasma were given. the patient was eventually successfully weaned and explanted. | [5 empty] | The investigation of hemolysis, vf/vt/af/at, and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. as per the data log review and clinical information provided, the root cause of the hemolysis issue was most likely patient condition related as it was resolved by giving blood volume. as the necessary information was not provided, the root cause of the vt/vf and vascular damage peripheral vessel was not determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9961 | 1220648-2025-27867 | 21912667 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21912667&pc=OZD | 04/05/2025 | 04/25/2025 |  | 04/25/2025 |  |  | 04/05/2025 | 04/25/2025 |  |  |  |  | 04/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025583014 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a 68 year old male patient had a cp pump placed for support due to an st-elevated myocardial infarction. patient was on multiple resuscitation meds and placed on the cp and ecmo via femorals that had pad. the pump was placed and ischemia occurred. the team placed a fem-fem to support the flow but, this did not resolve. the pump was explanted; however, the patient continued on ecmo support. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9962 | 1220648-2025-27868 | 21913042 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21913042&pc=PYX | 04/05/2025 | 04/25/2025 |  | 04/25/2025 |  |  | 04/05/2025 | 04/25/2025 |  |  |  |  | 04/25/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2026635487 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the45 year old patient on bipella support had suction during ambulation that resolved on its own. the next day, intermittent suction events occurred overnight while the patient laid flat. the rp flow was reduced. on day six of impella support, the rp was explanted. echo the day before showed the inlet pointed toward the mitral but no suction events. the patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
9963 | 1035166-2025-00027 | 21913058 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21913058&pc=DYB | 03/28/2025 | 06/30/2025 |  | 04/25/2025 |  |  | 03/31/2025 | 04/25/2025 |  |  |  |  | 04/25/2025 |  | INTRODUCER, CATHETER |  |  | INTRODUCER KIT FOR IMPELLA® |  |  | DYB |  |  | 0052-3006 |  |  | OSCOR INC |  |  | US |  | DP-19591 |  | 0052-3006 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Distributor |  |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that at the start of the impella 5.5 implant, oozing from axillary kit sheath without impella yet in place occurred. the oozing was located at the blue ring at pre perforated area. the sheath was split. the sheath was exchanged, and impella was inserted successfully. there was no harm to the patient. | [5 empty] | Conclusion not yet available, evaluation in process. a follow-up report will be submitted as soon as the investigation is complete. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report. | The following sections were updated in follow-up 1. b4, g3, g6, h2, h6 & h11. no product was provided for analysis. procedure completed successfully with different device. there was no adverse event with the patient reported. customer reported that patient had lost 10ml of blood and no surgical or medical intervention required to control the blood loss. the device history records were reviewed to confirm that the device passed all applicable in-process and final inspections. no product was returned to oscor inc. for evaluation; however, a complaint notification was provided. without the return of the actual device in question for evaluation, oscor inc. is unable to determine the exact cause for this incident. at this time, it is not possible to assign a definitive root cause for the event as reported. the complaint is non-verifiable as the product was not returned for evaluation. no corrections or corrective actions will be taken. device was not returned for analysis and review of the manufacturing documents did not reveal any discrepancies that would have contributed to this event. this complaint is non-verifiable. oscor inc. is submitting the above report to comply with 21 cfr part 803, the medical device reporting regulation. the report is based upon information obtained by oscor inc. which the company may not have been able to investigate or verify prior to the date the report was required by fda. this report does not constitute an admission that the device, oscor inc., or its employees, caused or contributed to the event described in this report. nor does this report reflect a conclusion by fda, oscor inc., or its employees, that the report constitutes an admission that the device, oscor inc., or its employees caused or contributed to the event described in this report.
9964 | 1220648-2025-27794 | 21898707 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21898707&pc=OZD | 03/25/2025 | 05/08/2025 |  | 04/24/2025 |  |  | 03/25/2025 | 04/24/2025 |  |  |  |  | 04/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026614242 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and experienced a stroke and subsequently expired. the patient was an athlete who coded while training for a triathlon and was initially implanted with an impella cp pump for mcs. the impella cp supported the patient for approximately six hours before being removed and escalated to an impella 5.5 device at the tertiary center for increase in support. two days after start of the impella 5.5 support, a brain bleed prompted the team to stop systemic anticoagulation. however, two days later, the patient expired. care was withdrawn and organs harvested after a total of four days on support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of stroke/cva has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
9965 | 1220648-2025-27795 | 21898856 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21898856&pc=OZD | 04/03/2025 | 04/24/2025 |  | 04/24/2025 |  |  | 04/03/2025 | 04/24/2025 |  |  | 01/31/2026 |  | 04/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025509682 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that during attempted implantation of an impella cp pump, the 14 french peel-away sheath was placed in the right femoral artery, and later was found to be occlusive via an angiogram. the doctor performed an angiogram of the left femoral artery and determined it to also occlusive and unable to use. the implant was aborted. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
9966 | 1220648-2025-27796 | 21899068 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21899068&pc=OZD | 11/04/2024 | 04/24/2025 |  | 04/24/2025 |  |  | 11/04/2024 | 04/24/2025 |  |  |  |  | 04/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025562139 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  |  | Black Or African American | Malfunction | No | Yes | Positioning Failure | Detachment of Device or Device Component | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient that was found to be in cardiogenic shock on inotropes was taken to the operating room for emergent impella 5.5 placement. during support it was noted that the impella position on a echocardiogram was not optimal (rotated up toward mitral valve). during the pump reposition in the unit the impella came out of the left ventricle and the doctor was unable to recross. additionally, the posterior part of the sterile sleeve detached. the decision was made to replaced with a new impella 5.5. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed.
9967 | 1220648-2025-27797 | 21899169 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21899169&pc=OZD | 03/13/2025 | 04/24/2025 |  | 04/24/2025 |  |  | 03/13/2025 | 04/24/2025 |  |  |  |  | 04/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026612951 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that the decision was made for impella cp placement on a patient in cardiogenic shock. the patient went in to ventricular fibrillation, after 15 minutes of resuscitation the decision was made to end all therapies and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The device has been received for evaluation. upon completion of the investigation a supplemental mdr will be filed. section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9968 | 1220648-2025-27798 | 21899259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21899259&pc=OZD | 04/03/2025 | 05/05/2025 |  | 04/24/2025 |  |  | 04/03/2025 | 04/24/2025 |  |  |  |  | 04/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026612792 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | EXTRACORPOREAL MEMBRANE OXYGENATION | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella cp along with extracorporeal membrane oxygenation (ecmo) was implanted in a patient scheduled for transcatheter aortic valve replacement. venoarterial ecmo was cannulated and then an impella cp was placed for venting. the patient was taken to the unit to recover. later, the patient was taken to the operating room for ecmo decannulation. after decannulation and removal of the distal perfusion catheter, there was concerned for distal perfusion to the right foot. upon review, vascular recommended removing the impella cp after returning to the unit. impella cp weaned to p2 and pulled across the valve at p1. repo sidearm was wired with .035¿. the impella cp was turned to p0 and pulled out maintaining arterial access with the sidearm wire. the site used perclosed twice. the patients outcome is stable. | [5 empty] | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9969 | 1220648-2025-27799 | 21899260 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21899260&pc=OZD | 03/15/2025 | 04/24/2025 |  | 04/24/2025 |  |  | 03/15/2025 | 04/24/2025 |  |  |  |  | 04/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026622624 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the 55-year-old male patient on impella 5.5 support had a small hematoma at the insertion site. the site was monitored, and anticoagulant was restarted in the afternoon. it was further reported that the patient was walking daily and impella was repositioned on day five of support. the patient remained on support. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿.
9970 | 1220648-2025-27800 | 21899313 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21899313&pc=OZD | 04/14/2025 | 07/10/2025 |  | 04/24/2025 |  |  | 04/21/2025 | 04/24/2025 |  |  |  |  | 04/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2025594844 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that a 66-year-old male patient on 5 day support with the cp and ecmo with cardiogenic shock and cardiomyopathy. the pump support was then explanted and upon explant the team observed thrombus burden. there was no intervention noted. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-27800. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27800 as it is unknown if the initial reporter also sent the report to the food and drug administration.
9971 | 1220648-2025-27801 | 21899386 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21899386&pc=OZD | 04/03/2025 | 06/10/2025 |  | 04/24/2025 |  |  | 04/03/2025 | 04/24/2025 |  |  |  |  | 04/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026631978 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | TRANSCATHETER AORTIC VALVE REPLACEMENT. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that upon implantation of an impella cp, the pump stopped due cannula kinking at the aorta. due to another company's transcatheter aortic valve replacement valve they were not able to be insert any more to straighten out the cannula. attempts were made at repositioning and titrating the p level but the motor current was not recovered. the impella cp was explanted and a balloon pump was inserted. the patient survived. | [5 empty] | Device inserted over tavr valve easy with no resistance and the kink was observed after the pump was started so we don't know exactly what caused the kink. despite repositioning pump and adjusting p- level flow issue not resolved. complaint device not returned for analyzing. data log not available to review. low pump flow: the cause of the low pump flow was a reported kinked cannula. cause of the kink was unknown due to insufficient clinical details. product damaged (cannula kink): the cause of the product damaged (cannula kink) was a reported kinked cannula. cause of the kink was unknown due to insufficient clinical details. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of product damage (cannula kink) and low pump flow has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
9972 | 1220648-2025-27802 | 21899463 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21899463&pc=OZD | 03/31/2025 | 05/27/2025 |  | 04/24/2025 |  |  | 03/31/2025 | 04/24/2025 |  |  |  |  | 04/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026615876 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was admitted following stress test and echocardiogram. decision was made for impella cp after they were unable to stent during percutaneous coronary intervention to right coronary artery. the impella was placed successfully, however during support it was reported of slight ooze from impella site, redressed site for angle matching. the urine was pink tinged and an echocardiogram was obtained for repositioning. the impella was pulled back 5cm, the urine cleared and oozing improved. the impella was successfully weaned and explanted. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of positioning issues and access site bleeding has been completed. the impella device was not returned for evaluation. access site bleeding - major: on the same day as device insertion, oozing from repositioning sheath site. perclose sutures tightened. angle matching gauze changed to match angle of insertion. probelm was noted as resolved the next day. product was not returned. the cause of the access site bleeding - major was most likely use related as the tightening of the perclose sutures and angle matching resolved the issue. positioning issue: on the night after device insertion, roughly 12 hours hours later, impella was pulled back 5cm. urine cleared. pump was not returned. data log review showed pump position alarms that resolve the night impella repositioning took place. the cause of the positioning issues was not determined since it is not known how the pump became deep necessitating the repositioning. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9973 | 1220648-2025-27803 | 21899501 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21899501&pc=OZD | 04/18/2025 | 07/16/2025 |  | 04/24/2025 |  |  | 04/18/2025 | 04/24/2025 |  |  |  |  | 04/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1666136 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that when replacing the purge cassette, the automated impella controller (aic) screen displayed "purge not detected" along with the "procedure not completed alarm.¿ the purge air removal procedure was completed, but the problem did not improve. the aic was exchanged, and a new purge cassette was used. priming was performed, and the pump was connected to the new aic and operated normally. there was no patient harm. | [5 empty] | The impella console and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. data analysis: cp pump was connected to the console on 4/14. on 4/18, the user performs a cassette change from the new purge cassette can be seen inserted multiple times along with the purge cassette but the console does not continue through the purge wizard. eventually this triggers consecutive log entries of "(b)(6) 2025 12:12:29 imcgui: updateguipurgechange: pc_display_purge_fluid_detect_error 5" then "(b)(6) 2025 12:12:34 imcgui: processpurgechangeerrorfluidnotdetectedoksoftkey: ok soft key: 4024". the user continues to get fluid not detected entries and was unable to complete the procedure on this console. the real time (rt) logs for this period confirm that during this period of purge fluid not detected, the purge flow ticks are fast primed 7 times but the purge flow is 0ml/hr and purge pressure is 0 mmhg. device analysis: the failure mode was not reproduced during testing. the purge unit was taken apart and examined, but there was nothing that would have prevented pressure and flow (motor1 and motor2). root cause: there was no purge disc/flag issue found during the case. the cause of the purge issues could not be determined because the failure mode was not reproduced.
9974 | 1220648-2025-27804 | 21899656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21899656&pc=OZD | 04/03/2025 | 05/14/2025 |  | 04/24/2025 |  |  | 04/03/2025 | 04/24/2025 |  |  |  |  | 04/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026656037 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was implanted with an impella cp and exhibited hemolysis and limb ischemia. the impella cp was removed and the patient was implanted with an impella 5.5. the patient survived. | [5 empty] | The investigation of hemolysis and limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9975 | 1220648-2025-27805 | 21899671 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21899671&pc=OZD | 03/31/2025 | 05/19/2025 |  | 04/24/2025 |  |  | 03/31/2025 | 04/24/2025 |  |  |  |  | 04/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025597731 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following insertion, an impella 5.5 pump was unable to advance through the patient's vasculature. after multiple failed attempts, the implant was aborted. it was noted that the surgeon was aware of the patient's smaller caliber vessels with calcification prior to the attempts. there was no noted patient injury. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since the patient had small vessels and calcification. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
9976 | 1220648-2025-27806 | 21899744 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21899744&pc=OZD | 04/03/2025 | 04/24/2025 |  | 04/24/2025 |  |  | 04/03/2025 | 04/24/2025 |  |  |  |  | 04/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025604076 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella cp and exhibited increased leg pain and loss of pulses in the foot. the impella was weaned and explanted. the patient survived. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9977 | 1220648-2025-27807 | 21899818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21899818&pc=OZD | 03/31/2025 | 05/08/2025 |  | 04/24/2025 |  |  | 03/31/2025 | 04/24/2025 |  |  |  |  | 04/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026634126 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during the implant, intra-aortic balloon pum/ (iabp) was removed over a 0.018 wire and the 14fr x 13cm sheath was placed. while the patient was in the unit their groin was bleeding and a large hematoma was noted. the doctor made the decision to replace the impella. groin bleeding was controlled. a new impella was implanted in the left femoral. the doctor felt that upon removing iabp, a vessel was struck and broke off some calcium which caused the bleeding. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation has been completed. no product was returned. the root cause of vascular damage was most likely use issue since upon removing iabp, a vessel was struck and broke off some calcium which caused the bleeding. | Investigation results were provided in the initial manufacturer device report 1220648-2025-27807. the following codes were added to reflect the investigation findings h6 codes 3221 and 4315 were reported incorrectly on manufacturer device report 1220648-2025-27807. new code was added to type of investigation, investigation findings, and investigation conclusions.
9978 | 1220648-2025-27808 | 21899975 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21899975&pc=OZD | 04/03/2025 | 05/22/2025 |  | 04/24/2025 |  |  | 04/03/2025 | 04/24/2025 |  |  |  |  | 04/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026635366 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that an impella 5.5 pump was placed to support a patient admitted in with cardiomyopathy, valvular disease, and other cardiac risk factors. upon advancing into the left ventricle, the catheter kinked and failed to deliver. the pump was removed and a new impella 5.5 pump was placed. the patient survived. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues and product damage (cannula kink) has been completed. the returned product was visually inspected and a cannula kink and catheter kink under the motor housing was observed. delivery issue: extremely difficult to advance wire to left ventricle (lv). second pump able to be delivered with second wire and buddy wire. the root cause of the delivery issue was most likely patient condition related to the patient's vessel tortuosity. product damage (cannula kink): impella cannula and catheter kinked/buckled at turn during both insertion attempts. the root cause of the product damage cannula kink issue was most likely patient condition related to the delivery issues requiring multiple attempts related to the patient's vessel tortuosity. product damage catheter kink: failure mode added based on returned product observation of catheter kink under the motor housing. the root cause of the product damage catheter kink issue was most likely patient condition related to the delivery issues requiring multiple attempts related to the patient's vessel tortuosity. d9 device returned to manufacturer date added.
9979 | 1220648-2025-27809 | 21900059 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21900059&pc=OZD | 03/28/2025 | 05/02/2025 |  | 04/24/2025 |  |  | 03/28/2025 | 04/24/2025 |  |  |  |  | 04/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026631942 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that a patient experienced some oozing at their access site during impella 5.5 support. bival was held temporarily and manual pressure was applied to manage the condition. a few days later, the patient noted some pain at their site. no intervention was required to address the pain. support continued with the implanted pump. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27809 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smart assist for use during cardiogenic shock, section: potential adverse events (unites states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿
9980 | 1220648-2025-27810 | 21900067 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21900067&pc=OZD | 04/01/2025 | 04/24/2025 |  | 04/24/2025 |  |  | 04/01/2025 | 04/24/2025 |  |  |  |  | 04/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026629691 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient developed ischemia in their right leg following impella cp implant. the patient was escalated to impella 5.5 support. however, the condition remained. the patient's outcome is critical. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9981 | 1220648-2025-27811 | 21900199 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21900199&pc=OZD | 03/31/2025 | 06/03/2025 |  | 04/24/2025 |  |  | 03/31/2025 | 04/24/2025 |  |  |  |  | 04/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026633555 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support, and during support, a pump stop occurred. the pump was restarted, but the pump immediately stopped again. the pump was then removed and there was no patient harm. | [5 empty] | The data logs from the field were unable to be recovered. the pump was returned and inspected. inspection revealed biomaterial found on the impellor blades. the pump was attempted to be restarted but impella stopped -motor current (mc) high alarms were triggered. the biomaterial was then removed and the pump was able to be run for 10 minutes at nominal parameters. the cause of the pump stop was most likely biomaterial due to the biomaterial found on the impellor blades that recreated the impella stopped - mc high alarms reported in the field. after removing the biomaterial, the pump was able to be run for 10 minutes at nominal parameters. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. impella 5.5 with smartassist for use during cardiogenic shock, section: direct aortic insertion: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
9982 | 1220648-2025-27812 | 21900305 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21900305&pc=PYX | 04/03/2025 | 05/02/2025 |  | 04/24/2025 |  |  | 04/03/2025 | 04/24/2025 |  |  |  |  | 04/24/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1268753 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA RP FLEX SN (B)(6) | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient implanted with an impella rp flex support was transferred from one hospital to a second hospital. at the second hospital a new aic was started, and it alarmed that the placement signal was not reliable. hemodynamics were assessed and positioning was assessed. after repositioning the rp flex, the alarm remained. the impella rp flex was switched to a new console and placement signal resolved. | [5 empty] | The investigation of optical signal issue has been completed. the console logs from the reported day of event are consistent with complaint. the real time log data shows that placement signal was pulsatile (despite invalid optical signal) however, its pulsatility came from the electrical pressure sensor, and overall placement signal was inaccurate. characteristics of optical data - low signal not reliable (snr), low contrast, lamp at 100%, maxed-out gain of 2.6 - are consistent with an air gap. console ic4779 was tested and initial measurements of ¿5s¿ parameters using a fiber-coupled test cavity showed parameters within specification. when using a pump (attached via pump connector), the parameters - in particular snr ¿ did not meet specification. inspection of the optical pump connector revealed a piece of debris (presumably a broken piece of pump plug) lodged in the blue receptacle, preventing pump plug from being inserted fully, therefore creating a large air gap. when the piece of debris was removed, optical ¿5s¿ parameters were measured to be within specification (using a pump). the cause of the issue was a defective blue receptacle. d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2025-27812. | The automated impella controller was returned for investigation. should the device or any new information be received, a supplemental mdr will be filed.
9983 | 1220648-2025-27813 | 21900432 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21900432&pc=OZD | 04/04/2025 | 04/24/2025 |  | 04/24/2025 |  |  | 04/21/2025 | 04/24/2025 |  |  |  |  | 04/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025574230 |  | 1000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient on impella 5.5 expired on support. the cause of death was circulatory failure due to infection. the relationship between death and impella is unknown. although the possibility of infection due to the device was considered, the patient's underlying heart function was very poor and the patient was in a state where he was susceptible to infection. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components: "the infusion filter prevents bacterial contamination and air from entering the purge lumen." section: warnings and cautions: warnings: "to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers."
9984 | 1220648-2025-27814 | 21900554 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21900554&pc=OZD | 04/03/2025 | 04/24/2025 |  | 04/24/2025 |  |  | 04/03/2025 | 04/24/2025 |  |  |  |  | 04/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026631892 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was implanted with an impella 5.5 as bridge to transplant. during support it was noted that the patient had mild shoulder pain. the impella was functioning well at p5. the physician increased the impella to p6 however, the patient had short run of non-sustained ventricular tachycardia so it was decreased back to p5. additionally, the patient's urine was pink tinged. the team believed it was from a traumatic insertion. the patient successfully received the heart transplant. | [5 empty] | The impella 5.5 was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist system. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9985 | 1220648-2025-27815 | 21900618 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21900618&pc=OZD | 03/31/2025 | 05/19/2025 |  | 04/24/2025 |  |  | 03/31/2025 | 04/24/2025 |  |  |  |  | 04/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026619204 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. after three days of support, there were signs of hemolysis. the hemolysis prompted a call to the customer support center for support. the team observed urine color change and lactate dehydrogenase labs rising. the echocardiogram imaging had poor windows. the pump remained on for support despite the signs of hemolysis. the signs were noted to be improving and the patient is stable. | [5 empty] | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b2 was inadvertently left blank on the initial manufacturer device report number 1220648-2025-27815. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
9986 | MW5169498 | 21900892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21900892&pc=OZD | 04/19/2025 | 04/21/2025 |  | 04/24/2025 |  |  |  | 04/24/2025 |  |  |  |  | 04/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | 1000100 |  |  | ABIOMED, INC. |  |  | US | [5 empty] | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Voluntary report | Yes | Nurse | [6 empty] | 36 YR |  | Female | 79 KG |  |  | Malfunction | No | Yes | Fluid/Blood Leak | Defective Component | [5 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | [9 empty] | Pt received impella 5.5 in operating room. at the conclusion of the operation, it was noted that there was bleeding at the interface of the cable and the red plug. due to concern about damage to the inner mechanics of the pump, the device was replaced. rep notified. suspected to be a manufacturer defect. extended surgery time by over an hour.
9987 | 1220648-2025-27816 | 21901005 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21901005&pc=OZD | 03/31/2025 | 05/23/2025 |  | 04/24/2025 |  |  | 03/31/2025 | 04/24/2025 |  |  |  |  | 04/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026612532 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that a patient had a coronary artery bypass graft and a ventricular tachycardia (vt) ablation the day before. he was increasingly hemodynamically unstable and showed reduced left ventricle (lv) function. the decision was made for impella 5.5. the impella was implanted without complications. during support it was reported that two days after the impella implantation, a left ventricle thrombus was seen on computed tomography (ct). this was not present/not visible/not known before implantation. additional information was received that later indicated the following: the physician reported about a new ct scan. there was no lv thrombus and they stated that the first diagnosis was wrong. the patient had prolonged vt episodes the day before. ultimately, the patient expired on support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of vt/vf/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
9988 | 1220648-2025-27817 | 21901182 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21901182&pc=OZD | 03/29/2025 | 05/19/2025 |  | 04/24/2025 |  |  | 03/29/2025 | 04/24/2025 |  |  |  |  | 04/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025610766 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the pump was explanted and replaced in under a half an hour due to sensor failure, that may have taken place with pump insertion. the team replaced the pump and support was continued. there was no patient harm. | [5 empty] | The investigation of optical signal issue has been completed. no product or data logs were returned for evaluation. clinical details noted that shortly after pump was inserted and started, placement signal and lv waveforms were not reliable and giving inaccurate numbers. the root cause of the optical signal issue could not be definitively determined as no product or data logs were returned for analysis. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9989 | 1220648-2025-27818 | 21901376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21901376&pc=OZD | 03/11/2025 | 05/13/2025 |  | 04/24/2025 |  |  | 03/11/2025 | 04/24/2025 |  |  |  |  | 04/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026612685 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient experienced bleeding from their access site during impella cp support. an additional suture was placed at the site to manage the condition. support continued uninterrupted and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the access site bleeding issue was most likely use related, an additional suture (impella groin side) was placed because of bleeding to achieve hemostasis. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9990 | 1220648-2025-27819 | 21901398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21901398&pc=OZD | 03/29/2025 | 05/28/2025 |  | 04/24/2025 |  |  | 03/29/2025 | 04/24/2025 |  |  |  |  | 04/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026619295 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | Septic Shock | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during support, oozing was observed at the proximal most portion of the repositioning sheath. no crack or damage was observed. the staff monitored and impella will be replaced if needed. additionally, the patient had diminished left leg pulses. a femorofemoral bypass was placed. however, the pump was still explanted. furthermore, the patient also experienced septic shock and the cause is unknown. | [5 empty] | The investigation of access site bleeding, limb ischemia, and infection/sepsis has been completed. the pump was received for investigation. the reported oozing from the proximal most potion of the repositioning sheath where the t-lock connects. there were no cracks or damage seen. pump was returned and leak was reproduced. the root cause of the access site bleeding - major was most likely a damaged valve since the leak was reproduced on the returned pump. the root cause of the limb ischemia was not determined due to insufficient clinical details provided. the root cause of the infection was not determined due to insufficient clinical details provided. h6 type of investigation 4114 was erroneously entered on the initial manufacturer device report number. h10 narrative erroneously stated that the device was not received on the initial manufacturer device report number. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system; section: general patient care considerations; ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist: section: table 3.2 impella catheter components, ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
9991 | 1220648-2025-27820 | 21901676 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21901676&pc=OZD | 09/24/2024 | 04/24/2025 |  | 04/24/2025 |  |  | 09/24/2024 | 04/24/2025 |  |  |  |  | 04/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532940 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in cardiogenic shock from an acute myocardial infarction was placed on extracorporeal membrane oxygenation and was implanted with an impella cp device for mechanical circulatory support to unload the left ventricle. however, after implant of the impella cp device, it was indicated that the pump failed. the impella cp device was explanted in the procedural area and replaced with another impella cp device, which supported the patient for approximately two days before being explanted and escalated to an impella 5.5 device. the patient was noted to have survived the explant and was noted to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
9992 | 1220648-2025-27821 | 21901681 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21901681&pc=OZD | 04/03/2025 | 06/02/2025 |  | 04/24/2025 |  |  | 04/03/2025 | 04/24/2025 |  |  |  |  | 04/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1175558 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Decrease in Pressure | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during start up, the automated impella controller skipped the priming step. no patient was involved in this event. | [5 empty] | The automated impella controller was returned for investigation. should any new information be received, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - a/c power issues has been completed. the real time (rt) log data associated with pump 557054 on ic3502 indicates that at this point purge line was already primed, which explains why steps 4 and 5 of case start wizard were skipped. most likely because operator expected to see prompts to prime the pump, a purge change wizard was manually launched and pump was unplugged. console was then power-cycled, case start wizard was launched and pump was connected. the cause for skipping purging prompts of case start wizard was the fact that purge line was already primed, and pump was initialized. there was no ac power issue found during the case. the device functioned/operated to specification. d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2025-27821.
9993 | 1220648-2025-27822 | 21901861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21901861&pc=OZD | 04/03/2025 | 06/02/2025 |  | 04/24/2025 |  |  | 04/03/2025 | 04/24/2025 |  |  |  |  | 04/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1175559 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Decrease in Pressure | Application Program Problem | Device Sensing Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during start up, the automated impella controller skipped the priming step. no patient was involved in this event. | [5 empty] | The automated impella controller was returned for investigation. should any new information be received, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - purge issue ¿ other has been completed. when pump (b)(6) was connected to (b)(6) for the first time, the purge line was most likely already primed, since case start wizard was previously launched and priming had started (there was no pump connection recorded, and wizard was closed, then launched again). during, and in-between subsequent attempts to use the pump, console presented with controller error alarm (defective optical sensor lamp), and its logs recorded multiple communication issues with optical bench: events 1035, 1043 (due to optical bench resetting), and later events 1036 (due to momentary loss of light) and invalid_bad_pressure_sample_mask errors. calibration of g0 was started during 1st attempt but failed (due to unreliable placement signal), and again during 2nd attempt where it succeeded. at this point the pump was initialized, and calibrated g0 was recorded in pump eeprom. the 3rd and 4th attempt failed, due to optical system errors and the fact that purge line was already primed. after 5 failed attempts, pump (b)(6) was transferred to console (b)(6). it is most likely that optical signal issues were the primary reason for transferring the pump to backup console. the cause for skipping purging prompts of case start wizard was the fact that purge line was already primed, and pump was initialized. the cause of optical signal issues and controller alarm was defective optical bench, intermittently failing to provide light to the optical system. d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2025-27822. h6 medical device problem code 2880 was erroneously entered on the initial manufacturer device report number 1220648-2025-27822.
9994 | 1220648-2025-27823 | 21902009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21902009&pc=OZD | 04/03/2025 | 05/20/2025 |  | 04/24/2025 |  |  | 04/03/2025 | 04/24/2025 |  |  |  |  | 04/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026623813 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male | 61 KG | Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | Followup | [8 empty] | It was reported that placement of an impella cp was noted to be shallow. attempts at repositioning the device were unsuccessful due to patient anatomy. additionally, the patient had limb ischemia and subsequently, the pump was explanted and the patient survived. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation of positioning issues and limb ischemia has been completed. the impella device was not returned for evaluation. the data logs confirmed pump was in aortic area in early and middle of the case but resolved. no other issues were found on the logs. the root cause of positioning issue was most likely patient condition related since the patient had torturous aorta and heart appears horizontal. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
9995 | 1220648-2025-27824 | 21902082 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21902082&pc=OZD | 04/06/2025 | 05/19/2025 |  | 04/24/2025 |  |  | 04/06/2025 | 04/24/2025 |  |  |  |  | 04/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025569580 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was on impella cp support. a large hematoma and bleeding at the impella site was noted. it was resolved by changing the dressing, holding pressure and pausing heparin for some time. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of access site bleeding has been completed. it was unknown whether the hemostasis was achieved at the bleeding site. hence, the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
9996 | 1220648-2025-27825 | 21902578 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21902578&pc=OZD | 04/03/2025 | 06/16/2025 |  | 04/24/2025 |  |  | 04/03/2025 | 04/24/2025 |  |  |  |  | 04/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025585996 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the impella cp pump was kinked. while changing the sheath the impella went into the ventricle and kinked. the pump was exchanged for another impella cp and the patient survived. | [5 empty] | The investigation of positioning issues and product damage (cannula kink) has been completed. the impella device was not returned for evaluation. the pump was implanted with no issues. while changing the peel away sheath and inserting the repo sheath, the pump went into the ventricle and kinked the cannula, requiring the pump to be explanted and replaced. the physician reported friction while inserting the repo. no vessel conditions noted. excessive force stated to have been applied. no product returned. the cause of the product damage cannula kink was most likely use related as excessive force was used during the exchange of the repo sheath leading to malpositioning and cannula kink. the cause of the positioning issue was most likely use related to the pump going into the ventricle during the exchange of the repo unit with excessive force applied. h6 medical device problem code 3009 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27825. | The impella cp was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
9997 | 1220648-2025-27826 | 21902962 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21902962&pc=OZD | 04/03/2025 | 05/22/2025 |  | 04/24/2025 |  |  | 04/03/2025 | 04/24/2025 |  |  |  |  | 04/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025589377 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  |  | White | Death | Yes | Yes | Restricted Flow rate | Pumping Stopped | Malposition of Device |  |  |  |  | Atrial Fibrillation | Cardiac Arrest | [11 empty] | Life Threatening | Required Intervention | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient (unknown ethnicity) with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support and experienced a positioning issue, suction events, atrial fibrillation with cardioversion by implantable device and a pump stop with a restart. the patient would eventually expire. the patient presented with acute on chronic heart failure exacerbation status post right heart catheterization. patient was on inotropic support and afterload reduction drips now with intra-aortic balloon pump (iabp) being evaluated for left ventricular assist device (lvad). there were elevated pulmonary artery pressures despite iabp, and the decision was made to implant an impella 5.5 device for improved left ventricle off-loading and pulmonary artery pressures optimization. approximately seven days after start of the support, the impella 5.5 pump required repositioning due to placement signal low alarm as well as one suction, secured at 45cm. the patient had been cardioverted due to atrial fibrillation. the following day there was another suction event the patient encountered when leaning forward. the performance level was decreased from p-9 to p-8 with resolution. two days later, it was noted the patient was unfortunately no longer a candidate for advanced therapies given persistent fever and multi-organ dysfunction. the patient remained under supported despite being on performance level p-8 on near maximal pharmacological support. impella support continued per the plan. however, later that afternoon, the impella 5.5 device pump stopped alarm triggered three times, the pump stopped and restarted at performance level p-8. it was further explained that the event occurred while the patient was sitting in the chair being wheeled outside, which was aborted; patient taken back to the room the patient was alert and talking visiting with family. all drips remain where they were all day. the physician noted there is no back up plan if the pump stops again and fails to restart as the patient is a do not resuscitate status and there is no escalating care other than maintaining the current chemical support. the following day it was noted no additional pump alarms; however additional pressors were added to help maintain blood pressure while the patient managed affairs prior to plan for comfort care later in the day, which subsequently occurred, and the patient expired. | [5 empty] | The impact to performance and outcome as a result of the impella related events (e.g., pump stop alarms and stopping of the pump) is unknown. there is not enough to exclude the impella device used as an associated factor. the impella device was returned, and an evaluation is underway. upon completion of the analysis, a supplemental report will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: direct aortic insertion, ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of positioning issues, temporary pump stop, and vf/vt/af/at has been completed. there was no problem found during the case with pump relative to pump stop event. the device functioned/operated to specification. the root cause of the positioning issue was not determined due to insufficient clinical information. the root cause of the arrhythmia was unable to be determined due to insufficient clinical details.
9998 | 1220648-2025-27827 | 21903288 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21903288&pc=OZD | 03/31/2025 | 04/24/2025 |  | 04/24/2025 |  |  | 03/31/2025 | 04/24/2025 |  |  |  |  | 04/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025574650 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Disability | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with a impella cp to provide mechanical circulatory support. while on impella cp support, the patient experienced limb ischemia. the decision was made to percutaneously place a new cp through the axillary artery. both legs were amputated due to chronic occlusions while being supported by the second impella cp. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. there are to mdrs for the same patient however this report is for pump one (b)(6) and the other is for the second pump (b)(6). as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
9999 | 1220648-2025-27828 | 21903426 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21903426&pc=PYX | 04/05/2025 | 05/28/2025 |  | 04/24/2025 |  |  | 04/05/2025 | 04/24/2025 |  |  | 07/31/2025 |  | 04/24/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024390086 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported the physician attempted to place an impella rp in the left femoral vein. the right venticle was extremely dilated. the impella rp catheter kept kinking on itself in the right ventricle. the medical team tried a buddy to advance the catheter but were unable to advance into the pulmonary artery. the decision was made to remove and try an impella rp flex catheter. the impella rp flex was advanced without issue. there was no subsequent harm as a result. | [5 empty] | The investigation of delivery issues and product damage (cannula kink) has been completed. the impella device was not returned for evaluation. the cause of the product damage issue was most likely kinked cannula due to patient's extremely dilated right ventricle. the cause of the delivery issue was most likely related to the patient's anatomical condition. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella rp smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ impella rp smartassist system with automated impella controller section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ impella rp smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
10000 | 1220648-2025-27829 | 21903817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21903817&pc=OZD | 04/03/2025 | 04/24/2025 |  | 04/24/2025 |  |  | 04/03/2025 | 04/24/2025 |  |  |  |  | 04/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026617658 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support as a bridge to heart and kidney transplant and continued work up. a cutdown of right axillary artery was made to isolate vessel. graft anastomosis with hemashield 8mm graft. axillary insertion kit utilized. the aortic valve was "easily" crossed with 0.035 guidewire and diagnostic catheter and wire exchanged for abiomed 0.018. the impella 5.5 was "easily "inserted, initially directed towards septum, rotated into left ventricle free space with inflow to aortic valve measurement of 4.8 cm. the graft trimmed and impella secured and dressed. external fixation was placed while in the operating room and the patient was then transferred to the unit for ongoing management. post implant, same day, it was noted the patient experienced bleeding from the right axillary site. manual pressure was applied, and protamine was given to counteract the condition. support continued with minimal oozing going forward. no blood products were given as a result of the event. the status of the patient as a result of the event was indicated as unknown. however, the support continued for an additional 14 days until the impella 5.5 removal and cardiac transplantation. the device was successfully weaned and explanted. | [5 empty] | The impella device used was not returned for evaluation and analysis. a product history review was completed and noted the pump passed all post sterile inspection checks. the investigation was completed, and the root cause of access site bleeding was noted to be most likely a use issue since non-abiomed product was used which is not recommended per the instruction for use. instructions for use for the related event are as follows: impella 5.5 with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
10001 | 1220648-2025-27747 | 21889584 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21889584&pc=OZD | 04/03/2025 | 05/22/2025 |  | 04/23/2025 |  |  | 04/03/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026642632 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Pumping Stopped | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that upon insertion of the impella cp into the left ventricle, there was a ¿pump stopped¿ alarm. a fluoroscopy revealed that the cannula proximal to the outlet was kinked at the aorta. another company's transcatheter aortic valve replacement valve was not able to be inserted any more to straighten out the cannula. therefore, the decision was made to place another impella cp. the patient remained in stable condition. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of positioning issues and product damage (cannula kink) has been completed. the impella device was not returned for evaluation. the clinical states that upon insertion of impella cp into left ventricle (lv) pump had a pump stop alarm. fluoroscopy revealed cannula proximal to outlet was kinked at aorta. due to medtronic transcatheter aortic valve replacement (tavr) valve, was not able to be inserted any more to straighten out the cannula. decision was made to place another cp. the second cp had similar issues. the physician suspected that the device was unable to get fully past the aortic valve and maintain appropriate position in the ascending without kinking the cannula. the pump was not returned. data logs not available. positioning issue: the root cause of the positioning issue was most likely patient condition related to interaction with the patient's tavr as stated by the physician the pump was not able to be fully inserted past the aortic valve to maintain appropriate position. product damage (cannula kink): the root cause of the product damage cannula kink issue was most likely patient condition related to the patient's tavr as stated by the physician the pump could not be inserted more to straighten out the cannula due to tavr. pump stop: the root cause of the pump stop issue was not determined since product and data logs were not returned for investigation. instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: entering cardiac output. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
10002 | 1220648-2025-27748 | 21889698 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21889698&pc=OZD | 04/03/2025 | 05/29/2025 |  | 04/23/2025 |  |  | 04/03/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1319825 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Crack | Pumping Stopped | Application Program Problem | No Pressure |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller (aic) had two controller errors. the aic was replaced. this issue occurred during patient use and there was no patient harm. | [5 empty] | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. data analysis: cp pump 524918 was connected to ic5698 on 4/3 after 856 minutes of use on ic8828. (reviewing the cloud logs for ic8828, there was nothing notable.) 4 hours later, there is a pump stop on ic5698, the first of seven pump stops as the console attempted to restart the motor, including resetting the pump motor driver (pmd) twice as a response to controller failure event #138 for a pump speed deviation that was most likely what the complaint was for. rt and lt logs confirm pump stops at these times. the pump was disconnected and reconnected, and after the pump was restarted, there were occasional suction alarms but no more pump stops until the pump was moved to ic8828 later on 4/3. the device continued to perform without issue on the backup console. device analysis: ic5698 was returned for investigation but the pump stops were not reproduced. however, immediately upon booting up there was a controller error #1039 for no light to the ccd array despite the lamp confirmed on and outputting at the optical connector. the controller error was found to be due to a fiber break. additionally, the purge flag was broken. furthermore there was found to be cracks on the front and rear housings. root cause: the damaged purge flag is most likely due to the console being dropped or similar blunt force after the complaint date since the symptoms were not present in the logs from that date. the cause of the controller error on boot up during testing is most likely due to the console being dropped or similar blunt force after the complaint date since the symptoms were not present in the logs from that date. the cause of the pump stops during the case that triggered the two alarms from the complaint was not determined since the failures were not reproduced. instructions for use: (pump stop) impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
10003 | 1220648-2025-27749 | 21889785 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21889785&pc=OZD | 04/03/2025 | 05/05/2025 |  | 04/23/2025 |  |  | 04/03/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026616490 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bacterial Infection | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella 5.5 pump to support a patient who needed care escalated from an intra-aortic balloon pump and transferred to a tertiary center. the pump was supporting when antibiotics were needed due to methicillin-susceptible staphylococcus aureus bacteremia gram positive labs. it was further reported that care was withdrawn and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation of infection/sepsis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock, section: table 3.2 impella catheter components: ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
10004 | 1220648-2025-27750 | 21889859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21889859&pc=OZD | 01/27/2025 | 05/05/2025 |  | 04/23/2025 |  |  | 03/06/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1356082 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller (aic) triggered and displayed an advisory alarm ¿impella position unknown¿ despite having "great" pulsatility. the physician commented, the optical sensor was likely damaged. approximately two months later, abiomed's investigation engineer identified that the automated impella controller had intermittent low signal-to-noise ratio the clinical procedures. consequently, a request was made for the return of the aic for evaluation and analysis, which has been received. there was no report or indication of any patient harm because of this issue. | [5 empty] | The investigation of optical signal issue has been completed. the data analysis showed the console had numerous placement signal entries "processsampleforopminvalidsignalerror: calculator placement signal 1036" as well as frequent impella position unknown alarms as mentioned in the complaint. despite these alarms , the pump remained on this console for the duration of the case. the console was returned for investigation. a 5s test was run on the console and all values were passing. the ccd array was analyzed and was found to have no abnormalities in the envelope with a pump and without a pump. a pump was run overnight with no placement signal issues. there was most likely no malfunction of the console during this case as the console performed significantly better on neighboring cases, the failure mode was not reproduced during testing, and the benchtop testing revealed an optical bench performing as expected. | The automated impella controller device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. this is one of two devices associated with the event. refer to manufacturing report number 1220648-2025-26202 for the report on the impella 5.5 device (reported for possible positioning issue given the ¿impella position unknown¿ display).
10005 | 1220648-2025-27751 | 21890003 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21890003&pc=OZD | 04/02/2025 | 04/23/2025 |  | 04/23/2025 |  |  | 04/02/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026625333 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, the patient had oozing with hematoma at the access site. two units of blood products were given to the patient. anticoagulation was noted in target range; a plan was noted to wean and explant the impella cp device the following day, which was successfully completed with a survived outcome. the patient status post the event was indicated as unknown. | [5 empty] | Patient-specific information is unavailable at the time of this mdr writing; therefore, respective fields reflect "unknown" or left blank accordingly. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10006 | 1220648-2025-27752 | 21890004 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21890004&pc=OZD | 04/02/2025 | 05/15/2025 |  | 04/23/2025 |  |  | 04/02/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026634517 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient undergoing high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support and bleeding from site overnight and a hematoma formed. two units of packed red blood cells were administered in addition to femostop and manual pressure. subsequently, on echocardiogram, the impella cp device showed the position slightly deep. the impella was pulled back during under echocardiogram, and the placement alarms resolved. the following morning, it was noted that the pump ingested a clot. flows dropped to less than one. however, after several hours, the patient remained hemodynamically stable. therefore, the decision was made to remove the impella cp device without replacement. after explant of the impella, a large clot burden was visible in the inlet cage. the patient was reported to be stable following the device explant. | [5 empty] | The pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The pump was returned for investigation. a data log was not provided for this case. no evidence of physical damage was observed on the returned product. biomaterial was noted on the inlet cage of the returned product. further investigation revealed the broken impeller blade. the pump was tested in a bucket of water, and low pump flow was reproduced after removing the biomaterial. the cause of the positioning issue was most likely use-related, as the pump was pulled during use, based on the returned product investigation. the cause of the access site bleeding issue was not determined since sufficient clinical information was not provided. the broken impeller blade and the biomaterial could impact the low pump flow, and without logs, the field event could not be confirmed the root cause of low pump flow. the cause of the low pump flow issue was not determined without data log investigation. as the necessary information was not provided, the root cause of the thrombus was not determined.
10007 | 1220648-2025-27753 | 21890117 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21890117&pc=OZD | 04/02/2025 | 05/21/2025 |  | 04/23/2025 |  |  | 04/02/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026620844 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. approximately six days after start of the support, the impella anti-contamination sleeve ripped. however, support continued for four additional days before the patient was successfully weaned and the device was explanted. there was no harm to the patient as a result of the event. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of product damage (sleeve damage)has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
10008 | 1220648-2025-27754 | 21890239 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21890239&pc=OZD | 04/02/2025 | 05/20/2025 |  | 04/23/2025 |  |  | 04/02/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1504720 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the pump was supporting for 51 hours when explanted and escalated to an impella 5.5 device. the purge cassette on the day of explant failed to prime ("purge cassette defective¿ was displayed on the automated impella controller). to troubleshoot, the purge cassette and automated impella controller were both replaced. there was no harm to the patient as a result of the event. | [5 empty] | The automated impella controller device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. there was no issue found during the case. the device functioned/operated to specification. d9 date device returned to manufacturer added.
10009 | 1220648-2025-27755 | 21890268 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21890268&pc=OZD | 04/02/2025 | 05/19/2025 |  | 04/23/2025 |  |  | 04/02/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026638763 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 88 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient undergoing high-risk percutaneous coronary intervention was planned for impella cp mechanical circulatory support. however, the impella cp pump was unable to advance through the patient's vasculature during attempted delivery. the implant was subsequently aborted. the status of the patient after the event was unknown. however, there was no report or indication of harm to the patient. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since the implant of pump was aborted due to patient's anatomy. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
10010 | 1220648-2025-27756 | 21890513 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21890513&pc=OZD | 04/02/2025 | 04/23/2025 |  | 04/23/2025 |  |  | 04/02/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025579681 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient undergoing high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the right coronary artery was intervened on using rotational atherectomy to modify the calcific plaques. flows averaged about 3.4 liters per minute during the atherectomy and mean arterial pressure was maintained above 70 mmhg. the lesions were then ballooned, and intravascular ultrasound was used to size the vessel, and stents were inserted to secure the vessel. the impella was weaned and removed in the catheterization lab. the groin site could not be closed percutaneously, and surgical closure was required. there was access site bleeding requiring an unplanned surgical intervention for hemostasis. the patient was reported to be stable following the event. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10011 | 1220648-2025-27757 | 21890522 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21890522&pc=OZD | 04/02/2025 | 05/22/2025 |  | 04/23/2025 |  |  | 04/02/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025574214 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | Ventricular Fibrillation | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an unsuccessful delivery attempt of an impella 5.5 device followed by ventricular fibrillation requiring intervention. the patient was brought to operating room for escalation to an impella 5.5 insertion and impella cp removal. the physician was made aware that the impella cp did not have perclose and axillary anatomy had not been imaged. the physician elected to do plan for cut down and repair of femoral artery and wanted to proceed with 5.5 insertion. the axillary artery was exposed, 10mm graft beveled and anastomosed to vessel. the graft tunneled out laterally with appropriate distance/angle. the physician was asked if he wanted to obtain vessel imaging prior to the impella 5.5 being opened and primed and decided to proceed. the impella 5.5 was prepped and primed in standard fashion. a 23fr introducer in place with 2x graft locks. access to left ventricle was obtained with 0.035 j and 4fr non-taper angle glidecath. a 0.035 guidewire was exchanged for a 0.018 guidewire. activated clotting time was 263. glidecath was out, the impella 5.5 device was backloaded onto 0.018" and advanced until "2-3cm past anastomosis" and would not advance further. troubleshooting attempted; however, the decision to abort made after about approximately two hours of troubleshooting. during the event, the patient experienced ventricular fibrillation arrythmia and required defibrillation three times with the final shock at 200j. the procedure was completed, and the patient was transferred back to the unit in critical condition. the patient was planned for transfer to a higher-level care facility. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the analysis, a supplemental report will be filed. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues and vf/vt/af/at has been completed. product damage(cannula kink) and product damage(catheter kink) added since a cannula kink and catheter kink were found on the returned device. patient was being escalated from cp to 5.5 and reported having issues advancing the impella 5.5 more than 2-3cm past the anastomosis. troubleshooting was attempted but delivery was ultimately aborted. impella 5.5, purge cassette and guidewire were returned and a cannula kink and catheter kink were found on the pump. the root cause of the delivery issue was most likely patient condition related since reported issues advancing pump past the anastomosis. a cannula kink was found near the motor housing and in the middle of the cannula on the returned pump. the root cause of the product damage was a kinked cannula. a catheter kink was found near the motor housing on the returned pump. the root cause of the product damage was a catheter kink. as the necessary information was not provided, the root cause of the vt/vf was not determined.
10012 | 1220648-2025-27758 | 21890631 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21890631&pc=OZD | 04/02/2025 | 05/12/2025 |  | 04/23/2025 |  |  | 04/02/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026628637 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient undergoing a high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. post implant the impella site was bleeding despite angle matching and forward tension sutures. the decision was made to explant the impella cp device. coronaries were imaged again with contrast and showed no issues. the groin site was closed with perclose twice. the patient was transported to the unit and was noted to be in stable condition following the event. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10013 | 1220648-2025-27759 | 21890795 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21890795&pc=OZD | 04/02/2025 | 04/23/2025 |  | 04/23/2025 |  |  | 04/02/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025576325 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp (and an impella rp flex for bipella) mechanical circulatory support. during the support, same day post implant, there were placement signal low alarm with the impella cp device. the pump was repositioned multiple times and additionally, the patient experienced a ventricular tachycardia arrhythmic event on performance level less than p-4. the decision was made to maintain the support level at p-2. the following day the impella cp device along with the impella rp flex device was successfully weaned and explanted. the patient was reported to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
10014 | 1220648-2025-27760 | 21890816 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21890816&pc=OZD | 04/02/2025 | 04/23/2025 |  | 04/23/2025 |  |  | 04/02/2025 | 04/23/2025 |  |  | 01/31/2026 |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025485670 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported prior to the impella placement the patient failed to come off the pump twice post coronary artery bypass graft surgery and the patient failed to come off the pump twice to the intra-aortic balloon pump mechanical circulatory support (mcs). the decision was made to implant an impella 5.5 device for mcs. subsequently, a left ventricle perforation was observed. emergent closure of perforation was completed. the impella 5.5 device was pulled into the aorta and was placed in surgical mode for intervention with cross clamp and on cpb. the patient had friable vessels per the operating room staff. intra procedural complications with sewing grafts on and thought to have had some sort of connective tissue disorder. the left ventricle perforation was undetermined if caused by impella or other procedural complications. the perforation was closed as noted but the patient expired after approximately two hours of support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
10015 | 1220648-2025-27761 | 21891170 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21891170&pc=OZD | 04/03/2025 | 05/19/2025 |  | 04/23/2025 |  |  | 04/03/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026642611 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following impella cp implant, the device immediately triggered suction alarms coinciding with low flows. the p-level was adjusted, but the alarms continued. a kink was subsequently noted under fluoroscopy and the decision was made to replace the pump. there was no resulting patient harm. | [5 empty] | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of product damage (cannula kink) and low pump flow has been completed. the clinical details state that following ¿normal insertion of impella cp via lfa¿ utilizing ultrasound guidance and micro puncture to obtain access, device started in auto and begin having suction alarms and low flows. the physician tried to adjust p-level to resolve suction and had continued alarms. the physician noticed visible kink in cannula under fluoroscopy. the data logs confirm low flow and suction. the pump was returned and there was a kink in the cannula. the root cause of the low pump flow is due to a cannula kink but the cause of the kink is unknown. the root cause of the product damage (cannula kink) cannot be determined to due limited clinical details provided.
10016 | 1220648-2025-27762 | 21891250 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21891250&pc=OZD | 03/27/2025 | 05/07/2025 |  | 04/23/2025 |  |  | 03/27/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025599969 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  | White | Death | Yes | Yes | Positioning Failure | Restricted Flow rate | Increase in Pressure |  |  |  |  | Septic Shock | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support, there were suction alarms and the pump was repositioned. support continued. additionally, the patient had septic shock while on support. the cause and treatment of the septic shock is unknown. furthermore, high purge pressure was noted. tissue plasma activator was added to the purge. the patient remains on support. | Additional information received that the patient expired while on support. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: table 3.2 impella catheter components. ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | Additional information received from the clinical site informing that the patient expired while on support. the following fields were updated to reflect the new outcome. b2 death and date of death. b5 additional information added. d6b explant date. h1 type of report. h6 code 1802 added.
10017 | 1220648-2025-27763 | 21891384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21891384&pc=OZD | 04/01/2025 | 05/22/2025 |  | 04/23/2025 |  |  | 04/01/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026633632 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Material Twisted/Bent | [5 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp pump was placed along with an extracorporeal membrane oxygenation to support a patient admitted in with cardiac history and st - elevated myocardial infarction. the coronaries were intervened upon and the patient was sent for computed tomography due to free fluid in the abdomen, which prompted many blood products. during support, there was suction and limb ischemia issues. the team and family made the patient do not resuscitate and the patient expired. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of low pump flow and limb ischemia has been completed. product damage (cannula kink) was added since a cannula kink was found on the returned device. low flow - reported extended cardiopulmonary resuscitation (cpr) and hypoperfusion reported as contributing factors. the pump was returned and a cannula kink near the motor housing was observed. data logs show pump fluctuated between p7 and p2 on auto for the first hour. there were suction alarms from when the pump started and impella flow reduced alarms in response to suction. pump was then set to p-3, briefly turned up to p9 and back down but with flows below 0.5l/min. flows then fluctuate between 0.3l/min and 1l/min with more suction alarms when pump at p-2 for the next 4 hours. after turning the p-level to p-3, then back to p-2, flows stabilized around 0.6l/min for 4 hours. pump was ran at p1 for the last 21 hours with flows between 0.5l/min and 0.7l/min. the root cause of the low pump flow was most likely a kinked cannula as observed on the returned product due to extended cpr reported as a contributing factor. cannula kink - reported extended cpr and hypoperfusion reported as contributing factors to the failure modes. impella cp was returned and a cannula kink was found. the root cause of the product damage was a kinked cannula as observed on the returned product due to extended cpr reported as a contributing factor. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
10018 | 1220648-2025-27764 | 21891742 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21891742&pc=OZD | 04/02/2025 | 05/27/2025 |  | 04/23/2025 |  |  | 04/03/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636877 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump support placed to support the patient admitted in with a st-elevation myocardial infarction. it was reported that at the end of case, patient being reversed with protomine, impella flows dropped suddenly with spikes in motor current. a clot was suspected. once the impella was removed a clot was seen on outlet cage and motor housing. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of low pump flow and thrombus has been completed. the impella device was not returned for evaluation. as per clinical impella flows dropped suddenly with spikes in motor current. impella removed and clot seen on outlet cage and motor housing. the data log was reviewed and the flow was steady at p-9 until suctions occurred at time of reported flow issue on 02-apr followed by impella position alarms in aorta and ventricle which confirm unsuccessful attempts to solve the flow issue with repositioning. no placement signal issue observed. motor current (mc) spike observed with no p-level change and drop in flows at the end of the support and before decision of impella removal. the root cause of the low pump flow issue most likely was biomaterial ingestion due to low activated clotting time (act)'s based on log and clinical review. as the necessary information was not provided, the root cause of the thrombus was not determined. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-27764. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
10019 | 1220648-2025-27765 | 21892089 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21892089&pc=PYX | 04/03/2025 | 09/23/2025 |  | 04/23/2025 |  |  | 04/03/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026625513 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Positioning Failure | Device Sensing Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient on impella rp flex support and was transferred to a second hospital. upon observation via x-ray, the impella rp flex looked deep. the impella rp flex was repositioned. it was further reported that the patient plasma-free hemoglobin on first draw was 380. second draw was 350. the physician was concerned for hemolysis due to the patient only having one kidney. the next day the impella rp flex was removed and the patient continued on support with extracorporeal membrane oxygenation and protek support. the patients outcome is stable. | The complainant also reported that a placement signal unreliable alarm was also present. the impella was repositioned. the placement signal unreliable alarm remained despite repositioning. |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of hemolysis and positioning issues has been completed. the impella device was not returned for evaluation. the clinical states that on (b)(6) 2025: patient arrived from outside facility for transfer, impella rp flex noted to be deep. due to lack of clinical details, the root cause of the positioning issue of the impella being too deep cannot be determined. the data logs show that there was no trend in the placement signal, changes in motor current or positioning alarms. due to the lack of concrete information, the root cause of the hemolysis cannot be determined. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. there was no problem found pertaining to the pump¿s optical signal issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include placement signal issue description. corrections that were made from the initial manufacturer device report number 1220648-2025-27765. d10 concomitant medical products and therapy dates updated.
10020 | 1220648-2025-27766 | 21892268 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21892268&pc=OZD | 04/01/2025 | 05/13/2025 |  | 04/23/2025 |  |  | 04/01/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026640365 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient had an attempted impella cp implant during a high risk cardiac procedure when the patient went into cardiac arrest. there were multiple rounds of resuscitation performed and access was obtained for impella cp insertion. during the implant, the staff was unable to advance past the calcified iliac artery. the patient required continued cardiopulmonary resuscitation and subsequently expired. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since the pump was unable to advance past iliac artery due to calcification. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
10021 | 1220648-2025-27767 | 21892452 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21892452&pc=OZD | 04/01/2025 | 05/12/2025 |  | 04/23/2025 |  |  | 04/01/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026629030 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella 5.5 pump was unable to advance across the aortic valve during attempted placement. the implant was aborted as a result of the noted resistance. there was no noted patient injury resulting from the delivery issue. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, physician unable to position impella 5.5 across aortic valve and aborted 5.5 implantation. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
10022 | 1220648-2025-27768 | 21892547 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21892547&pc=OZD | 04/01/2025 | 04/23/2025 |  | 04/23/2025 |  |  | 04/01/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026634737 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient had an impella cp placed for support during a high-risk procedure. the pump was placed and supported for 4 days when there was bleeding observed at the site. the site was treated with blood and platelets. the pump was explanted and escalated to an impella 5.5 on the same day. the patient survived the procedure. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10023 | 1220648-2025-27770 | 21893002 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21893002&pc=OZD | 03/29/2025 | 04/23/2025 |  | 04/23/2025 |  |  | 03/31/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025611727 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support and after the explant, the patient was noted to have right side weakness and not speaking. a computed tomography (ct) scan was performed showing a left internal carotid artery terminus occlusion. a mechanical thrombectomy was performed. a ct scan post procedure shows left hemispheric infarct with cerebral edema, and midline shift. it was noted heparin or other form of anticoagulant was never started. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
10024 | 1220648-2025-27771 | 21893083 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21893083&pc=OZD | 04/01/2025 | 04/28/2025 |  | 04/23/2025 |  |  | 04/01/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1692374 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an automated impella controller (aic) was not detecting the purge disc despite the ¿click¿ being audibly heard. the aic was replaced. there was no patient harm reported. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. the logs from the complaint date revealed that the console issued purge disc not detected alarm several times. the console was examined, and a broken purge disc flag was identified. the purge disc flag was observed to be broken and was the root cause of the reported purge disc not detected alarm issue
10025 | 1220648-2025-27772 | 21893145 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21893145&pc=OZD | 03/31/2025 | 05/23/2025 |  | 04/23/2025 |  |  | 03/31/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026623878 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support, and during support, waveforms became decoupled and persistent suction on every power level occurred despite confirmed position, preload, and tv function. the pump was replaced and there was no patient harm. | [5 empty] | The investigation of optical signal issue and low pump flow has been completed. the impella device was not returned for evaluation. based on the clinical details provided and the nature of the motor current shift seen in data logs, the root cause of the optical signal issue was determined to most likely be due to a suction condition. it is unknown what caused the pump to go into a suction condition based on the clinical details provided. since clinical details provided were limited, data logs did not show a definitive cause, and product was not returned for investigation, the root cause of the low pump flows could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10026 | 1220648-2025-27773 | 21893411 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21893411&pc=OZD | 03/31/2025 | 05/15/2025 |  | 04/23/2025 |  |  | 03/31/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025607762 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient was on impella 5.5 support when the patient was moved to reposition impella in bed, immediately there was suction, and impella in ventricle alarms. after the physician repositioned, flows were better. the physician tried to reposition again for perfect placement under echo, and a motor current high alarm and pump stopped occurred. troubleshooting was done but the impella did not restart. the 5.5 was removed and a new pump was replaced. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of pump stop has been completed. the pump was returned for investigation. the data log shows that the pump stop with alarm 13 (motor current high) after 2 hours of pump run. the pump was unable to restart. no physical damage was observed on the returned product. the pump passed electrical testing. pump was then soaked in a bucket of water, where white calcium-like biomaterial was observed on the impeller. the pump was also unable to start during the test run. the white, calcium-like biomaterial was removed using dental picks. finally, the pump was run in a bucket of water as per specifications. the cause of the pump stop issue was biomaterial, based on returned product investigation. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
10027 | 1220648-2025-27774 | 21893765 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21893765&pc=OZD | 04/04/2025 | 05/27/2025 |  | 04/23/2025 |  |  | 04/04/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025606348 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient began to bleed from their impella cp site following implant. manual pressure was applied, gauze was added to the dressing and the repositioning sheath was adjusted but the bleed persisted. eight units of packed red blood cells and two units of platelets were given to treat the condition. the decision was ultimately made to explant the pump as a result of the bleed. the patient outcome is unknown. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the access site bleeding issue was most likely patient movement, bleeding noted from impella site upon arrival to room. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) . ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10028 | 1220648-2025-27775 | 21893860 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21893860&pc=OZD | 03/19/2025 | 04/23/2025 |  | 04/23/2025 |  |  | 03/19/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026617613 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Positioning Failure | [6 empty] | Tachycardia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported an impella 5.5 was inplanted in a patient admitted with cardiac history. the patient was being supported on two pumps for over 59 days. on day 26 of the being on the second pump, there were concerns noted for ventricular tachycardia (vt) due to poor position within the left ventricle. the vt triggered the automatic implantable cardioverter-defibrillator to shock the patient. the pump position was assessed by computed tomography imaging. however, despite all the measures, the care team determined that there were no options for the patient and the patient was transitioned to comfort care and subsequently expired. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10029 | 1220648-2025-27776 | 21894061 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21894061&pc=OZD | 04/04/2025 | 05/23/2025 |  | 04/23/2025 |  |  | 04/04/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636405 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | Material Twisted/Bent |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and after the implant, low pump flow was observed. there was a notable kink/bend located at impella cannula closest to outlet. the pump was explanted and the patient remained stable. | [5 empty] | The investigation of low pump flow has been completed. product was not returned for review. based on clinical description, the root cause of low flow was most likely kinked cannula. the root cause of cannula kink was most likely patient condition related to patient's anatomical anomalies as stated by the physician. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10030 | 1220648-2025-27777 | 21894525 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21894525&pc=OZD | 04/02/2025 | 04/23/2025 |  | 04/23/2025 |  |  | 04/02/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025607183 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [12 empty] | Disability | Life Threatening | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 device as an elective placement for mechanical circulatory support and the following day the patient experienced a stroke and expired. the patient presented with st-segment elevation myocardial infarction. the patient had multi-vessel disease and was planned for coronary artery bypass grafting with intraoperative placement of the impella 5.5 device. the patient had diffused post-op bleeding. heparin was withheld due to the bleeding with plans for washout and evaluation of heparin drip the next morning. the impella 5.5 device was successfully implanted for mechanical circulatory support. transesophageal echocardiogram showed the inlet 5cm to aortic valve in the operating room. the following day, the patient's pupils were blown. the patient was taken for a computed tomography scan, which showed an ischemic stroke with multiple emboli. the patient subsequently expired this same day. further information noted the impella was primed for 20 minutes before insertion, the graft and sheath were de-aired twice before insertion, and prior to impella placement, the patient had air bubbles in the left ventricle. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: potential adverse events, ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
10031 | 1220648-2025-27778 | 21895340 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21895340&pc=OZD | 04/03/2025 | 04/23/2025 |  | 04/23/2025 |  |  | 04/03/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025575888 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that delivery issues were experienced when placing an impella 5.5 into the introducer. the pump was advanced into the artery to the point near the innominate take off. with multiple attempts, the pump was unable to advance completely into the aorta. the surgeon was concerned about causing any damage to the artery with forcing it any harder. the surgeon believes there are some anatomical abnormalities causing this and was concerned even if it passes with additional attempts, this could cause a complication with explanting the pump. the patient was hemodynamically stable on minimal pressors and off extracorporeal membrane oxygenation. the decision was made by the surgeon that the risk outweighed the benefits and to explant the impella 5.5. the patient was sent up to the intensive care unit. the case was aborted due to the patient's anatomy and the patients outcome is unknown. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist: section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
10032 | 1220648-2025-27779 | 21895408 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21895408&pc=OZD | 04/03/2025 | 05/16/2025 |  | 04/23/2025 |  |  | 04/03/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026615321 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp pump was unable to pass through the iliac artery during attempted delivery. the implant was aborted and extracorporeal membrane oxygenation was implanted for patient support. the patients outcome is unknown. | [5 empty] | The root cause of delivery issue is determined to be patient condition because the pump was unable to pass through the iliac artery and the insertion was aborted. based on the clinical details the root cause of delivery issue is determined to be patient condition because the pump was unable to pass through the iliac artery and the insertion was aborted. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
10033 | 1220648-2025-27780 | 21895436 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21895436&pc=OZD | 04/03/2025 | 05/15/2025 |  | 04/23/2025 |  |  | 04/03/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1597593 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an automated impella controller was damaged during transport back to the hospital following a patient transfer. there was no patient involvement. | [5 empty] | There were no console log to review for this complaint. there was obvious damage to the front and rear housings. the console was able to perform as expected during testing as all the damage was to the exterior. root cause: the damage to the console was most likely a use issue. | The controller was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
10034 | 1220648-2025-27781 | 21895519 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21895519&pc=OZD | 04/01/2025 | 05/19/2025 |  | 04/23/2025 |  |  | 04/01/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026619345 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that the device migrated out of the left ventricle and into the aorta during patient support. attempts to re-advance the pump were unsuccessful and the cannula body subsequently kinked. the device was removed and flushed in a sterile bowl, however adequate flows could not be re- obtained at support level p-9. the pump was subsequently replaced as a result of the failures. there was no resulting patient harm. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of product damage (cannula kink), positioning issues, and low pump flow has been completed. the impella device was not returned for evaluation. patient had a small vertical left ventricle (lv). shortly after implant the cp migrated from lv to aorta and was unable to re-advance across the valve. wireless re-access was attempted. catheter kinking in body occurred while troubleshooting. pump was removed and ran in new sterile bowl with sterile water with low flows. the pump was then replaced. positioning issues: the root cause of the positioning issue was most likely patient condition related due to the patient's small vertical lv. product damaged (cannula kink): the cause of the product damaged (cannula kink) was most likely use related as the kinking was stated to have occurred during positioning troubleshooting when wireless re-access was attempted. the impella device is not indicated for wireless re-access per instructions for use. low pump flow: complaint device not returned for analyzing. data log reviewed: suction alarm occurred suddenly followed by drop in flow at the time of reported issues. position in aorta alarm occurred after suction. no motor current (mc) spike or placement signal issue observed. the flow was stated to be low when it run in the bucket of water outside the body. the cause of the low pump flow issue most likely was cannula kink.
10035 | 1220648-2025-27782 | 21895632 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21895632&pc=OZD | 04/01/2025 | 05/14/2025 |  | 04/23/2025 |  |  | 04/01/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025582348 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | Yes | Yes | Fluid/Blood Leak | Inaccurate Information | Wrong Label |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Patient representative reported left side patient ¿was shocked and afflicted with the insecurity caused by the news, regarding the possibility of developing cancer. a new surgery is already generating anxiety, anguish and fear in patient¿. patient is presenting capsular contracture baker grade iii (unspecified side) and cysts. device has been explanted and replaced. | [5 empty] | The investigation of product damage (hole in catheter) and packaging issues has been completed. product damage (hole in catheter). the clinical stated that when attempting to reposition the device, there was bleeding noted from the red pump plug. the product was returned, and there were 3 different needle puncture holes seen in the catheter around the 35cm mark, causing blood to leak from the red pump plug. the root cause of the product damage (hole in catheter) is most likely due to user error during insertion. packaging issues: an email was received from the representative that stated the impella 5.5 was labeled as a cp on the red pump plug. the product was returned and it was confirmed that the 5.5 was mislabeled as a cp. the sn was looked up in sap and it was confirmed to have left product as a 5.5. additionally, the pump was connected to the lab console and the console recognized it as a 5.5. the root cause of the packaging issue is due to an incorrect label. b1 adverse event was erroneously selected on the initial manufacturer device report number 1220648-2025-27782. b2 should have been left blank on the initial manufacturer device report number 1220648-2025-27782. h1 type of report was erroneously selected on the initial manufacturer device report number 1220648-2025-27782. h6 codes 1888 and 4051 were erroneously entered on the initial manufacturer device report number 1220648-2025-27782. new codes were added to health effect - clinical code and health effect - impact code. h10 instructions for use were erroneously entered on the initial manufacturer device report number 1220648-2025-27782.
10036 | 1220648-2025-27783 | 21895759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21895759&pc=OZD | 04/01/2025 | 05/21/2025 |  | 04/23/2025 |  |  | 04/01/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025612283 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the 64 year old female patient was on impella cp support and during support, the pump was started into auto mode once it reached 1.6 l min flows it would not go any higher. once taken out of auto mode there were suction alarms. the staff turned down the pump to p-0 then back up to p-2 and p-3 with same result. the doctor pulled the catheter back across the aorta and got normal flows. staff were unable to recross the valve and the impella was removed. the patient survived the procedure. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The investigation of low pump flow and positioning issues has been completed. low pump flow: clinical states that upon pump implant, low flows were observed at p-9 with suction alarms. same result was observed at lower p-levels. pump was pulled out for repositioning but unable to recross valve. pump was returned for investigation. after running the pump in the lab, saturated flows were observed along with degraded biomaterial stuck on the impeller of the pump. no cannula kink, damaged impeller blades were observed. data logs were investigated and there were suction alarms observed at p-level change and due to low flow, pump was explanted. no other abnormalities observed in data logs. therefore as per the product review, the root cause of the low pump flow issue was most likely biomaterial found on the impeller of the pump. positioning issue: clinical states that due to suction alarms and low flow, the pump was troubleshooted and pulled back into aorta and better flows were observed but the pump was unable to recross the valve leading to impella removal. the guidewire was removed before pump start and most likely the re-access was carried without a guidewire. therefore, the root cause of the positioning issue was most likely wireless re-access which is an improper technique. the impella device is not indicated for wireless re-access per instructions for use.
10037 | 1220648-2025-27784 | 21895813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21895813&pc=OZD | 04/01/2025 | 05/21/2025 |  | 04/23/2025 |  |  | 04/01/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025579707 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the cannula became kinked during a high risk cardiac procedure, after interaction with guide. the case was completed with limited support of impella due to kinked cannula. after the procedure, the cp was removed. there was no patient harm reported. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The investigation of product damage (cannula kink) has been completed. clinical states that the pump was kinked while manipulating the pump during insertion due to interaction with the guidewire. pump was returned for investigation. a kink was confirmed in the cannula near the motor housing side of the pump. no other physical abnormalities were observed. therefore, the root cause of the product damage issue was a kinked cannula that was a result of interaction with guidewire. low pump flow: clinical states that the pump had flows of 3.3l/min due to the kinked cannula. pump was returned and a cannula kink was observed near the motor housing of the pump. therefore, the root cause of the low pump flow issue was a kinked cannula which was a result of interaction with other device.
10038 | 1220648-2025-27785 | 21895921 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21895921&pc=OZD | 04/13/2025 | 07/16/2025 |  | 04/23/2025 |  |  | 04/14/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025601281 |  | 1000459 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support, there was a leak from the purge sidearm. the purge pressure and flow rate remained at the same values. staff decided to not replace the pump, but to monitor. the patient remained on support and there is no patient harm. | [5 empty] | The investigation of purge leak ¿ pump has been completed. the returned pump was inspected and there were no abnormalities observed. there were no damages observed to the pump sidearm. the pump was purged for over 24 hours in the lab and no leak was observed. purge flow and purge pressure values were stable within the expected range. clinical detail noted that during the case the purge system continued working without any issues. data log analysis does not reveal any leak indication or leak location. the root cause of the purge leak - pump was not determined as the leak did not reproduce on the returned product and purge flow and pressure values were stable during reproduction testing. the investigation of purge leak ¿ pump has been completed. d6b explant date was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27785. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27785 as it is unknown if the initial reporter also sent the report to the food and drug administration. g2 report source; foreign was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27785. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
10039 | 1220648-2025-27786 | 21895932 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21895932&pc=OZD | 04/10/2025 | 07/10/2025 |  | 04/23/2025 |  |  | 04/16/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2025594052 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that an implanted impella cp pump stopped unexpectedly during patient support. the pump was removed as a result of the failure. the patient remained supported via drug control following explant. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the impella cp with smartassist for use during cardiogenic shock states the following: section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation of pump stop has been completed. the root cause of the pump stop was not determined since no product was returned. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-27786. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27786 as it is unknown if the initial reporter also sent the report to the food and drug administration.
10040 | 1220648-2025-27787 | 21895971 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21895971&pc=OZD | 04/17/2025 | 07/10/2025 |  | 04/23/2025 |  |  | 04/17/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026655007 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 pump was unable to advance through the patient's vasculature during attempted delivery. it was noted that the right subclavian was calcified and that resistance was met at the brachiocephalic artery. after multiple failed attempts, the implant was aborted. the patient remained supported with extracorporeal membrane oxygenation. there was no patient harm. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since the implant was aborted as patient had small vessels and calcification. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27787 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
10041 | 1220648-2025-27788 | 21896037 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21896037&pc=OZD | 04/17/2025 | 07/10/2025 |  | 04/23/2025 |  |  | 04/21/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 180905 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [11 empty] | Required Intervention | Death | [23 empty] | R |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and post implant, same day, it was noted when the p value was increased to performance level p-8 or higher, the aortic position waveform overshot. consequently, the automated impella controller was exchanged, and the issue resolved. however, following, approximately six days later the patient expired noted from circulatory failure. | The user facility reported during use with a 56 year old male patient, the ao position waveform on the console (aic) was overshot when the p value was increased to p8 or higher. the aic was replaced and the problem was resolved. there was no patient harm reported. |  |  |  |  | Medical safety review of the event was completed and noted there is not enough information to exclude the automated impella controller (aic) as a contributing factor in the outcome. optical signal issue required exchange in aic device, and there was a brief period of interrupted support to the patient. updates were made to respective sections: b1, b2, h1, and h6 coding (health/product problem) accordingly given the medical safety review outcome. additionally, updates were made to section d9 and h6 coding (investigation) accordingly, as the automated impella controller device was returned, and an evaluation is underway. upon completion of the evaluation/analysis, a supplemental report will be filed. section b5 was updated to include complete details of the reported event. | The investigation of optical signal issue has been completed. based on the device analysis, data analysis, and the clinical details; the root cause for the placement signal not reliable alarm (optical signal issue) was due to the defective optical bench with low signal not reliable (snr). b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-27788. d4 lot number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27788. d6a should have been left blank on the initial manufacturer device report number 1220648-2025-27788. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27788 as it is unknown if the initial reporter also sent the report to the food and drug administration. h5 was erroneously selected on the initial manufacturer device report number 1220648-2025-27788. h8 was erroneously selected on the initial manufacturer device report number 1220648-2025-27788. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
10042 | 1220648-2025-27789 | 21896714 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21896714&pc=OZD | 04/01/2025 | 05/22/2025 |  | 04/23/2025 |  |  | 04/01/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025551671 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that significant bleeding was encountered at the axillary site leading to the development of a hematoma. the bleeding occurred when the graft was being sewn on for the planned implant. manual pressure and a sandbag were applied to manage the condition and packed red blood cells were transfused. support continued uninterrupted following the intervention. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. reported having issues getting good control with vessel loops and sewing on the graft. hematoma following procedure managed with manual pressure and sandbag to site. impella cp pump was returned and no damages or abnormalities were observed. the root cause of the access site bleeding issue was not determined due to insufficient clinical information.
10043 | 1220648-2025-27790 | 21896715 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21896715&pc=OZD | 04/01/2025 | 05/23/2025 |  | 04/23/2025 |  |  | 04/01/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025539824 |  | 0048-0044 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | Material Split, Cut or Torn | [5 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additionally, the patient was subsequently taken to the operating room to remove the impella and have vascular repair for the femoral artery for its surgical closure. | The user facility reported that the impella cp sterile sleeve was torn. the sterility was maintained by sterile tape sealed onto sterile sheath and toughy. there was no patient harm reported. |  |  |  |  | The investigation of product damage - sleeve has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. should any new investigation results be received regarding vascular damage - peripheral vessel and supplemental report will be submitted. b1 adverse event was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27790. b2 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27790. b5 additional information that was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27790. d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2025-27790. g2 report source; foreign was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27790. h1 type of report was erroneously selected on the initial manufacturer device report number 1220648-2025-27790. h6 health effect - clinical code 4582 was erroneously selected on the initial manufacturer device report number 1220648-2025-27790. new codes added to health effect - clinical code and health effect - impact code h6 component code 4723 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27790. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
10044 | 1220648-2025-27791 | 21896724 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21896724&pc=OZD | 04/16/2024 | 05/05/2025 |  | 04/23/2025 |  |  | 04/01/2025 | 04/23/2025 |  |  | 10/31/2025 |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024422689 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that experienced left bundle branch block with a "possible" relationship to the impella device used: ¿subject with baseline interventricular conduction delay. had slight increase in conduction delay following impella placement. the patient was successfully weaned and explanted. | [5 empty] | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-27791. g2 report source; study was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27791. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
10045 | 1220648-2025-27792 | 21896730 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21896730&pc=OZD | 04/01/2025 | 05/22/2025 |  | 04/23/2025 |  |  | 04/01/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025574879 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Stroke/CVA | Tachycardia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. there were continued position concerns with associated ventricular arrhythmias. the device was repositioned three times since insertion. support continued. additionally, the patient had a code stroke called for increased lethargy and left sided weakness. right mca calcified clot was found and was evacuated. furthermore, the patient had a hemorrhagic subarachnoid stroke. systemic heparin was withheld due to the bleeding. the patient was successfully weaned. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ | The investigation of stroke/cva, vf/vt/af/at, positioning issues, and thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. instructions for use: section: warnings and cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ h10 instructions for use were incomplete on the initial manufacturer device report number 1220648-2025-27792.
10046 | 1220648-2025-27793 | 21896737 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21896737&pc=OZD | 02/24/2025 | 05/09/2025 |  | 04/23/2025 |  |  | 04/02/2025 | 04/23/2025 |  |  |  |  | 04/23/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025610845 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, the catheters crossed in the left ventricle and when the field medical catheter was ablating, the impella would drop to p0 and then the impella was manually brought up to the previous p level. this occurred several times throughout the case. the patient was explanted in the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high-risk cpi section: entering cardiac output ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation of the low pump flow and temporary pump stop has been completed. no product was returned for investigation. no logs were returned for analysis. low pump flow: the root cause of the low pump flow was not determined as no product was returned for investigation and no logs were returned for analysis. temporary pump stop: the root cause of the temporary pump stop was not determined as no product was returned for investigation and no logs were returned for analysis. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-27793.
10047 | 1220648-2025-27718 | 21878931 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21878931&pc=OZD | 03/28/2025 | 04/22/2025 |  | 04/22/2025 |  |  | 03/28/2025 | 04/22/2025 |  |  |  |  | 04/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026614107 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Malposition of Device | [6 empty] | Ischemia | Obstruction/Occlusion | Implant Pain | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and the 14fr. sheath was 100% occluding the artery. the patient was screaming in pain from the right lower extremity. the patient was given pain medication and eventually intubated as during the procedure the patient was thrashing and almost fell off the table. the wire position was lost; however, after intubation, the physician was able to complete the revascularization. post stenting peel-away sheath was removed. right lower extremity iliac angiogram was performed with repositioning sheath side access. the flow was restored to extremity. following, in the morning overnight, the patient's urine blood was tinged and the echocardiogram showed impella placement deep. the physicians decided to wait until later in the morning to address the positioning and possibly reposition. pulsatility was returning to impella. also, no pulses could be found distally; therefore, the plan was noted to have a vascular ultrasound later in the morning. that afternoon it was noted the patient went asystole four times and chemical code was given to achieve electrical activity. patient was having signs of multi-system organ failure. the patient's feet remain cold bilaterally. the patient would ultimately expire the following day. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
10048 | 1220648-2025-27719 | 21878934 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21878934&pc=OZD | 03/29/2025 | 05/19/2025 |  | 04/22/2025 |  |  | 03/29/2025 | 04/22/2025 |  |  |  |  | 04/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026624544 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Shock | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device via the right axillary artery for mechanical circulatory support. the patient was also cannulated for extracorporeal membrane oxygenation. the following day it was noted the patient was bleeding from cutdown site and a developed a "massive" hematoma from the upper chest to shoulder and neck. three units of blood products were administered. of note, the patient was also bleeding from gastrointestinal with over 300cc suctioned from the nasogastric tube. consults requested from cardiothoracic surgery and gastrointestinal physician. it was noted the team believed the patient was entering hemorrhagic shock. the patient was noted not a candidate for advance options. following the event the patient was noted to be in critical condition. withdraw of care discussion was planned in the event the patient did not wake up. prognosis remained very poor, neurological status was unknown, and the patient would expire approximately nine days later. care was withdrawn. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the access site bleeding issue was most likely patient condition related, as there was bleeding from the access site and patient also gastrointestinal bleeding.
10049 | 1220648-2025-27720 | 21878941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21878941&pc=OZD | 03/29/2025 | 05/23/2025 |  | 04/22/2025 |  |  | 03/29/2025 | 04/22/2025 |  |  |  |  | 04/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026641437 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Malposition of Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient (unknown race) with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the impella cp pump subsequently migrated into the aorta during patient support. attempts to re-advance the pump were unsuccessful and the decision was made to replace the device. there was no harm to the patient as a result of the event. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. based on the clinical details the cause of the positioning issues most likely was patient condition related due to large left ventricle. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10050 | 1220648-2025-27721 | 21878944 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21878944&pc=OZD | 03/29/2025 | 05/21/2025 |  | 04/22/2025 |  |  | 03/29/2025 | 04/22/2025 |  |  |  |  | 04/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026613018 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Pump | Malposition of Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in postcardiotomy cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support and during the support, the patient grabbed the impella red plug, became agitated and ultimately the impella had an "impella stopped" alarm. noted 2cm change in position per the marker. the staff was unable to restart the pump after multiple attempts and even after an automated impella controller exchange. the physician at bedside pulled back and explanted the impella 5.5 device. possible placement of another impella 5.5 was being considered; physician was noted to be formulating the plan forward. it was noted the patient survived the explant. however, status of the patient following the event was unknown. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: direct aortic insertion: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of pump stop and positioning issues has been completed. the impella device was not returned for evaluation. positioning issue: the cause of the positioning issue was most likely patient movement related as the clinical stated the patient was delirious and accidentally pulled on the pump causing it to come out of position. pump stop: data logs show sudden motor current (mc) spike with alarm pump stop 119. the pump attempts to restart but is unable to with more mc spikes and alarm 119 and 115. the cause of the pump stop was most likely an electrical open due to excessive force by the patient as the clinical stated the patient was delirious and pulled on the pump leading to pump stop.
10051 | 1220648-2025-27722 | 21878963 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21878963&pc=OZD | 03/29/2025 | 06/27/2025 |  | 04/22/2025 |  |  | 03/29/2025 | 04/22/2025 |  |  |  |  | 04/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1389732 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and during support, the following day, a controller error alarm triggered with loss of placement signal. the placement signal did return, and controller error alarm resolved. however, the hospital team recommended removing the automated impella controller (aic) from support and send it in for investigation. the aic was exchanged, and there was no harm to the patient as a result of the event. | [5 empty] | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. the logs confirmed the reported failure mode given there was a controller error alarm (opm comm crc error ¿ event set 1045) triggered during the clinical procedure. real time (rt) logs confirmed that during that time all signals received from optical bench (placement, snr, contrast, lamp, gain) were represented as -16384 values due to the crc error. console was, but the issue was not reproduced. the cause of the aic issue was a software issue. h10 narrative erroneously stated the device was not received on the initial manufacturer device report number 1220648-2025-27722. | The automated impella controller device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10052 | 1220648-2025-27723 | 21881066 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21881066&pc=OZD | 03/29/2025 | 05/23/2025 |  | 04/22/2025 |  |  | 03/29/2025 | 04/22/2025 |  |  |  |  | 04/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026622014 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 92 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs). the guidewire was removed and the pump attempted to start. however, a pump failure alarm triggered. an attempt was made to start/restart without success. as well, the console was exchanged the pump still did not start. ultimately, the impella cp was explanted and a new impella cp device was implanted and successfully started. there was no harm to the patient as a result of the event. latest update noted the patient currently remains on impella mcs. | [5 empty] | The investigation of pump did not start has been completed. data log review showed pump unable to start with pump stop alarm. pump was returned and inspected. inspection revealed damage to the outflow cage with one of the connecting struts angled in towards the impellor blades, preventing them from rotating freely. the pump stop reproduced in the lab. the pump passed electrical testing. the cause of the pump did not start was likely a damaged outflow cage due to the deformed outflow struts found on the returned pump which did not allow free rotation of the impellor blade. lack of clinical information relating to how the cage was damaged. d9 device returned to manufacturer date added. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: using the automated impella controller with the impella catheter. ¿achieving an act 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
10053 | 1220648-2025-27724 | 21881200 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21881200&pc=OZD | 03/29/2025 | 05/30/2025 |  | 04/22/2025 |  |  | 03/29/2025 | 04/22/2025 |  |  |  |  | 04/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025611592 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and during the support there was high motor current, and the pump stopped. the staff attempted to restart the pump multiple times without success. the impella cp device was explanted and replaced. the patient status as a result of the event was indicated as unknown. however, latest information indicated the patient continued on impella mcs. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the analysis, a supplemental report will be filed. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ | The investigation of pump stop has been completed. impella cp was returned for evaluation. upon visual inspection, a large purge gap was present under the impeller blade. no data logs were returned for evaluation. clinical details noted that a pump stop occurred on day 7 of support with a spike in motor current. pump was attempted to be restarted multiple times without success. the root cause of the pump stop was most likely due to bearing wear within indicated design life of 8 days per design trace matrix.
10054 | 1220648-2025-27725 | 21881425 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21881425&pc=OZD | 04/02/2025 | 05/22/2025 |  | 04/22/2025 |  |  | 04/02/2025 | 04/22/2025 |  |  | 09/30/2025 |  | 04/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024399464 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 10 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Crack | Fluid/Blood Leak | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was on impella cp support. it was reported that the impella cp had a purge leak alarm. upon inspection, the purge side was fractured at the base of the reservoir just before the junction of the leur lock to the filter. the impella cp was removed and the patient survived the procedure. | [5 empty] | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: cleaning: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of purge leak ¿ pump has been completed. the data logs were reviewed and there were no abnormalities noted with purge pressures or flows. purge pressure never dropped below 400 mmhg throughout the case, and there were no purge related alarms. returned product was investigated and the purge sidearm reservoir was completely broken off from the sidearm filter at the joint. based on data logs, this was not the state of the sidearm at the time of the reported leak or at any point in the case, as there would have been a severe drop in purge pressure and corresponding alarms. however, it is the location that the field reported the leak to originate from, so the initially reported crack likely propagated to fully fracture after the case or during return of the product. based on the clinical details provided and returned product, the root cause of the purge leak was most likely joint reservoir to filter. the cause of this damage is unclear as insufficient clinical details were provided.
10055 | 1220648-2025-27726 | 21881668 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21881668&pc=OZD | 04/02/2025 | 04/22/2025 |  | 04/22/2025 |  |  | 04/02/2025 | 04/22/2025 |  |  |  |  | 04/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026620819 |  | 1000100 |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Tamponade | Ectopic Heartbeat | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was on impella 5.5 support and on day one in the intensive care unit, the impella p-level was increased to p8, but the patient had a suction event along with ventricular ectopy, so it was decreased back to p7. it was further reported that on day five of support, the impella 5.5 was explanted urgently. the patient was found to have pericardial tamponade from the suspected impella graft leak. the patient's outcome at time of explant was noted as survived. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
10056 | 1220648-2025-27727 | 21881970 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21881970&pc=OZD | 04/02/2025 | 05/28/2025 |  | 04/22/2025 |  |  | 04/02/2025 | 04/22/2025 |  |  |  |  | 04/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025582301 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | Malposition of Device | [5 empty] | Tachycardia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp was not able to go above p3 without suction alarms. the impella cp was repositioned the next day as multiple positioning alarms occurred overnight. on day four of impella support, the impella cp was repositioned again. on day seven of impella cp support, the patient experienced torsade requiring cardiopulmonary resuscitation with return of spontanous circulation achieved. the impella appeared malposition in the chest x-ray and was repositioned. it was further reported that the patient expired on impella support. there is not enough evidence to exclude the impella cp as an associated factor in the patient's outcome. | [5 empty] | The investigation of vf/vt/af/at and positioning issues has been completed. the impella device was not returned for evaluation. the cause of the positioning issues most likely was use related due to cardiopulmonary resuscitation (cpr). as the necessary information was not provided, the root cause of the torsad arrhythmia was not determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10057 | 1220648-2025-27728 | 21882290 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21882290&pc=PYX | 03/27/2025 | 04/22/2025 |  | 04/22/2025 |  |  | 04/02/2025 | 04/22/2025 |  |  |  |  | 04/22/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025483246 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Pulmonary Embolism | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was on impella 5.5 support. echocardiogram showed the right ventricle was dilated and general hypokinesis. the patient decompensated, biventricular failure with frequent ventricular tachycardia on dual inotropes. cardiogenic shock was called and the decision was made to escalate care. the patient was taken for impella rp flex placement. the impella rp flex was successfully placed. long-term outcome and quality indicator (loqi) impella registry database reported upon two adverse events on 2025-apr-02 both with "unknown/undetermined" relationship to the impella use. adverse event three was for sustained ventricular arrhythmia (impella 5.5 pump) and adverse event four for right pulmonary embolus (impella rp flex pump). loqi noted that following: "(b)(6) 2025 presented with chest pain. ctpa showed a subsegmental pe and he was started on heparin. his symptoms have improved and he was transitioned to eliquis. echo showed ef 50-55%, rv mildly dilated with function low normal. fk level remained within appropriate limits. (b)(6) 2025 discharged in stable condition". this report represents the impella rp flex adverse event. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10058 | 1220648-2025-27729 | 21882432 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21882432&pc=OZD | 03/21/2025 | 06/03/2025 |  | 04/22/2025 |  |  | 03/26/2025 | 04/22/2025 |  |  |  |  | 04/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1719481 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Other | Health Professional | User facility | Company representation |  |  | 65 YR |  | Female | 68 KG | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | Yes | Initial submission | Followup | [8 empty] | It was reported that while a patient was on impella support, the purge pressure increased to over 1000. multiple nurses changed tubing, de-aired line, and purge cassette. however, the issue was not resolved. abiomed representative was called and asked that they exchange the automated impella controller. after switching the controller, the purge pressure returned within defined limits and the impella was no longer alarming. there was no patient harm. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - purge issue ¿ other has been completed. the data analysis confirmed the reported failure mode. the console was tested and the purge pressure accuracy was within specifications per the preventative maintenance procedure. unable to reproduce the ¿purge pressure high¿ (event #408) or ¿purge system blocked¿ (event #409) alarms. upon performing the de-air procedure and the cassette change procedure while the optical test pump was running, the purge pressure and purge flow returned to the appropriate values without triggering events #408 and #409." the root cause of the ¿purge pressure high¿ and ¿purge system blocked¿ alarms were most likely due to a use issue. it could be related to a kink in the purge line.
10059 | 1220648-2025-27730 | 21882473 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21882473&pc=OZD | 01/03/2025 | 04/22/2025 |  | 04/22/2025 |  |  | 04/02/2025 | 04/22/2025 |  |  |  |  | 04/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025564013 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported through abiomed¿s long-term outcome and quality indicator (loqi) impella registry, that a patient experienced sustained ventricular arrhythmia during impella 5.5 support. the patient was defibrillated in response to the condition, after which, the patient was in and out of ventricular tachycardia (vt). amiodarone and lidocaine bolus, 40 meq potassium, albumin, and lidocaine drops was administered at this time. that night, the patient once again experienced sustained vt coinciding with suction events. the condition was said to have an unknown/undetermined relationship to the use of the impella 5.5. the patient survived the event,. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10060 | 1220648-2025-27731 | 21882590 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21882590&pc=OZD | 04/14/2025 | 07/16/2025 |  | 04/22/2025 |  |  | 04/15/2025 | 04/22/2025 |  |  |  |  | 04/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025597974 |  | 1000459 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 41 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the patient was on impella 5.5 support and during support, there was high purge pressure. the purge cassette and the automated impella controller were replaced to no avail. the pump was then replaced. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The pump was returned for investigation. the data log shows small motor current spike while running on a steady p-level, followed by saturated pump flow for a short period of time. high purge pressure began to gradually rise after the motor current spike, culminating in a purge system high alarm within 15 minutes. no physical damage was observed on the returned product. the pump was primed, and high purge pressure was noted. the pump was started in a bucket of water, and purge pressure gradually returned to normal specification limits. the pump was further tear down to check the purge gap, and biomaterial was found in the purge gap. the cause of the high purge pressure issue was biomaterial, based on data log review and returned product investigation. d9 device returned to manufacturer date added.
10061 | 1220648-2025-27732 | 21882894 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21882894&pc=OZD | 02/14/2025 | 04/22/2025 |  | 04/22/2025 |  |  | 04/01/2025 | 04/22/2025 |  |  |  |  | 04/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1084911 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that a patient was placed on impella 5.5 support for hemodynamic support during bridge to transplant. during support it was reported that the pump stopped after the patient stepped on the cord. the patient coded and required cardiopulmonary resuscitation, shocked and emergent extracorporeal membrane oxygenation placement at bedside. after stabilizing the patient a new impella 5.5 was placed. furthermore, biomedical engineer advised that the console was involved with the incident with the patient and wanted the console to be looked at to double check if the console was properly working. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation has been completed. the impella controller was returned for investigation. six days into running, there was a pump stop followed by continuous inability to run the pump, even after having the plug disconnected and reconnected twice. the console was returned but performed as expected. 25-41574-1 describes how there was an electrical short among other pump issues stemming from being stepped on by the patient confirming that this console performed as expected. there was no problem found during the case with the console as it functioned/operated to specification.
10062 | 1220648-2025-27733 | 21883532 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21883532&pc=OZD | 04/14/2025 | 10/06/2025 |  | 04/22/2025 |  |  | 04/17/2025 | 04/22/2025 |  |  |  |  | 04/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025584374 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 39 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella cp was used on a patient and there was a motor current waveform issue. there was also suspicion of thrombus aspiration. the issue did not improve but the patient was monitored for a few days. however, the aorta position waveform disappeared, but there was no problem with the pump's operation, so the patient was switched to a left ventricle assist device. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the pump was returned for investigation. data logs were not provided for this case run. the returned pump failed the laser continuity test, and light was leaking at kink at the red handle's kink protector. additionally, an signal not reliable (snr) of 0 was observed during optical bench testing. the pump was able to run in a bucket of water as per specifications but did not display the placement signal value. the cause of the optical signal issue was most likely damaged optical fiber line, based on returned product investigation. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
10063 | 1220648-2025-27734 | 21883544 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21883544&pc=OZD | 02/02/2025 | 11/03/2025 |  | 04/22/2025 |  |  | 04/03/2025 | 04/22/2025 |  |  |  |  | 04/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1461907 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | Yes | Device Sensing Problem | [6 empty] | Insufficient Information | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [5 empty] | IMPELLA 5.5. | [18 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was on impella support and during support, there was a placement signal issues. the investigating engineer found that the placement signal issues were likely related to the automated impella controller. care was withdrawn and the patient expired. | The medical safety officer (mso) has reviewed the complaint and determined that the patient expired three days later after aic device issue for issue unrelated to the automated impella controller (aic). there was no consult exchange. no hemodynamic compromise, delay or interruption in therapy. procedure/therapy was continued with same aic. therefore, unlikely for chance of death or serious injury. |  |  |  |  | The investigation has been completed. no product was returned. eighteen days after the pump was started, the console displayed the message ¿placement signal not reliable (opm invalid signal, optical pump connected) ¿ alarm.¿ according to the long term logs during the peak signal-to-noise ratio alarm, the pump flow remained unchanged, confirming that there was no impact on pump functionality. after 22 hours, a ¿controller error (opm comm crc invalid) [optical pump connected] ¿ alarm¿, occurred but was resolved within a second, with no impact on pump functionality as per the ong term log. support continued for an additional one day and 16 hours. after the pump was disconnected from the automated impella controller (aic) on (b)(6) 2025, the log recorded abnormal optical parameters, especially light: 0.45v and snr: 0.5. the root cause for the controller error and loss of placement signal is due to an optical bench firmware communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic software operated as designed. | The following information has been revised according to the medical safety officer (mso) review on manufacturer device report 1220648-2025-27734. b1 revised to reflect product malfunction only. b2 should have been reflected as blank. b5 mso statement. h1 type of report revised to reflect malfunction. h6 revised to reflect appropriate coding for malfunction. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data analysis: ¿ the case associated with the reported complaint was initiated on (b)(6) 2025, with the pump (sn: (b)(6)). ¿ eighteen days after the pump was started, the console displayed the message ¿placement signal not reliable (opm invalid signal, optical pump connected) ¿ alarm,¿ event #1036. ¿ according to the lt log during the psnr alarm, the pump flow remained unchanged, confirming that there was no impact on pump functionality. ¿ after 22 hours, a ¿controller error (opm comm crc invalid) [optical pump connected] ¿ alarm,¿ event #1045, occurred but was resolved within a second, with no impact on pump functionality as per the lt log. ¿ support continued for an additional 1 day and 16 hours. ¿ after the pump was disconnected from the aic on (b)(6) 2025, the log recorded abnormal optical parameters, especially light: 0.45v and snr: 0.5. the console successfully passed the preventive maintenance procedure on march 3, 2025. the optical parameters were good. device analysis: this console was tested after arriving at fs on (b)(6) 2025. although there were no optical signal issues at the customer site between (b)(6) and (b)(6) 2025, the analog output from the optical bench was found to be distorted. after running the console for 12 hours, the console displayed ¿controller error (defective optical sensor lamp) ¿ alarm,¿ event #1039. although there was a typical 4.8v average pwm input voltage at the optical bench lamp connector p4, the lamp was off, and the digital multimeter confirmed infinite resistance across the lamp. root cause: ¿ the root cause of the psnr issue and oscillating contrast was most likely a malfunction in the optical bench lamp. ¿ the root cause for the controller error (event #1045) and loss of placement signal is due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d9 updated. h3 updated. h6 updated. d10 concomitant medical products and therapy dates updated.
10064 | 1220648-2025-27735 | 21883776 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21883776&pc=OZD | 02/04/2025 | 04/22/2025 |  | 04/22/2025 |  |  | 04/01/2025 | 04/22/2025 |  |  | 12/31/2025 |  | 04/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025468856 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella 5.5 pump supporting in (b)(6) 2025 for 23 days and was then weaned and explanted. the spanish agency reported to abiomed that the medical team reported upon an ischemic stroke and that there was a thrombus/biomaterial observed in the patient or on the device. further comments from the reporter were translated as quoted "moderate-severe disability, requires assistance walking. measures taken: end of product use. in recovery." | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death". in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
10065 | 1220648-2025-27736 | 21883835 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21883835&pc=OZD | 04/14/2025 | 07/16/2025 |  | 04/22/2025 |  |  | 04/15/2025 | 04/22/2025 |  |  |  |  | 04/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1325615 |  | 0042-0000-JP |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 41 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support, there was high purge pressure. the purge cassette and the automated impella controller were replaced to no avail. the pump was then replaced and the issues resolved. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - purge issue ¿ other has been completed. the impella device was not returned for evaluation. the console log confirmed the reported failure mode. the case started on the (b)(6) 2025 at 20:59:41 but with the console and within 10 minutes of starting the pump, the "purge pressure high event 408" alarm was triggered, prompting an attempt to change the purge cassette. despite this, the "purge pressure high" alarm persisted and was not resolved. subsequently, the aic was replaced on same day, however, the "purge pressure high" alarm continued to be active. ultimately, the resolution of the issue was achieved only after replacing the pump. the console was not returned to field service for further investigation. root cause: the root cause for the purge pressure high was not related to any problems with the aic. no issues were found given the console.
10066 | 1220648-2025-27737 | 21884200 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21884200&pc=OZD | 04/14/2025 | 07/15/2025 |  | 04/22/2025 |  |  | 04/16/2025 | 04/22/2025 |  |  |  |  | 04/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1581781 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient on impella cp experienced a console (aic) controller error alarm during use in the icu room which automatically cleared after 30 seconds.. the procedure was completed with the impella. there was no health hazards to the patient. | [5 empty] | Console logs from the reported day of event are consistent with complaint and show console unexpectedly shut down and rebooted. after the reboot, connected pump was recognized and restarted, data streaming had been disabled (by design), and console presented with advisory alarm ¿unexpected shutdown¿ (#603). interruption of hemodynamic support was no longer than 30 seconds due to pump motor driver reinitialization during boot-up. console was tested at fs but issue was not reproduced, and console operated within specification. root cause of the unexpected shutdown was not determined. most likely component that could have contributed to the issue is 4-in-1 assembly, including cf cards (no file system corruption was observed). repair: replace 4-in-1 assembly (including cf cards) as a precaution and perform functional check of the console. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
10067 | 1220648-2025-27738 | 21884684 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21884684&pc=OZD | 03/29/2025 | 04/22/2025 |  | 04/22/2025 |  |  | 03/29/2025 | 04/22/2025 |  |  |  |  | 04/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026634760 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  |  |  | Death | Yes | Yes | Failure to Pump | [6 empty] | Cardiac Arrest | [12 empty] | Death | Required Intervention | [23 empty] | A |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in cardiogenic shock from acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and after placement, the physician attempted to start the pump on auto. however, there was an immediate motor current high alarm noted, and the pump failed to start. the impella cp device was removed and reinserted. the second attempt to start the pump met with the same motor current high alarm and the pump failed to start. the decision was made to place a new impella cp device which was successfully completed via the right femoral artery. however, during the percutaneous intervention to the 100% occluded right coronary artery, the patient coded. the team was unable to obtain return of spontaneous circulation, and the patient expired. | [5 empty] | Medical safety review of the event noted there is not enough evidence to exclude the impella cp pump 1 as an associated factor in the patient's outcome. delay in needed support; may have add to an already complex situation with the patient. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10068 | 1220648-2025-27739 | 21884956 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21884956&pc=OZD | 03/14/2025 | 07/14/2025 |  | 04/22/2025 |  |  | 04/16/2025 | 04/22/2025 |  |  |  |  | 04/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026622088 |  | 1000459 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 27 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported that while a patient was on impella 5.5 support they noticed purge fluid was leaking from the sidearm. there was no change in the purge pressure, so the impella 5.5 was continued to be used for several days, considering that there was no problem with its continued use. there was no patient harm. | [5 empty] | The pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | H.3 device not returned to manufacturer for evaluation should of reflected the device was available to be evaluated on the initial manufacturer device report number 1220648-2025-27739. if the device was not evaluated has been revised to reflect device evaluation anticipated, but not yet begun. | The investigation of purge leak ¿ pump has been completed. the pump was returned for investigation. data log review was not required for his case. the returned pump was inspected, and stress mark was found on the back of the reservoir. leak test was performed, revealing a leak from the membrane radial seal of the purge reservoir during priming. the root cause of the purge leak - pump was determined to be due to a damage sidearm reservoir as evidenced by the results of the leak test.
10069 | 1220648-2025-27740 | 21885359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21885359&pc=OZD | 04/03/2025 | 04/22/2025 |  | 04/22/2025 |  |  | 04/03/2025 | 04/22/2025 |  |  |  |  | 04/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026620195 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had an impella cp pump support ongoing for a patient admitted in with a non-st-elevation myocardial infarction. it was reported that the impella was positioned towards mitral. unable to achieve flows greater than p2 and unable to reposition pump at this time. the decision was made to remove the impella. the patient remained stable and survived the procedure. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed.
10070 | 1220648-2025-27741 | 21886502 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21886502&pc=OZD | 04/03/2025 | 05/22/2025 |  | 04/22/2025 |  |  | 04/03/2025 | 04/22/2025 |  |  |  |  | 04/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026635392 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was admitted for acute myocardial infarctions/cardiogenic shock was implanted with an impella cp. the patient was infused with one unit of red blood cells to address hemolysis. after eight days of support care was withdrawn and the patient expired. the relationship between the impella cp and cause of death is unknown. | [5 empty] | Impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states), acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). the impella cp was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. instructions for use: section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2025-27741. h6 health effect - impact code 1802 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27741. h10 instructions for use were incomplete on the initial manufacturer device report number 1220648-2025-27741.
10071 | 1220648-2025-27742 | 21886513 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21886513&pc=OZD | 04/03/2025 | 04/28/2025 |  | 04/22/2025 |  |  | 04/03/2025 | 04/22/2025 |  |  |  |  | 04/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025607089 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after an impella cp was implanted, limited flow was noted. fem-to-fem bypass sheaths were placed to improve flow to the impella leg. the distal sheath was exchanged several times while on impella support. the last time the sheath was exchanged, clots were observed and the sheath was removed. a clot traveled down the leg, and the leg became molded and cold. the impella cp was exchanged for a new impella cp. a femoral cut down was performed to remove the thrombus. the patient was noted to survive device explant. | [5 empty] | Impella cp with smart assist system instructions for use for the related event are as follows: section: potential adverse events (united states), acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). the impella cp was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of limb ischemia and thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2025-27742. d6b explant date was erroneously entered on was erroneously entered on the initial manufacturer device report number 1220648-2025-27742. h6 health effect - impact code 4624 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27742. h6 component code 3039 was erroneously entered on the initial manufacturer device report number 1220648-2025-27742.
10072 | 1220648-2025-27743 | 21886560 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21886560&pc=OZD | 04/04/2025 | 04/22/2025 |  | 04/22/2025 |  |  | 04/04/2025 | 04/22/2025 |  |  |  |  | 04/22/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | IMPELLA CONTROLLER, PACKAGED, US |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Yes |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | It was reported that the plastic piece that holds the purge cassette air sensor in place on an impella aic was broken. the aic was removed from service. no patient was involved. | [5 empty] | The investigation for the reported console purge disc/flag issues has been completed. the console was returned for evaluation and it was confirmed that the top left area of the retainer had a large crack. the logs were not relevant to the reported issue. the blue purge disc retainer, purge disc detect flag, and purge flag spring were replaced. after a full functional check the device was returned to the customer for use. the root cause of this issue was a broken purge retainer.
10073 | 1220648-2025-27744 | 21886587 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21886587&pc=OZD | 04/04/2025 | 09/24/2025 |  | 04/22/2025 |  |  | 04/04/2025 | 04/22/2025 |  |  |  |  | 04/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026636240 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Malposition of Device | Device Sensing Problem | [5 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | It was reported that a patient implanted with an impella 5.5 experienced ventricular tachycardia and an echocardiogram showed the impella 5.5 was positioned slightly deep. the device was pulled back into position and the patient was treated with a bolus of amiodarone and lidocaine. eventually the impella 5.5 was removed and the patient receive cardiopulmonary bybass. | The customer reported that the device exhibited an optical signal issue. |  |  |  |  | The impella 5.5 was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of vf/vt/af/at and positioning issues has been completed. the pump was returned for investigation. some clamping marks were found near the motor housing, which do not affect the product's performance during the test. the pump passed the laser continuity test and the optical bench test. the data log shows many intermediate spikes to 200 in the placement signal (ps) during the reported discrepancy in ps numbers by the doctor. there were several pump position in aorta alarms present during the ps spike event. the signal not reliable (snr) was lower than the acceptable range throughout the case. the real time (rt) log was closely reviewed, and it was found that the placement signal and raw optical signal barcodes were independent and did not show any trends in motor current or pump flow during that time. the cause of the positioning issue was not determined since sufficient clinical information was not provided. cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. timing of the arrythmia event in relation to impella support (during or post-implant) electrocardiogram (ekg) recordings of the arrhythmia episodes. evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. response to any antiarrhythmic treatments or interventions during the event. any documented device-related issues or complications during impella support evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. d9 date device returned to manufacturer added. d4 model number was erroneously entered on the initial manufacturer device report number. e4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration. h5 was erroneously entered on the initial manufacturer device report number. h6 health effect - clinical code 2130 was erroneously entered on the initial manufacturer device report number. new health effect - clinical code entered. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the pump was returned for investigation. some clamping marks were found near the motor housing, which do not affect the product's performance during the test. the pump passed the laser continuity test and the optical bench test. the pump also underwent pressure and water bucket tests for the optical sensor, and no abnormalities were recorded on the placement signal or the 5s optical signal parameters. the data log shows many intermediate spikes to 200 in the placement signal during the reported discrepancy in ps numbers by the doctor. there were several pump position in aorta alarms present during the ps spike event. the snr was lower than the acceptable range throughout the case. the rt log was closely reviewed, and it was found that the placement signal and raw optical signal barcodes were independent and did not show any trends in motor current or pump flow during that time. the cause of the optical signal issue was not determined since it could not be reproduced on the returned product. please refer to (b)(4) for aic-related optical issues. the cause of the positioning issue was not determined since sufficient clinical information was not provided. cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. ¿timing of the arrythmia event in relation to impella support (during or post-implant) ¿electrocardiogram (ekg) recordings of the arrhythmia episodes ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. ¿response to any antiarrhythmic treatments or interventions during the event. ¿any documented device-related issues or complications during impella support. ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. ¿review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical signal issue "2917" as part of a retrospective review of optical signal issues in accordance with fda recommendation.
10074 | 1220648-2025-27745 | 21886602 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21886602&pc=OZD | 04/04/2025 | 06/06/2025 |  | 04/22/2025 |  |  | 04/04/2025 | 04/22/2025 |  |  |  |  | 04/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026629577 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | It was reported that after explantation of an impella 5.5 pump the patient required shock twice and medical management of cardiac arrthymia with a "possible" relationship to the impella use. reportable harm. later, the patient exhibited acquired vonwillebrand disorder with "probable" relationship to the impella 5.5. it was reported that the patient was successfully weaned from the impella. | [5 empty] | The impella 5.5 was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of access site bleeding and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d4. model number was erroneously entered on the initial manufacturer device report number 1220648-2025-27745. e4. should have been left blank on the initial manufacturer device report number 1220648-2025-27745 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1. manufacturing site name, address, country, and fax number were erroneously entered on the initial manufacturer device report number 1220648-2025-27745.
10075 | 1220648-2025-27746 | 21886605 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21886605&pc=OZD | 04/04/2025 | 05/14/2025 |  | 04/22/2025 |  |  | 04/04/2025 | 04/22/2025 |  |  |  |  | 04/22/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025561456 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Cardiac Perforation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | It was reported that while implanted with an impella 5.5 the patient required ventricular septal defect (vsd) repair surgery. it was suggested to pull the impella back into the aorta, however the physician proceeded to perform the procedure without shallowing the impella. after the vsd repair, a perforation was noted in the inferior lateral wall of left ventricle. the perforation was patched and support continued. additionally, signals were lost in the patient's right arm. vascular surgery found the right brachial artery signal, but no signal was found distal. the right radial artery catheter was removed and vascular cutdown was performed at the brachial artery. an attempted thrombectomy toward radial and vascular ultimately concluded this was likely not a embolic event, but a combination of decreased flow, low cardiac output and vasoconstriction. the impella 5.5 was weaned successfully. | [5 empty] | The investigation of ventricle perforation and limb ischemia has been completed. the impella device was not returned for evaluation. based on the clinical details the cause of the ventricle perforation issue was most likely use related, based on given clinical details. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2025-27746. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27746 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella 5.5 was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10076 | 1220648-2025-27699 | 21871767 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21871767&pc=OZD | 03/25/2025 | 04/21/2025 |  | 04/21/2025 |  |  | 03/25/2025 | 04/21/2025 |  |  |  |  | 04/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026627621 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that there was concern for hemolysis with a patient on impella cp support. the hemolysis was resolved by reduction of patients after load. no change was made to the impella, positional or otherwise, and the impella was not removed for hemolysis. the impella was weaned and removed as the patients heart had recovered. the pump was discarded. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella pump was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." section: use of echocardiography for positioning of the impella catheter: "evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it." impella catheter too far into the left ventricle: "obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine." section: hemolysis: "hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis." "patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis." section: suction: "suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure."
10077 | 1220648-2025-27700 | 21871871 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21871871&pc=OZD | 03/31/2025 | 04/21/2025 |  | 04/21/2025 |  |  | 03/31/2025 | 04/21/2025 |  |  |  |  | 04/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025610360 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on impella 5.5 support and during the implant, the abiomed 0.018 inch guidewire kinked and the impella 5.5 was unable to make the turn. the doctor used a boston science platinum plus wire to implant the pump. support was continued and the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10078 | 1220648-2025-27701 | 21872007 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21872007&pc=OZD | 04/01/2025 | 04/24/2025 |  | 04/21/2025 |  |  | 04/01/2025 | 04/21/2025 |  |  |  |  | 04/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026629259 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient developed a hematoma from the axillary access site during impella 5.5 support due to heart failure. a washout was performed to manage the condition and support was continued with the implanted device. the patient was bridged to transplant. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
10079 | 1220648-2025-27702 | 21872159 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21872159&pc=OZD | 04/14/2025 | 04/21/2025 |  | 04/21/2025 |  |  | 04/14/2025 | 04/21/2025 |  |  |  |  | 04/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025595057 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient on impella cp support developed lower limb ischemia. antegrade sheath was inserted but sufficient perfusion was not achieved. support continued until successful weaning and explant. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
10080 | 1220648-2025-27703 | 21872320 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21872320&pc=OZD | 04/01/2025 | 05/16/2025 |  | 04/21/2025 |  |  | 04/01/2025 | 04/21/2025 |  |  |  |  | 04/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1670339 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the automated impella controller (aic) did not prime purge fluid when starting a new case. there was some difficulty inserting the cassette into the aic. patient impact is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - purge issue ¿ other has been completed. the console logs were consistent with what was reported in the complaint. the reported purge priming issue was unable to be reproduced. console was able to prime without any issues and operated as intended while testing with a pump and purge cassette kit. no issues were found with the accuracy of the purge pressure sensor during testing. the reported purge cassette insertion issue was also unable to be reproduced. the purge cassette was able to be inserted without any issue. the console was not able to prime the purge because insufficient pressure was building up in the purge disc during the priming process, likely due to a kink in the purge line. the real time log data showed that the purge pressure did not increase from 0 while the console was fast priming. the root cause of the purge fluid priming issue was most likely a kink of the purge line. the root cause of the issue with inserting the purge cassette was unable to be determined due to the inability to reproduce. d9 date device returned to manufacturer added.
10081 | 1220648-2025-27704 | 21872466 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21872466&pc=OZD | 03/29/2025 | 04/21/2025 |  | 04/21/2025 |  |  | 03/29/2025 | 04/21/2025 |  |  |  |  | 04/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568936 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. approximately three after start of the support purge pressure and purge system blocked alarm occurred. lysis protocol was implemented, and the impella purge values returned to normal. no patient harm was reported. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
10082 | 1220648-2025-27705 | 21872532 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21872532&pc=OZD | 03/31/2025 | 05/19/2025 |  | 04/21/2025 |  |  | 03/31/2025 | 04/21/2025 |  |  |  |  | 04/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025607588 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp aborted due to patient's anatomy, unable to pass through vasculature. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, aborted insertion due to patient's anatomy.
10083 | 1220648-2025-27706 | 21872821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21872821&pc=OZD | 03/31/2025 | 04/21/2025 |  | 04/21/2025 |  |  | 03/31/2025 | 04/21/2025 |  |  |  |  | 04/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025605332 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hematoma | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support and during support, the patient had an access site hematoma. manual pressure was held. two units of blood products were given for low hemoglobin. support continued. additionally, the pump was turned to p2 and the patient had ventricular tachycardia. no shocks were given, the p level was increased and the position was checked. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10084 | 1220648-2025-27707 | 21873905 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21873905&pc=OZD | 03/25/2025 | 04/21/2025 |  | 04/21/2025 |  |  | 03/25/2025 | 04/21/2025 |  |  |  |  | 04/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025587279 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  | White | Injury | Yes | Yes | Failure to Advance | [6 empty] | Ectopic Heartbeat | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on impella 5.5 support and during the implant, the impella 5.5 could not advance past the impella cp in the ascending aorta with the .018 inch guidewire. a .027 inch wire was pulled from impella rp kit to support the case. the impella was implanted successfully. additionally, the patient had ectopy while on support. the pump was pulled back 2 centimeters. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
10085 | 1220648-2025-27708 | 21874035 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21874035&pc=PYX | 03/26/2025 | 05/20/2025 |  | 04/21/2025 |  |  | 03/26/2025 | 04/21/2025 |  |  | 10/31/2025 |  | 04/21/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024423950 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella rp support and after the implant, there was high purge pressure observed. the purge cassette was replaced with no avail. tissue plasma activator was added to the purge. support continued until care was withdrawn and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella rp with smartassist system: section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ | The investigation of high purge pressure has been completed. the impella device was not returned for evaluation. the data logs confirm high purge pressure alarms. the logs show a shift in motor current and placement signal coinciding with a drop in pump flows. after this, there is a 12 minute rise in purge pressure until high purge pressure alarm is triggered. clinical mentions tissue plasminogen activator (tpa) was added and logs show purge pressure values returned to normal after about 4-5 hours. the root cause of the high purge pressure was most likely biomaterial build up as evidenced by the gradual rise in purge pressure after motor current and placement shift.
10086 | 1220648-2025-27709 | 21874309 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21874309&pc=OZD | 03/31/2025 | 09/24/2025 |  | 04/21/2025 |  |  | 03/31/2025 | 04/21/2025 |  |  |  |  | 04/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026615084 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Device Sensing Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, there was access site bleeding. blood products were given and heparin was withheld due to the bleeding. additionally, the pump was showed deep on echocardiogram and the doctor pulled back 1cm. the patient remains on support. | Us complainant reported the patient was on impella cp support and during support the placement signal was not correlating with arterial line pressures. motor current and flow were within limits for p level. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. no product was returned for evaluation. access site bleeding - major: clinical details noted that the patient had increased jp output from impella access site. heparin was stopped and patient received multiple units of blood products. additionally, pressure dressing was applied and bleeding eventually resolved. the root cause of the access site bleeding could not be definitively determined was limited clinical details were provided. positioning issues: clinical details noted that "placement signal low" alarms were being triggered throughout support. echo showed that pump position was slightly deep and was pulled back twice to be measured at 3.7cm. clinical details noted that pump was secured and touhy-borhst was locked after pump insertion. the exact time of the positioning issues was not noted. the root cause of the positioning issues could not be definitively determined as limited clinical details were provided on the time of initial mispositioning of the pump. optical signal issue: data logs were reviewed and lt log showed "placement signal low" alarms being triggered during support. signal not reliable (snr) is low and variable during support, contrast is variable with all other optical parameters remaining stable during support. clinical details noted that pump position was slightly deep in ventricle with "placement signal low" alarms were triggered with ao signal not matching patient's arterial line. pump was repositioned and measuring at 3.7cm in ventricle and positioning alarms resolved and ao waveforms were matching arterial line. the root cause of the optical signal issue could not be definitively determined as pump repositioning did improve suspected optical issue, however, no product was returned to confirm no issues with the optical sensor. upon review of data logs, it is apparent that the console is a potential contributing factor of the issue as low snr was present from the start of the case and repositioning the pump improve the reported discrepancy of the ao waveform and patient's arterial line as stated in the clinical details. please refer to complaint # (b)(4) for an investigation into the console. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical signal issues "2917" as part of a retrospective review of optical signal issues in accordance with fda recommendation.
10087 | 1220648-2025-27710 | 21874532 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21874532&pc=OZD | 03/31/2025 | 05/13/2025 |  | 04/21/2025 |  |  | 03/31/2025 | 04/21/2025 |  |  |  |  | 04/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025595590 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support and during support, the patient had oozing at the access site. fresh frozen plasma was given to the patient. support was continued until successful explant. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10088 | 1220648-2025-27711 | 21874878 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21874878&pc=OZD | 01/20/2025 | 09/25/2025 |  | 04/21/2025 |  |  | 04/01/2025 | 04/21/2025 |  |  |  |  | 04/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025579395 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  | Hispanic | White | Death | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | Sepsis | Thrombocytopenia | Thrombosis/Thrombus | [8 empty] | Death | Required Intervention |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant also reported that the impella 5.5 optical sensor stopped working. flows and motor current were confirmed to be accurate. | On april 1st, 2025, the vigilance unit of medical devices of the spanish agency of medicines and medical devices has been informed through renacer registry about an incident. on (b)(6) 2025 the patient on impella 5.5 support developed heparin induced thrombocytopenia and was switched from heparin to argatroban. on (b)(6) 2025 non-central nervous system arterial thromboembolism and major intra-abdominal hemorrhage occurred. anticoagulation was discontinued and left bicarbonate in the purge. there was spontaneous bleeding from branches of the hepatic artery with hypovolemic shock. attempted embolization by interventional radiology was noted. on (b)(6) 2025 the patient suffered an ischemic stroke with symptoms but there was no significant disability. anticoagulation was resumed. occipital stroke without sequelae. recovered without sequelae. on (b)(6) 2025 the patient had klebsiella pneumoniae infection and septic shock and expired two days later. |  |  |  |  | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue: product was returned and evaluated. the pump's 5s parameters were within specification. the pump passed laser continuity testing. the root cause of the optical signal issue was not determined since data logs were not available to analyze the sensor's behavior during the case. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-27711. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: table 3.2 impella catheter components. ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation of infection/sepsis, stroke/cva, vascular damage - peripheral vessel, thrombus, and thrombocytopenia have been completed. data logs are not available. product was returned and evaluated. the pump's 5s parameters were within specification. the pump passed laser continuity testing. as the necessary information was not provided, the root cause of infection/sepsis, stroke/cva, vascular damage - peripheral vessel, thrombus, and thrombocytopenia was not determined.
10089 | 1220648-2025-27712 | 21875359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21875359&pc=OZD | 03/27/2025 | 06/11/2025 |  | 04/21/2025 |  |  | 03/27/2025 | 04/21/2025 |  |  |  |  | 04/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636203 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Malposition of Device | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Rupture | [11 empty] | Death | Disability | Required Intervention | Life Threatening | [21 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with undergoing high-risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support (mcs) and experienced a vessel rupture and positioning issues. the patient would subsequently expire. while inserting the 14fr peel away sheath, the vessel ruptured, and the groin was treated with a covered stent to close the damaged vessel. the puncture was sealed with manta closing device. the physician decided to switch to the right side for access, and the impella cp device was implanted. during the high-risk pci, after intervention of the left anterior descending artery, the patient became critically unstable. the impella cp device kept patient on a low-level stable. left ventricular ejection fraction was 10% prior to the procedure and became worse. it was further noted that during the impella mcs, suction was present due to lack of preload and beginning right heart failure. several positioning alerts were present due to the patient's movement, and the pump had to be repositioned several times. subsequently, the was transferred to a university hospital for treatment but would ultimately expire. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | Vascular damage - peripheral vessel: during initial implantation attempt from left groin, the vessel ruptured and the groin was treated with a covered stent to close the damaged vessel. puncture was sealed with manta closing device. the pump was not returned. the cause of the vascular damage - peripheral vessel was not determined since product was not returned and insufficient clinical details. positioning issues: after placement of the impella cp and following hrpci procedure, suction was present due to lack of preload and beginning rhf. several positioning alerts were present due to patient movement. the pump was not returned. the cause of the positioning issues was most likely patient movement due to the clinical information stating positioning alarms occurred during patient movement.
10090 | 1220648-2025-27713 | 21875440 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21875440&pc=OZD | 03/31/2025 | 05/08/2025 |  | 04/21/2025 |  |  | 03/31/2025 | 04/21/2025 |  |  |  |  | 04/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026617682 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a decision for the impella cp was made after a patient was found to be in cardiogenic shock secondary to peripartum cardiomyopathy and ejection fraction had decompensated. the impella cp was successfully placed. there were no impella cp issues during support. the pump was successfully weaned and explanted. the medical doctor attempted to use the one perclose and the patient lost distal pulses at that time so surgical cutdown. vascular repair was completed and pulses were regained. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of the limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
10091 | 1220648-2025-27714 | 21875522 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21875522&pc=OZD | 03/27/2025 | 05/05/2025 |  | 04/21/2025 |  |  | 03/27/2025 | 04/21/2025 |  |  |  |  | 04/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026623933 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 57 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [10 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support and during support the patient developed a nosebleed approximately 13 days after start of the support. heparin was stopped in response to the condition. three days later, a reduction in flow was noted. the pump was repositioned to troubleshoot the issue, after which the patient went into a three-minute run of ventricular tachycardia as reported under the abiomed's long-term outcome and quality indicator impella registry noted with a definite relationship to the impella device used. bedside echo performed. impella measuring 5.1 and appears to possible be suctioning on mitral valve. impella rep notified and interventional fellow on call came to bedside as well as impella rep. during this time, patient motor current wave form flattened, and flows reduced from 2.5 to 1. nipride stopped. during repositioning and evaluation patient went into sustained ventricular tachycardia, given a push of amiodarone with eventual chemical conversion back to normal sinus rhythm right before anesthesia arrived at bedside. unable to get a cuff pressure during this time and r radial arterial line was placed by anesthesia. a line and impella pressures seem to be correlating. the patient remained alert and oriented throughout the event. ldh and lactic acid after event were normal. the patient was transported to operating room for impella replacement due to the low flow with flat motor current. a clot was visualized on the device. the patient recovered with no sequelae. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the analysis, a supplemental report will be filed. section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The pump was returned for investigation. no physical damage was observed on the returned product, other than a cannula kink near inlet cage. the pump was then soaked in a bucket of water to remove dried blood. biomaterial was found inside the cannula at the inlet cage after soaking the pump. the pump was then tested in the bucket of water and ran according to specifications. the pump was also primed, and purge-related issues were noted. as per the given clinical information, the pump flow was reduced, and the doctor confirmed that the pump was away from the mitral valve and some way interacted with papillary muscle. pump flow was not improved after repositioning. the data log showed no signs of motor current spikes during the low pump flow event. it was confirmed that flow dropped gradually on a steady pressure level of p5. additionally, there was a cannula kink, and biomaterial was observed on the returned product, but pump flow was within specification limits during test. the cause of the low pump flow issue was not determined since the issue could not be reproduced on the returned product, and the data logs were not sufficient to derive the root cause without adequate clinical information. the cause of the vascular damage issue was not determined since sufficient information was not provided. the relation between the nosebleed and the impella was unknown. as the necessary information was not provided, the root cause of the thrombus and vt/vf was not determined. b2 life threatening was erroneously selected on the initial manufacturer device report number 1220648-2025-27714.
10092 | 1220648-2025-27715 | 21875959 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21875959&pc=OZD | 03/27/2025 | 04/21/2025 |  | 04/21/2025 |  |  | 03/27/2025 | 04/21/2025 |  |  |  |  | 04/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025548322 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. patient was also cannulated for extracorporeal membrane oxygenation. following the implant, the impella cp. device was noted to be possibly in the papillary muscle and was repositioned. flows post repositioning were noted to be "good." the patient received two units of packed red blood cells (prbcs) and was administered albumin. it was subsequently noted the patient developed a hematoma at the impella access site. the following day, there were multiple suctions alarms, but issue was responsive to volume. another one unit of prbcs was administered. a computed tomography performed revealed right groin extravasation around impella insertion site but currently stable. the next day the patient underwent left heart catheterization with percutaneous coronary intervention to the left anterior descending artery. however, the next day systemic anticoagulation was withheld due to stable the right groin hematoma. the unit of platelet were transfused overnight and with transfusing one unit of prbc this day. however, the following day, the impella cp device was removed after washout. the impella cp was replaced, patient escalated to an impella 5.5 device. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.. instructions for use for the related event are as follows: impella cp with smartassist system. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
10093 | 1220648-2025-27716 | 21876182 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21876182&pc=OZD | 03/27/2025 | 10/29/2025 |  | 04/21/2025 |  |  | 03/27/2025 | 04/21/2025 |  |  |  |  | 04/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026631177 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock with critical left main, left anterior descending and left circumflex arteries disease was implanted with an impella cp device for mechanical circulatory support (mcs). approximately two days after start of the support, the patient developed a hematoma at the access site. heparin was withheld and mcs continued. it was reported the patient remained critical but stable condition following the event. the patient was noted as a possible transfer for higher level of care but did undergo percutaneous coronary intervention at current facility. latest update noted the patient was successfully weaned and the device was explanted after approximately six days of total support with a survived outcome. | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured loss of pulses in right lower extremity with a "definite" relationship to the impella device used: ¿on (b)(6) 2025, the patient underwent impella device removal in the catheterization lab due to loss of pulses in the right lower extremity. the procedure was initially well tolerated, with the patient remaining hemodynamically stable. peripheral angiography confirmed successful hemostasis at the impella access site.¿ |  |  |  |  | The investigation of access site bleeding and limb ischemia has been completed. the impella device was not returned for evaluation. access site bleeding - major: the root cause of access site bleeding is determined to be anticoagulation management because the hematoma was stabilized after reducing the heparin. limb ischemia - reversible: the root cause of the limb ischemia was unable to be determined due to insufficient clinical details. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. b5 updated to include additional information received from the clinical site h6 updated to reflect additional information received from the clinical site g2 report source; study was missing from manufacturer device report 1220648-2025-27716.
10094 | 1220648-2025-27717 | 21877848 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21877848&pc=OZD | 03/25/2025 | 09/25/2025 |  | 04/21/2025 |  |  | 03/25/2025 | 04/21/2025 |  |  |  |  | 04/21/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026629043 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Malposition of Device | Device Sensing Problem | [5 empty] | Hematoma | Hemolysis | Tachycardia | Thrombosis/Thrombus | Implant Pain | [8 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 device via the right auxiliary artery for mechanical circulatory support (mcs) and the following day, the patient experienced ventricular tachycardia requiring a shock by the implantable cardioverter defibrillator. an ultrasound was completed to verify placement. the impella 5.5 device was deep under transthoracic echocardiogram and was pulled back. two days later, a non-occlusive thrombus in right radial artery was noted. subsequently, approximately four days later on a computed tomography angiogram, it showed a right pectoral muscle enlargement concerning intramuscular hematoma. the patient had increasing pain overnight. the patient remains status two for heart transplant. following four days. lactate dehydrogenase were hemolyzed. heparin was withheld two days later, and right arm pain was much lower. the patient remained status 2 for transplant and continued on impella mcs. latest update noted approximately 12 days later the patient was planned to undergo the heart transplant. | Us complainant reported that a patient was on impella 5.5 for hemodynamic support. it was also noted of placement signal failure that spontaneously recovered after purge cassette change. |  |  |  |  | As last reported, the device remains implanted, supporting the patient at the time of the mdr writing, and therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. optical signal issue - (b)(6) 2025: placement waveforms accurate on impella. over the weekend, placement signal failure. spontaneously recovered after purge cassette change. data logs confirm the clinical narrative showing psnr alarms on (b)(6) with snr drop to 0, contrast oscillating, light and gain drop. lamp steps up to 100. upon review of data logs, it is apparent that the console is the potential cause of the issue. please refer to complaint # (b)(4) for an investigation into the console. on (b)(6) 2025 reports placement signal drift, 20pts lower on impella than on cuff pressure. then on (b)(6) 2025 pt map on impella reading 10 points lower on impella and pump position was confirmed via echo and placement signal wave form was visible. data logs for these dates confirm the clinical narrative and show variable ps. the mc is stable with no spikes. a few suction alarms on (b)(6) that resolve. the cause of the optical signal issue, ps lower than patient's cuff pressure, was unable to be determined due to lack of clinical information and no product returned. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical signal issues per a retrospective review of optical signal issues in accordance with fda recommendation.
10095 | 1220648-2025-27695 | 21869906 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21869906&pc=OZD | 03/28/2025 | 05/30/2025 |  | 04/19/2025 |  |  | 03/28/2025 | 04/19/2025 |  |  |  |  | 04/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025567735 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that an implanted impella cp pump slipped into the patient's aorta at the time that cpr was being performed on the patient. the pump was replaced in response to the positioning issue. the patient reportedly passed away the following day. there is not enough evidence to disassociate the patient's passing with impella support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ a1-a5 are unknown. | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. clinical states that the patient required cardiopulmonary resuscitation (cpr) which resulted in impella slipping into aorta. pump was not returned for investigation. as per the clinical information provided, the root cause of the positioning issue was use related to the cpr performed while pump still inside the patient. e4 should have been left blank on the initial manufacturer device report number: 1220648-2025-27695 as it is unknown if the initial reporter also sent the report to the food and drug administration.
10096 | 1220648-2025-27696 | 21869907 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21869907&pc=OZD | 03/28/2025 | 05/15/2025 |  | 04/19/2025 |  |  | 03/28/2025 | 04/19/2025 |  |  | 12/31/2025 |  | 04/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025446064 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility had a cp pump placed for the patient admitted in cardiogenic shock. the cp was placed but, the patient experienced signs of hemolysis and limb ischemia. the hemolysis did not prompt intervention. the ischemia was seen and had distal perfusion placed. the patient survived and was weaned off after 97 hours of support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia and hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27696 as it is unknown if the initial reporter also sent the report to the food and drug administration. h6 component code 3038 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27696.
10097 | 1220648-2025-27697 | 21869908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21869908&pc=OZD | 03/28/2025 | 05/05/2025 |  | 04/19/2025 |  |  | 03/28/2025 | 04/19/2025 |  |  |  |  | 04/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026629851 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that a patient was on 5.5 pump support initiated for a cardiomyopathy and was deemed not to be a transplant candidate prior. the patient had signs of infection at the site of insertion and was treated with antibiotics. the patient bridged to left ventricular assist device (lvad). | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (unites states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿ | The investigation of infection/sepsis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27697 as it is unknown if the initial reporter also sent the report to the food and drug administration.
10098 | 1220648-2025-27698 | 21869909 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21869909&pc=OZD | 03/28/2025 | 05/02/2025 |  | 04/19/2025 |  |  | 03/28/2025 | 04/19/2025 |  |  |  |  | 04/19/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026622698 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that a 83 year old male patient experienced heavy bleeding following repositioning during impella 5.5 support. the pump was noted to have been implanted via direct access and roughly 700ml of blood was removed from their chest tube after repositioning during the night. five units of packed red blood cells were given to manage the condition. the patient was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (unites states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of access s site bleeding is determined to be patient movement because the patient was repositioned last night in icu. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27698 as it is unknown if the initial reporter also sent the report to the food and drug administration.
10099 | 1220648-2025-27651 | 21864247 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21864247&pc=OZD | 03/21/2025 | 06/11/2025 |  | 04/18/2025 |  |  | 03/21/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1102527 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the biomed requested a change of console due to an alarm on their circuit board. the device was set aside for inspection. there was no patient harm. | [5 empty] | Data analysis: the case associated with the reported complaint was initiated on march 20, 2025, with the pump (sn: (b)(6)). the case was performed for 17 hours and 29 minutes. multiple alarms were displayed during the case (imc_logs_ic1603.pdf). in the rlm logs, there were multiple rlc process reboots and rlm reboots. during these reboots, any alarms on the aic were not displayed on the impella connect (mar_20_rlc&rlm_reboot.pdf). this confirms the reported failure mode. note: the aic was misreported as a "circuit board." device analysis: this console was tested after arriving at fs. the rlm reboot was reproduced, and the logs show potential micro sd card corruption (ic1603_during_testing.png & reproduced_rlm_reboot.pdf). the envelope (without the optical test cavity) has a couple of minor dents (envelope_ic1603.pdf), which are unrelated to the reported failure mode. the front panel optical connector was found to be contaminated, and the debris could not be removed upon cleaning (unrelated to the reported failure mode). root cause: the root cause for the alarm not being displayed on the ¿impella connect¿ was due to the rlc and rlm reboots. the cause of the reboots was not determined, it could be attributed to intermittent micro-sd card corruption there was no controller failure alarm in the case associated to the reported complaint. it was misreported. ad hoc task: before returning this console (ic1603) back to the customer, fs was instructed to perform the following: replace the microsd card (0042-2000) due to intermittent corruption. replace the optical bench (004201012) due to an abnormal envelope. replace the front panel optical connector (0042-2736) due to debris. perform sections 23 & 24 of the pm procedure (0042-7309). perform a full functional check (0042-7316). | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
10100 | 1220648-2025-27652 | 21864459 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21864459&pc=OZD | 03/27/2025 | 06/05/2025 |  | 04/18/2025 |  |  | 03/27/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025611000 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Death | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the doctor wanted to optimize the pump positioning under the premise that it was too deep, but the catheter appeared to be stuck. the physician was unable to move the pump forward or backward into the desired location. the automated impella controller (aic) subsequently triggered a placement signal low alarm and the patient's blood pressure began to drop. the staff was advised not to increase the support level until the positioning issue could be resolved. upon follow up, it was discovered that during the repositioning attempts, the t-lock had been turned by the staff rather than the touhy-borst valve, which was why the catheter remained firmly in position. repositioning was performed successfully following the correction. the patient later passed away that same day. there is not enough information to disassociate the use of the impella with the patient outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings, cautions and precautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ understanding and managing impella catheter position alarms. ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ correct impella catheter position: ¿for optimal positioning of the impella catheter, the inlet area of the catheter should be 3.5 cm below the aortic valve annulus and well away from papillary muscle and subannular structures. the outlet area should be well above the aortic valve.¿ | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the positioning issue was use related to improper technique of using the touhy-borst lock.
10101 | 1220648-2025-27653 | 21864532 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21864532&pc=OZD | 03/27/2025 | 04/18/2025 |  | 04/18/2025 |  |  | 03/27/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1874262 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the impella aic displayed a control error message. there was no report of patient involvement. | [5 empty] | The investigation for the reported controller error issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the cause of the aic issue was a defective pbm board.
10102 | 1220648-2025-27654 | 21864630 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21864630&pc=OZD | 03/27/2025 | 09/16/2025 |  | 04/18/2025 |  |  | 03/27/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025580904 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an implanted impella 5.5 pump on a 53-year-old female began to experience high purge pressure during patient support. a purge system blocked alarm appeared shortly thereafter and tissue plasminogen activator was run through the purge to manage the condition. there was no resulting patient injury. | The family decided to withdraw care and the patient expired. there was report of high purge pressure; purge line was changed. there was no report or indication given information at hand there was low pump flow. per medical safety officer review based on information there was no interruption or unwanted change in therapy. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smartassist states the following: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of high purge pressure has been completed. the impella device was not returned for evaluation. data log review showed high purge pressure build up gradually over the course of 2 hours while purge flows trended down. the cause of the high purge pressure was most likely biomaterial due to the gradual build up of the purge pressure occurring, purge flow trending down, and clinical details that tissue plasminogen activator (tpa) resolved the issue. b5 mso statement was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27654.
10103 | 1220648-2025-27655 | 21864721 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21864721&pc=OZD | 03/27/2025 | 05/12/2025 |  | 04/18/2025 |  |  | 03/27/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025539838 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during the impella cp insertion on a 79-year-old male, the cp could not pass through the stented iliac. the impella inlet would buckle on fibrotic area. cp cannula observed to be completely occlusive around illiac stent lumen. physician at this time opted to explant cp given patient hemodynamic stability. the patient survived the procedure. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since the implant of pump was aborted due to resistance met because of small vessel size. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states) - ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions - ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings - ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings - ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
10104 | 1220648-2025-27656 | 21864892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21864892&pc=OZD | 03/28/2025 | 04/18/2025 |  | 04/18/2025 |  |  | 03/28/2025 | 04/18/2025 |  |  | 08/31/2025 |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024386082 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that at the start of the impell 5.5 implant, oozing occurred from the axillary kit sheath without impella yet in place. the oozing was located at the blue ring at pre perforated area. damage was observed on the sheath after removal. the sheath was exchanged, and impella was inserted successfully. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10105 | 1220648-2025-27657 | 21865001 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21865001&pc=OZD | 03/27/2025 | 05/19/2025 |  | 04/18/2025 |  |  | 03/28/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025561413 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 5.5 pump support was initiated to support a patient admitted in with known coronary artery disease and cardiomyopathy. the 5.5 was placed but there was a hematoma and after 6 days the team observed a purge sidearm leak. the pump remained on for support despite the leakage. the purge failure and purge pressure remain stable. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of purge leak ¿ pump has been completed. the pump was not returned for investigation. the data log was available until march 24th. according to the clinical details a purge side arm leak was reported on march 27th without affecting any purge parameters. the cause of the purge leak issue was most likely leak from the sidearm but the exact location of the leak was unknown without provided picture or returned product. d6b explant date added.
10106 | 1220648-2025-27658 | 21865207 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21865207&pc=PYX | 03/28/2025 | 04/18/2025 |  | 04/18/2025 |  |  | 03/28/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025503148 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that an implanted impella rp flex pump triggered high purge pressure alarms roughly five minutes following placement. the purge cassette and bags were replaced but the issue persisted. blood was then seen backing up in the purge line and the decision was made to replace the pump. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use states the following: impella rp with smartassist system section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
10107 | 1220648-2025-27659 | 21865438 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21865438&pc=OZD | 03/28/2025 | 05/21/2025 |  | 04/18/2025 |  |  | 03/28/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026613044 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Vascular Dissection | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient had an attempted impella 5.5 implant and during the implant, the 0.035 glide wire and al1 were inserted through peel-away sheath and attempted to cross aortic valve. as access was being attempted with the 0.035 wire, anesthesia noticed an aortic dissection with unknown origin. origin appeared to be in aortic arch and continued into descending. no ascending aortic dissection present. wire and catheter removed. the procedure was aborted and patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | The investigation of delivery issues and vascular damage - great vessel has been completed. the impella device was not returned for evaluation. the clinical states that the pump implant was aborted since there was an aortic dissection observed during use of 0.035 non-abiomed guidewire. no issue stated with introducer sheath or impella 5.5 pump itself. pump and guidewire were not returned. delivery issue: the cause of the delivery issue was most likely use related to the dissection with the non abiomed guidewire as that issue prevented delivery of the pump. vascular damage - great vessel: the cause of the vascular damage great vessel issue was most likely use related to the implant of the 0.035 non abiomed guidewire.
10108 | 1220648-2025-27660 | 21865521 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21865521&pc=OZD | 03/28/2025 | 04/25/2025 |  | 04/18/2025 |  |  | 03/28/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025572083 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Delivered as Unsterile Product | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a cp was placed for the support of a high risk cardiac patient. the support was for an hour and was then explanted. the team did report upon a sterility/packaging issue. there was a "work in progress" signage within the the sealed cp box. no noted harm or injury/infection noted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Upon review of the provided information, there is currently no allegation of a malfunction or serious injury associated with the implanted pump. a regulatory report is no longer necessary.
10109 | 1220648-2025-27661 | 21866393 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21866393&pc=OZD | 03/28/2025 | 04/18/2025 |  | 04/18/2025 |  |  | 03/28/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026627608 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Female |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient presented cath lab on two pressers for evaluation of mcs. upon arrival due to hemodynamic compromise md decided to proceed with impella cp support. the impella was placed and support was then initiated, however after a few minutes of runtime, continuous suction alarm displayed. troubleshooting and repositioning were attempted with no change at which point physician wanted to replace pump for a new one. the pump was removed, and replaced with another impella cp. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10110 | 1220648-2025-27662 | 21866442 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21866442&pc=OZD | 02/06/2025 | 04/18/2025 |  | 04/18/2025 |  |  | 02/06/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025608622 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Delivered as Unsterile Product | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the packaging of an impella cp device arrived with a broken seal. there was no patient involvement. | [5 empty] | The investigation has been completed. no product or images of issue were returned for evaluation. clinical details noted that the seal of the packaging was broken upon arrival to customer. the root cause of the packaging issues - sterility cannot be determined as no product or photos of issue in the field were returned and limited clinical details were provided.
10111 | 1220648-2025-27663 | 21866912 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21866912&pc=OZD | 03/26/2025 | 04/18/2025 |  | 04/18/2025 |  |  | 03/26/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026620092 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Deformation Due to Compressive Stress | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | In preparation for an impella 5.5 implant to a patient with cardiomyopathy, the scrub technician pulled the guidewire out of the plastic packaging and noticed the tip was damaged/bent. a platinum pro plus guidewire was used to implant the impella 5.5 device which was successfully completed. the patient was hemodynamically stable at end of case and transferred to unit on impella support there was harm to the patient as a result of the event. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10112 | 1220648-2025-27664 | 21867034 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21867034&pc=OZD | 02/01/2025 | 09/24/2025 |  | 04/18/2025 |  |  | 02/01/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025534530 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Arrhythmia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | We received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured recurrent ventricular tachycardia with a "possible" relationship to the impella device used: the user facility reported the patient was on impella cp support and during support, the patient had ventricular tachycardia (vt). the patient was cardioverted and the pump position was checked. additionally, there were persistent placement signal low alarms and intermittent ventricular waveform via placement signal. the position was verified and support continued. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. optical signal issue: it was stated that there were multiple ps low alarms but positioning was confirmed. the data logs were observed and there were ps low alarms but all the 5s parameters were stable. there were many positioning alarms seen, but they were accurately triggered. the motor current was also sightly flat throughout the case. due to the clinical stating that positioning was confirmed and there was no product returned, the root cause of the optical signal issue could not be determined. vt/vf: cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. timing of the arrythmia event in relation to impella support (during or post-implant). electrocardiogram (ekg) recordings of the arrhythmia episodes. evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. response to any antiarrhythmic treatments or interventions during the event. any documented device-related issues or complications during impella support. evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical signal issues "2917" as part of a retrospective review of optical signal issues in accordance with fda recommendation.
10113 | 1220648-2025-27665 | 21867296 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21867296&pc=OZD | 03/31/2025 | 05/27/2025 |  | 04/18/2025 |  |  | 03/31/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026622812 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 was removed due to low flows and concern for thrombus. the impella was having low flows with flat motor current. the impella position was confirmed with echocardiogram, but suction alarms occurred above p-5. the volume status was adequate. thrombus was never confirmed. there was no patient harm. | [5 empty] | The investigation of low pump flow has been completed. the returned complaint 5.5 pump visually inspected. biomaterial observed in purge gap and round impeller blades. no cannula kink or broken blade found. dried dextrose and blood washed for better investigation with no observation then impella 5.5 connected to console to reproduce reported low pump flow issue. pump run for more than 10 minutes. flow was with specified limit at p-9. after impella run in investigation lab still biomaterial observed on top of the impeller blades. insufficient data log returned only for first 2 days of support; so motor current (mc) spike observed no placement signal issue seen. suctions occurred and position in aorta alarms observed with no evidence the reported issue and date issue reported. low pump flow: the cause of the low pump flow issue most likely was biomaterial as observed on the returned product. failure mode thrombus added as a result of biomaterial observation. as the necessary information was not provided, the root cause of the thrombus was not determined. b1 updated to both adverse event and product problem based on investigation results. b2 updated based on investigation results. d9 device returned to manufacturer date added. h1 type of report updated to serious injury based on investigation results. h6 health effect - clinical code 4438 was added based on investigation results. instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10114 | 1220648-2025-27666 | 21867450 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21867450&pc=OZD | 03/26/2025 | 05/15/2025 |  | 04/18/2025 |  |  | 03/26/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025566817 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was being prepped for impella cp support and there was a priming issue with the pump. de-air was attempted with purge solution and visual purge solution out of outlet cage was not seen. purge flow on automated impella controller was not there. the pump was replaced and the patient remained stable. there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Additional information provided in h3, h6 and h10. the impella cp was returned for evaluation. upon visual inspection, no damage on pump was observed. pump was connected to a lab console and purged for approximately 15 minutes without issue. purge was observed exiting purge gap when connected to console. data logs at time of case start show that case start was completed and purge pressure and purge flows were gradually normalizing after case start. no purge or priming alarms noted. pump was disconnected and reconnected to console and purge pressure was continuing to normalize. no issue with purge flow ticks noted on logs. clinical details noted that pump was not priming correctly and no purge fluid was visualized at the outflow cage. additionally, no purge flow was present on the aic. upon review of the returned product and data logs, no problem was found with the pump during the case.
10115 | 1220648-2025-27667 | 21867490 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21867490&pc=OZD | 03/26/2025 | 05/14/2025 |  | 04/18/2025 |  |  | 03/26/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025597511 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, there was a sudden pump stop. the pump was explanted and the procedural outcome was the patient survived surgery. there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock. section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of pump stop has been completed. there were no data logs returned. the cp stopped after running for 216 hours. the pump was returned and no damages or biomaterial was seen. there was a large purge gap which is indicative of bearing wear. the pump was attempted to be restarted in the lab and ran for a short time before stopping again. the impella cp's design life is 8 days which is indicated in the design trace matrix 0046-9910. the root cause of the pump stop is most likely due to bearing wear based on the returned product. d6b explant date was erroneously entered on the initial manufacturer device report number 1220648-2025-27667.
10116 | 1220648-2025-27668 | 21867527 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21867527&pc=OZD | 03/26/2025 | 04/18/2025 |  | 04/18/2025 |  |  | 03/26/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026634123 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. hematuria was noted in indwelling catheter bag and dusky impella extremity was also noted. additionally, the impella looked deep and needed to be repositioned. the impella cp was replaced with a impella 5.5 and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
10117 | 1220648-2025-27669 | 21867980 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21867980&pc=OZD | 03/27/2025 | 05/15/2025 |  | 04/18/2025 |  |  | 03/27/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026617830 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Thrombocytopenia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that the 66 year old female patient was on impella 5.5 support and during support, the patient was profoundly thrombocytopenic requiring platelet transfusions. decision was made by surgeon not to start heparin. staff noticed dampening of the motor current, impella position unknown alarms, separation of the ao/lv waveforms and frequent suction events. echo confirmed completely collapsed left ventricle with the pump sitting deep in the ventricle. the pump was explanted. upon explant an extensive clot was seen within the inlet cage as well as almost completely occlusive at the outlet. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Additional information provided in h6 and h10. the investigation into the reported issue has been completed.no product was returned for evaluation. thrombocytopenia: the cause of the thrombocytopenia was not determined due to insufficient clinical details. low pump flow: data logs were reviewed and at time of reported low pump flow shows persistent suction alarms occurring with flows lower than expected range for p-level. motor current becomes variable and increases in amplitude without changes in p-level. no motor current spikes observed at time of reported low flows. clinical details noted that overnight motor current was observed to be dampening with "impella position unknown" and 'suction" alarms. patient was hemodynamically stable at the time with flows of 1,4l/min at p-4. additionally, it was noted that a large thrombus was present on the inflow and outflow cage at explant. the root cause of the low pump flow could not be definitively determined as no product was returned for evaluation. positioning issues: data logs were reviewed and at beginning of case shows pump was initiated then turned down to p-0 and then restarted. ¿impella position unknown¿ alarms were triggered persistently at beginning of case. the root cause of the positioning issues could not be definitively determined as limited clinical details were provided at time of reported issue. thrombus: extensive clot was seen on both inlet and outlet cage upon explant. thrombus can be observed in the body or on the pump upon explant. to determine the true root cause of the thrombus, sufficient clinical information mentioned would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined.
10118 | 1220648-2025-27670 | 21867984 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21867984&pc=OZD | 03/27/2025 | 05/29/2025 |  | 04/18/2025 |  |  | 03/27/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026613886 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that after positioning the 78 year old male patient, the placement signal of an impella cp pump became unreliable. there were no kinks at the time the signal was lost. the positioning of the pump was verified to be correct and the staff was advised to monitor the motor current if support was to continue. the decision was ultimately made to explant the pump. there was no resulting patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of optical signal issue has been completed. the impella cp was not returned. data logs were not recovered. the cause of the optical signal issue was not determined since product was not returned, data logs were not recovered, and insufficient clinical details. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27670 as it is unknown if the initial reporter also sent the report to the food and drug administration.
10119 | 1220648-2025-27671 | 21867989 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21867989&pc=PYX | 03/27/2025 | 05/15/2025 |  | 04/18/2025 |  |  | 03/27/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026623795 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 35 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that a 35 year old female patient with severe right ventricle failure was put on impella rp for hemodynamic support. during support, it was noted that due to ct surgeon against systemic heparin the care team decided not to start it. it was also reported that there was a motor uptick and decrease flows noticed on aic. they noted that crrt was clotting off. care team notified of clot ingestion suspicion and bolus of heparin given along with heparin gtt started. the patient was taken to cvor to perform one vessel cabg to rca. the impella had ingested clot so plan was to explant impella and place new one at end of case. the physician pulled impella without issue and placed new suture with good hemostasis. large clot noticed on exit cage of impella. one vessel cabg went well and rv thought to have recovered enough that team made decision not to place new impella. the patient survived the procedure. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental mdr will be filed. | The investigation of low pump flow, hemolysis, and thrombus has been completed. upon visual inspection, a large biomaterial was present on the impeller blade and around the purge gap. pump was connected to lab console and run for approximately 10 minutes at p-9. low pump flows did reproduce with pump flows running lower than expected p-level range for p-9. biomaterial was still present on pump after pump run. additionally, a dried clot was present at outflow cage of pump and remained in pump during pump run. clinical details noted that patient was not on systemic anticoagulation and at time of low flows, a uptick in motor current and decrease in flows was seen on the console. the root cause of the low pump flow was most likely due to biomaterial ingestion based on returned product investigation. hemolysis: biomaterial present on impeller blade and wrapped around purge gap on returned pump. data logs were reviewed and lt log of case showed a motor current spike with speed deviation, drop in pump flows, and step up in placement signal. clinical details noted that labs showed signs of hemolysis day after low flows and uptick in motor current was observed on the console. clinical team suspected that the ingested clot was causing the hemolysis. the root cause of the hemolysis was most likely due to biomaterial ingestion based on returned product and data log analysis. as the necessary information was not provided, the root cause of the thrombus was not determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27671 as it is unknown if the initial reporter also sent the report to the food and drug administration.
10120 | 1220648-2025-27672 | 21867994 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21867994&pc=OZD | 03/30/2025 | 04/18/2025 |  | 04/18/2025 |  |  | 03/30/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025544059 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | Gastrointestinal Hemorrhage | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient suffered a suspected gi bleed during impella 5.5 support. heparin was being run through the purge at the time of the noted issue. support was increased and heparin was removed from the purge shortly after. two days later, the patient was diagnosed with heparin induced thrombocytopenia. support continued uninterrupted. the pump was successfully weaned and explanted. | [5 empty] | The impella pump was not returned for investigation. should new information be received, a supplemental manufacturer device report will be filed. impella 5.5 for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella 5.5 with smartassist for use during section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
10121 | 1220648-2025-27673 | 21867998 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21867998&pc=OZD | 03/30/2025 | 04/18/2025 |  | 04/18/2025 |  |  | 03/30/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636916 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that during support it was noted that the insertion site to the left groin was oozing, it was dressing and re-enforced. additionally, the next day, it was noted that the site was stable. however, the patient received 1 unit of packed red blood cells (prbc). the pump was successfully weaned and explanted. | [5 empty] | The impella pump was not returned for investigation. should new information be received, a supplemental manufacturer device report will be filed. instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
10122 | 1220648-2025-27674 | 21868000 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21868000&pc=OZD | 03/31/2025 | 09/17/2025 |  | 04/18/2025 |  |  | 03/31/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1399060 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Display or Visual Feedback Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | Followup | [7 empty] | In the companion manufacturer device report number 1220648-2025-26400 during the investigation of the pump data logs it was discovered the automated impella controller (aic) console was experiencing issues/flickering lamp during support. | Additional information indicated that, during patient support, the pump experienced placement signal unreliable alarm. |  |  |  |  | The investigation into the reported issue has been completed. pump (b)(6) was run on (b)(6) from 1/20 to 2/23. starting on 2/15, there are unreliable placement signal alarms. the logs show that placement signal went to 3000 mmhg as the contrast oscillated and logs further confirmed this optical behavior. elements of the failure mode were reproduced during testing. when the console was turned on, there was a controller error for bad lamp confirmed by and the optical bench lamp was found to be off during this time. this alarm intermittently resolved and re-triggered during testing. the cause of the loss of placement signal and oscillating contrast was a malfunctioning lamp on the optical bench. | The console was returned for investigation. should new information be received, a supplemental manufacturer device report will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated with additional information. h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2025-27674.
10123 | 1220648-2025-27675 | 21868005 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21868005&pc=OZD | 03/30/2025 | 05/19/2025 |  | 04/18/2025 |  |  | 03/31/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026625409 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that according to the physician, the patient cut the cable on the red plug-in. the patient was stable. a new impella implantation was not planned at this time. | [5 empty] | The impella pump was returned for investigation. should new information be received, a supplemental manufacturer device report will be filed. instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of pump stop has been completed. the clinical details state that according to the physician, the patient cut the cable on the red plug-in. the patient is stable. the returned product shows the catheter cut and the kink protector removed from the red pump plug. the data logs also confirm an electrical open pump stop and ps going out of range at the same time. the root cause of the pump stop is due to open electrical lines due to excessive force. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-27675. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27675 as it is unknown if the initial reporter also sent the report to the food and drug administration.
10124 | 1220648-2025-27676 | 21868019 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21868019&pc=OZD | 04/01/2025 | 04/18/2025 |  | 04/18/2025 |  |  | 04/01/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025580982 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [5 empty] | INFUSION SET TYPE: AUTOSOFT XC INSULIN TYPE: HUM | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant had a impella cp pump placed to support a patient admitted in with a non-st elevation myocardial infarction and in need of support. during support it was reported that the site was bleeding and a hematoma developed. the team troubleshot the bleed by explanting the pump. the patient remained stable throughout the procedure and survived. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10125 | 1220648-2025-27677 | 21868131 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21868131&pc=OZD | 04/01/2025 | 04/18/2025 |  | 04/18/2025 |  |  | 04/01/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026616532 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient on impella cp support. it was reported that an echocardiogram showed left ventricular thrombus with a history of atrial fibrillation. the pump was removed the following day. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: impella cp with smart assist catheter insertion, section: axillary insertion of the impella cp with smart assist catheter, section: use of impella with transcatheter aortic valves, ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
10126 | 1220648-2025-27678 | 21868195 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21868195&pc=OZD | 03/31/2025 | 04/18/2025 |  | 04/18/2025 |  |  | 03/31/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025606218 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  |  | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that the decision for impella cp was made to support a patient. during explant of the cp the perclose was attempted to use for hemostasis, however upon removal of cp significant bleeding noted- concerns for fem artery tear. vascular was called to the operating room for intervention. cutdown was performed and left fem artery repaired and pedal signals were noted. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation has been completed. no product was returned. the root cause of vascular damage peripheral vessel cannot be determined since the product was not returned and insufficient clinical details were provided. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10127 | 1220648-2025-27679 | 21868227 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21868227&pc=OZD | 04/14/2025 | 04/18/2025 |  | 04/18/2025 |  |  | 04/14/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025553385 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that when replacing the purge system on the impella cp, the automated impella controller (aic) displayed a message, the screen went black, and impella temporarily stopped. the aic started up again and then started supporting the patient. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: entering cardiac output ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿
10128 | 1220648-2025-27680 | 21869635 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21869635&pc=OZD | 08/09/2024 | 04/18/2025 |  | 04/18/2025 |  |  | 03/26/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Other | [5 empty] | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | Yes | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the following was reported by medwatch: "impella placed and dad never recovered and got extremely worse". no further information was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10129 | 1220648-2025-27681 | 21869639 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21869639&pc=OZD | 03/26/2025 | 04/18/2025 |  | 04/18/2025 |  |  | 03/26/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026616512 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Respiratory Insufficiency | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a impella 5.5 pump placed to support a patient admitted in with underlying cardiac and systemic comorbidities. the patient was enrolled in the impact study. the patient was supported for on the 5.5 for just under 17 days when weaned/explanted. afterwards the impact study reported upon an ae with "definite" relationship to the impella procedure of tracheostomy. the field rep was aware of tracheostomy but noted it was performed to optimize patient after being on ventilator support for multiple weeks. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed.
10130 | 1220648-2025-27682 | 21869640 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21869640&pc=OZD | 03/26/2025 | 04/18/2025 |  | 04/18/2025 |  |  | 03/26/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1824431 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) was not recognizing the purge disc. new aic used successfully. there was no patient harm. | [5 empty] | The investigation has been completed. the impella pump was returned for investigation. data review: there was no evidence found in the logs to confirm the reported failure. device analysis: the console was examined after arriving. a broken purge disc flag was identified. the purge disc flag was observed to be broken and was the root cause of the reported purge disc not detected alarm issue.
10131 | 1220648-2025-27683 | 21869646 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21869646&pc=OZD | 03/26/2025 | 04/18/2025 |  | 04/18/2025 |  |  | 03/26/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025607011 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a impella 5.5 pump that had been placed to support a patient awaiting transplant. the pump was supporting for 8 days when there were low flows and suction. the team troubleshot with meds/volume and pump position assessment. the pump remains on for support despite the persistent low flows and the patient remains stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10132 | 1220648-2025-27684 | 21869647 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21869647&pc=OZD | 03/26/2025 | 04/18/2025 |  | 04/18/2025 |  |  | 03/26/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025583047 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient developed ischemia during impella cp support. there were concerns for rhabdomyolysis due to the poor perfusion. a right leg fasciotomy was performed and a distal perfusion catheter was placed to manage the condition. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
10133 | 1220648-2025-27685 | 21869648 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21869648&pc=OZD | 03/26/2025 | 05/15/2025 |  | 04/18/2025 |  |  | 03/26/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025511481 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 21 YR |  | Male |  | Hispanic |  | Malfunction | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | Required Intervention |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient (unknown race) with heart failure cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support as a bridge to heart transplant. patient was currently listed status two for a repeat heart transplantation. the patient was returned to the cardiac catheterization lab for a biopsy and a right heart catheterization with a ramp study to assess the hemodynamics under echogram and with a swan at various performance levels to test the performance of the impella pump. it was demonstrated that the device was not adequately unloading, and the patient required additional support during wait for a redo heart transplant. the patient was reported to be therapeutic on systemic anticoagulation. a decision was made to explant and replace the device. the impella 5.5 was explanted without issue and the right axillary graft was used for extracorporeal membrane oxygenation arterial cannulation. when the impella 5.5 device was explanted, there was thrombus confirmed on the inlet, and this was felt to be a factor. the patient was reported to be in critical condition following the event. | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. a 5.5 impella was implanted to replace for the previous 5.5 pump. once the second pump appeared to be in position per fluoroscopy, a transthoracic echo was done to confirm placement. adjustments were made, although at the end of the procedure, it was still thought that the pump may be pointed towards the mitral valve. under fluoroscopy and transthoracic echo, pump appeared to be in the free left ventricle space away from the septum and measuring 5.04 cm, at 44 cm at the exit site. the pump was explanted shortly afterwards. |  |  |  |  | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. product and data were not returned for review. the root cause of positioning issue was unable to be determined as limited clinical details were provided and no product was returned for review. corrected information provided in b5, h1 and h6.
10134 | 1220648-2025-27686 | 21869649 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21869649&pc=OZD | 03/26/2025 | 04/18/2025 |  | 04/18/2025 |  |  | 03/26/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026614761 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 49 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient (unknown race and ethnicity) with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. after placing the impella cp device, oozing from access site after positioning the repositioning sheath was noted. additional compression with added hemostatic sponge was used to treat the bleeding the patient was transported to the unit on impella support and noted to be in stable condition following the event. the patient was supported for approximately five days before being successfully weaned and the impella cp device explanted with a survived patient outcome. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10135 | 1220648-2025-27687 | 21869652 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21869652&pc=OZD | 03/26/2025 | 07/30/2025 |  | 04/18/2025 |  |  | 03/26/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025599971 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 88 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Malposition of Device | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Tachycardia | Cardiogenic Shock | [8 empty] | Death | Required Intervention |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Additional notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured bleeding with a "possible" relationship to the impella device used: ¿patient was indicated for red blood cell transfusion with a hgb less than 9 or a hct of less than 27 with symptomatic anemia. patient's labs were drawn and a hgb was found to be 6.9 g/dl the morning of (b)(6) 2025. patient was given 1 unit or rbc's and blood re-draw was found to be 8.3 g/dl, further rbc's were not needed at that time.¿. the patient recovered with no sequelae. | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and approximately one day after start of the support, the patient experienced tachycardic events for which cardioversion was required. the physician arrived and repositioned the pump to optimal placement and decreased the support level to assist with management of the tachycardia events. the following day, the impella was inadvertently pulled back into the aorta when the patient was being turned; it was noted the nurse forgot to check the touhy borst. the physician was unable to visualize the issue. the felt the patient was hemodynamically stable and the decision was made to explant the impella cp device. the patient survived the explant and was reported to be in stable condition following the event. | An additional loqi notification was received and indicated that the patient developed acute anemia secondary to hemolysis believed to have a probable relationship to impella cp support. a blood transfusion was performed to treat the condition. it was then noted that the patient suffered from an acute kidney injury with a probable relationship to impella support. the condition was believed to have stemmed from cardiogenic shock, impella-related hemolysis, and contrast nephropathy. the patient was started on a bumetanide infusion and continuous renal replacement therapy in response to condition. it was reported that this was ongoing at the time of the patient's passing. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. |  |  |  | The investigation of positioning issues and vf/vt/af/at has been completed. the pump was not returned for investigation, and a data log review was not required for this case run. the cause of the positioning issue was most likely use related since the pump was pulled back to aorta during turn. as the necessary information was not provided, the root cause of the vt/vf was not determined. b5 revised with additional information received from the clinical site. h6 health effect - clinical code and health effect - impact code revised to reflect the additional information received. the investigation vascular damage - peripheral vessel has not been completed. should the device or any new information be received, a supplemental mdr will be filed. g2 report source; study was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27687. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | Instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ b2 death was inadvertently omitted on manufacturer device report number 1220648-2025-27687-1. date of death is unknown. b5 additional information and patient outcome was inadvertently omitted on manufacturer device report number 1220648-2025-27687-1. h1 type of report was erroneously entered on manufacturer device report number 1220648-2025-27687-1. h6 health effect - clinical code 1886, 2041, 2262 and health effect - impact codes 4644 and 1802 were inadvertently omitted on manufacturer device report number 1220648-2025-27687-1. h10 instructions for use were inadvertently omitted on manufacturer device report number 1220648-2025-27687-1.
10136 | 1220648-2025-27688 | 21869654 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21869654&pc=OZD | 03/28/2025 | 05/15/2025 |  | 04/18/2025 |  |  | 03/28/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025589201 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella 5.5 was placed for the bipella support of a patient being evaluated for a transplant. after 2 weeks, the pump had purge pressure high alarms that prompted tissue plasmingen activator (tpa) and pump removal/replacement. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: the instructions for use for the impella 5.5 with smartassist states the following: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of high purge pressure has been completed. data logs were not available for review. returned product was investigated, and after soaking the pump, biomaterial residue - not a clot/thrombus - was visualized inside the purge gap and in the outflow cage. there were no other visual abnormalities. the pump and returned purge cassette were connected to a console and run in a bucket. high purge pressure immediately reproduced and pump flows were saturated for the respective p-level. a window was cut in the catheter just proximal to the motor housing, and no blood ingress was noted in the purge line. the catheter was then cut in that same location, and high purge pressure resolved, isolating the blockage to the motor housing/purge gap. though biomaterial residue was seen in the purge gap and high purge pressure reproduced, the root cause of the high purge pressure could not be determined, as data logs were not available for review to investigate the nature of the rise in purge pressure. d9 date device returned to manufacturer added.
10137 | 1220648-2025-27689 | 21869660 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21869660&pc=OZD | 03/28/2025 | 04/24/2025 |  | 04/18/2025 |  |  | 03/28/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025592673 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella cp pump was inadvertently placed too low in the patient right groin. despite being able to implant the device, the staff had concerns with the access point and opted to place a new impella cp pump via the left femoral artery. the pump was explanted and a hematoma developed at the site. no intervention was required to address the hematoma. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings. ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely use issue since the pump was placed too low and was in superficial femoral artery (sfa).
10138 | 1220648-2025-27690 | 21869661 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21869661&pc=OZD | 03/28/2025 | 09/27/2025 |  | 04/18/2025 |  |  | 03/28/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026618333 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | Device Sensing Problem | [5 empty] | Tachycardia | [12 empty] | Death | Required Intervention |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the impella cp was placed and was taken to icu, however the patient was in sustained ventricular tachycardia. during support there was a call regarding ongoing placement signal not reliable (psnr) alarm. the nurse reported active alarm with audio disabled. placement signal waveforms remained. additionally, the there was sporadic wrong position alarms. motor current (mc) was also narrow. the physician was advised to confirm position. and they are working on an art line. ultimately the patient was decompensating. the family did not want cpr and the patient expired. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of positioning issues has been completed. after a few hours on support, sporadic wrong position alarms occurred. prevailing care team theory is patient is non-pulsatile. the patient began to decompensate in icu. product was not returned. the cause of the positioning issue was not determined since product was not returned and insufficient clinical details. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue and ventricular tachycardia. pump was not returned. data log review showed that during placement signal not reliable alarms, signal-to-noise ratio (snr) decreases with contrast, gain increases, and lamp and light remain constant. the cause of the optical signal issue was not determined since product was not returned, data logs showed no known optical failure patterns, and insufficient clinical information. based on clinical review, patient was in sustained ventricular tachycardia, which had started prior to implantation. the root cause of the vt/vf/at/af was unable to be determined due to insufficient clinical details. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated to include placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-27690. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated.
10139 | 1220648-2025-27691 | 21869662 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21869662&pc=OZD | 03/28/2025 | 05/30/2025 |  | 04/18/2025 |  |  | 03/28/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025540846 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 87 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Component Missing | Packaging Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that upon going to place a 13cm 14fr introducer for impella cp implant, it was noted that the opened kit was missing the introducer. another kit was opened for introducer placement. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of packaging issues has been completed. the impella device was not returned for evaluation. the root cause of the packaging issue was determined to most likely be a missing component based on the clinical details provided. however, it could not be verified without returned product. h6 medical device problem code 3007 was erroneously entered on the initial manufacturer device report number 1220648-2025-27691.
10140 | 1220648-2025-27692 | 21869675 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21869675&pc=OZD | 03/28/2025 | 04/18/2025 |  | 04/18/2025 |  |  | 03/28/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026623841 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, it was noted that with impella at p3, flows were 1.7. there were intermittent ¿preventing retrograde flow alarms¿ when map above 85. additionally, apical laminated thrombus was noted on am bedside echo. position at that time was felt to be deep and pump was repositioned 1cm. the patient was returned to or for pump removal due to thrombus and intraaortic balloon pump was inserted. the patient outcome is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
10141 | 1220648-2025-27693 | 21869677 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21869677&pc=OZD | 03/28/2025 | 10/02/2025 |  | 04/18/2025 |  |  | 03/28/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025586987 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | Embolism/Embolus | No Clinical Signs, Symptoms or Conditions | [17 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The user facility reported the patient was on impella cp support and during support, there were high purge pressure alarm present. the purge cassette was working and it seem that the problem was in the sidearm. additionally, there were placement signal issues. there were impella in aorta alarms that came up 4-5 times when the impella was properly placed. no patient harm was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental mdr will be filed. | The investigation of high purge pressure has been completed. the returned product was evaluated. no data logs were returned for analysis. high purge pressure: the returned product had no abnormalities observed. no biomaterial on impeller or catheter kinks. the pump was run in the lab and high purge pressure did not reproduce. the root cause of high purge pressure was not determined as no issue reproduced on the returned product and no logs were returned for analysis with limited clinical information related to failure. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. the returned product had no abnormalities observed. the 5s parameters were within range and the pump passed the light continuity test. the data logs show a motor current spike with speed deviation before the onset of the impella in aorta alarms indicating possible biomaterial interaction. there were also suction alarms during the same time frame. the root cause of the optical signal issue was most likely due to biomaterial causing false 'impella in aorta' alarms as evidenced by the data log analysis and clinical evidence of correct placement. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated to include placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2024-11809. d10 concomitant medical products and therapy dates updated. | The returned product was evaluated. thrombus failure mode added based on log evidence of ingestion. high purge pressure: the returned product had no abnormalities observed. no biomaterial on impeller or catheter kinks. the pump was run in the lab and high purge pressure did not reproduce. the data logs show a gradual rise in purge pressure after the pump was started until high purge pressure alarms were triggered. purge system blocked alarm was also triggered. high purge pressure was sustained. about 30 minutes into support there was a motor current spike with speed deviation indicative of an ingestion. about an hour later there was another ingestion event with a motor current spike and speed deviation followed by low pump flows. high purge pressure was still sustained. the root cause of the high purge pressure was determined to be biomaterial as evidenced by data log analysis and returned product testing. thrombus: as the necessary information was not provided, the root cause of the thrombus was not determined.
10142 | 1220648-2025-27694 | 21869678 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21869678&pc=OZD | 03/28/2025 | 05/21/2025 |  | 04/18/2025 |  |  | 03/28/2025 | 04/18/2025 |  |  |  |  | 04/18/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported via a publication entitled "cholecystectomy during impella cp support in cardiogenic shock", that a patient developed heparin induced thrombocytopenia during impella cp support. heparin was removed from the purge in response to the condition. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ a1-a5 are unknown. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. e.1 name prefix/title, first name and last name are unknown. h.4 device manufacture date is unknown. | D6: implant/explant date unknown. the investigation of thrombocytopenia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. citation: mikolaj blaziak, jarocki, m., weronika wietrzyk, katarzyna modrzejewska, izabela swierczek, & wiktor kuliczkowski. (2024). cholecystectomy during impella cp support in cardiogenic shock. advances in interventional cardiology, 20(4), 509¿510. https://doi.org/10.5114/aic.2024.144781 g2 report source; literature was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27694. h10 article citation was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27694. article attachment was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27694.
10143 | 1220648-2025-27598 | 21853715 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21853715&pc=OZD | 12/13/2024 | 04/17/2025 |  | 04/17/2025 |  |  | 12/13/2024 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025536682 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 39 year old male patient on impella 5.5 support. during support it was noted that the pump kept turning to mitral valve and they were unable to properly position. the p level was reduced. the pump was found to be facing mitral apparatus and unable to flow higher than p-5. surgery tried to reposition and turn p-level higher and impella came out of lv. patient was taken to the or for 5.5 replacement. the patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
10144 | 1220648-2025-27585 | 21853728 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21853728&pc=OZD | 03/20/2025 | 04/17/2025 |  | 04/17/2025 |  |  | 03/20/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025575778 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 51 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient on impella cp experienced a marbled leg without foot pulses, however fem fem bypass was not being considered for the time being. additionally, it was noted that the impella was removed due to hemolysis, the lactate dehydrogenase and indirect bilirubin was not measurable, the day before it was strongly increased, however no echocardiogram was performed. position was noted in good position the day before. the patient was noted as stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
10145 | 1220648-2025-27610 | 21854052 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21854052&pc=OZD | 03/24/2025 | 05/15/2025 |  | 04/17/2025 |  |  | 03/24/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026623457 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the pump had a faulty optical sensor. the staff attempted to reposition the pump in response to the irregular readings, but the device was pulled out of the ventricle and into the aorta. when attempts to readvance the pump failed, the decision was made to explant the device due to the patient's hemodynamic stability. there was no patient harm. | [5 empty] | The pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | Additional information provided in h3, h6 and h10. the investigation into the reported issue has been completed. positioning issues the clinical stated that there was a placement signal low alarm so they repositioned the device. it was never confirmed they were having issues with positioning or the pump was in the incorrect position. echo images were poor so the impella was pulled completely out of the ventricle and was not able to be advanced back across the aortic valve. product was returned and there were no damages or abnormalities seen. due to lack of clinical information, the root cause of the positioning issues could not be determined. optical signal issue the clinical stated that the ao placement signal did not match the aortic pressure line. the data logs show that the 5s parameters are stable throughout the case. there was a slight trend in the ps and the flows around 5:15 pm. there are suction alarms and placement signal low alarms seen throughout the whole run. two impella position in ventricle alarms are confirmed towards the end of the run. they are in suction conditions for almost the whole case and confirmed to be in altering positions. there was no issue with the returned product. the pump was ran and no alarms reproduced. additionally, it was pressure tested using the uson and the sensor was able to hold pressure at the respected pressure settings. due to lack of clinical information provided, the root cause of the optical signal issue cannot be determined.
10146 | 1220648-2025-27608 | 21854266 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21854266&pc=OZD | 03/25/2025 | 05/02/2025 |  | 04/17/2025 |  |  | 03/25/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1251864 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the impella aic battery charge dropped to half when patient was ambulatory. no additional information on patient outcome was provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. e.1 name prefix/title, first name and last name are unknown. | The investigation of automated impella controller (aic) - battery charging/other issues has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
10147 | 1220648-2025-27602 | 21854492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21854492&pc=OZD | 03/25/2025 | 06/03/2025 |  | 04/17/2025 |  |  | 03/25/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025593522 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Renal Impairment | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the patient had a craniotomy for a subdural hematoma evacuation. the patient had heparin in the purge prior to this issue. | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured acute kidney injury with a "possible" relationship to the impella device used: ¿admitted with aki secondary to cardiac arrest. started on crrt.¿ the patient outcome recovered with sequelae. |  |  |  |  | D6 explant date: unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of vascular damage peripheral vessel and renal failure has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b5 updated with additional information received from the clinical site. d6b explant date added. h6 revised to reflect the additional information received from the clinical site. instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ g2 report source; study was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27602. h6 health effect - clinical code 1884 and health effect - impact code 4624 were inadvertently omitted on the initial manufacturer device report number 1220648-2025-27602.
10148 | 1220648-2025-27601 | 21854889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21854889&pc=OZD | 03/27/2025 | 04/30/2025 |  | 04/17/2025 |  |  | 03/27/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026623617 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, there was a small leak noted from the purge sidearm. the leak was not impacting the purge flow or purge pressure with plans to place bone wax around the side arm. no patient harm was reported. | [5 empty] | The investigation of purge leak ¿ pump has been completed. the pump was not returned for investigation. data logs show no trend related to the purge leak on purge parameters. the cause of the purge leak issue was most likely a leak from the sidearm, but the exact location of the leak is unknown. h5 was was erroneously selected on the initial manufacturer device report number 1220648-2025-27601. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
10149 | 1220648-2025-27613 | 21855126 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21855126&pc=OZD | 03/24/2025 | 04/17/2025 |  | 04/17/2025 |  |  | 03/24/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1411995 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that an automated impella controller experienced a battery failure while in use in the catheterization lab. there was no reported involvement. | [5 empty] | The investigation for the reported battery failure (red alarm) issue has been completed. the console logs were consistent with what was reported in the complaint. multiple battery failure (transport connection blown/missing) alarms were observed in the logs. the xonsole was tested. the reported battery failure was able to be reproduced. results of running batttest revealed that cell 1 of battery 1 had a voltage that was significantly less than the rest of the cell stacks in the battery. isolative testing was performed by swapping battery 1 with a known good one which resolved the alarm. the root cause of the battery failure was determined to be an internal cell imbalance in battery 1.
10150 | 1220648-2025-27609 | 21855324 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21855324&pc=OZD | 03/22/2025 | 04/17/2025 |  | 04/17/2025 |  |  | 03/22/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1273768 |  | 0042-0000-EU |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a purge disc not detected alarm appeared on the automated impella controller (aic) during patient mechanical circulatory support. the purge cassette was replaced. however, the issue persisted. the aic was exchanged to resolve the failure. there was no report of patient harm as a result of the event. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10151 | 1220648-2025-27607 | 21855432 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21855432&pc=OZD | 03/25/2025 | 06/03/2025 |  | 04/17/2025 |  |  | 03/25/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1153946 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an automated impella controller (aic) was turned on during preparation for patient support and displayed a controller error message. the aic was swapped in response to the noted failure. there was no patient harm. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. the data analysis confirmed the reported failure mode. this console was tested and upon power cycling the console 10 times, unable to reproduce the controller error. the root cause for the intermittent keyboard communication issue could not be determined due to the inability to reproduce.
10152 | 1220648-2025-27606 | 21855552 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21855552&pc=OZD | 03/25/2025 | 05/15/2025 |  | 04/17/2025 |  |  | 03/25/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026625256 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the decision for impella cp was made when a 80 year male patient continued to decompensate. during support they reported suction events that were related to volume and resolved after albumin and were able to increase p-level to 9. the patient was successfully weaned, however, had oozing around the impella site after a ¿turn¿ of the patient. manta was used to close the site. patient was successfully weaned. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and low pump flow has been completed. the product was returned and evaluated. no abnormalities were found with the repositioning sheath. data logs are not applicable to access site bleeding. clinical details state that the patient experienced bleeding at the impella site after being rolled onto their side after they had been weaned for explant. the root cause of the access site bleeding was most likely patient movement since clinical details noted that the bleeding began after rolling the patient onto their side. the returned pump catheter had been cut in the field, meaning the pump could not be run in the lab. visual inspection showed biomaterial near the purge gap and on the impeller. there were no cannula kinks found. data log review showed no motor current spikes/trends indicating biomaterial interaction. suction alarms occurred over the first night of support mostly resolving within 2 hours. there was 20 mmhg drop in placement signal prior to the suction alarms. clinical details reported suction alarms during the first night of support that they attributed to volume issues. the suction alarms resolved after giving 2x albumin based on clinical details. the root cause of the suction was most likely patient condition related since there was a drop in placement signal in the data logs and clinical details note the suction alarms resolved after giving volume.
10153 | 1220648-2025-27586 | 21855719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21855719&pc=OZD | 03/20/2025 | 04/17/2025 |  | 04/17/2025 |  |  | 03/20/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026634368 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was attempted to be implanted with an impella device for mechanical circulatory support. the physician was unable to advance the pump through the distal portion of the sheath. the decision was made to abort the procedure. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings; section: pre-support evaluation; section: axillary insertion of the impella 5.5 catheter; section: direct aortic insertion; section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
10154 | 1220648-2025-27593 | 21855817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21855817&pc=OZD | 03/27/2025 | 05/14/2025 |  | 04/17/2025 |  |  | 03/27/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1548488 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the automated impella controller (aic) had a loose rga port. no further details were provided. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic): damage before therapy has been completed. data analysis: not informative for this investigation. the logs did not show any anomalies. device analysis: the console was sent in for investigation. it was determined that the external vga port on the impella connect module/radio link module was defective. although the external vga port was not loose, the monitor displayed a blank screen when connected via the external vga port, making it impossible to extend the display from the aic to the external monitor. root cause: the root cause the external vga port on rlm, not connecting was due to the defective external vga port on impella connect module (rlm).
10155 | 1220648-2025-27600 | 21856011 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21856011&pc=OZD | 02/13/2025 | 04/17/2025 |  | 04/17/2025 |  |  | 02/13/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025605968 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 80 year old female patient with an impella cp placed for high risk cardiac procedure. the impella cp was placed and flows up to 3.6lpm. successful ivus guided poba of lm/lad with penumbra used for mechanical thrombectomy. decision made to leave peel-away sheath in place. contrast injection performed through sidearm with no flow distally. 14f peel-away exchanged for repo sheath and contrast injection was repeated and revealed no flow distally. decision was made to place swan and evaluate for tolerance of wean. the impella was successfully weaned and removed with 3 total perclose devices used. femstop was applied. the patient survived the procedure. | [5 empty] | The investigation for the limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details or the device return, the root cause of the reported limb ischemia could not be determined. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10156 | 1220648-2025-27611 | 21856065 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21856065&pc=OZD | 03/25/2025 | 05/12/2025 |  | 04/17/2025 |  |  | 03/25/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026620640 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella 5.5 pump on a 57 year old female patient was unable to advance through the axillary artery into the ascending aorta during implant. calcification was noted in the artery and the decision was made to abort the implant. an impella cp pump was subsequently placed for patient support. there was no resulting patient harm. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since the implant of pump was aborted due to calcification in artery and hence unable to advance the pump. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.
10157 | 1220648-2025-27590 | 21856123 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21856123&pc=OZD | 03/20/2025 | 05/13/2025 |  | 04/17/2025 |  |  | 03/20/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026633153 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Embolism/Embolus | [10 empty] | Other | Required Intervention |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following impella cp implant, the patient developed a bleed at their access site. it was noted that a hematoma developed shortly have the repositioning sheath was placed for support. manual pressure and a femostop was applied, but the issue persisted. the perclose sutures were eventually tightened and the sheath was re-sutured to resolve the bleed. the procedural outcome was the patient survived surgery. | Fibrin clot was discovered on the impella catheter. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and thrombus has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the access site bleeding issue was most likely use related, perclose suture limbs were tightened to create intra vascular tamponade, and sheath was re sutured. bleeding and hematoma successful resolved. as the necessary information was not provided, the root cause of the thrombus was not determined". instructions for use: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." b5 updated with information inadvertatnly omitted on the initial manufacturer device report number 1220648-2025-27590. h6 health effect - clinical code 4438 and health effect - impact code 4624 were inadvertently omitted on the initial manufacturer device report number 1220648-2025-27590. h10 instructions for use (thrombus) were inadvertently omitted on the initial manufacturer device report number 1220648-2025-27590.
10158 | 1220648-2025-27603 | 21856138 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21856138&pc=OZD | 03/26/2025 | 04/17/2025 |  | 04/17/2025 |  |  | 03/26/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026617468 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient undergoing high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support (mcs). impella mcs was successful for under one hour and then the device was explanted, which revealed a vascular dissection to the femoral access site that prompted a covered stent for hemostasis post dry seal patch. the patient was reported to be stable following the event without further complications. | [5 empty] | The impella has been completed. the product was not returned for evaluation. the explant revealed vascular damage to the femoral access site. the root cause of vascular damage peripheral vessel could not be determined since the product was not returned and insufficient clinical details were provided.
10159 | 1220648-2025-27588 | 21856365 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21856365&pc=OZD | 03/20/2025 | 05/12/2025 |  | 04/17/2025 |  |  | 03/20/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026623565 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Female |  | Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 pump was unable to advance through the patient's vasculature during attempted delivery. the implant was subsequently aborted. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, impella insertion aborted due to patient's anatomy.
10160 | 1220648-2025-27604 | 21856375 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21856375&pc=OZD | 03/25/2025 | 05/29/2025 |  | 04/17/2025 |  |  | 03/25/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025603413 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Female |  |  |  | Malfunction | No | Yes | Material Separation | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during impella cp implant, the covering on a .018 guidewire came loose. following successful implant, high resistance was met when attempting to pull back the guidewire. upon removal, it was noted that the outer wire sheath had come loose at the tip. no wire remnants were detectable at the site. there was no resulting patient harm. it was noted that the survival outcome at explant was that the patient expired. medical safety review of the event noted that there was no injury to the patient as a result of this device issue and there was no action taken to preclude...(no piece separation or remnant). the patient was taken off support and expired and based on information at hand death is unrelated to the impella device. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the pump was not returned. the guidewire was returned and the tip was uncoiled, however, based on the packaging of the returned wire, it is most likely not the abiomed 0.018 wire which was returned. delivery issue - the root cause of the delivery issue was not determined due to insufficient clinical details. guidewire damage- the root cause of the guidewire damage was not determined due to insufficient clinical details.
10161 | 1220648-2025-27594 | 21856420 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21856420&pc=OZD | 03/27/2025 | 05/23/2025 |  | 04/17/2025 |  |  | 03/27/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1218086 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a console with damage. the rga port was loose and not connecting. there was no reported patient involvement. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The complainant intended to complain about the vga port not the rga port. the incorrect console was reported for this complaint, as the correct console with a vga port issue was ic8532 as captured in manufacturer device report number 1220648-2025-27593. there was no malfunction on ic4074. the console did not need to be returned for investigation. thus, the complaint does not meet the criteria for mdr reportability as defined by 21 cfr 803.50. the issue will be tracked and trended internally.
10162 | 1220648-2025-27589 | 21856572 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21856572&pc=OZD | 03/20/2025 | 04/17/2025 |  | 04/17/2025 |  |  | 03/20/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025605045 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump stopped unexpectedly during patient support. the pump was restarted successfully, but the decision was ultimately made to replace the device. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: entering cardiac output ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
10163 | 1220648-2025-27605 | 21856611 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21856611&pc=OZD | 03/26/2025 | 05/08/2025 |  | 04/17/2025 |  |  | 03/26/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025575848 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the insertion of an impella 5.5 device to a patient in postcardiotomy cardiogenic shock was unsuccessful due to the patient's anatomy. during the procedure, the 10 millimeter hemashield platinum graft anastomosed to the axillary artery and 0.035¿ j wire was placed in the left ventricle preoperatively. coronary artery bypass graft surgery three times was performed with a plan to place the impella 5.5 device to come off bypass. however, the physician was unable to advance the pigtail and the j wire popped out of the left ventricle. after multiple attempts and various wires and catheters, the physician was still unable to get past calcification in innominate artery. the impella 5.5 device was opened and priming during this process. however, it was unable to be inserted into patient due to anatomical issues, even with manual manipulation while the chest was open. computed tomography scan was reviewed for alternative options, but none were viable. ultimately, the physician decided to support the patient with intra-aortic balloon pump. the patient was noted to be relatively stable with improved contractility on transesophageal echocardiogram at the end of the case. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, after multiple attempts and various wires and catheters, still unable to get past calcification in innominate artery.
10164 | 1220648-2025-27597 | 21856674 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21856674&pc=OZD | 03/25/2025 | 04/28/2025 |  | 04/17/2025 |  |  | 03/27/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026620065 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Other | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, there was non-sustained ventricular tachycardia. the procedure was successful and the patient's outcome is stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
10165 | 1220648-2025-27587 | 21856888 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21856888&pc=OZD | 03/20/2025 | 04/17/2025 |  | 04/17/2025 |  |  | 03/20/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636795 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the decision for impella cp placement with assisted percutaneous coronary intervention (pci) was made on a patient with multi vessel disease. it was reported that impella was inserted but due to patient anatomy, was unable to cross aortic valve after several attempts. the impella was removed and pci was done without support. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high-risk pci. section: potential adverse events (united states) - ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.)¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves - ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions - ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings - ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves - ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings - ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
10166 | 1220648-2025-27599 | 21856893 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21856893&pc=OZD | 03/27/2025 | 04/28/2025 |  | 04/17/2025 |  |  | 03/27/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1738105 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an automated impella controller failed to recognize that it was plugged in. errors "ac power disconnected, running on battery power" appeared even though the battery was charging while plugged in. the biomed engineer confirmed that the battery was reengaged upon return from preventative maintenance service and the issue persisted while plugged in. there was no reported patient involvement. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller - a/c power issues has been completed. it can not be determined from the logs whether or not the console was plugged in during these uncharging times, only that the console did not see ac power. lt power logs confirmed that the batteries were not charging when the imc logs indicated there was no ac power connection. device analysis: the console again was returned for investigation but the symptoms were not reproduced during any testing. root cause: the cause of the console allegedly not charging when plugged into wall power was not determined since failure mode was not reproduced.
10167 | 1220648-2025-27591 | 21857162 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21857162&pc=OZD | 03/21/2025 | 05/12/2025 |  | 04/17/2025 |  |  | 03/21/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026622868 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Required Intervention | Death |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the 70 year old female patient was on impella cp support and during support, an echo was performed which revealed the device had flipped and inlet was near papillary bed. echo was brought back to bedside and several attempts were made to reposition cp. the doctor believed the pigtail of the cp was wrapped/wedged into chordae tendineae. additionally, the pump was only able to achieve flows at p6. the pump was explanted and an impella 5.5 was implanted. the patient expired. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of low pump flow and positioning issues has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10168 | 1220648-2025-27614 | 21857256 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21857256&pc=OZD | 03/25/2025 | 05/14/2025 |  | 04/17/2025 |  |  | 03/25/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026629968 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the 61 year old patient was on impella 5.5 support and during support, the patient was being transferred when the pump came across the aortic valve. the pump was advanced. the patient was eventually bridged to transplant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of positioning issues has been completed. the cause of the positioning issue was not determined due to lack of clinical information and no product returned.
10169 | 1220648-2025-27621 | 21857442 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21857442&pc=OZD | 03/26/2025 | 04/17/2025 |  | 04/17/2025 |  |  | 03/26/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026620647 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and approximately two days after start of the support, within a few hours of starting heparin, the patient developed hematoma at insertion site with 1pt hemoglobin drop. heparin was paused and pressure applied to insertion site. the patient began having pain at insertion site. surgery consulted; however, it was noted there was no plan for washout or intervention at such time. the patient was reported to be in stable condition following the event and continues on impella mcs. | [5 empty] | The device remains implanted, supporting the patient at the time of the mdr writing, and therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: potential adverse events. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
10170 | 1220648-2025-27596 | 21857702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21857702&pc=OZD | 03/27/2025 | 06/09/2025 |  | 04/17/2025 |  |  | 03/27/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1162662 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) was turned on and a controller failure alarm occurred. there was no report of patient involvement. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | Data analysis: on the complaint date, march 27, 2025, immediately after the initial boot-up, the console displayed the ¿controller failure (purge pressure sensor defective) ¿ alarm,¿ event #418. the console was manually shut down and rebooted twice, each time, the console displayed event #418 (imc_logs_ic3243.pdf). this confirms the reported failure mode. device analysis: this console was tested after arriving at fs. although the ¿controller failure¿ event #418 was initially reproduced, it could not be replicated after power cycling the console. found the front panel optical connector contaminated, unable to remove the debris upon cleaning (unrelated to the reported failure mode and the clinical procedure). root cause: the root cause of the controller failure was not determined due to the inability to reproduce it consistently. it could be due to either a pud board malfunction or a purge pressure sensor malfunction. ad hoc task: before returning this console (ic3243) back to the customer, fs was instructed to perform the following. replace the purge pressure sensor (5500-0036) due to an intermittent controller failure alarm. replace pud pca (0042-1105) due to an intermittent controller failure alarm. replace the front panel optical connector (0042-2736) due to debris. · perform sections 10, 11, 12, and 23 of the pm procedure (0042-7309). · perform a full functional check (0042-7316).
10171 | 1220648-2025-27592 | 21857792 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21857792&pc=OZD | 04/09/2025 | 04/17/2025 |  | 04/17/2025 |  |  | 04/10/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025594215 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 44 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Ventricular Fibrillation | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient was on impella cp support due to heart failure. it was reported that while on support, the patient experienced torsades de point and expired. the cause of death is unknown, but it is believed to be septic shock, which resulted in ventricular fibrillation. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
10172 | 1220648-2025-27615 | 21858348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21858348&pc=OZD | 03/25/2025 | 05/23/2025 |  | 04/17/2025 |  |  | 03/25/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025597894 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 38 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Valvular Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support for eight days, the patient developed a chordae rupture resulting in mitral valve failure. the following day, the impella cp was removed, and a mitral valve repair was performed. the patient's outcome is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: impella cp with smart assist catheter insertion, section: axillary insertion of the impella cp with smart assist catheter, section: use of impella with transcatheter aortic valves: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: impella cp with smartassist catheter insertion: ¿impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve.¿ section: use of echocardiography for positioning of the impella catheter impella catheter inlet to the aorta: ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms: ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ | The investigation of heart valve damage has been completed. the pump was returned for investigation. the returned pump had a cut cannula in two separate pieces. the pump was visually inspected, and no abnormalities were reported on the returned product. there was some biomaterial present on the pigtail. the pump was run as per specifications in a bucket of water. the data log was not provided for this case run. according to the clinical report, several patient movements were reported during the time leading up to the chordae rupture and mitral valve damage, which could have caused the pump to be in an improper position during patient movement. the cause of the heart valve damage issue was not determined since no abnormalities were reported on returned product and the sufficient clinical information was not provided. d9 device returned to manufacturer date added.
10173 | 1220648-2025-27616 | 21858745 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21858745&pc=OZD | 03/25/2025 | 05/27/2025 |  | 04/17/2025 |  |  | 03/25/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026619606 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Fluid/Blood Leak | [6 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella cp support and during support, there was a purge leak from the sidearm at the luer-connector. a purge bypass was created and support was continued. it was further reported that based upon conversation with the family, care was withdrawn and the patient subsequently expired. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: cleaning: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of purge leak ¿ pump has been completed. the pump was returned for investigation. sidearm bypass was observed on the returned product. the severed sidearm was inspected, and a crack along the length of the yellow luer was noted. the sidearm was connected to the purge fluid via the purge cassette, and the leak was reproduced. data log was not provided for case run. the root cause of the purge leak - pump was determined to be due to a damaged sidearm yellow luer due to high tension in combination with disinfectant and certain drug ingredients such as oil, alcohol, lipids etc.
10174 | 1220648-2025-27625 | 21858755 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21858755&pc=OZD | 03/25/2025 | 05/14/2025 |  | 04/17/2025 |  |  | 03/25/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026625130 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on bipella support. on day one of impella cp support, the access site was bleeding around the repositioned sheath. lidocaine with epinephrine was injected. the reposition sheath was pushed back to the level of the skin and manual pressure was done. bleeding was not resolved so new sutures were applied and angle match dressing was done. two units of blood cells were administered. additionally, the impella cp measuring was shallow. the physician advanced the impella cp but the pigtail was in the papillary. the physician was not able to free the pigtail. no active alarms currently. the impella was still into the papillary structure but was free of mitral leaflets. no repositioning was done per the medical doctor. the pump was eventually weaned and explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and positioning issues has been completed. the product was not returned for review. based on clinical description, the root cause of access site bleeding was most likely use related since the repo-sheath was not fixed to the skin correctly. the root cause of the positioning issue was most likely patient condition related since repositioning was not considered due to small left ventricle (lv) based on the clinical review.
10175 | 1220648-2025-27595 | 21858768 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21858768&pc=PYX | 03/27/2025 | 05/23/2025 |  | 04/17/2025 |  |  | 03/27/2025 | 04/17/2025 |  |  | 08/31/2025 |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024392240 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella rp device for mechanical circulatory support. while on support, the motor current on the implanted impella rp pump suddenly spiked coinciding with a drop in flows. at that point in time, the patient's labs were noted to have hemolyzed. the flow issues continued into the following day at which point the flow rates finally began to normalize. despite the improvement, the patient's pathology results prompted the decision to withdraw care. the patient expired following removal of the pump. there are no notes that the patient's hemolysis had resolved prior to their passing. there is not enough evidence to disassociate the use of the impella with the patient's passing. | [5 empty] | The investigation of low pump flow and hemolysis has been completed. the data logs confirmed failure mode with motor current (mc) spiked followed low pump flow that triggered auto suction response and suction alarms. this event only occurred about 5 minutes on 3/27 and 15 minutes on 03/28 then back to normal range. no other issue was found on the logs. product was returned for review. visual inspection found biomaterial inside the of cage. no other issues were found on the rest of the pump. pump passed hemolysis test. the root cause of low flow was most likely due to biomaterial ingestion since the biomaterial was found on the pump, the data logs meet the pattern of low flow and hemolysis and the pump had no issue under inspection visually. based on the investigation findings the failure mode thrombus was added and as the necessary information was not provided, the root cause of the thrombus was not determined. d9 device returned to manufacturer date added. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10176 | 1220648-2025-27622 | 21858846 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21858846&pc=OZD | 03/25/2025 | 05/01/2025 |  | 04/17/2025 |  |  | 03/25/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025548466 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella 5.5 support and during support, the pump had a spot on the red plug where white outer sheath on driveline catheter pulled back exposing inner lines. the spot was secured with tape/tegaderm. also, the driveline securement devices were loose and the red plug was able to move back and forth. securements were reinforced. no issues or faults were reported for the pump function. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of product damage (sleeve damage)has been completed. the pump was not returned for investigation, and a data log review was not required for this case run. according to the clinical details, the sterile sleeve was separated and exposing the lumens and wires beneath. it was secured with tape. the cause of the sleeve damage issue was most likely use related.
10177 | 1220648-2025-27629 | 21859048 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21859048&pc=OZD | 03/25/2025 | 05/13/2025 |  | 04/17/2025 |  |  | 03/25/2025 | 04/17/2025 |  |  | 09/30/2025 |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024396432 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 insertion was aborted after several unsuccessful attempts to advance the impella 5.5 through vasculature. angiogram of left axillary revealed stenosis. the patient remained in stable condition. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, impella unable to pass through vasculature. | The impella device was discarded by the customer and therefore, ¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
10178 | 1220648-2025-27627 | 21859098 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21859098&pc=OZD | 03/25/2025 | 04/17/2025 |  | 04/17/2025 |  |  | 03/25/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026621984 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | Ischemia | [11 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that a patient was on impella cp support. during support it was reported that the right leg was mottled. hypotension continued the next day and the right leg was very mottled. additionally, it was noted that the patient's lactate continued to rise. fulminant organ failure, and acidosis inhibiting mechanism of action of medications. ultimately, care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer as the patient expired on support. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10179 | 1220648-2025-27619 | 21859131 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21859131&pc=OZD | 03/25/2025 | 04/17/2025 |  | 04/17/2025 |  |  | 03/25/2025 | 04/17/2025 |  |  | 01/31/2026 |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025524038 |  | 0048-0047 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Prime | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that during support from a impella cp on a 78 year old male patient, a purge set triggered multiple air-in-purge alarms. troubleshooting was performed, but the issue persisted. the staff opted to replace the cassette, but the case was completed prior to the new replacement. as such, a definitive cause for the problem could not be determined. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10180 | 1220648-2025-27626 | 21859205 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21859205&pc=OZD | 03/26/2025 | 05/21/2025 |  | 04/17/2025 |  |  | 03/26/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025594882 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with difficulty with an impella cp device for mechanical circulatory support; however, the device had to be immediately explanted. there were known aortic valve contraindications to impella cp implant; the patient had critical aortic stenosis and aortic insufficiency. however, the team decided to proceed after balloon aortic valvuloplasty (bav). after inflation, the impella cp device had difficulty crossing valve and had an unusual trajectory into left ventricle. the entire aorta was extremely tortuous and calcified. there was also severe dilation of aortic root and arch. the cannula developed a kink just below the outlet which would not resolve with manipulation. the impella cp pump was implanted/started, had suction and low flow. the decision was made to explant the impella cp device, do balloon aortic valvuloplasty with a larger balloon and attempt new impella cp device placement. the patient was reported to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of delivery issues, low pump flow, and product damage (cannula kink) has been completed. the root cause of the delivery issues was most likely patient condition related as the patient had critical aortic stenosis and severe ai which is contraindicated for impella cp use per instructions for use. the root cause of the product damage was most likely due to a cannula kink per clinical details provided. the root cause of the low pump flow was most likely patient condition related as low from was present right after support was initiated, a cannula kink was noted, and patient had diseased vasculature. instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system with automated impella controller section: warnings & cautions: warnings: section: pre-support evaluation section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ h10 instructions for use were inadvertently omitted on the initial manufacturer device report number 1220648-2025-27626.
10181 | 1220648-2025-27631 | 21859208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21859208&pc=OZD | 03/01/2025 | 04/17/2025 |  | 04/17/2025 |  |  | 03/25/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563082 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured hemolysis with a "probable" relationship to the impella device used: ¿the patient developed hemolysis after impella cp and va ecmo placement. ldh peak 3535 (b)(6) 2025. additionally, urinalysis on (b)(6) 2025 was positive for blood, red and white blood cells. the patient received several transfusions for drops in hemoglobin without overt bleeding. admission hemoglobin 14.4 g/dl, nadir 7.0 g/dl (b)(6) 2025.¿ the patient recovered with no sequelae. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
10182 | 1220648-2025-27623 | 21859227 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21859227&pc=OZD | 03/25/2025 | 05/14/2025 |  | 04/17/2025 |  |  | 03/25/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026629966 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported during support of an impella 5.5 on a patient for cardiomyopathy, proper positioning could not be obtained. despite attempts to reposition the device, the pump could not be directed away from the mitral valve. support was continued with the implanted pump despite the compromised positioning. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of positioning issues has been completed. product was not returned for review. the root cause of positioning issue was unable to be determined as limited clinical details were provided and no product was returned for review.
10183 | 1220648-2025-27632 | 21859340 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21859340&pc=OZD | 03/25/2025 | 05/21/2025 |  | 04/17/2025 |  |  | 03/25/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025518526 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a thrombus developed in the left ventricle during impella 5.5 support. the apical thrombus was found near the inlet of the impella 5.5. cardiology discussed options and the decision was ultimately made to explant the pump. the patient's status was updated to expired following explant. per medical safety officer review, the pump was prematurely explanted because of the thrombus which may or may not be related to the use of the impella device. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | Per medical safety officer review, the pump was prematurely explanted because of the thrombus which may or may not be related to the use of the impella device. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation of thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2025-27632. h1 type of report was erroneously selected on the initial manufacturer device report number 1220648-2025-27632.
10184 | 1220648-2025-27620 | 21859372 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21859372&pc=OZD | 03/07/2025 | 04/30/2025 |  | 04/17/2025 |  |  | 03/07/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026622752 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 50 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the complainant reported that the patient had access site bleeding and ventricular tachycardia. on arrival to the intenstive care unit, the impella groin site started briskly oozing so pressure was held by the nurse. the blue suture hub was advanced slightly, re-sutured by intensivist with forward tension sutures and then the bleeding stopped. additionally, a notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry indicting that the patient experienced recurring ventricular tachycardia with a "possible" relationship to the impella device used: ¿the patient experienced recurring ventricular tachycardia. one episode of ventricular tachycardia was refractory and required direct current cardioversion with 200 j with successful conversion to sinus rhythm with biventricular pacing, ventricular ectopy including runs of non-sustained ventricular tachycardia. underlying rhythm is sinus with biventricular pacing (pacemaker settings ddd)¿. the patient recovered with sequelae. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and vf/vt/af/at has been completed. the impella device was not returned for evaluation. access site bleeding- the root cause of the access site bleeding issue was most likely use related, re-sutured by intensivist with forward tension sutures and then the bleeding stopped. as the necessary information was not provided, the root cause of the vt/vf was not determined. g2 report source; study was inadvertently omitted on the initial manufacturer device report number.
10185 | 1220648-2025-27634 | 21859552 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21859552&pc=OZD | 03/25/2025 | 04/17/2025 |  | 04/17/2025 |  |  | 03/25/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025546186 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient on day one of impella 5.5 support was taken back to the operating room for pocket washout due to hematoma formation. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation has been completed. the impella device was not received from the customer. the root cause of access site bleeding could not be determined since the product was not returned and insufficient clinical details were provided.
10186 | 1220648-2025-27624 | 21859736 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21859736&pc=OZD | 03/26/2025 | 05/23/2025 |  | 04/17/2025 |  |  | 03/26/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025606076 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device via the right femoral artery for mechanical circulatory support and during support there was high purge pressure, low pump flow, and the patient developed hemolysis. activated clotting time was not therapeutic prior to the impella cp device insertion. the patient was transferred to the operating room to proceed with the coronary artery bypass graft (cabg) surgery. the physician pulled the impella cp pump back into descending aorta and left at support level p-1 for the cabg case. the physician proceeded with the saphenous vein graft (svg) to the left anterior descending artery and the svg to first diagonal. the impella was re-advanced and utilizing transesophageal echocardiogram back across aortic valve. the support was increased as bypass was weaned. purge pressure high/ blocked alarms started at this point. after weaning from bypass, the alarm persisted. the team discussed options extensively including escalation to impella 5.5, replacing the impella cp, or extracorporeal membrane oxygenation (ecmo). given the patient's support needs and anatomy considerations, the decision was made to leave the current impella in, transfer to cardiovascular intensive care unit (cvicu) and decide on just impella cp replacement with or without ecmo placement in the catheterization lab. on arrival to the cvicu, urine color was noted to be very dark (cherry coke). the patient was waiting for the catheterization lab for the impella replacement. impella cp device the motor current started to trend up and suction was unable to break even at support level p-1. there was concern the impella cp was failing or soon to fail. the impella cp device was eventually replaced and ecmo support added. the patient continued support for an additional five days before being successfully weaned and replacement impella cp device explanted with a survived patient outcome. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the analysis, a supplemental report will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. hemoylsis section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction : ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ high purge pressure: section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation of high purge pressure and hemolysis has been completed. the data logs were not available. hemolysis: clinical details state that hemolysis was suspected during the case due to dark urine. the impella was placed under tee guidance. it was suspected that the clot at the outlet area resulted in hemolysis. product was returned and evaluated. there was some biomaterial near the purge gap, but no other abnormalities. the pump passed hemolysis testing. the root cause of the hemolysis was not determined since data logs are not available and the hemolysis did not reproduce in the lab. high purge pressure: clinical details state that there were purge pressure high/blocked alarms after the pump was re-advanced post-cabg. the act was not therapeutic prior to insertion. the purge cassette was replaced with no improvements. no purge line kinks were observed. the returned product had some biomaterial in the purge gap, but it did not reproduce high purge pressure in the lab. there was no evidence of blood ingress in the purge line. the root cause of the high purge pressure was not determined since data logs are not available and the high purge pressure did not reproduce in the lab.
10187 | 1220648-2025-27618 | 21859941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21859941&pc=OZD | 03/25/2025 | 09/23/2025 |  | 04/17/2025 |  |  | 03/25/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025610869 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, it was noted that the patient had increased oozing and a hematoma from the right chest insertion and packed red blood cells (prbc's) were given. surgery team was notified. plan of care to address bleeding with sandbag, blood products and holding heparin infusion. additionally, impella on echocardiogram showed 4.3 and tilted towards the mitral valve, which caused suction when centravl venous pressure is <7. this also causes the left ventricle signal to be damped. placement signal was also approximately 15mmhg less than left radial art line. it was also noted that there were purge pressure low alarms, which they identified a leak within the purge cassette. the patient's outcome is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | Additional information provided in h6 and h10. access site bleed: product was not returned for investigation and data logs are not relevant to the issue. since limited clinical detail was provided regarding the complication and product was not returned for investigation, the root cause of the access site bleed could not be determined. low pump flow: data logs were reviewed and suction alarms and low pump flows were confirmed throughout the entire day. there were no positioning alarms to confirm malposition. there was an instance early in the day where the motor current modulation dropped, dampening the lvp, and was followed by a suction event which aligns with the report. product was not returned for investigation. since it's unclear from clinical details if the pump was repositioned or not, data logs showed intermittent low flows and suction alarms throughout the entire day, and product wasn't returned, the root cause of the low flows could not be determined. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. product was not returned for investigation. data log review confirmed the reported trends seen in the placement signal during a suction event, however, optical signal metrics were stable suggesting the sensor was performing to specification. for these reasons, the root cause of the optical signal issue could not be determined. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-27618. d10 concomitant medical products and therapy dates.
10188 | 1220648-2025-27612 | 21859999 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21859999&pc=OZD | 03/23/2025 | 05/06/2025 |  | 04/17/2025 |  |  | 03/23/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026626904 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was in cardiogenic shock from an acute myocardial infarction and was implanted with an impella cp device for mechanical circulatory support. during support, the patient experienced access site bleeding, low pump flow, and device malpositioning with repositioning. the patient would subsequently expire. the patient presented to the hospital from jail in a declining condition. the patient was noted to have stopped eating two days prior. the decision was made to implant an impella cp device in the setting of cardiogenic shock, which was successfully completed. the patient was prepared and transferred out to a high-level care facility the following day. once on the unit, the patient was bleeding at the impella access site and all other access sites. the patient was intubated and sedated and placed on a triple lumen central venous catheter and arterial line. arterial blood gas showed combined acidosis, but was improving. the patient was paced via transvenous pacemaker and resting on impella cp support at performance level p-2 to p-4. however, there was intermittent suction alarms. the performance level was unable to go above p-4 without suction. the patient was transfused with one unit of packed red blood cells (prbcs). the impella site still was noted with constant bleeding. pressure was held for 30 minutes with no change. angle was supported but approximately two inches of the repositioning sheath was exposed. at bedside, the physician re-sutured the site and repositioned the impella cp pump. the bleed was noted to have stopped at all sites. it was noted that the patient was not a candidate for extracorporeal membrane oxygenation. the patient remained in critical condition but stable after rewarming, repositioning and volume administration. the following day the patient remained in critical condition. overnight the patient experienced bleeding from the endotracheal tube. the volume was resuscitated now with two units of prbcs. anticoagulation was put on hold. the patient was too unstable for pan scan. a discussion with the family was planned for plan of care that may include palliative care. the following night the patient was taken to the operating room for a washout. sternal wires was the culprit of the chest bleeding. the patient received "lots" of blood products. the bleeding subsided, and it was noted the patient looked markedly better. the patient was noted to be hemodynamically stable. the plan was to wean the impella with hopes the patient wakes up for extubating. however, later in the evening, the patient would expire. and care was withdrawn. | [5 empty] | There is not enough information to exclude the impella cp device as an associated factor. the patient was very ill and the impact to the outcome as a result of the impella related event is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding, low pump flow, and positioning issues has been completed. the impella device was not returned for evaluation. access site bleeding - major: the cause of the access site bleeding - major most likely was patient condition related as the patient was bleeding at all access sites. low pump flow: complaint device not returned for analyzing. no data log available to review. not enough evidence provided to confirm cause of low flow. the cause of the low pump flow issue cannot be determined since complaint device and data logs were not returned for analysis and insufficient clinical data provided. positioning issues: the cause of the positioning issues most likely was patient condition related to the patient's small left ventricle.
10189 | 1220648-2025-27617 | 21860084 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21860084&pc=OZD | 04/10/2025 | 07/11/2025 |  | 04/17/2025 |  |  | 04/11/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025611222 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 39 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, the patient had severe hemolysis. the pump was explanted. there were no problems with the pump position or volume management. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of hemolysis has been completed. the pump was returned for investigation. the cannula and inflow cage had no biomaterial. no impeller or any physical damage observed. pump also passed hemolysis testing with mih = 7.71. data logs were returned and investigated. logs shows numerous ¿impella in aorta¿ alarms along with ¿impella position unknown¿ alarms starting early into the case. no placement signal trend, motor current spike or suction alarms observed. clinical states no issues with positioning or volume management. the root cause of the hemolysis issue was not determined due to lack of clinical information, returned product passed hemolysis testing, and data log investigation was inconclusive. d9 device returned to manufacturer date added | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
10190 | 1220648-2025-27633 | 21860216 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21860216&pc=OZD | 03/26/2025 | 04/17/2025 |  | 04/17/2025 |  |  | 03/26/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026624584 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. when the patient self-explanted the 5.5 and placed back on va ecmo. the patient was then implanted with a new impella 5.5. while on support, after 10 days of the 5.5 support a cva bleed was observed. the team removed heparin from the purge and used sodium bicarb. the pump remained on for 5 more days and then the care was withdrawn and organs harvested. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system & impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10191 | 1220648-2025-27636 | 21860357 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21860357&pc=OZD | 03/25/2025 | 05/14/2025 |  | 04/17/2025 |  |  | 03/25/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026634520 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient had an attempted impella cp implant and it was aborted due to patient anatomy and the long peel away sheath kinked. the impella cp was unable to track to the left ventricle due to anatomy. there was no patient harm. | [5 empty] | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions - ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings - ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system - section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves - ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings - ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues and introducer damage ¿ mechanical has been completed. the impella cp pump was returned, but introducer was not returned for investigation. delivery issues: device attempted implantation on (b)(6) 2025. during procedure long peel away sheath kinked and cp unable to track to lv. md noted unable to track due to anatomy. case noted indicate patient vessel condition was tortuous and location of resistance was axillary artery. case aborted. the pump was returned and inspected. 2 kinks were observed at the 25 and 32cm marks of the catheter. the cause of the delivery issue was patient condition related due to reported calcification and tortuosity of the patients vessels introducer damage - mechanical: device attempted implantation on (b)(6) 2025. during procedure long peel away sheath kinked and cp unable to track to lv. md noted unable to track due to anatomy. case noted indicate patient vessel condition was tortuous and location of resistance was axillary artery. case aborted. the introducer was not returned. the cause of the introducer damage - mechanical was most likely patient condition related due to reported calcification and tortuosity of the patients vessels.
10192 | 1220648-2025-27630 | 21860404 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21860404&pc=OZD | 03/26/2025 | 04/17/2025 |  | 04/17/2025 |  |  | 03/26/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025601128 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with postcardiotomy cardiogenic shock implanted with an impella cp device for mechanical circulatory support developed ischemia the following day after start of the support. the peel-away sheath was removed, but the patient's pulse failed to return in the lower extremity. the condition was monitored until explant two days later. pulse returned after explant. the patient was reported to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
10193 | 1220648-2025-27635 | 21860774 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21860774&pc=OZD | 05/27/2024 | 04/17/2025 |  | 04/17/2025 |  |  | 03/26/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025485778 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured severe hypotension with a "possible" relationship to the impella device used: ¿during the impella-supported high-risk pci, the patient developed hypotension. baseline systolic blood pressure was >90 mmhg, nadir sbp 81 mmhg. norepinephrine was added starting at 2 mcg/min with a peak of 3 mcg/min with good result.¿ the patient recovered with no sequelae. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10194 | 1220648-2025-27628 | 21861682 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21861682&pc=OZD | 03/26/2025 | 04/17/2025 |  | 04/17/2025 |  |  | 03/26/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025541973 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient undergoing a high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. post procedure, a groin complication occurred after removal of the impella cp device. the peel-away sheath was accessed with .035 wire to maintain access for perclose. the sheath was removed after contra-lateral access and femoral artery balloon inflation. perclose x2 deployed, second perclose failed and third perclose was deployed, protomine was given and there were no visible signs of bleeding at said moment. a moderate soft hematoma was noted after drape removed. manual pressure held while the patient was moved to holding area for recovery. after arriving to holding area, femstop was applied. it is unknown if blood products were administered. the status outcome of the patient following the event was indicated as unknown. | [5 empty] | The device was not returned for analysis. the investigation was completed, however, and noted the root cause of access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided. instructions for use for the related event are as follows: impella cp with smartassist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
10195 | 1220648-2025-27637 | 21862653 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21862653&pc=OZD | 03/24/2025 | 06/13/2025 |  | 04/17/2025 |  |  | 03/30/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025525454 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 18 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | Leak/Splash | Increase in Pressure | Material Split, Cut or Torn |  |  |  | Unspecified Infection | Tachycardia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that roughly twenty days into support, the patient had white ooze around their impella site. although it was believed to be potentially related to the adhesive from the dressing, the patient was started on prophylactic antibiotics. six days later, the sterile sleeve on the pump was torn. tegaderm was used to secure the sleeve. two days later, a purge leak was observed. a connection was found to be loose and the issue was fixed. thirteen days later, the patient had runs of ventricular tachycardia (vt) during support. the pumps positioning was found to be heavily reliant on the patient's position. in that regard, intermittent suction was encountered throughout the course of support. repositioning was performed and volume was given with varying results. the patient was reported to be in stable condition as support continued with the implanted pump. | [5 empty] | The impella pump was not returned for investigation. should new information be received, a supplemental manufacturer device report will be filed. | The investigation of product damage (sleeve damage), infection/sepsis, vf/vt/af/at, purge leak ¿ pump, high purge pressure, and low pump flow has been completed. the impella device was not returned for evaluation. purge leak - pump: as per the given clinical details, leaking from the purge was noted from loose connection, which has since been fixed. the cause of the purge leak issue was most likely user-related, as it was resolved after fixing the loose connection. high purge pressure: the data log shows a 3-hour gradual rising trend in purge pressure, followed by purge system block alarm. this trend resolved after 11 hours, with a gradual decrease in purge pressure. following that day, a slight rising trend in purge pressure was observed, with purge flow slightly decreasing; however, this normalized within days without any purge-related alarms. during the high purge pressure event, the doctor switched back to heparin in the purge solution, and tissue plasminogen activator (tpa) was also administered. the cause of the high purge pressure issue was most likely biomaterial, based on provided clinical details. low pump flow: several suction alarms were reported throughout the case run. pump flow was at the lower specification range and dropped low while running on steady high p levels (p7 and p8). no motor current spike, major placement signal trend, or positioning issue was confirmed from the data log review. according to clinical details, the patient was experiencing volume-related issues, and multiple blood products were given during the case run. additionally, there were observations noted by the doctor regarding position-based suction alarms. the cause of the low pump flow issue was most likely patient condition related. product damage (sterile sleeve damage): as per the given clinical details, sterile sleeve on the impella was torn and it was secured with tegaderm. the cause of the sterile sleeve damage issue was most likely use related. infection/sepsis: an infection/sepsis can occur during or after impella support. the root cause of the infection cannot be determined. vf/vt/af/at: as the necessary information was not provided, the root cause of the vt/vf was not determined. thrombus: thrombus failure mode was added as an investigated failure mode based on the high purge pressure investigation. as the necessary information was not provided, the root cause of the thrombus was not determined. d6b added instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ h6 medical device problem code 1491 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27637. h10 instructions for use were inadvertently omitted on the initial manufacturer device report number 1220648-2025-27637. | E4 should have been left blank on the initial manufacturer device report number 1220648-2025-27637 as it is unknown if the initial reporter also sent the report to the food and drug administration.
10196 | 1220648-2025-27638 | 21862654 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21862654&pc=OZD | 03/27/2025 | 05/15/2025 |  | 04/17/2025 |  |  | 03/27/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026624458 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient undergoing high-risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support and experienced low pump flow. the impella cp device was placed across the aortic valve and turned on to auto. however, flows would not go above 1.6l/m. the performance level was reduced to p-6 with suction alarms. the performance level was reduced further to p-4 and suction alarms did not resolve. support level then reduced to p-2 and volume was added; however, suction alarms were still present. consequently, the physician decided to explant the impella cp device and proceeded with the pci. the report noted there was no harm to the patient as a result of this event. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the analysis, a supplemental report will be filed. | The investigation of low pump flow has been completed. returned complaint cp pump visually inspected. no biomaterial observed. no cannula kink or broken blade found. dried dextrose washed for better investigation first with no observation. impella cp ran in the lab and flow was with specified limit at p-9. no low flow can be reproduced and no other issue observed on returned pump. data logs reviewed: no motor current (mc) spike observed. suctions occurred after a few minutes of support as mentioned in the clinical. a quick resolved impella in aorta alarm observed before explant. clinical states repositioning performed with no issue resolve. volume given which also did not resolve the issue. low pump flow: the cause of the low pump flow cannot be determined due to lack of clinical information, issue not reproduced in the lab, and no abnormality found on the returned pump. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27638 as it is unknown if the initial reporter also sent the report to the food and drug administration.
10197 | 1220648-2025-27639 | 21862655 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21862655&pc=OZD | 03/27/2025 | 04/17/2025 |  | 04/17/2025 |  |  | 03/27/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026627670 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Life Threatening | Death | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient (unknown race and ethnicity) with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and some time thereafter implant, same day, the patient had access site bleeding. heparin was withheld and two units of blood products were administered. the patient would subsequently expire. care was withdrawn. | [5 empty] | There are limited details and the impact from the impella-related event is unknown. although care was withdrawn, there is not enough information to exclude the impella cp device as an associated factor. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10198 | 1220648-2025-27640 | 21862658 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21862658&pc=OZD | 03/29/2025 | 05/19/2025 |  | 04/17/2025 |  |  | 03/29/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025561168 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was to be implanted with an impella 5.5 device for mechanical circulatory support. it was reported that impella 5.5 placement was aborted due to the patient vascular anatomy. the impella cp was placed instead. no patient harm was reported. | [5 empty] | The impella pump was not returned for investigation. should new information be received, a supplemental manufacturer device report will be filed. instructions for use as follows: & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on clinical details the root cause of delivery issue is determined to be patient condition because the pump was unable to pass through the due to calcification in the vessels and the insertion was aborted. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27640 as it is unknown if the initial reporter also sent the report to the food and drug administration.
10199 | 1220648-2025-27641 | 21862659 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21862659&pc=OZD | 03/29/2025 | 04/17/2025 |  | 04/17/2025 |  |  | 03/29/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025597602 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 92 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that that oozing was noted at the impella cp access site during patient support. a mild drop in hemoglobin levels was noted and two units of packed red blood cells were given. support continued uninterrupted. | [5 empty] | The impella pump was not returned for investigation. should new information be received, a supplemental manufacturer device report will be filed. instructions for use as follows: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10200 | 1220648-2025-27643 | 21862663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21862663&pc=OZD | 03/30/2025 | 05/23/2025 |  | 04/17/2025 |  |  | 03/30/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025583031 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Problem | [6 empty] | Hemolysis | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that hemolysis was suspected, and the inlet appeared to be close to the left ventricle wall. the doctor did not want to reposition and the patient remained on support through the duration of therapy. | [5 empty] | The impella pump was not returned for investigation. should new information be received, a supplemental manufacturer device report will be filed. instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿. impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿. section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions, including catheter position, pre-existing medical conditions, and small left ventricular volumes, may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿. | The investigation of hemolysis and positioning issues has been completed. the impella device was not returned for evaluation. hemolysis: clinical states that the pump was found to be against the mitral apparatus and the pump was chosen to not be repositioned. pump was explanted. pump was not returned for investigation. data logs show all slight decrease in pump flow at a steady p-level followed by suction alarm and decrease in p-level. data logs suggest positioning issue as a reason for hemolysis. therefore, the root cause of the hemolysis issue was most likely improper positioning of the device due to improper technique where the pump was facing mitral apparatus. positioning issues: clinical states that red tinged urine was observed and echocardiogram was done. positioning was observed to be against the mitral apparatus. pump was not returned. therefore, the root cause of the positioning issue was not determined since insufficient clinical information about the positioning issue was provided. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27643 as it is unknown if the initial reporter also sent the report to the food and drug administration.
10201 | 1220648-2025-27642 | 21862665 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21862665&pc=OZD | 03/29/2025 | 04/17/2025 |  | 04/17/2025 |  |  | 03/29/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636872 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Embolism/Embolus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that there was oozing and hematoma at the access site. hematoma to the right groin pressed out with only bruising remaining. the site was redressed and angle matched. additionally, vascular was consulted for intermittent common femoral artery obstruction on the right leg seen on computed topography scan. support continued uninterrupted. | [5 empty] | The impella pump was not returned for investigation. should new information be received, a supplemental manufacturer device report will be filed. instructions for use as follows: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
10202 | 1220648-2025-27644 | 21862666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21862666&pc=OZD | 03/30/2025 | 04/17/2025 |  | 04/17/2025 |  |  | 03/30/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025604066 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that after impella cp placement the right foot became cold, and loss of pulse noted. the heart team decided to remove impella. post impella cp removal, flow returned to lower leg. the foot was warm with distal pulses. | [5 empty] | The impella pump was not returned for investigation. should new information be received, a supplemental manufacturer device report will be filed. instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10203 | 1220648-2025-27645 | 21862668 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21862668&pc=OZD | 03/30/2025 | 05/23/2025 |  | 04/17/2025 |  |  | 03/30/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025609417 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | Positioning Problem | [5 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the pump was placed, and limb was seen to be ischemic, and an external bypass was placed. the patient was then transferred to the tertiary center and the pump was seen to be deep and showing signs of hemolysis. the team observed the blood urine and a pfhg was drawn. the pump was attempted to be repositioned. ultimately the decision was made to explant the impella cp and escalate to 5.5 pump. | [5 empty] | The impella pump was not returned for investigation. should new information be received, a supplemental manufacturer device report will be filed. instructions for use as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis and limb ischemia has been completed. the returned compliant device visually inspected. no biomaterial observed. no broken blade or sharp edge found. data log reviewed: impella position in ventricle alarms observed. no suction alarms seen during support. no motor current (mc) spike observed outside of p-level changes. no placement signal issue observed. clinical stated: patient received as transfer with bloody urine output. echo showed impella cp was deep at about 6cm measurement to av annulus. repositioning performed and cp was pulled back to 4.9cm. the cause of the hemolysis issue most likely was improper positioning related and impella required repositioning after transferring. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. the failure mode positioning issues was added the root cause of the positioning issue was most likely use related with pump retracting after patient transfer and was deep about 6cm to aortic annulus. d9 device returned to manufacturer date was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27645. h3 was erroneously selected on the initial manufacturer device report number 1220648-2025-27645. h6 type of investigation 4114 was erroneously entered on the initial manufacturer device report number 1220648-2025-27645. h10 narrative erroneously stated, " the pump was not returned for investigation" on the initial manufacturer device report number 1220648-2025-27645.
10204 | 1220648-2025-27646 | 21862669 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21862669&pc=OZD | 03/30/2025 | 09/23/2025 |  | 04/17/2025 |  |  | 03/30/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026629362 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that there was placement signal unreliable, and nursing staff educated on watching motor current and flow. it was also noted that the impella was repositioned, and bolus was given. the impella was successfully weaned and explanted. | [5 empty] | The impella pump was not returned for investigation. should new information be received, a supplemental manufacturer device report will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. there was no product returned. the data logs show that the placement signal (ps) goes out of range about 21 minutes after starting the pump. the signal-to-noise ratio (snr) drops to 0, light increases, lamp increases, contrast decreases and gain increases. the ps remains out for the duration of the case. due to lack of product returned, it cannot be confirmed the true root cause. the investigation has been completed for positioning issue. there were no clinical details regarding why the impella was repositioned. due to lack of sufficient information returned, the root cause of the positioning issues cannot be determined. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-27646. d10 concomitant medical products and therapy dates updated. e4 erroneously entered in the initial report.
10205 | 1220648-2025-27647 | 21862671 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21862671&pc=OZD | 03/30/2025 | 04/17/2025 |  | 04/17/2025 |  |  | 03/30/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026618327 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that a patient was implanted with impella cp pump for unknown indications. it was reported that the impella cp catheter kinked along the cannula causing suction alarms. the pump was explanted in procedural area. there was no patient harm reported. | [5 empty] | The impella pump was not returned for investigation. should new information be received, a supplemental manufacturer device report will be filed.
10206 | 1220648-2025-27648 | 21862817 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21862817&pc=OZD | 03/20/2025 | 05/05/2025 |  | 04/17/2025 |  |  | 03/20/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025601437 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemodynamic instability | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in cardiogenic shock from acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support and experienced low pump flow resulting in device replacement. however, the patient would subsequently expire. the patient came in septic and found to have endocarditis from liver disease. echocardiogram showed aortic insufficiency and possible ventricular septal defect. coronaries were noted as clean. ejection fraction was 25% and left ventricular end-diastolic pressure was 32. the decision was made to place an axillary impella cp device which was successfully placed. flows were 3.6 liters per minute at performance level p-9. the peel-away sheath was removed, and repositioning sheath was placed with device moved in ventricle. the device was repositioned, and the flows would not go above 1.6 liters per minute consequently, the device was removed, cleaned and ran in fluids with flows still not going above 1.6 liters per minute. consequently, the device was replaced with a new impella cp device. it was noted there was no clot observed on the impella cp device (pump 1) after removal. the patient would expire approximately two days later. the surgical team and family decided that aortic valve surgery would be "extremely" risky, and the decision was made to remove life-sustaining measures including impella cp support. | [5 empty] | Medical safety review of the event noted there was a delay in the needed mechanical circulatory support which may have added to the already poor prognosis situation. although the impella is not the cause of death, the impact to the outcome is unknown. there is not enough evidence to exclude th impella cp device pump 1 as an associated factor to the event outcome. the impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of low pump flow has been completed. the returned complaint device visually inspected. damaged impeller blades found. no biomaterial, no cannula kink observed. low pump flow: the cause of the low pump flow issue was damaged impeller blades due to interaction with other device.
10207 | 1220648-2025-27650 | 21862954 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21862954&pc=OZD | 03/26/2025 | 04/28/2025 |  | 04/17/2025 |  |  | 03/26/2025 | 04/17/2025 |  |  |  |  | 04/17/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025604099 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with difficulty an impella cp device pump 1 for mechanical circulatory support. however, the device had to be immediately explanted. there were known aortic valve contraindications to the impella cp implant. the patient had critical aortic stenosis and aortic insufficiency. however, the team decided to proceed after balloon aortic valvuloplasty (bav). after inflation, the impella cp device pump 1 had difficulty crossing the valve and had an unusual trajectory into the left ventricle. the entire aorta was extremely tortuous and calcified. there was also severe dilation of the aortic root and arch. the cannula developed a kink just below the outlet which would not resolve with manipulation. the impella cp pump was implanted/started, had suction and low flow. the decision was made to explant the impella cp device pump 1, do a balloon aortic valvuloplasty with a larger balloon and attempt to implant a new impella cp device. the balloon aortic valvuloplasty was performed with a larger balloon (22 millimeters). the new impella cp device pump 2 was prepped. the wire was able to cross the valve and exchanged to a new v-18. however, the impella cp device pump 2 was unable to cross the aortic valve with the cannula; shaft and wire buckling occurred in dilated aortic root/arch. attempted buddy catheter and several changes of trajectory were made without success. the patient had stabilized with the balloon aortic valvuloplasty. therefore, the decision was made to abort the impella cp device pump 2 implant. the patient was planned as the first case, urgent transcatheter aortic valve replacement for the following day. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention & impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this is one of two devices related to the patient's implant experience. refer to manufacturing report number 1220648-2025-27626 for the report on the impella cp device pump 1. | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be patient condition because of severely dilated, calcified and tortuous aortic root/arch and the pump was unable to be delivered.
10208 | 1220648-2025-27519 | 21847188 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21847188&pc=OZD | 03/19/2025 | 05/06/2025 |  | 04/16/2025 |  |  | 03/19/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026626848 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Delivered as Unsterile Product | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A physician reported that the packaging seal was already broken. therefore, the pump should be replaced. there was no patient involvement. | [5 empty] | The investigation of packaging issues- sterility has been completed. no product or photos were returned. the root cause of the packaging issue - sterility was not determined as no product or images were returned to confirm issue. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10209 | 1220648-2025-27520 | 21847271 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21847271&pc=OZD | 03/19/2025 | 09/22/2025 |  | 04/16/2025 |  |  | 03/19/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026615231 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  |  | White | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | Device Sensing Problem |  |  |  |  | Atrial Fibrillation | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | Embolism/Embolus | [9 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant also reported that the impella cardiac output was jumping around. the patient was on ecpella with a very apulsatile aortic placement signal. | The user facility reported a patient was found to have complex tachycardia and was intubated and placed on extracorporeal membrane oxygenation and impella cp. during support, it was noted that the patient had positive blood cultures and was started on antibiotics. the impella cp appeared deep so it was repositioned under echocardiogram guidance. additionally, the patient had an episode of atrial fibrillation and was treated with intravenous bolus. it was noted that the oozing from the access sites had slowed down after the stitch was placed. the patient survived the procedure. |  |  |  |  | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the 5s parameters on the returned product were within specification. the pump had no abnormalities and passed the light continuity test. the data logs confirm 'update cardiac output' alarms. the alarms triggered when pump position was changed and when the pump was disconnected. based on the trigger criteria for the alarms, it seems like the alarm triggers were met. the root cause of the optical signal issue was unable to be definitively determined as insufficient information related to timing of clinical events in relation to alarm triggers was provided. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-27520. d10 concomitant medical products and therapy dates updated. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: table 3.2 impella catheter components, ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of infection/sepsis, access site bleeding, vt/vf, positioning issues, thrombus and low pump flow. thrombus and low pump flow was added by the investigator. has been completed. infection/sepsis: an infection/sepsis can occur during or after impella support. when diagnosing the cause of an infection in a patient, the following clinical evidence would need to be considered holistically: -patient symptoms and medical history obtained by the treating clinician -physical examination findings -results of special investigations : laboratory test results & imaging studies -microbiological culture and sensitivity results -response to previous treatments and therapies -any relevant epidemiological information or exposure history -concomitant devices in use -preexisting patient condition such as underlying infection, remote septic sources such as indwelling lines/catheters or other bioprosthetic material in situ. -care of access site because this evidence has not been provided, the root cause of the infection cannot be determined ¿access site bleeding: there were no abnormalities on the returned product. clinical mention oozing in groin access sites overnight (ecmo and impella) which slowed down the following day. the root cause of the access site bleeding - minor was most likely patient condition as evidenced by the report of the bleeding from multiple access sites. vt/vf: cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease ¿timing of the arrythmia event in relation to impella support (during or post-implant) ¿electrocardiogram (ekg) recordings of the arrhythmia episodes ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities ¿response to any antiarrhythmic treatments or interventions during the event ¿any documented device-related issues or complications during impella support ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand ¿review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. positioning issues: there were no abnormalities on the returned product. clinical reported impella also appeared deep, so was repositioned under echo guidance. no additional information was provided on how pump became mispositioned. the root cause of the positioning issues was not determined as there were no abnormalities on the returned product and insufficient information related to positioning failure was provided. thrombus: the failure mode of ¿thrombus¿ is to be used when a thrombus is noted but cannot be connected to any other adverse events or decrease in pump performance. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions. ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency). ¿relevant laboratory test results, such as coagulation profiles and platelet counts, ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). ¿medications or treatments that the patient was receiving at the time of thrombus formation. ¿surgical or procedural details if applicable (e.g, cardiac catheterization) ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. low pump flow: there were no abnormalities on the returned product. the pump was run in the lab and the low pump flow was not reproduced. the data logs show the drop in flows occurred when the pump was in the aorta, an incorrect position. clinical confirmed pump was in incorrect position via echo. the root cause of the low pump flow was most likely due to improper positioning of the pump as clinical imaging and log analysis suggest improper positioning of the pump during low pump flow onset and low pump flow issue did not reproduce on the returned product. d6b explant date was erroneously entered on the initial manufacturer device report number 1220648-2025-27520.
10210 | 1220648-2025-27521 | 21847362 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21847362&pc=OZD | 03/19/2025 | 05/07/2025 |  | 04/16/2025 |  |  | 03/19/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025612242 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | Embolism/Embolus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient on impella 5.5 support developed hematuria. the impella level was dropped to p4 and one unit of packed red blood cells was given. the next day, hematuria was resolved and hemoglobin was stable. the patient survived the procedure. | [5 empty] | The investigation of hemolysis has been completed. complaint device not returned for analyzing. data log reviewed: no mc spike outside of p-level change observed. many unknown imeplla positioning alarms seen. a few impella in aorta alarms observed after insertion. suctions occurred at the time hemolysis issue reported. no placement signal issue seen. impella deep mentioned in the clinical and repositioned, blood given to the patient. clinical stated hematuria resolved but ldh rose again next day, unclear if hemolysis completely resolved after repositioning and blood given. hemolysis: the cause of the hemolysis issue cannot be determined since complaint device not returned for analysis and insufficient clinical data provided. failure mode: thrombus added since noted large clot burden in graft. as the necessary information was not provided, the root cause of the thrombus was not determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
10211 | 1220648-2025-27522 | 21847454 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21847454&pc=OZD | 03/18/2025 | 05/06/2025 |  | 04/16/2025 |  |  | 03/18/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026626072 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | Gastrointestinal Hemorrhage | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information was further reported that stated that while the patient had continued hemolysis, the patient was escalated to the impella 5.5. later on, the patient developed a massive gastrointestinal bleed which required multiple units of blood. however, due to the patient's condition, care was withdrawn and the patient expired. | The user facility reported a patient still had significant hemolysis due to impella pigtail being wrapped in mitral apparatus. plasma free hemoglobin remained elevated. the patient was escalated to an impella 5.5. |  |  |  |  | The investigation of hemolysis and positioning issues has been completed. the impella device was not returned for evaluation. should additional information become available regarding the gastrointestinal hemorrhage a supplemental report will be submitted hemolysis: clinical states that old cp was swapped out and the new cp has significant hemolysis due to the pigtail being wrapped in the mitral apparatus. no other information was provided. pump was not returned. partial data logs were provided. therefore, the root cause of the hemolysis issue was most likely improper positioning of the device as mentioned in the clinical information. positioning issue: clinical states that the imaging done for initial placement, confirmed that the pump was not in the correct position and was facing the mitral apparatus. repositioning was unsuccessful. pump was not returned for investigation. therefore, the root cause of the positioning issue was not determined since product was not returned and the reason for the improper positioning is unknown. additional information: b2: selected death and added date of death. b5-additional event details leading to patient death. h6-added clinical code 4476, impact code 1802, 4643. mdr manufacturer report #1220648-2025-27846 was filed on the impella 5.5 for the gastrointestinal bleed. mdr manufacturer report # 1220648-2025-27850 was filed on the impella rp for the gastrointestinal bleed. both the impella 5.5 and impella rp were in-situ at the time of gi bleed. medical safety officer review: the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. the patient's significant hemolysis while using the impella cp device (first implanted impella) and continued with the escalation to the impella 5.5, could have been one of multiple reasons that the massive bleed occurred days later when on bipella (impella 5.5 and rp flex). the patient's demise cascaded from the cp malfunction causing significant hemolysis and could not be dissociated from the outcome of death. there was additional intervention and the patient required an escalation because of higher therapy. the patient was in poor condition and subsequently met his/her demise due to significant gi bleed in which it cannot be said it is not associated with the impella device. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter, ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle, ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis, ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions, including catheter position, pre-existing medical conditions, and small left ventricular volumes, may also play a role in patient susceptibility to hemolysis.¿ section: suction, ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
10212 | 1220648-2025-27523 | 21847489 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21847489&pc=OZD | 03/18/2025 | 05/20/2025 |  | 04/16/2025 |  |  | 03/18/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025609905 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Death | Required Intervention |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a critical patient was implanted with an impella device for mechanical circulatory support. bleeding was suspected, but staff could not see any active bleed under fluoro, considering dissection. giving epi pushes, fluid, blood product and intermittently coding patient for low blood pressure. bleeding was from the right femoral artery (where transcatheter aortic valve replacement (tavr) sheath was previously). the 14fr peel away was removed and repo sheath was advanced. staff placed sutures and matched angle entry, significantly slowing the bleeding. the patient was stabilized and sent to icu. the patient was still having sudden drops in blood pressure, having to intermittently code, give blood products and fluid. right flank hematoma was noted in icu and it kept getting bigger as time went on. the patient was sent to get ct scan - showed retroperitoneal bleed. gen surg consulted to fix the bleed (he was willing to fix). the patients family decided to withdraw care. the patient subsequently expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation of access site bleeding and vascular damage: peripheral vessel has been completed. the impella device was not returned for evaluation. access site bleeding: based on the clinical details the root cause of access site bleeding is determined to be use issue because the bleeding was resolved by matching the angle and placing sutures. vascular damage: the bleed from right femoral artery (rfa) has been mentioned but no details related to the cause of damage. hence, the root cause of the vascular damage was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10213 | 1220648-2025-27524 | 21847580 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21847580&pc=OZD | 03/19/2025 | 04/21/2025 |  | 04/16/2025 |  |  | 03/19/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025594734 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device via the right femoral artery for mechanical circulatory support and experienced bleeding from the access site and limb ischemia requiring intervention. the following day after impella implant, the patient was taken to the operating room for cannulation and difficulty was encountered. the decision was made to open the chest and centrally cannulate. extracorporeal membrane oxygenation (ecmo) flows were increased to 3.9 and impella performance level was p-2 with appropriate flows. subsequently, it was noted the patient began bleeding from multiple sites due to coagulopathy. massive transfusion pack was given along with cell saver and crystalloid. the patient's chest was left open with central cannulation. right groin decannulated and repaired. later this same day, it was noted the patient was more stable. however, bleeding continued from the chest. additionally, the patient's right leg was stiff and had no pulse. vascular surgeon was consulted. the plan was to remove the impella and distal profusion catheter later in the day which was completed. the patient survived the explant and was noted to be in stable condition following the event. | [5 empty] | The investigation of access site bleeding and limb ischemia has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the access site bleeding issue was most likely patient condition related, the patient had bleeding from multiple sites due to coagulopathy. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10214 | 1220648-2025-27525 | 21847670 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21847670&pc=OZD | 03/19/2025 | 05/12/2025 |  | 04/16/2025 |  |  | 03/19/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025591473 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and developed oozing from the impella access site post implant same day. the existing perclose was tightened and the condition was resolved. the patient continued on support until being successfully weaned and the device was explanted with a survived patient outcome. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the access site bleeding issue was most likely patient movement, groin oozing after patient attempted to sit up.
10215 | 1220648-2025-27526 | 21847790 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21847790&pc=OZD | 03/19/2025 | 04/22/2025 |  | 04/16/2025 |  |  | 03/19/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026623546 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support while being worked up for transplant or left ventricular assist device (lvad). the patient developed a hematoma at the impella 5.5 access site during support. a washout was performed, and several clots were subsequently removed from the site. support was continued following the intervention. latest updated, as of day 31 of support, the patient was planned for the operating room for a lvad placement and impella 5.5 removal. | [5 empty] | The investigation of access site bleeding and thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
10216 | 1220648-2025-27527 | 21847831 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21847831&pc=OZD | 03/18/2025 | 04/16/2025 |  | 04/16/2025 |  |  | 03/18/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026618291 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Other | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that the patient experienced bleeding from the impella cp insertion site after arrival to icu despite angle matching and forward tension sutures. manual pressure was applied, surgiseal applied, and pressure dressing placed. the patient outcome was noted to be stable. | [5 empty] | The investigation into the reported issue has been completed. the pump was not returned for investigation; however, clinical information was sufficient to determine the root cause. the root cause of the reported issue was most likely patient movement since the bleeding started occurring after arrival to icu. as noted in the impella ifu: impella cp with smartassist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10217 | 1220648-2025-27528 | 21847899 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21847899&pc=OZD | 03/19/2025 | 05/08/2025 |  | 04/16/2025 |  |  | 03/19/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025484408 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and developed hemolysis following the implant. the staff worked to increase the patient's central venous pressure while decreasing their systemic vascular resistance in response to the condition. the hemolysis was noted to have resolved overnight and support was continued uninterrupted. at the end of support, approximately five days later, the patient was successfully weaned, and the device was explanted with a survived patient outcome. | [5 empty] | The investigation of hemolysis has been completed. product was not returned for review. based on clinical description and data review, the root cause of hemolysis was most likely patient condition related since the patient had very low central venous pressure (cvp) when hemolysis occurred & resolved after fluid compensation. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
10218 | 1220648-2025-27529 | 21847916 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21847916&pc=OZD | 03/18/2025 | 04/25/2025 |  | 04/16/2025 |  |  | 03/18/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026629216 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  |  | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Material Split, Cut or Torn | [5 empty] | Arrhythmia | Hematoma | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Additional information was received. it was reported that the patient is had incision and drain of axillary pocket due to serous drainage and swelling of right upper extremity. the physician suspects hematoma causing swelling and interstitial tissue drainage. the old hematoma was evacuated and sent for culture and sensitivity, no infectious concerns, new drain was placed in addition to receiving 1 unit on packed red blood cells. also reported that the patient did require 3 boluses to equal 1 liter of intravenous fluid for maps 50s. patient was noted to have a lot of ectopy. additionally, it was unsure if small tear to anticontamination sleeve close to distal lock, appears to have tegaderm wrapping around and dried betadine. | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced significant bleeding from the axillary site, pressure dressing and sandbag were applied. anticoagulation was withheld. hemoglobin levels dropped, so the patient was administered 1 unit of packed red blood cells. the bleeding resolved with an insignificant hematoma. the pump was increased to p6; however, suction events occurred. albumin was administered and pump was reduced to p5. the patient was noted to be stable. |  |  |  |  | B5 revised to reflect additional information received, which was inadvertently omitted from the initial manufacturer device report 1220648-2025-27529. h6 revised revised to reflect additional information received, which was inadvertently omitted from the initial manufacturer device report 1220648-2025-27529. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10219 | 1220648-2025-27530 | 21848047 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21848047&pc=OZD | 03/18/2025 | 09/30/2025 |  | 04/16/2025 |  |  | 03/19/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025580179 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Tachycardia | [12 empty] | Required Intervention | Death |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant also reported that the implanted impella 5.5 pump experienced intermittent placement signal issues during patient support. despite positioning being correct, it was noted that the provided readings were not accurate. the issues resolved shortly after each irregularity. | Additional information received that eighteen days after initiating of support, care was withdrawn and the patient's status was updated to expired. there is not enough evidence to disassociate the use of the impella with the patient's outcome. | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support and approximately four days after start of the support, the patient experienced runs of ventricular tachycardia. the pump was repositioned in response to the condition. support was able to continue without interruption. the impella 5.5 device was successfully weaned and explanted with a survived patient outcome. |  |  |  | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue and cardiac arrythmia. there was no product returned. the data logs show tha there was an increase in the placement signal (ps). the 5s parameters are relatively stable, but the signal-to-noise ratio (snr) slightly trends down. without sufficient evidence, and the product not being returned, the root cause of this optical signal issue cannot be determined. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include placement signal issue description. h6 updated to include placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-27530. d10 concomitant medical products and therapy dates updated. | B2 outcome death and date of death was inadvertently omitted on the initial manufacturer device report number. 1220648-2025-27530. b5 patient outcome of death was inadvertently omitted on the initial manufacturer device report number. 1220648-2025-27530. g3 date received by manufacture was 4/6/2025. h1 type of report was erroneously entered on the initial manufacturer device report number. 1220648-2025-27530. h6 health effect - impact code 1802 was inadvertently omitted on the initial manufacturer device report number. 1220648-2025-27530. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10220 | 1220648-2025-27531 | 21848153 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21848153&pc=OZD | 03/14/2025 | 05/21/2025 |  | 04/16/2025 |  |  | 03/21/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2025604577 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Patient-Device Incompatibility | Adverse Event Without Identified Device or Use Problem | [5 empty] | Perforation of Vessels | Cardiac Tamponade | Cardiac Perforation | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that the patient had an attempted impella cp implant and after the sheath and guidewire were inserted, the patient had a cardiac tamponade with a left ventricle (lv) rupture. the implant was aborted. the impella package was opened, but due to lv rupture impella was not inserted. instead, iabp+ecmo was inserted and a surgical procedure was performed. it is unknown if the guidewire or attempted implant contributed to the lv rupture. the patient expired. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of delivery issues and ventricle perforation has been completed. the impella device was not returned for evaluation. ventricle perforation: during insertion of 16fr insertion sheath, a non abiomed product, and operation of the guidewire, blood pressure decreased, and cardiac tamponade and left ventricle (lv) rupture (free wall rupture) were confirmed by angiography. the pump was not returned. it is unclear whether the 16 fr sheath, guidewire, or pump contributed to the perforation and how. the cause of the ventricle perforation was not determined since product was not returned and insufficient clinical details. delivery issues: the impella package was opened, but due to cardiac tamponade and lv rupture, impella was not inserted. the cause of the delivery issue was most likely use related as per the clinical details, the pump was not implanted due to cardiac tamponade and lv perforation. instructions for use: impella cp with smartassist system for use during cardiogenic shock: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist system for use during cardiogenic shock, section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ b1 product problem was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27531. b2 date of death was erroneously entered on the initial manufacturer device report number 1220648-2025-27531. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-27531. d2 common device name was erroneously entered on the initial manufacturer device report number 1220648-2025-27531. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27531 as it is unknown if the initial reporter also sent the report to the food and drug administration. h6 health effect - clinical code 2135 was erroneously entered on the initial manufacturer device report number 1220648-2025-27531. h6 medical device problem code 2682 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27531 h10 instructions for use were erroneously entered on the initial manufacturer device report number 1220648-2025-27531. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
10221 | 1220648-2025-27532 | 21848267 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21848267&pc=OZD | 04/08/2025 | 07/16/2025 |  | 04/16/2025 |  |  | 04/09/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025574574 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Defective Device | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the 0.018 guidewire required for inserting the impella had an initial defect. the floppy part at the tip of the guidewire was kinked before use. a new guidewire was used. there was no patient harm. | [5 empty] | The investigation of product damage (guidewire damage - peripheral vessels) has been completed. the guidewire was returned for investigation. data log review was not required for this case run. the floppy end of the guidewire was bent, and a kink was reported on it. no images or evidence were provided to confirm an out-of-box packaging issue. the cause of the out of box guidewire damage is unknown and the product will be returned to the vendor for further analysis. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10222 | 1220648-2025-27533 | 21848335 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21848335&pc=OZD | 03/19/2025 | 04/16/2025 |  | 04/16/2025 |  |  | 03/19/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025608565 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | A |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in postcardiotomy cardiogenic shock was implanted with an impella 5.5 device for mechanical circulatory support and after seven days of support, the device was successfully explanted. the patient was speaking clearly and some minutes later the patient began to have trouble speaking, trouble word finding, and was aphasic. stat head computed tomography revealed an acute left m2 and m3 occlusions of the middle cerebral artery. neurology was consulted, and heparin was initiated. leading up to this, the patient had been subtherapeutic on activated clotting time (act), and the physician had been encouraged by the hospital staff as well as abiomed staff to initiate anticoagulation to which the physician refused. the patient was very coagulopathic post coronary artery bypass graft surgery and had to go back to the operating room for a redo sternotomy and clot evacuation. therefore, the physician chose to withhold anticoagulation post operative for this reason although the physician was repeatedly informed of the goal act of 160-180. after the occlusion was identified, the physician acknowledged and attributed the reluctance to start heparin. the patient survived the event and was noted to be in stable condition. | [5 empty] | The investigation has been completed. as this clinical information was not provided, the true root cause of the stroke/cerebrovascular accident could not be determined. as the necessary information was not provided, the root cause of the thrombus was not determined. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
10223 | 1220648-2025-27534 | 21848492 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21848492&pc=OZD | 03/19/2025 | 04/21/2025 |  | 04/16/2025 |  |  | 03/19/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636747 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and following the implant, an access site hematoma developed. angiogram showed no extravasation. source of the hematoma arterial versus venous was unknown. the decision made to explant the impella cp and reimplant it on the opposite, left, side. the right femoral artery was closed using two perclose along with balloon tamponade, and hemostasis was achieved. impella cp device was then reimplanted successfully via left femoral artery for continued mcs. the patient was reported to be stable following the event. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10224 | 1220648-2025-27535 | 21848505 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21848505&pc=OZD | 03/23/2025 | 04/16/2025 |  | 04/16/2025 |  |  | 03/23/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636781 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the 62-year-old male patient on impella cp support had no dopplerable pulses to impella leg. additionally, the physician repositioned the cp as the patient was having brief episodes of ventricular tachycardia (vt) which have since been resolved. the cp was exchanged for impella 5.5 due to limb ischemia. the patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10225 | 1220648-2025-27536 | 21848512 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21848512&pc=OZD | 04/02/2025 | 07/15/2025 |  | 04/16/2025 |  |  | 04/04/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1838368 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | Unintended Application Program Shut Down | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) was being serviced and an abnormal noise occurred. the aic could not be powered off, so a forced shutdown was performed. after restarting, controller error alarm occurred, so it was decided to refrain from clinical use. there was no patient involvement. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic): controller error (yellow alarm) has been completed. based on that data and device analysis there were no issues were found with the aic for the reported failure mode of "aic-controller error". e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27536 as it is unknown if the initial reporter also sent the report to the food and drug administration.
10226 | 1220648-2025-27537 | 21848706 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21848706&pc=OZD | 03/21/2025 | 05/28/2025 |  | 04/16/2025 |  |  | 03/22/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025594796 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and post implant, same day, the patient had bleeding problems at the groin and a hematoma developed. blood products were administered as a result of the event. support was continued uninterrupted until the end of the support approximately three days later. the patient was successfully weaned and the device was explanted with a survived patient outcome. | [5 empty] | The investigation of access site bleeding has been completed. impella cp was returned and no abnormalities were found. the root cause of the access site bleeding was not determined due to insufficient clinical details were provided. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10227 | 1220648-2025-27538 | 21848738 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21848738&pc=OZD | 03/23/2025 | 04/16/2025 |  | 04/16/2025 |  |  | 03/23/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636245 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the 72-year-old male patient on impella cp support developed limb ischemia. absent pulses distal to cp, but present via doppler on contralateral side. it was further reported that care was withdrew and the patient expired. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10228 | 1220648-2025-27539 | 21848793 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21848793&pc=OZD | 03/22/2025 | 05/12/2025 |  | 04/16/2025 |  |  | 03/22/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1738177 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient newly diagnosed with ischemic heart failure with 10% ejection fraction and multivessel disease was planned for coronary artery bypass graft surgery and implanted with an impella 5.5 device for mechanical circulatory support (mcs). approximately one day after start of the impella mcs, the automated impella controller (aic) shut down and then turned back on. the nurse indicated "yes" to resume and noted that "everything looks okay right now." there were no alarms triggered. it was recommended that te aic be exchanged in case of reoccurrence. however, the patient continued support with the aic and the staff monitored the situation. the outcome as a result of the event was indicated as unknown. however, nine days later it was noted support concluded. the patient was bridged to a left ventricular assist device and the impella 5.5 device was successfully explanted. | [5 empty] | The automated impella controller device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - a/c power issues has been completed. the console performed as expected until around 18:36:11 at which point the real time (rt) log prematurely ends and the imc logs show the console rebooted. after the reboot, the console asked the user to resume pump at previous p-level by design at the same time as the unexpected shutdown alarm triggered. support for the patient was resumed and continued through the end of the case on (b)(6). device analysis: the console was returned for testing but the failure mode was not reproduced as the console was able to run a pump uninterrupted and all voltages checked were as expected. root cause: the cause of the unexpected reboot is not determined as it was unable to be reproduced.
10229 | 1220648-2025-27540 | 21848931 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21848931&pc=OZD | 03/22/2025 | 05/13/2025 |  | 04/16/2025 |  |  | 03/22/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025564826 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and high purge pressure and purge system blocked alarms were encountered. tissue plasminogen activator protocol was initiated and infused overnight due to persistent low purge flow/high purge pressure. the following morning, the flows were noted to have "greatly" improved. support was continued uninterrupted for an additional five days as the patient was successfully weaned and the device was explanted with a survived patient outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of high purge pressure has been completed. the cause of the high purge pressure was not determined since product was not returned, data logs were not recovered, and insufficient clinical details.
10230 | 1220648-2025-27541 | 21848944 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21848944&pc=OZD | 03/24/2025 | 04/16/2025 |  | 04/16/2025 |  |  | 03/24/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636261 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, it was difficult finding pulses on all four extremities. there was a concern for distal limb flow and the pump was explanted. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
10231 | 1220648-2025-27542 | 21849133 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21849133&pc=OZD | 03/21/2025 | 04/16/2025 |  | 04/16/2025 |  |  | 03/21/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026622866 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on impella cp support. an echocardiogram was performed and impella cp was noted to be flipped and near the papillary bed. several attempts were made to reposition but was not successful. the patient was brought to operting room. when the patient was turned to plan rolling board to transfer the patient, impella flow dropped with suction. several attempts were made to resolve suction but was unsuccessful. the decision was made to replace the existing impella cp with a new impella cp to restore some flow to the patient. the patients outcome is critical. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10232 | 1220648-2025-27543 | 21849277 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21849277&pc=OZD | 03/24/2025 | 06/05/2025 |  | 04/16/2025 |  |  | 03/24/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026617288 |  | 0550-0002 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 46 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, shortly after insertion (with an optimal position confirmed both by ultrasound and angiography), it started signaling position alarms (impella position in the aorta) and purge system alarms (high purge pressure). the staff tried to de-air the purge cassette to ensure glucose was flowing out. since the alarm persisted, it was decided to replace both the purge bag and the purge cassette. the console started showing the warning ¿purge fluid not detected¿ and the pump was replaced. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of high purge pressure has been completed on 5/23/2025. the pump was not returned. data log review showed high purge pressure. mc spike occurred without a corresponding p-level change that also caused the purge pressure to trend up and the purge flow to trend down. the cause of the high purge pressure was most likely biomaterial ingestion due to a motor current (mc) spike outside of a p-level change in conjunction with the pump flow indicating temporary saturated flows and rising purge pressures. d9 device returned to manufacturer date on 5/30/2025, should any new information become available after investigating the device a supplemental report will be submitted.
10233 | 1220648-2025-27544 | 21849411 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21849411&pc=OZD | 03/21/2025 | 04/16/2025 |  | 04/16/2025 |  |  | 03/21/2025 | 04/16/2025 |  |  | 07/31/2025 |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024370575 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a 64-year-old male patient on impella 5.5 and venoarterial extracorporeal membrane oxygenation for support during high-risk surgery. after the chest was closed and the site was secured, the patient had to be taken back to the operating room due to bleeding in the chest cavity. ultimately the patient expired on support. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10234 | 1220648-2025-27545 | 21849416 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21849416&pc=OZD | 03/22/2025 | 05/12/2025 |  | 04/16/2025 |  |  | 03/22/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025551869 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that following impella cp placement the patient abruptly sat up and the pump came back across the valve. attempts to wirelessly reposition the device were unsuccessful and the decision was made to replace the pump. there was no patient harm. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. positioning issues: the cause of the positioning issues most likely was due to patient movement since the issue occurred while the cath lab staff moved patient over from cath table to bed and patient abruptly sat up/bent legs and pump began alarming. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10235 | 1220648-2025-27546 | 21849435 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21849435&pc=OZD | 03/21/2025 | 04/30/2025 |  | 04/16/2025 |  |  | 03/21/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636789 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that following implant on a 62 year old female patient, an impella cp pump experienced persistent suction due to a kinked cannula. attempts to unkink the cannula were unsuccessful, however the decision was made to proceed with the case. as support continued, the patient lost pedal pulses in their impella leg. the pump was subsequently explanted in response to the condition. the patient survived the procedure. | [5 empty] | The investigation of limb ischemia, low pump flow, and product damage (cannula kink) has been completed. the pump was not returned for investigation. the data log shows that the low pump flow remains low from the start to the end of the case, accompanied by an active suction alarm. according to the clinical report, a cannula kink was noted on the implanted pump, and the doctor was not able to unkink it. the cause of the low pump flow issue was most likely kinked cannula. the cause of the cannula kink issue was not determined since sufficient clinical information was not provided. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. h6 component code 3038 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27546. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10236 | 1220648-2025-27547 | 21849466 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21849466&pc=OZD | 03/21/2025 | 05/08/2025 |  | 04/16/2025 |  |  | 03/21/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026615094 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 82 year old female that was implanted with impella cp pump experienced suction events after it was powered on for patient support. repositioning was attempted, but the issue persisted. the pump was removed and then re-inserted, but flow rates continued to remain low. the staff opted to proceed with the scheduled the cardiac procedure with the restricted flow rate. following the procedure, the pump was quickly weaned and explanted. upon inspection of the device, a kink on the cannula was discovered near the motor housing. the patient survived the procedure. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The investigation of low pump flow and product damage (cannula kink) has been completed. returned complaint cp pump visually inspected. cannula kink observed. no biomaterial or broken blade seen. cannula soaked in water to remove the dried dextrose for better inspection. physician tried to manipulate pump position and ultimately pulled it out of the left ventricle (lv) so impella was pulled out of the body and after re-insert flow still was low and noted the cannula kink near motor housing. low pump flow: the cause of the low pump flow issue was cannula kink due to improper technique. product damage (cannula kink): the cause of the product damage (cannula kink) was cannula kink due to improper technique.
10237 | 1220648-2025-27548 | 21849473 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21849473&pc=OZD | 03/24/2025 | 05/27/2025 |  | 04/16/2025 |  |  | 03/24/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1384051 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 46 YR |  | Male |  |  |  | Malfunction | No | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, shortly after insertion (with an optimal position confirmed both by ultrasound and angiography), it started signaling position alarms (impella position in the aorta) and purge system alarms (high purge pressure). the staff tried to de-air the purge cassette to ensure glucose was flowing out. since the alarm persisted, it was decided to replace both the purge bag and the purge cassette. a new purge cassette was used with the replacement pump and immediately the console repeated the warning ¿purge fluid not detected.¿ staff then replaced the automated impella controller (aic) to resolve the issue. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - purge issue ¿ other has been completed. based on the data and device analysis the root cause for the purge related issue was most likely due to the glucose residue. g1 reporting contact address and country was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27548. g1 manufacturing site name, address, country and fax number were erroneously entered on the initial manufacturer device report number 1220648-2025-27548.
10238 | 1220648-2025-27549 | 21849484 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21849484&pc=OZD | 03/21/2025 | 04/19/2025 |  | 04/16/2025 |  |  | 03/21/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026617488 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella cp support and during support, there was bleeding from the impella access site. during assessment, it was noted that a forward tension suture was quite loose. the doctor put in another forward tension suture. the impella site was redressed and still had slight ooze. a femstop was placed above the impella site and hemostasis was obtained. four units of blood products was given to the patient. the patient survived the procedure. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of access site bleeding was most likely use issue since hemostasis was achieved by forward tension suture tightening. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10239 | 1220648-2025-27550 | 21849615 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21849615&pc=OZD | 03/21/2025 | 12/11/2025 |  | 04/16/2025 |  |  | 03/21/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1683914 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | Hemodynamic instability | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported the automated impella controller (aic) had a yellow alarm displaying 50% battery charge after the patient took a lap. the console was then plugged in and showed 100% in about 15 seconds. the aic was replaced. there was no patient harm. | [5 empty] | Corrected information provided in b1, b2, h1 and h6. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - battery charging/other issues has been completed. data analysis: the console logs from the reported day of the event align with the complaint. on (b)(6) 2025 the battery level low (#23) alarm was triggered at 11:54:16. prior to that, the ac power was disconnected at 11:47:10 i.e. before 7mins of the battery low alarm. the battery level low (#24) alarm self-resolved one minute, at 11:55:16. battery data embedded in the log file indicated a single sample of the battery remaining capacity parameter for battery 1, which had a value of 0, leading to the activation of the battery level low alarm. there was no evidence to suggest that communication with battery 1 was interrupted. device analysis: the console was returned for further investigation. the reported complaint could not be reproduced, even after the console was power cycled and operated with the pump and purge cassette for a couple of hours. communication between the pbm and the batteries was monitored, revealing no issues. additionally, a visual inspection of the internal circuitry showed no abnormalities. root cause: the root cause of the ¿battery level low (#23)" was not determined as was not able to reproduced.
10240 | 1220648-2025-27551 | 21849642 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21849642&pc=OZD | 03/21/2025 | 05/12/2025 |  | 04/16/2025 |  |  | 03/21/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025579705 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  |  | White | Injury | Yes | Yes | Restricted Flow rate | Increase in Pressure | [5 empty] | Embolism/Embolus | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an implanted impella 5.5 pump developed high purge pressure during patient support. heparin was added to the purge and the pressure began to trend downwards. support continued with the implanted pump following the event. the patient was bridged to left ventricular assist device and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smartassist states the following: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The product was returned and evaluated. there was biomaterial wrapped around the purge gap and on the impeller. no kinks were found. high purge pressure did not reproduce and there was no blood ingress in the purge line. the investigation of high-pressure purge, pump stop, and thrombus has been completed. data log review shows the purge pressure rising from 500 to greater than 1000 mmhg within 3 minutes after a mc spike, speed deviation, and shift in pump flow. this occurred about 20 hours into the case and triggered high purge pressure alarms. the purge pressure remained high for 20 minutes before gradually recovering over the next two hours. heparin was added to the purge concentration while the purge pressure recovered. upon further data log review, a high motor current pump stop was found at the end of the case ((b)(6) 2025) with multiple motor current spikes and a slight increase in purge pressure prior to the pump stop, which continued increasing after the pump stop. this event was not mentioned in clinical details. the failure modes thrombus and pump stop were added to account for this event. high purge pressure: clinical details reported high purge pressure alarms on (b)(6) 2025. the root cause of the high purge pressure was most likely biomaterial ingestion as data log review showed a motor current spike, speed deviation, and shift in pump flow prior to the rise in purge pressure. pump stop: a pump stop with "impella stopped - motor current high" alarm was found at the end of the data logs, which was not reported in clinical details. there was biomaterial found on the returned product, which is most likely related to the pump stop event rather than the high purge pressure event on (b)(6) 2025 that resolved. the root cause of the pump stop was most likely biomaterial ingestion since the returned product had biomaterial wrapped around the purge gap and on impeller and data logs show multiple motor current spikes prior to the pump stop. thrombus: as the necessary information was not provided, the root cause of the thrombus was not determined. instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.". in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ b1 revised to reflect both adverse event and product problem based on investigation result findings. d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2025-27551.
10241 | 1220648-2025-27552 | 21849743 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21849743&pc=OZD | 03/21/2025 | 05/07/2025 |  | 04/16/2025 |  |  | 03/21/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636786 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | Material Twisted/Bent |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the 62-year-old female patient had an attempted impella cp implant and during the implant, the cannula kinked near the aortic valve. multiple attempts to advance, pull back, torque, and reposition but none successful in removing the kink. the pump was removed and replaced. the patient survived the procedure. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk pci states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. delivery issue - reported cannula kink occurred near the aortic valve during delivery. pump was removed after unsuccessful attempts to reposition. replacement pump had the same issue. suspected delivery issues due to patient anatomy and patient had aortic calcification. the root cause of the delivery issue was most likely patient condition related since clinical reported the patient had aortic calcification and the replacement pump had the same delivery issues. cannula kink - reported cannula kink occurred near the aortic valve during delivery. kink was found on the cannula near the motor housing on the returned pump. the root cause of the product damage was most likely a kinked cannula due to patient condition since the patient had aortic calcification.
10242 | 1220648-2025-27553 | 21849790 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21849790&pc=OZD | 03/21/2025 | 04/16/2025 |  | 04/16/2025 |  |  | 03/21/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025561453 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a 47-year-old female patient was on impella 5.5 support and during support, the patient had arrhythmias. impella¿s positioning was checked twice and appeared to be in the correct position. patient had consistent ventricular fibrillation for hours despite interventions. care was withdrawn and the patient expired. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10243 | 1220648-2025-27554 | 21849801 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21849801&pc=OZD | 03/21/2025 | 04/16/2025 |  | 04/16/2025 |  |  | 03/21/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025608598 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a groin bleed developed during impella cp support for a patient. manual pressure was initially held, but once lifted, the bleed returned instantly. best practices were used to troubleshoot the condition but the bleed persisted. the decision was ultimately made to explant the pump. following removal, the perclose failed prompting application of manual pressure and a femostop. upon additional analysis, the bleed was found in the left femoral circumflex artery, for which a stent was placed and three units of blood were given. the patients outcome is stable. | [5 empty] | The investigation for the vascular damage has been completed. groin bleeding around repo sheath site was noted. manual pressure was held and when lifted, the bleed was instantly. bleeding did not improve with angle matching and following best practices. 8.0 covered stent was placed and bleeding was stopped. the root cause of vascular damage peripheral vessel could not be determined since the product was not returned and insufficient clinical details were provided.
10244 | 1220648-2025-27555 | 21849879 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21849879&pc=OZD | 03/21/2025 | 05/19/2025 |  | 04/16/2025 |  |  | 03/21/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026612720 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the 63 year old male patient had an attempted impella cp implant and the catheter came back across the valve. the repo sheath was in and the doctor elected to advance wirelessly. the cannula kinked and the doctor elected to replace pump. the patient survived the procedure. | [5 empty] | The investigation of positioning issues has been completed. the complaint device not returned for analyzing. clinical stated: while exchanging the sheath, the catheter came back across the valve. the repo sheath was in and physician elected to advance wirelessly. the cannula kinked and physician elected to replace pump. positioning issues: the root cause of the positioning issue most likely was wireless re-access due to improper technique. the impella device is not indicated for wireless re-access per instructions for use. failure mode: product damage- cannula kink added since it occurred during reported positioning issue. product damaged (cannula kink): the cause of the product damaged (cannula kink) was cannula kink due to improper technique. the impella device is not indicated for wireless re-access per instructions for use. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
10245 | 1220648-2025-27556 | 21850074 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21850074&pc=OZD | 03/22/2025 | 05/02/2025 |  | 04/16/2025 |  |  | 03/22/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026624535 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella 5.5 support and during support, the patient had access site bleeding. several blood products were given. support was continued. care was withdrawn and the patient expired. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of access site bleeding is determined to be patient condition because the bleeding is occurring from all the access sites. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10246 | 1220648-2025-27557 | 21850085 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21850085&pc=OZD | 03/24/2025 | 04/22/2025 |  | 04/16/2025 |  |  | 03/24/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1458318 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 79 YR |  | Female |  |  |  | Death | Yes | Yes | Decrease in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during preparation of the purge system for impella cp placement, an error occurred on the automated impella controller (aic) saying no purge fluid. after troubleshooting it seems that the drive shaft in the aic was blocked and therefore no purge fluid could be pumped. the aic was exchanged and impella was successfully placed. it was further reported that the patient expired soon after arriving at icu. the minimal interruption in therapy cannot be excluded from the patient¿s outcome. | [5 empty] | The investigation of automated impella controller (aic) - purge issue ¿ other has been completed. the console logs were consistent with what was reported in the complaint. further inspection of the console revealed that there was dried glucose on the purge motor. the observed dried glucose likely caused the purge motor to get stuck to the purge housing. because the purge motor was stuck, it was unable to rotate and purge the line. the root cause of the purge fluid detection issue was a seized purge motor due to dried glucose ingress. g1 reporting contact address was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27557. g1 manufacturing site name and address was erroneously entered on the initial manufacturer device report number 1220648-2025-27557. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
10247 | 1220648-2025-27558 | 21850090 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21850090&pc=OZD | 03/21/2025 | 05/19/2025 |  | 04/16/2025 |  |  | 03/21/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026638668 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the 52 year old male patient was on impella cp support and during support, the pump position was being optimized. when nurse was adjusting patient in the bed, impella was pulled back into aorta. attempts to reposition device under echo in icu unsuccessful. patient brought to cath lab for explant and replacement of impella cp. there was no harm to the patient. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of positioning issues has been completed. product was not returned for review. based on clinical description, the root cause of positioning issue was most likely patient movement since the pump pulled out of the left ventricle (lv) while patient was being adjusted in the bed.
10248 | 1220648-2025-27559 | 21850093 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21850093&pc=OZD | 03/20/2025 | 05/07/2025 |  | 04/16/2025 |  |  | 03/20/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025563830 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after initiating impella cp support on auto, flows ramped up with appropriate waveforms. after exchanging repositioning sheath for peel-away, catheter movement occurred, and the physician had to reposition. repositioning resulted in a kink in the impella cannula and impella flow was reduced. further repositioning resolved the kink in cannula but did not improve impella flows. decision was made to explant the impella and replace with new cp. there was no harm to the patient. | [5 empty] | The investigation of low pump flow, positioning issues, and product damage (cannula kink) has been completed. clinical states that pump flows were acceptable at first. then after exchanging repositioning sheath for peel-away, catheter movement occurred and the physician had to reposition. repositioning resulted in cannula kink and flows dropped. stated that repositioning resolved the kink in cannula but did not improve impella flows. lack of clinical details regarding patient conditions or anatomy. returned complaint cp pump visually inspected. cannula kink observed near outlet. no biomaterial, no broken blade found. cannula soaked in water for better inspection to dissolve dried dextrose. positioning issues: the cause of the positioning issues most likely was use related as it occurred during the exchange of the repo sheath. low pump flow: though the clinical states repositioning resolved the cannula kink in the field (as seen on echo), a kink was observed on the returned product. the cause of the low pump flow issue was likely a cannula kink as observed on the returned product. product damage (cannula kink): the cause of the product damage cannula kink issue was unable to be determined due to lack of clinical details. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental mdr will be filed.
10249 | 1220648-2025-27560 | 21850169 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21850169&pc=OZD | 03/20/2025 | 04/16/2025 |  | 04/16/2025 |  |  | 03/20/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026633358 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support and during support, persistent oozing was noted. manual pressure was held for 20 minutes to try to resolve oozing around sheath. after manual pressure was held persistent oozing still noted. the doctor decided the safest option was to wean and explant the pump. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10250 | 1220648-2025-27561 | 21850209 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21850209&pc=OZD | 03/22/2025 | 05/13/2025 |  | 04/16/2025 |  |  | 03/22/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026629956 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an impella 5.5 pump was placed for a patient undergoing coronary artery bypass graft due to coronary microvascular disease and coronary artery disease. the pump was supporting with no swan and sodium bicarbonate in the purge. the site had to be washed out and surgically explored. the team also gave blood products. the blood products also may have benefitted the suction/low flows that were observed. the pump supported for just under 10 days before explanted. the patients outcome is stable. | [5 empty] | The investigation of access site bleeding and low pump flow has been completed. product was not returned for review. based on clinical description and data review, the root cause of low flow was most likely patient condition related since the low flow occurred follow severe access site bleeding and resolved by blood & fluid compensation. the cause of the bleeding was most likely patient condition related since the bleeding was reported to be resolved with washout and multiple blood products. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10251 | 1220648-2025-27562 | 21850212 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21850212&pc=OZD | 03/20/2025 | 05/08/2025 |  | 04/16/2025 |  |  | 03/20/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026612774 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the decision for impella cp was made after a patient was taken to the cath lab for right heart catheterization. it was noted that the insertion was difficult due to tortuous right common femoral artery, requiring multiple impella sheaths, however it was inserted with difficulty. the wire was removed and the device attempted to start, however would not start. a new cp was inserted and worked well. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of pump did not start has been completed. no product was returned. data log review showed successful priming and just after ps becomes visible, the pump is turned up to p-9 but immediately has a motor current spike and pump stop with motor current high alarm. there was no change in motor speed at this time. the pump was removed and replaced with another cp. the root cause of the pump not starting was not determined since product was not returned.
10252 | 1220648-2025-27563 | 21850258 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21850258&pc=OZD | 03/24/2025 | 04/16/2025 |  | 04/16/2025 |  |  | 03/24/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1103719 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on impella support when there was an automated impella controller (aic) to aic transfer. the aic recognized the pump and resumed flows at p7. the purge cassette, bag and purge disc were transferred over to the new aic. there was a purge disc was not detected alarm. the alarm did not resolve despite attempts to reposition the purge disc, ¿burping¿ the tubing, and performing a cassette and bag change with a new purge cassette. the new purge disc expressed the same alarm. ultimately, the decision was made to put the patient back on the original aic and transport the patient. | [5 empty] | The investigation for the reported purge disc/flag issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device logs from the complaint date showed multiple occurrences of the "purge disc not detected" alarm. initially the pump was inserted before connecting the purge cassette, and the pump was then resumed at its previous p-level. shortly after the purge cassette and disc were inserted, the "purge disc not detected" alarm was triggered and persisted even after the console was rebooted. the console was tested. the issue of purge disc not detected was reproduced by connecting the purge cassette. it was confirmed that the purge flag was missing/broken. the root cause for the purge disc not detected alarm was a broken purge flag.
10253 | 1220648-2025-27564 | 21850367 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21850367&pc=OZD | 03/21/2025 | 04/16/2025 |  | 04/16/2025 |  |  | 03/21/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026624340 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the 85 year old female patient was on impella cp support and at the start of the case, 1 perclose didn't take and a 2nd was deployed that took. the patient was complaining of right leg pain. the pump was explanted, and staff couldn't get a right pedal pulse and leg was starting to mottle. placed another sheath in the left femoral artery and shot the right groin - no flow down the right leg. doctor said they either preclosed the artery shut or it is the angioseal. they ballooned it opened enough to get some flow down the leg, but had multiple perfs in the process. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10254 | 1220648-2025-27565 | 21850393 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21850393&pc=OZD | 03/23/2025 | 04/16/2025 |  | 04/16/2025 |  |  | 03/23/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026633360 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient on impella cp support had a plasma free hemoglobin of 173. the impella cp was weaned and removed. the patient's outcome is stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
10255 | 1220648-2025-27566 | 21850398 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21850398&pc=OZD | 03/21/2025 | 04/16/2025 |  | 04/16/2025 |  |  | 03/21/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025585549 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the 74 year old male patient developed some oozing from the groin. no systemic anticoagulant was running. dressing was changed, site re-matched angle, and some manual pressure given. the patient was given one unit of packed red blood cells overnight was administered. the next day the groin was clean, dry, and intact. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10256 | 1220648-2025-27567 | 21850421 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21850421&pc=OZD | 03/24/2025 | 04/16/2025 |  | 04/16/2025 |  |  | 03/24/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025527332 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, when the procedure was almost completed, the patient started complaining of bad left leg pain. hand injection was done via re-access port on left femoral cp. occlusive flow noted to left leg. the physician tried to get antegrade sheath into left superficial femoral artery for distal perfusion, but it was unsuccessful. the decision was made to remove the impella cp. left groin perclosed and manual pressure held, and hemostasis achieved. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
10257 | 1220648-2025-27568 | 21850432 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21850432&pc=PYX | 03/24/2025 | 05/15/2025 |  | 04/16/2025 |  |  | 03/24/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025487291 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella rp device for mechanical circulatory support. while on support, there was oozing from all puncture sites. heparin was removed from the purge and bicarb was placed. additionally, the impella flows dropped and placement signal was off the screen. unable to re-zero pas. the pump was explanted and the patient expired. there is not enough evidence to dissociate the impella pump as an associated factor to the outcome given bleeding event and pump flow therapy suboptimal. | [5 empty] | The investigation of vascular damage - peripheral vessel, access site bleeding, and low pump flow has been completed. the impella device was not returned for evaluation. vascular damage - peripheral vessel: clinical details noted that bleeding was occurring from multiple sites. it was not noted if any repairs were performed for bleeding. the root cause of the vascular damage - peripheral vessel could not be definitively determined as limited clinical details were provided. access site bleeding the clinical details noted that patient was oozing from all puncture sites. the root cause of the access site bleeding was most likely patient condition related as patient was bleeding from multiple sites per clinical details. low pump flow: data logs were reviewed and lt log at time of issue showed a large motor current spike with a speed deviation and sudden drop in flows. additionally, a step up in placement signal was observed at time of motor current spike. pump flows remained lower than p-level range for remainder of case. clinical details noted that patient was not on any forms of anticoagulation at time of reported issue. the root cause of the low pump flow was most likely due to biomaterial ingestion based on returned data log analysis. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella rp with smart assist system with automated impella controller section: warnings and cautions instructions for use for the related event are as follows: ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ section: using the automated impella controller with the impella rp flex smartassist system catheter, use of the repositioning sheath and the 23 fr peel-away introducer: ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ impella rp smartassist system with automated impella controller & rp flex with smart assist system with automated impella controller, section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: potential adverse events (united states): ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ rp flex with smartassist system with automated impella controller, section: warnings & cautions: warnings: ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿
10258 | 1220648-2025-27569 | 21850602 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21850602&pc=OZD | 03/23/2025 | 04/16/2025 |  | 04/16/2025 |  |  | 03/23/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026634551 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Hispanic | White | Death | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient arrested at home, remained pulseless until arrival to emergency room. return of spontaneous circulation was achieved, and the patient was taken to catheterization lab. left anterior descending artery was found occluded. it was opened and stented. the impella cp was placed post this intervention. the patient then had pulseless electrical activity arrest. cardiopulmonary resuscitation was performed for 20 minutes with multiple rounds of epinephrine. return of systemic circulation was attained again. echocardiogram showed the impella cp was placed across the mitral valve. the physician attempted to reposition the impella cp away from the mitral valve, but it was pulled into the aorta. the physician was unable to get back the impella cp into the ventricle. impella cp was removed from the body. the ventricle was reassessed and the same impella cp was inserted using a v18 wire. after restarting the impella cp, the physician was concerned about the placement signal and motor current waveforms and low flow. the impella cp was removed and a new impella cp was inserted. the patient coded again immediately after placing the new pump and expired in the procedural area. | [5 empty] | The impella device was not received from the customer as the pump stayed with the expired patient. therefore, the investigation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
10259 | 1220648-2025-27571 | 21850724 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21850724&pc=OZD | 03/23/2025 | 05/12/2025 |  | 04/16/2025 |  |  | 03/23/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025594030 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that when implanting the impella cp in a patient, the vasculature was tortuous. the physician felt like there was calcification to work through and then felt the same when crossing the aortic valve. toward the end of the case, it was noted that the patients impella site leg was cool, dusky, and unable to find a dopplerable pulse. the patient¿s hemodynamics was maintained so ultimately the decision was made to remove the impella cp after percutaneous coronary intervention. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues and limb ischemia has been completed. the impella device was not returned for evaluation. based on the clinical details, the root cause of delivery issue was most likely patient condition since the implant of pump was aborted due to resistance met because of tortuous / small vessel size. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined.
10260 | 1220648-2025-27572 | 21850725 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21850725&pc=OZD | 03/24/2025 | 05/08/2025 |  | 04/16/2025 |  |  | 03/24/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531723 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  |  | White | Death | Yes | Yes | Positioning Failure | Restricted Flow rate | [5 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, there was ongoing suction alarms. echo showed 4.6cm inlet to aortic valve annulus. the echo also showed inlet close to mitral valve: not impeding movement, but leaflet may be close to inlet while the valve is open. the doctor declined to reposition. care was withdrawn and the patient expired. | [5 empty] | The pump was not returned for investigation. clinical observations indicated that the pump appeared to be positioned deeply, with the inlet located close to the mitral valve. despite this positioning, the pump was kept in place. there were intermittent suction alarms during the case, which prompted a reduction in the pump's p-level to manage the suction concerns. the data log did not record any position-related alarms or motor current spike. the pump had several suction alarms and low pump flow while running on a higher p-level (p9). reducing the p-level to p7 resolved the low pump flow, but the suction alarm persisted for some time. later during the case, the p-level was further reduced to resolve the suction condition. the cause of the low pump flow issue was most likely patient condition related, based on data log review. the cause of the positioning issue was not determined since sufficient clinical information was not provided. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings, cautions and precautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ understanding and managing impella catheter position alarms: ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ correct impella catheter position: ¿for optimal positioning of the impella catheter, the inlet area of the catheter should be 3.5 cm below the aortic valve annulus and well away from papillary muscle and subannular structures. the outlet area should be well above the aortic valve.¿
10261 | 1220648-2025-27570 | 21850752 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21850752&pc=OZD | 03/24/2025 | 09/29/2025 |  | 04/16/2025 |  |  | 03/24/2025 | 04/16/2025 |  |  | 07/31/2025 |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024370685 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Stroke/CVA | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant also reported that following impella 5.5 implant, the aorta (ao) placement signal gave irregular readings. the difference between the ao placement signal and arterial line increased and decreased coinciding with the rise and fall in pressure. | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the patient had a massive stroke. care was later withdrawn and the patient expired. there is not enough evidence to disassociate the use of the impella and the patient's outcome. |  |  |  |  | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue and stroke. no product was returned. data log review showed no abnormalities with the 5s optical parameters. no abnormalities were found with the imc logs. the root cause of the optical signal issue was not determined since product was not returned and no abnormalities were seen with the optical parameters in the data logs. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include placement signal issue description. h6 updated to include placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-27570. d10 concomitant medical products and therapy dates updated. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smart assist for use during cardiogenic shock, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10262 | 1220648-2025-27573 | 21850869 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21850869&pc=OZD | 03/25/2025 | 05/14/2025 |  | 04/16/2025 |  |  | 03/25/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026634518 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a patient admitted in with known coronary artery disease and plan for a high-risk percutaneous coronary intervention. the pump failed to deliver and the 13 french x 25 centimeter sheath had kinked with the attempted delivery. the medical doctor reported that the patient had a heavily calcified access site. the pump was aborted and there was no patient harm. | [5 empty] | The pump and introducer were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | Clinical details report that difficult anatomy and heavily calcified arteries caused the introducer sheath to kink during access. the impella cp was unable to be delivered for this reason. it is unclear from clinical details how far the impella cp was delivered when the team aborted. the investigation of introducer damage: mechanical and delivery issues has been completed. introducer damage mechanical: returned product was investigated and there was a kink in the sheath as reported. based on the clinical report that the patient had heavily calcified anatomy and the returned introducer had a kink, the root cause of the introducer damage was determined to be a sheath kink. delivery issues: the returned pump was investigated and there were no abnormalities noted, particularly to the cannula or catheter. based on clinical details provided that the introducer sheath kinked due to diseased vasculature, and the delivery of the cp was unsuccessful, the root cause of the delivery issue was determined to be patient condition.
10263 | 1220648-2025-27574 | 21850870 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21850870&pc=OZD | 03/24/2025 | 04/16/2025 |  | 04/16/2025 |  |  | 03/24/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026615805 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient experienced oozing at the site, pressure was held, placed pressure dressing in addition to receiving one unit of red blood cells. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation for the reported access site bleeding issue has been completed. the impella device was not received from the customer and therefore, the root cause of the access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10264 | 1220648-2025-27575 | 21850873 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21850873&pc=OZD | 03/24/2025 | 05/23/2025 |  | 04/16/2025 |  |  | 03/24/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1117231 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella aic had a controller error alarm. the aic will be returning for investigation. there was no report of patient involvement. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. during the data analysis there was no controller error in the logs on or around the complaint. the failure mode was not reproduced as the console was able to successfully shutdown on every attempt. there was no controller error with this console. the cause of the issue was most likely user related as they could have inadvertently rotated from ok to cancel without realizing it, therefore cancelling shutdown. e2 was erroneously selected on the initial manufacturer device report number 1220648-2025-27575.
10265 | 1220648-2025-27576 | 21850903 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21850903&pc=OZD | 03/24/2025 | 04/16/2025 |  | 04/16/2025 |  |  | 03/24/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025609862 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Positioning Failure | Increase in Pressure | [5 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the pump left deep at 4 cm after repositioning from 5 cm. support continued and staff monitored. additionally, low purge flows and high purge pressures were noted. tissue plasma activator was added to the purge and support continued. care was withdrawn on support and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10266 | 1220648-2025-27577 | 21850968 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21850968&pc=OZD | 03/24/2025 | 04/16/2025 |  | 04/16/2025 |  |  | 03/24/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025572015 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed hematoma on the right groin. the site was redressed, and oozing was stabilized. heparin was turned off for about six hours and then re-started. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10267 | 1220648-2025-27578 | 21851006 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21851006&pc=OZD | 03/24/2025 | 09/25/2025 |  | 04/16/2025 |  |  | 03/24/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025601463 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, on day 26 of 5.5 support there was a high spike in motor current and then the impella stopped. the impella restarted and later replaced. there was no patient harm. | The complainant reported that the patient is waiting for a heart transplant on impella 5.5 support. on day 21 of 5.5 support the impella lost placement signal overnight. upon inspection there was still good motor currents and flows at p-5. echo confirmed correct positioning. alarms and signal were disabled, and decision was made to monitor. it was further reported that on day 26 of 5.5 support there was a high spike in motor current and then the impella stopped. the impella restarted but the plan is to exchange the 5.5. not reportable for the placement signal as this issue does not affect the pump performance. |  |  |  |  | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of pump stop has been completed. the cut pump was sent for investigation. a severe kink was noted near the kink protector. the laser continuity test failed at the kink protector. additionally, a significant amount of biomaterial was noted on the impeller. there was no sign of blood ingress revealed after slicing the catheter near the motor housing. further electrical testing confirms that one out of three electrical lines was open at the noted kink at the kink protector. the data log confirmed that on the 21st day of support. also, several pumps stop with motor current spike to 1200 and alarm 119-impella stop. the cause of the pump stop issue was an open electrical line at kink protector. | This report is being filed as part of a retrospective review of historical records. this mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the cut pump was sent for investigation. a severe kink was noted near the kink protector. the laser continuity test failed at the kink protector. additionally, a significant amount of biomaterial was noted on the impeller. there was no sign of blood ingress revealed after slicing the catheter near the motor housing. further electrical testing confirms that one out of three electrical lines was open at the noted kink at the kink protector. the data log confirmed that on the 21st day of support, a fiber break trend was seen in all 5s parameters with active psnr alarm, which persisted until the end of the case. also several pump stop with motor current spike to 1200 and alarm 119-impella stop. no abnormalities were reported on purge pressure parameters during the case. the cause of the pump stop issue was an open electrical line at kink protector. the cause of the optical signal issue was most likely damaged optical fiber line at kink protector due to kink in catheter. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: optical signal issue reported in effort of a retrospective review of optical signal issues in accordance with fda recommendation.
10268 | 1220648-2025-27579 | 21851038 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21851038&pc=OZD | 03/20/2025 | 04/16/2025 |  | 04/16/2025 |  |  | 03/20/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026612822 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Renal Failure | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that one day into impella cp support, the patient had a weak pulse in their lower left extremity. the decision was made to explant the pump in response to the condition. it was later reported via abiomed¿s long-term outcome and quality indicator (loqi) impella registry that the patient endured worsening renal failure with an unknown (undetermined) relationship to the impella device. the patient was said to have recovered from the condition. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
10269 | 1220648-2025-27580 | 21851043 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21851043&pc=OZD | 03/21/2025 | 04/16/2025 |  | 04/16/2025 |  |  | 03/21/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025604741 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that a patient was taken to cardiovascular operating room for escalation from an impella cp to a impella 5.5. the escalation was successful. during support on the 5.5 there was an aorta alarm. reposition was attempted multiple times under echocardiogram guidance however was unsuccessful. two drips added to maintain blood pressure. decision was made to either exchange 5.5 or add impella cp for quick immediate support. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10270 | 1220648-2025-27581 | 21851060 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21851060&pc=OZD | 03/21/2025 | 05/19/2025 |  | 04/16/2025 |  |  | 03/21/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025560848 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic |  | Injury | Yes | Yes | Restricted Flow rate | Infusion or Flow Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that upon starting up the pump, flows were reduced at higher p-levels and suctions was encountered above p5. upon rotating the device, a thrombus was noted on the inlet of the device. the pump was removed and run in sterile water. the left ventricle was rewired and the pump was successfully re-inserted for patient support. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation of thrombus, low pump flow, and saturated flow has been completed. no product was returned. data logs show low flows when pump is initially started at p5 with position unknown alarms. flow becomes saturated and increases to around 3l/min when p-level is lowered to p2. biomaterial is likely in the cannula causing low flows, then traveling to the impeller, adding additional load causing saturated flows. pump was then turned off and surgical mode was enabled. pump was then run briefly at p1, turned off, then turned back up to p8 with flows around 4.5l/min and stayed around 5l/min for the rest of the case. there was also a step up in placement signal after the pump was restarted. this aligns with clinical reporting removing the pump, running it in sterile water to removed biomaterial, then reimplanting and having better flows. the root cause of the low flow and saturated flow was most likely biomaterial since clinical reported biomaterial on the inlet cage, dampened motor current and logs show flows improved after the pump was removed, run in sterile water, then reimplanted. as the necessary information was not provided, the root cause of the thrombus was not determined. e4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration.
10271 | 1220648-2025-27582 | 21851217 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21851217&pc=OZD | 03/24/2025 | 04/16/2025 |  | 04/16/2025 |  |  | 03/24/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026613440 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Ventricular Fibrillation | [11 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, and during the implant, the patient experienced ectopy. lidocaine push and amiodarone drip administered. additionally, echo was performed, and a pericardial effusion was seen on echo. ultrasound also performed on the right popliteal artery showing no flow. patient was transferred back to the cath lab for urgent pericardiocentesis and impella explant due to no blood flow to distal leg. the clinical rep stated that the pericardial effusion was caused specifically from the infarct. the doctor believed that this was a left ventricle rupture. the clinical rep stated this likely occurred prior to the impella implant however, they were unsure if the impella contributed to the pericardial effusion. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
10272 | 1220648-2025-27583 | 21851620 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21851620&pc=OZD | 03/24/2025 | 05/14/2025 |  | 04/16/2025 |  |  | 03/24/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025593615 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had a hematoma at the access site. pressure was held and heparin was withheld. additionally, the pump initially measured at 1.10 cm, the physician advanced the pump to 2.12. the physician elected to leave it there even though it was shallow and continued support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings, cautions and precautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ correct impella catheter position ¿for optimal positioning of the impella catheter, the inlet area of the catheter should be 3.5 cm below the aortic valve annulus and well away from papillary muscle and subannular structures. the outlet area should be well above the aortic valve.¿ | The investigation of access site bleeding and positioning issues has been completed. the impella device was not returned for evaluation. the root cause of the positioning issue was most likely patient movement since the pump was accidentally pulled into the aorta when boosting the patient. the root cause of the access site bleeding was not determined since product was not returned and insufficient clinical details provided. h6 health effect - clinical code 1884 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27583.
10273 | 1220648-2025-27584 | 21851908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21851908&pc=OZD | 03/24/2025 | 05/13/2025 |  | 04/16/2025 |  |  | 03/24/2025 | 04/16/2025 |  |  |  |  | 04/16/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025601564 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 43 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, a coronary artery bypass grafting (cabg) was performed. retrograde flow alarms and high ao ps shown and flows less than 1 l/min. the physician decided to replace with a new impella 5.5 without issues. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of low pump flow has been completed. the root cause of low flow was unable to be determined as limited clinical details were provided and no product was returned for review.
10274 | 1220648-2025-27466 | 21839025 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21839025&pc=OZD | 03/18/2025 | 04/15/2025 |  | 04/15/2025 |  |  | 03/18/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025580853 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that they suspected minimal blood flow in left leg post impella cp insertion. patient has history of peripheral artery disease in both legs and a percutaneous catheter bypass was placed from right to left to help blood flow. the impella was successfully explanted post right coronary artery revascularization. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
10275 | 1220648-2025-27467 | 21839082 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21839082&pc=OZD | 03/18/2025 | 05/08/2025 |  | 04/15/2025 |  |  | 03/18/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025584987 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, the pump was slightly deep positioning but left where it is due to concerning, small aorta anatomy. support continued and the patient survived. | [5 empty] | The investigation of positioning issues has been completed. clinical states that the patient had bicuspid aortic valve disease and a calcified aorta due to which the previous pump was kinked. the new pump was left little deep due to patient's concerning small aorta anatomy. pump was not returned. the root cause of the positioning issue was most likely patient condition related to the patient's small aorta, diseased bicuspid aortic valve, and calcification. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10276 | 1220648-2025-27468 | 21839141 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21839141&pc=OZD | 03/18/2025 | 05/27/2025 |  | 04/15/2025 |  |  | 03/18/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1847282 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller (aic) black power cord was received damage and was not usable. this occurred on (b)(6). it broke when they unplugged it from the wall. the aic was replaced. the patient remained stable and survived. | [5 empty] | The investigation of automated impella controller (aic) - battery charging/other issues has been completed. console logs from the reported day of event did not contain any information relevant to this failure mode. at the time of complaint, the staff provided a photo showing ac wall plug separated from the console power cord. when console arrived for investigation, its power cord was either already repaired or replaced on-site (we were unable to ascertain which one it was). the root cause of the aic wall plug separation could not be determined. d9 device returned to manufacturer date added d6a and d6b should have been left blank on the initial manufacturer device report 1220648-2025-27468. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10277 | 1220648-2025-27469 | 21839509 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21839509&pc=OZD | 03/18/2025 | 04/15/2025 |  | 04/15/2025 |  |  | 03/18/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026646211 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Female |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support and during support, the patient had access site bleeding. additional compression was needed and blood products were given. the patient remained stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10278 | 1220648-2025-27470 | 21839749 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21839749&pc=OZD | 03/18/2025 | 04/15/2025 |  | 04/15/2025 |  |  | 03/23/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025531396 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | Yes | Positioning Failure | [6 empty] | Ventricular Fibrillation | Thrombocytopenia | Thrombosis/Thrombus | [10 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a st-segment elevation myocardial infarction patient coded upon arrival to the lab, return of spontaneous circulation was achieved and decision made for impella cp placement. during support it was noted that the impella was measuring deep on echocardiogram (echo), however, interventional cardiologist did not wish to reposition at that time. the rep spoke to physician about a new potential thrombus or infection attached to cp in left ventricle on echo. additionally, the patient coded however return of spontaneous circulation was achieved. the patient was reintubated during code- ventricular fibrillation arrest. echo showed cp deep and physician was guided through repositioning with echo. it was noted that medical doctor thinks heparin-induced thrombocytopenia. was the cause for the small thrombus that has appeared on the cannula in addition was unsure if the impella was tickling in the ventricle. ultimately the doctor wanted the cp removed. the cp was weaned and successfully explanted. the patient remained stable and survived. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: impella cp with smart assist catheter insertion, section: axillary insertion of the impella cp with smart assist catheter, section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
10279 | 1220648-2025-27471 | 21839835 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21839835&pc=PYX | 03/18/2025 | 09/26/2025 |  | 04/15/2025 |  |  | 03/18/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025551628 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Stroke/CVA | Hemolysis | [11 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient with right heart failure arrived at the cardiac catheterization lab from the intensive care unit on extracorporeal membrane oxygenation (ecmo) and ventilator. the impella rp was inserted. ecmo was decreased and working in conjunction with the impella rp flex. the nurse stated that the impella rp placement signal and pa signal were missing and that there could be a thrombus in the cannula. hemoglobin had increased and decreases of flows were noted. additionally, the patient developed hemolysis and due to intracranial hemorrhage, the patients care was withdrawn. the patient expired. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer as the patient passed away and the pump was not removed. therefore, the investigation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for hemolysis, low pump flow, stroke/cva, and placement signal issue. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. thrombus can be observed in the body or on the pump upon explant. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. on the 17th, placement signal was trending down as well as flows. the cvp is negative as well. there was spike in motor current around 8pm. flows and motor current remained variable after the fact. at 8:45am on the 18th there were spikes in motor current and placement signal with small decreases in flow. based on the timeline of the reported hemolysis and low pump flow, as well as with the data log review, the root cause of the hemolysis and low pump flow is most likely due to ingested biomaterial. the flow and placement signal are trending down. flows and motor current are also variable. the 5s parameters are stable. due to lack of product returned, and that is not enough evidence to determine if there is a volume issue or biomaterial influencing the optical sensor, the root cause of the optical signal issue cannot be determined. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. h6 updated to include placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-27471. d10 concomitant medical products and therapy dates updated. g1 reporting contact email updated.
10280 | 1220648-2025-27472 | 21839901 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21839901&pc=OZD | 03/18/2025 | 05/08/2025 |  | 04/15/2025 |  |  | 03/18/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025591553 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after explanting the impella cp it was difficult closing the access site. the physician attempted perclose four times with failure, needing to replace a 14 french short cook sheath for hemostasis. the staff was holding pressure on the site. one unit of packed red blood cells and a liter of fluid were given. vascular surgeon repaired the site the next day. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10281 | 1220648-2025-27473 | 21839990 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21839990&pc=OZD | 03/17/2025 | 08/25/2025 |  | 04/15/2025 |  |  | 03/19/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025608271 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Nerve Damage | Ischemia Stroke | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Upon later notification via impact study, it was additionally reported that the patient endured potential pericardial bleeding with a possible relation to the impella device, placement, and procedure. in addition, bleeding complications were reported during impella insertion. surgical intervention and blood products were given to manage the bleed. | It was reported via an impact study that a patient on impella 5.5 mechanical circulatory support experienced nerve damage during impella support. it believed that the nerve damage had a probable relationship to the impella procedure. additional information was not provided. seven days following explant, the patient experienced an ischemic stroke. an mri showed micro-embolic infarctions, subacute on both sides, no symptoms, no therapy. the stroke was believed to have a probable relationship to the impella procedure. the patient was noted to have recovered native heart function. the device was explanted, and the patient was noted to have survived the explant. |  |  |  |  | The investigation has been completed. no product was returned. per the clinical investigation, in order to make a root cause determination, all of the clinical details would have to be provided. the root cause of the stroke/cva was unable to be determined due to insufficient clinical details. the cause of the peripheral nerve injury was not determined due to insufficient clinical details. bleeding was reported during insertion. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. vascular damage was reported for pericardial bleed. the root cause of the vascular damage was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. | B5 updated with additional information from the clinical site h6 updated to reflect the additional failure mode should new information or investigation results for access site bleeding become available a supplemental report will be submitted. instructions for use for the related event are as follows: stroke/cva: "strokes/cvas can occur during or after impella support, and can be either ischemic or hemorrhagic. to make a determination as to the cause of the stroke, the following clinical information must be considered: -patient's medical history, including any previous strokes, cardiovascular disease, or risk factors such as hypertension, diabetes, or atrial fibrillation -timing of the stroke in relation to impella support (during or post-implant) -detailed description of the symptoms experienced by the patient -neurological examination findings and any focal deficits observed -diagnostic imaging results, such as ct scan or mri of the brain, to identify the type and location of the stroke (ischemic or hemorrhagic) -laboratory test results, including coagulation profiles and cardiac markers -any relevant procedural or device-related factors, such as duration and settings of impella support, presence of embolic protection devices, or any documented procedural complications -response to any previous anticoagulation or antiplatelet therapies -evaluation of potential alternative causes of stroke, such as arrhythmias, embolic sources, or underlying vascular pathology as this clinical information was not provided, the true root cause of the stroke/cva could not be determined.¿ peripheral nerve injury: the cause of the issue was not determined due to insufficient clinical details. g1 manufacturing site name, address, country and fax number were erroneously entered on the initial manufacturer device report number 1220648-2025-27473. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
10282 | 1220648-2025-27474 | 21840036 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21840036&pc=OZD | 04/01/2025 | 05/20/2025 |  | 04/15/2025 |  |  | 04/01/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | It has been reported from a literature titled: a case of abdominal aorta occlusion caused by impella removal. that a patient who underwent aortic valve replacement and mitral valve replacement for heart failure due to aortic and mitral stenosis needed impella cp support. during surgery, there were difficulties in weaning the patient from cardiopulmonary bypass, so an impella cp and venoarterial extracorporeal membrane oxygenation (va-ecmo) were inserted. after returning to the intensive care unit (icu), persistent bleeding from the drain was observed, and open-chest hemostasis was performed twice in the icu, on the 1st and 2nd postoperative days, and activated clotting time was controlled with the goal of reaching the lower limit of normal range. on the 5th postoperative day, va-ecmo was removed and weaning from impella was initiated. on the sixth postoperative day, the ic was removed under general anesthesia, and occlusion of both common iliac arteries from the distal abdominal aorta was observed. stent graft insertion was performed. the outcome of the patient is unknown. | [5 empty] | D6: implant/explant dates unknown. the investigation of vascular damage, peripheral vessel and thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: impella cp with smart assist catheter insertion, section: axillary insertion of the impella cp with smart assist catheter, section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
10283 | 1220648-2025-27475 | 21840252 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21840252&pc=OZD | 03/31/2025 | 07/03/2025 |  | 04/15/2025 |  |  | 04/04/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1393502 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that when priming the impella cp smartassist with an automated impella controller (aic), the instructions for the old model, "please connect the red plug," was displayed on the screen, and it was not possible to proceed. the aic was replaced with another aic and support was continued. | [5 empty] | Additional information provided in h3, h6 and h10. console logs are mostly consistent with complaint and show entries indicating same purge cassette being unexpectedly disconnected and reconnected. logs were also ¿flooded¿ with imc-ahp error messages, related to rfid communication, e.g.: log data indicates disruption in purge stats since start of case, until purge cassette was replaced on (b)(6) 2025. log data shows that each time purge cassette was erroneously detected, ¿homing¿ process started, which means purge cassette piston movement was stopped and piston fully retracted, momentarily interrupting purge flow, which affected purge stats calculation. logs also shows that the erroneous pattern of purge piston movements was no longer reproduced after purge cassette was replaced, and the case continued until its conclusion. the console was tested but the issue of intermittent purge cassette detection was not reproduced, therefore root cause could not be determined. as a precaution the rfid circuitry was recommended to be replaced as it was the most likely component to have intermittently malfunctioned. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
10284 | 1220648-2025-27476 | 21840332 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21840332&pc=OZD | 04/01/2025 | 04/15/2025 |  | 04/15/2025 |  |  | 04/01/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that information obtained from literature indicated that the patient encountered access site bleeding. a 74-year old patient was being provided impella support and on the third day of admission, swelling with a sense of fullness and subcutaneous bleeding spots suddenly appeared in the right forearm, so the peroneal artery catheter was removed and compression hemostasis was performed. after surgery, bleeding continued and surgical hemostasis was performed, but hemostasis was difficult, so a transfusion of 44 units of irradiated red blood cells (rbc), 40 units of fresh frozen plasma (ffp), and 40 units of platelets (pc) was required within 24 hours. on the fourth day of admission, hemostasis was performed again, but hemostasis was not achieved. therefore, a tourniquet was wrapped around the right upper arm and intermittent tourniquet was performed. as a result, within 24 hours, the amount of transfusion could be reduced to 28 units of rbc, 28 units of ffp, and 20 units of pc. on the fifth day of hospitalization, the impella was removed and the amount of heparin administered was reduced, and the amount of bleeding also decreased. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10285 | 1220648-2025-27477 | 21840500 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21840500&pc=OZD | 03/18/2025 | 04/21/2025 |  | 04/15/2025 |  |  | 03/18/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED INC. |  |  | US |  | 2025581521 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp was removed due to a patient heavily oozing from the impella site. hemostasis was achieved using the pre-close method with two perclose. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details, the root cause of the access site bleeding issue was anticoagulation management due to high patient act values. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10286 | 1220648-2025-27478 | 21840825 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21840825&pc=OZD | 03/22/2025 | 11/20/2025 |  | 04/15/2025 |  |  | 03/22/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1427009 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported during impella 5.5 support to a patient with cardiomyopathy, intermittent controller error alarms were encountered. the placement signal and left ventricle signal were intermittently dashed out and not present on the automated impella controller (aic), which appeared to be triggering the controller error alarms. it appears the aic remained in use. the latest update indicated the patient was stable on current support and status 2 for heart transplant. | [5 empty] | The automated impella controller device was not received from the customer. it appears the device remains in use. therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. console logs (obtained from the cloud) from the reported day of event are consistent with complaint and show multiple occurrences of controller error alarms #1045 (opm communication crc invalid) and entries ¿invalid_bad_pressure_sample_mask¿ (invalid sample due to ambient pressure out of range.) there are also numerous psnr events (1036 and 130) and one placement signal not reliable alarm. ltlog data showed unrealistic values of contrast, gail, lamp, light and invalid values of placement signal coinciding with controller error alarms and events. snr had been low (< 2000) throughout the case, and did not contribute to the controller error alarms, but caused all psnr events. rtlog data was not available. imc logs also indicated frequent interruptions of data streaming, and eventually complete loss of usb communication from aic to rlm (unrelated to complaint). console was initially not returned for investigation, but later arrived at fs. any data logs related to this complaint were overwritten on the device¿s cf cards. console testing and inspection revealed additional non-conformances: defective optical bench failing ¿5s¿ specification (causing psnr condition due to low snr), and some persistent contamination of the optical pump connector fiber. test cavity 1. test cavity 2. cavity length: 14334 nm. snr: 377. contrast: 45.5%. lamp: 90.6%. gain: 1.30. mano: 761.0 mmhg. light: 2.70 v. cavity length: 16499 nm. snr: 1599. contrast: 22.6%. lamp: 75.6%. gain: 1.33. mano: 761.0 mmhg. light: 2.78 v. monitoring of analog output of optical bench ccd detector revealed distortion of the signal, consistent with low snr. a compromised (misaligned) miniusb plug on the rlm was found to have intermittent electrical contact, which led to frequent disconnections and reconnections. although not reproduced during testing, field symptoms as well as the pattern of optical parameters in imc and ltlogs are consistent with loss of opm data due to bench firmware anomaly. distortion of the ccd output explained low snr, and compromised backstrap cable - interruptions of data streaming unrelated to complaint). a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d10 concomitant medical products and therapy dates updated.
10287 | 1220648-2025-27479 | 21840834 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21840834&pc=OZD | 03/18/2025 | 05/12/2025 |  | 04/15/2025 |  |  | 03/19/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025601570 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Vascular Problem | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted an impella 5.5 device for elective placement and following impella 5.5 implant via direct access, the patient experienced a large bleed from chest tubes. multiple blood products were given, and the patient was taken back to the operating room for a chest washout. upon examination, it was discovered that a branch vessel of the inferior mesenteric artery (ima) had not been sealed. bleeding continued for another two days and required additional blood products to be given. at that point, it was additionally found that the right marginal artery had been nicked by a sternal wire. support continued with the implanted pump and was noted to be doing well before successfully weaning and explant of the impella 5.5 device with a survived outcome. | [5 empty] | The investigation of vascular damage - peripheral vessel has been completed. the pump was not returned for investigation. a data log review was not required for this case. according to the clinical report, the pump was placed surgically using direct aortic insertion. the doctor reported that the right mammary artery had been nicked by a sternal wire, and the vessel was ligated. chest tube drainage improved after the patient was sent to the ccu for repair. the cause of the vascular damage issue was most likely related to the use of a non-abiomod device. the relation between vascular damage and the impella was unknown based on the provided clinical details. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
10288 | 1220648-2025-27480 | 21840955 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21840955&pc=PYX | 03/18/2025 | 05/08/2025 |  | 04/15/2025 |  |  | 03/18/2025 | 04/15/2025 |  |  | 02/28/2026 |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED INC. |  |  | US |  | 2025478763 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that four days into impella rp flex support, the patient developed hematuria. it was noted that the condition developed after support was changed to p5. the pump was explanted later that same day with no known resolution to the suspected hemolysis prior to removal. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of hemolysis has been completed. no product was returned for analysis. the hematuria complaint was entered on 3/18 at 6:42am. the data logs show a motor current spike on 3/18 a little after midnight with a speed deviation, drop in purge and pump flows characteristic of an ingestion. this occurred before the pump was turned to p5. the root cause of the hemolysis was most likely biomaterial as evidenced by the ingestion pattern in the data logs.
10289 | 1220648-2025-27482 | 21841149 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21841149&pc=OZD | 03/18/2025 | 05/06/2025 |  | 04/15/2025 |  |  | 03/19/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026617746 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  |  | White | Malfunction | No | Yes | Increase in Pressure | Malposition of Device | Infusion or Flow Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock implanted with an impella 5.5 device for mechanical circulatory support experienced positioning issues and high purge pressure. approximately one day after implant of the impella 5.5 implant, the pump appeared deep near the papillary muscle and was repositioned with improved flow. the following day the impella had high purge pressure and low purge flow. partial thromboplastin times throughout therapy was noted 51-58 with peripheral heparin drip. bicarb purge was changed to heparin 25u/ml purge. the purge flow continued to decrease to 2 ml/hr. and the purge pressure increase to greater than 1050. tissue plasma activator protocol was initiated and added to the purge. discussion was made regarding the possibility of pump saturation as well due to the elevated calculated flows and left ventricular waveform above the aortic placement signal. the patient continued on support with monitoring of the situation. it was indicated the patient underwent percutaneous while on impella mcs. however, date of pci is unknown. the patient support ended approximately 5 days later after the high purge pressure event. the impella 5.5 device was successfully weaned and explanted with a survived outcome. | [5 empty] | The investigation of high purge pressure (hpp) and positioning issues has been completed. high purge pressure: after 2 days on support high purge pressure was noted. troubleshooting efforts occurred with purge changed to heparin 25u/ml and tissue-type plasminogen activator (tpa) protocol initiated. data log review showed high purge pressure alarms occurring without motor current (mc) spike and 9 minutes later the impella was disconnected for 19 seconds and then plugged back in. the hpp alarms (purge flow low) continued 2 hours later without a mc spike. the purge pressure and motor current eventually trends down and resolved. the returned pump reproduced the high purge pressure and was torn down to show evidence of biomaterial found in the motor housing. the cause of the high purge pressure was biomaterial due to issue reproducing in lab, biomaterial found during pump teardown, and tpa resolving the issue in the field. positioning issues: device implanted (b)(6) 2025 without issue. after 1 day on support, md had to reposition device via echo as inlet was deep near pap. data log review showed position unknown alarms at the beginning of the case that line up with clinical details for physcian repositioning the pump. however, later in the case impella position alarms occur without known clinical cause. pump was returned and inspected. the returned pump showed no abnormalities related to pump positioning. the cause of the positioning issue was not determined due to insufficient clinical details. saturated flow: after 2 days on support, clinical states possibility of pump saturation discussion occurred with care team due to elevated calculated flows. data log review showed high pump flows at and above 5.0l/min when impella was at p-7. no mc spike prior to the issue, but mc trending down as the saturation is resolving that corresponds with the high purge pressure resolving. clinical team used tpa. pump was returned and inspected. returned pump continued to show elevated pump flows at p-7 flowed showing 5.0l/min. the pump was torn down and evidence of biomaterial was found in the motor housing. the cause of the saturated flow was biomaterial due to issue reproducing in lab, biomaterial found during pump teardown, and tpa resolving the issue in the field. | The impella device was returned, and an evaluation is underway. upon completion of the analysis, a supplemental report will be filed. section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
10290 | 1220648-2025-27483 | 21841205 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21841205&pc=OZD | 03/18/2025 | 04/21/2025 |  | 04/15/2025 |  |  | 03/18/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026615956 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that after multiple attempts, it was determined that the impella 5.5 was entering the vessel, but the vessel was simply too small to advance the impella 5.5 any further. the medical doctor was very nervous that he would risk causing damage to the vessel if he proceeded with the impella 5.5. at this point, the impella 5.5 was removed from the patient. the decision was made to switch to an impella cp. the patient remained stable. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since the pump could not be implanted due to patient's anatomy being small. h10 investigation results were inadvertently omitted on the initial manufacturer device report number 1220648-2025-27483. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist, section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
10291 | 1220648-2025-27481 | 21841243 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21841243&pc=OZD | 03/18/2025 | 05/12/2025 |  | 04/15/2025 |  |  | 03/18/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED INC. |  |  | US |  | 2025605319 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp pump was unable to cross the aortic valve during attempted delivery. after multiple unsuccessful attempts, the implant was aborted. there was no patient harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system: section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be patient condition because multiple wires were unable to cross the aortic valve because of vessel stenosis.
10292 | 1220648-2025-27484 | 21841533 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21841533&pc=OZD | 03/19/2025 | 04/15/2025 |  | 04/15/2025 |  |  | 03/19/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026623538 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  |  |  | Malfunction | No | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported the insertion of impella 5.5 device (escalation from an impella cp device) to a patient (unknown race and ethnicity) in cardiogenic shock from an acute myocardial infarction was unsuccessful. multiple attempts were made to pass impella 5.5 via left axillary without success. the left approach was aborted, and the right axillary accessed. multiple attempts were made to pass impella 5.5 via right axillary and also without success. consequently, the impella 5.5 implant was aborted and impella cp was successfully implanted via right axillary artery. the patient was transferred to the unit. patient outcome post the event was indicated as unknown. however, there was no report or indication of adverse effects to the patient as a result of this event. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
10293 | 1220648-2025-27485 | 21841540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21841540&pc=OZD | 03/18/2025 | 05/19/2025 |  | 04/15/2025 |  |  | 03/18/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025596780 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was on impella cp for hemodynamic support. during support, there was a "impella in ventricle" alarm that was reported due to patient agitation. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. clinical states that after 3 hours of implant, there were "impella in ventricle" alarms due to patient agitation. pump was not returned for investigation. as per the clinical information provided, the root cause of the positioning issue was use related to patient movement. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10294 | 1220648-2025-27486 | 21841639 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21841639&pc=OZD | 03/19/2025 | 04/15/2025 |  | 04/15/2025 |  |  | 03/19/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026616871 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient undergoing high-risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support. successful protected pci was performed. the 14fr peel away re inserted to left femoral artery and angiogram performed through side port. angiogram revealed no distal femoral artery flow. a 5fr antegrade sheath was distally placed to the 14fr and connected with right femoral artery 7fr. sheath. surgery was consulted for next steps (unknown details). the impella cp device was eventually, successfully weaned after four hours of support and explanted with a survived patient outcome. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
10295 | 1220648-2025-27487 | 21841739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21841739&pc=OZD | 03/19/2025 | 04/15/2025 |  | 04/15/2025 |  |  | 03/19/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1193360 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported that they encountered an automated impella controller (aic) battery failure, red alarm during prep of the device. no further information was provided. | [5 empty] | The investigation for the console battery failure "red alarm" has been completed. the console logs were consistent with what was reported in the complaint. multiple instances of battery failure (transport connection blown/missing) alarms were observed in the logs. the console was returned for evaluation. the reported battery failure was able to be reproduced. batttest results revealed that the cells in both batteries had voltages of below 2400mv. the battery failure alarm resolved after testing this console with two known good batteries. the root cause of the battery failure was identified as depleted batteries, which resulted in battery lockout.
10296 | 1220648-2025-27488 | 21841813 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21841813&pc=OZD | 03/19/2025 | 04/22/2025 |  | 04/15/2025 |  |  | 03/19/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1763601 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that there was an automated impella controller (aic) error alarm that lasted two minutes and self-resolved. this was associated with a ¿blank¿ placement signal. as a result, the aic was swapped out. there was no patient harm reported. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. the console logs from the reported day of event are consistent with complaint and show controller error alarm. console was tested, but reported issue was not reproduced. however, standard testing revealed deficiency of the optical system: defective optical pump connector and defective (noisy) optical bench, which caused the device to sometimes fail ¿5s¿ parameter specification: the cause of the aic issue was a software issue.
10297 | 1220648-2025-27490 | 21841943 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21841943&pc=OZD | 03/19/2025 | 05/27/2025 |  | 04/15/2025 |  |  | 03/19/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025596465 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Female |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient (unknown ethnicity) with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and same day some time post implant the patient was noted to have hemolysis. it was noted the position of the pump and the volume status were "good." the following day care was withdrawn from the patient, and the patient expired. additional details surrounding the demise outcome is unknown and the impact as a result of the hemolysis is unknown as well. | [5 empty] | The investigation of hemolysis has been completed. clinical states that hemolysis was observed after pump implant with pump in good position and good volume. pump was not returned for investigation. data logs were investigated and there was no motor current spike or positioning issue observed. the data logs were inconclusive. therefore the root cause of the hemolysis issue was not determined since the pump was not returned and data logs were inconclusive. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
10298 | 1220648-2025-27489 | 21842204 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21842204&pc=OZD | 03/20/2025 | 04/15/2025 |  | 04/15/2025 |  |  | 03/20/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2026622102 |  | 1000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Tamponade | Pericardial Effusion | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 81 year old female patient who developed acute mitral regurgitation (mr) was intubated without improving despite being inserted with intra-aortic balloon pump (iabp), so decision was made for impella 5.5. it was noted that immediately after insertion of 5.5, lv rapidly collapsed, pericardial effusion was observed, and cardiac tamponade occurred. oozing rupture occurrence. the 5.5 was explanted in the procedural area. the patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
10299 | 1220648-2025-27491 | 21842218 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21842218&pc=OZD | 03/16/2025 | 04/22/2025 |  | 04/15/2025 |  |  | 03/16/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026614970 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Atrial Fibrillation | Atrial Flutter | Low Blood Pressure/ Hypotension | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that there was product damage. the sterile sleeve ripped and detached on the distal segment of the catheter by the red impella plug. the middle and proximal segment of the sterile sleeve was still attached and in good condition. the nurse cleaned with betadine and reinforced with tegaderms. it was later reported that the patient decompensated and coded. the impella position was checked due to concerns of placement signal. additionally, the patient experienced atrial fibrillation and atrial flutter for which the p-level was adjusted for. however, the patient remained hypotensive and resuscitation was continued. the patient subsequently expired. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of product damage and vf/vt/af/at has been completed. the pump was not returned for investigation, and a data log review was not required for this case. the patient experienced ventricular tachycardia (vt) during pump use. additionally, the sleeve had ripped and detached, but this was managed with tegaderm. the cause of the sterile sleeve damage issue was most likely use related, based on given clinical details. as the necessary information was not provided, the root cause of the vt/vf was not determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10300 | 1220648-2025-27492 | 21842281 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21842281&pc=OZD | 03/23/2025 | 05/12/2025 |  | 04/15/2025 |  |  | 03/23/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026634655 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient (pt) critically ill, in cardiogenic shock, reportedly decompensated further once arriving to the icu. decision was made for impella cp. successful impella implant. initial intrinsic pulsatility very weak, however did improve after initiating support. pt¿s pressures also improved. during support it was reported that the patient was mechanically ventilated. the pt's clotting factors had been markedly abnormal and nursing staff stated he was ¿bleeding from every line.¿ with the dic, patient required packed red blood cells, cryo, and fresh frozen plasma. this was noted to not be impella related. additionally, the patient status had improved slightly; however, staff were concerned about his neurologic status-gcs of 5 (suspect possible cva). it was noted that in the last 24 hours, patient had received 17 units of blood products. they started hemodialysis as well. the nurse noted significant bleeding from the impella groin site. an intensivist adjusted the angle of insertion slightly and then resecured it with sutures. this did appear to alleviate the bleeding issue. ultimately care was withdrawn and the patient expired. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the access site bleeding issue was most likely use related, an intensivist adjusted the angle of insertion slightly and then resecured it with sutures to achieve hemostasis.
10301 | 1220648-2025-27493 | 21842596 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21842596&pc=OZD | 03/23/2025 | 04/15/2025 |  | 04/15/2025 |  |  | 03/23/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026617686 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient had no pulse on the right leg. fem to fem bypass was done successfully. | [5 empty] | D6 explant date unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
10302 | 1220648-2025-27494 | 21842806 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21842806&pc=OZD | 03/24/2025 | 09/23/2025 |  | 04/15/2025 |  |  | 03/24/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025601378 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | Device Sensing Problem | [5 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured left frontal ischemic stroke while on impella 5.5 support. | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the impella 5.5 white cable got stuck in the ct scanner while supporting the patient. flows stable and motor current pulsatile. additionally, there was oozing at the impella insertion site that was resolved with applying a clot. it was further reported that the patient was not following commands and had an inability to move right ue. ct scan performed showing no acute changes, but chronic lacunar infarcts was noted. | The complainant also reported that the placement signal not reliable alarm was present with a loss of placement signal. flows were stable, and motor current was pulsatile. |  |  |  | The investigation of access site bleeding, pump stop, and stroke/cva has been completed. the failure mode 'pump stop' was removed as investigation revealed there was no pump stop during case. no product was returned for analysis. access site bleeding - major: clinical reported oozing at impella site resolved with clot. the root cause of the access site bleeding - major was not determined as no product was returned and insufficient clinical information related to failure was provided. stroke/cva: strokes/cvas can occur during or after impella support, and can be either ischemic or hemorrhagic. to make a determination as to the cause of the stroke, the following clinical information must be considered: -patient's medical history, including any previous strokes, cardiovascular disease, or risk factors such as hypertension, diabetes, or atrial fibrillation -timing of the stroke in relation to impella support (during or post-implant) -detailed description of the symptoms experienced by the patient -neurological examination findings and any focal deficits observed -diagnostic imaging results, such as ct scan or mri of the brain, to identify the type and location of the stroke (ischemic or hemorrhagic) -laboratory test results, including coagulation profiles and cardiac markers -any relevant procedural or device-related factors, such as duration and settings of impella support, presence of embolic protection devices, or any documented procedural complications -response to any previous anticoagulation or antiplatelet therapies -evaluation of potential alternative causes of stroke, such as arrhythmias, embolic sources, or underlying vascular pathology as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. b5 updated with additional information received from the clinical site. g2 report source; study was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27494. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. optical signal issue: clinical reported placement signal not reliable (psnr) alarms occurred following white cable getting stuck in the computed tomography (ct) scanner. the data logs confirm psnr alarms and show placement signal going out of range, signal-to-noise ratio (snr) dropping to zero, lamp going up to 100, and a decrease in gain. contrast increased in amplitude. the pattern identified corresponds to a fiber break due to patient cable torque. there was no recovery for the remainder of support. the root cause of the optical signal issue was most likely a damaged fiber line due to an interaction with another device (ct scanner) as evidenced by clinical information and data log analysis. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include placement signal issue description. corrections that were made from the initial manufacturer device report number 1220648-2025-27494. d10 concomitant medical products and therapy dates.
10303 | 1220648-2025-27495 | 21843079 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21843079&pc=OZD | 03/14/2025 | 06/26/2025 |  | 04/15/2025 |  |  | 03/17/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026625099 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  |  |  | Death | Yes | Yes | Difficult to Remove | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. once impella was explanted, removal issues were noted one of the perclose sutures ripped. the physician decided to place angioseal and once angioseal was deployed, groin bleeding did not cease. femoral angiogram was performed and a stent was deployed, but did not resolve issue. the patient was taken to operating room, where he coded. a fem-fem bypass was performed. the patient was then taken back to the icu and ultimately expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The investigation of access site bleeding and removal issues has been completed. the impella device was not returned for evaluation. removal issues: pump was not returned. data logs are not applicable as the incident took place after impella removal. the cause of the removal issue was use related due to the tearing of the perclose suture that was placed following the procedure where the impella was removed. access site bleeding: pump was not returned. data logs are not applicable as the incident took place after impella removal. the cause of the access site bleeding was use related due to the tearing of the perclose suture that was placed following the procedure where the impella was removed. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-27495. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: ¿avoid manual compression of the inlet and outlet areas of the cannula assembly.¿ section: warnings, cautions, and precautions: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
10304 | 1220648-2025-27496 | 21843092 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21843092&pc=OZD | 03/21/2025 | 05/27/2025 |  | 04/15/2025 |  |  | 03/21/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026617945 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Malposition of Device | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, positioning issues were noted. the pump was repositioned multiple times. however, the pump could not achieve optimal positioning. additionally, the patient developed limb ischemia. manipulation of the repo sheath did not improve the outcome. a recommendation for a femorofemoral bypass was rejected by the physician as the patient had a very "bad" prognosis. the device was explanted. the patient's condition following the event and survival at explant was indicated as unknown. | [5 empty] | The investigation of positioning issues and limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10305 | 1220648-2025-27497 | 21843658 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21843658&pc=OZD | 03/20/2025 | 04/15/2025 |  | 04/15/2025 |  |  | 03/20/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1734780 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. during use in the icu, the automated impella controller (aic) had "purge disc not detected" red alarm during attempt to change purge cassette. the aic was swapped for another aic, and the device was inserted and running without issue. | [5 empty] | The investigation for the reported purge disc/flag issue has been completed. the impella device was received from the customer and an investigation regarding the returned device has been completed. the evaluation of the device noted that the logs from the complaint date revealed that the console issued purge disc not detected alarm multiple times. the console was examined after arriving and a broken purge disc flag was identified. the root cause of this issue was a broken purge disc flag.
10306 | 1220648-2025-27498 | 21843682 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21843682&pc=OZD | 01/30/2025 | 04/15/2025 |  | 04/15/2025 |  |  | 03/20/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025568513 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Low Blood Pressure/ Hypotension | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | It was reported via an impact study that a patient experienced hypotension and blood loss anemia during impella cp support. the conditions were believed to have a possible relationship to the impella device and procedure. support was maintained until successful completion of the planned percutaneous coronary intervention. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10307 | 1220648-2025-27501 | 21843719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21843719&pc=OZD | 03/20/2025 | 04/15/2025 |  | 04/15/2025 |  |  | 03/20/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025605544 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 29 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support and after the implant, flows of the impella would not climb higher than 0.5 l/min. physician ordered more heparin and decided to remove pump and regain ventricular access. once pump was removed, physician found some thrombus around the outlet, and decided to use a new pump to continue support. blood products were given to the patient. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
10308 | 1220648-2025-27500 | 21843725 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21843725&pc=OZD | 03/20/2025 | 04/15/2025 |  | 04/15/2025 |  |  | 03/20/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026636732 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that the decision for impella cp was made on a patient with coronary artery disease. during support it was reported of suction alarms x3. albumin was given. additionally, it was noted that they attempted to increase p level unsuccessfully and decision was made to exchange cp for fear of device malfunction. total of (b)(4) units given for the case. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
10309 | 1220648-2025-27502 | 21843736 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21843736&pc=OZD | 03/20/2025 | 04/15/2025 |  | 04/15/2025 |  |  | 03/20/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026613429 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 34 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support and during support, the patient had massive bleeding at the impella access site. a femstop was placed at the right groin. staff was unable to run systemic anticoagulation due to massive bleeding. abdominal wound vac placed and femoral bleeding resolved. blood products were given to the patient. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10310 | 1220648-2025-27503 | 21843794 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21843794&pc=OZD | 03/20/2025 | 05/09/2025 |  | 04/15/2025 |  |  | 03/20/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025512475 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, the patient had possible sepsis. antibiotics were started. the sepsis was not confirmed though. additionally, a thrombus was wrapped around the pigtail. the pump performed fine for another 36 hours and was explanted without issue. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist. section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of thrombus and infection/sepsis has been completed. thrombus - reported seeing thrombus wrapped around the pigtail under echo on (b)(6) 2025. pump and data stick were returned. biomaterial wrapped around the impeller on the returned pump. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions. ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency) ¿relevant laboratory test results, such as coagulation profiles and platelet counts. ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states) ¿medications or treatments that the patient was receiving at the time of thrombus formation ¿surgical or procedural details if applicable (e.g, cardiac catheterization) ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events) ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. infection/sepsis - reported concerns for sepsis on 21-mar-2025. an infection/sepsis can occur during or after impella support. when diagnosing the cause of an infection in a patient, the following clinical evidence would need to be considered holistically: -patient symptoms and medical history obtained by the treating clinician -physical examination findings -results of special investigations : laboratory test results & imaging studies -microbiological culture and sensitivity results -response to previous treatments and therapies -any relevant epidemiological information or exposure history -concomitant devices in use -preexisting patient condition such as underlying infection, remote septic sources such as indwelling lines/catheters or other bioprosthetic material in situ. -care of access site because this evidence has not been provided, the root cause of the infection cannot be determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27503 as it is unknown if the initial reporter also sent the report to the food and drug administration. h6 code 4114 was reported incorrectly on the initial manufacturer device report number 1220648-2025-27503.
10311 | 1220648-2025-27504 | 21843804 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21843804&pc=OZD | 03/20/2025 | 05/07/2025 |  | 04/15/2025 |  |  | 03/20/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1457200 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  |  | White | Malfunction | No | Yes | Failure to Read Input Signal | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) was unable to recognize the impella. the aic was exchanged, and the new aic worked successfully. | [5 empty] | During the data analysis the pump connection was not recognized, indicating a rarely occurring epm crystal issue this confirmed the reported failure mode. the console was tested by power cycling the console 10 times, upon connecting the optical test pump, every time the pump was recognized in the first attempt. unable to reproduce the pump recognition issue. root cause: the pump detection issue was most likely due to the rarely occurring epm crystal issue. the impellatronic pca does not have permanent damage, but this intermittent pump detection failure occurs infrequently. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27504 as it is unknown if the initial reporter also sent the report to the food and drug administration. h6 code 2880 was reported incorrectly on manufacturer device report 1220648-2025-27504. new code was added to medical device problem code. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
10312 | 1220648-2025-27505 | 21843818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21843818&pc=OZD | 03/20/2025 | 05/12/2025 |  | 04/15/2025 |  |  | 03/20/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026620826 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support had some bleeding at the impella access site. blood products were administered. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of access site bleeding is determined to be patient movement because the bleeding was started after transfer to icu. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27505 as it is unknown if the initial reporter also sent the report to the food and drug administration.
10313 | 1220648-2025-27506 | 21843954 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21843954&pc=OZD | 03/20/2025 | 10/17/2025 |  | 04/15/2025 |  |  | 04/05/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026619134 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem |  |  |  |  | Sepsis | Tachycardia | [11 empty] | Other | Required Intervention |  |  |  |  | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The patient was having some compensatory pauses and short ventricular tachycardia (vt) runs followed by impella in aorta alarms which lasted 5-10 seconds. it was discussed how pauses can briefly affect preload and decrease motor current differential to less than 50ma which can trigger false impella in aorta alarms. | The complainant reported continuous intermittent 'impella in aorta' alarms were encountered for which a discussion was given on how they could be triggered. the false impella in aorta alarms continued to appear whenever the patient coughed or moved but would self-resolve. additionally, the patient shaded for positive blood cultures on antibiotics and line holiday. however, a few days later the patient became unshaded post resolution of infection. the patient remains on support. | It was reported that the device exhibited optical sensor issue. |  |  |  | The investigation of infection/sepsis and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b1 product problem was erroneously selected on the initial manufacturer device report number 1220648-2025-27506. b2. other serious event selection was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27506. b3. date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-27506. b5. revised with information that was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27506. e4. should have been left blank on the initial manufacturer device report number 1220648-2025-27506 as it is unknown if the initial reporter also sent the report to the food and drug administration. h6. health effect - clinical code 2095 and health effect - impact code 4614 were inadvertently omitted on the initial manufacturer device report number 1220648-2025-27506. h6. medical device problem code 1158 was erroneously entered on the initial manufacturer device report number 1220648-2025-27506. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the pump was not returned for investigation. according to the given clinical information, several self-resolving impella position in aorta alarms were reported during the case, particularly with the patient coughing and moving. the data log recorded several impella position unknown alarms, indicating the impella position in the aorta. there were no suction alarms, and no trends in motor current or placement signals were observed during the case. signal-to-noise ratio (snr) was reported fluctuating within the specification range throughout the case run, regardless of the presence of false position in aorta alarms. no other 5s parameters were affected during the case. the cause of the optical signal issue was most likely patient condition related, based on clinical details and data log review. an infection/sepsis can occur during or after impella support. when diagnosing the cause of an infection in a patient, the following clinical evidence would need to be considered holistically: patient symptoms and medical history obtained by the treating clinician physical examination findings results of special investigations : laboratory test results & imaging studies microbiological culture and sensitivity results response to previous treatments and therapies any relevant epidemiological information or exposure history concomitant devices in use preexisting patient condition such as underlying infection, remote septic sources such as indwelling lines/catheters or other bioprosthetic material in situ. care of access site because this evidence has not been provided, the root cause of the infection cannot be determined. cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. timing of the arrythmia event in relation to impella support (during or post-implant) electrocardiogram (ekg) recordings of the arrhythmia episodes. evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. response to any antiarrhythmic treatments or interventions during the event any documented device-related issues or complications during impella support evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated to include placement signal coding. d10 concomitant medical products and therapy dates updated.
10314 | 1220648-2025-27507 | 21844010 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21844010&pc=PYX | 03/21/2025 | 09/24/2025 |  | 04/15/2025 |  |  | 03/21/2025 | 04/15/2025 |  |  | 01/31/2026 |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025462186 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  |  | White | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The user facility reported that the impella rp had intermittent placement signal unreliable alarms on a 61 yea rold female patient. placement signal appeared normal for several seconds and then disappeared for several seconds. motor current remained consistent/stable. additionally, the rp encountered some intermittent suction alarms. the patient was eventually successfully weaned and explanted. | [5 empty] | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the data logs confirm alarm #1053 placement signal out of range and alarm #1055 opm invalid bench service. there were continuous suction alarms during which the signal-to-noise ratio (snr) and contrast decreased slightly but were within normal range. lamp dropped abruptly and gain increased abruptly but were also within acceptable range. during this time the placement signal and motor current were noisy. there were no motor current spikes or ingestion patterns. this occurred for about an hour after which lamp increased abruptly, and gain decreased. the clinical information provided mentions that volume was being pulled continuous veno-venous hemodialysis (cvvhd). this procedure likely caused the placement signal issues and suction alarms. the issue resolves during the case as volume is no longer pulled. the root cause of the optical signal issue was most likely patient condition related as evidenced by the issue occurrence during the clinical procedure of volume being pulled cvvhd. the data logs match with the clinical timeline as supporting evidence. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. updated h6 to contain optical signal coding. d10 therapy dates. | The investigation of low pump flow has been completed. the impella device was not returned for evaluation. the data logs show continuous suction alarms occurred during the time period when volume was being pulled continuous venovenous hemo dialysis (cvvhd). the data logs do not show any motor current spikes. this procedure likely caused the placement signal issues and suction alarms. the issue resolves during the case as volume is no longer pulled. the root cause of the low pump flow was most likely patient condition relayed as evidenced by the issue occurrence during the clinical procedure of volume being pulled cvvhd. the data logs match with the clinical timeline as supporting evidence. e2 was erroneously selected on the initial manufacturer device report number 1220648-2025-27507. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27507 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
10315 | 1220648-2025-27508 | 21845062 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21845062&pc=OZD | 03/28/2025 | 04/15/2025 |  | 04/15/2025 |  |  | 03/28/2025 | 04/15/2025 |  |  | 09/30/2025 |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024403353 |  | P140003 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Product Quality Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that a 5.5 pump being prepared for a study/impact patient. during unpacking of the device, there was a noted "user error" and the device became unsterile. the device was replaced. there was no patient harm. | [5 empty] | The investigation into the reported issue has been completed. product was not returned for review. the cause of packaging issues - sterility was most likely use related since the pump became unsterile during unpacking of the device and was stated to be a user error.
10316 | 1220648-2025-27509 | 21845367 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21845367&pc=OZD | 03/19/2025 | 05/08/2025 |  | 04/15/2025 |  |  | 03/19/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025580894 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | Yes | Positioning Failure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, the pump fell out the left ventricle. staff were unable to confirm the position of the pigtail, and the pump was explanted and extracorporeal membrane oxygenation was implanted. the patient expired. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of positioning issues has been completed. the pump was not returned for investigation. data log review was not required for this case run. according to the clinical report, the impella was pulled into the aorta during the patient's transfer from the table to the bed. the cause of the positioning issue was most likely related to use, as the pump position was compromised during the patient's transfer. b1 product problem was inadvertently omitted on the initial manufacturer device report number. h6 codes 4582 and 4628 were reported incorrectly on manufacturer device report 1220648-2025-27509. new codes were added to health effect - clinical code and health effect - impact code. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
10317 | 1220648-2025-27510 | 21845447 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21845447&pc=OZD | 03/20/2025 | 04/15/2025 |  | 04/15/2025 |  |  | 03/20/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025601445 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was noted that there were significant events of low pulsatility in which patient remained supported due to impella flows sustaining 3.5 l. decision was made to put impella to p-7 post procedure. additionally, the patient developed a quick non-pulse in left limb. no pulses were found in distal limb. the cp was successfully explanted and put on intra-aortic balloon pump. there were no pulses in foot after removal. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
10318 | 1220648-2025-27511 | 21845451 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21845451&pc=OZD | 03/20/2025 | 04/15/2025 |  | 04/15/2025 |  |  | 03/20/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1161393 |  | 0042-0000 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the automated impella controller had a "system configuration not possible, replace console immediately" error. the console was replaced and there was no patient harm. | [5 empty] | The investigation of aic - controller failure (red alarm) has been completed. the console was returned for evaluation. data analysis: the console logs showed a communication issue with the pbm (power battery manager) during the initial bootup. due to the communication issue, the aic console rebooted automatically (as designed) to attempt another power on. after rebooting, the console displayed the "system self check failed" screen, and the console was then forcefully shut down. device analysis: the reported failure mode was reproduced during testing at the field service. upon bootup, the console rebooted automatically and showed the "system self check failure" screen. visual inspection confirmed the pbm contamination which has impacted a neighboring communication channel. the root cause of the "system self check failure" was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps.
10319 | 1220648-2025-27512 | 21845453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21845453&pc=OZD | 03/20/2025 | 04/15/2025 |  | 04/15/2025 |  |  | 03/20/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026634593 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Hispanic | Native Hawaiian Or Other Pacific Islander | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Vascular Problem | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the left access side was used for impella cp placement and pci procedure. at the post 14fr sheath explant the femoral vessel was perforated requiring covered stent and/or balloon tamponade and the patient expired in the procedural area. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system and impella 5.5 with smartassist for use during section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
10320 | 1220648-2025-27513 | 21845459 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21845459&pc=OZD | 03/20/2025 | 05/20/2025 |  | 04/15/2025 |  |  | 03/20/2025 | 04/15/2025 |  |  | 11/30/2025 |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024430082 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was being implanted with an impella cp device for mechanical circulatory support and the impella cp was unable to be implanted due to the patient¿s calcified and tortuous vessel. after passing the long peel away sheath, the impella could not be advanced any further, and the case was aborted. there was no patient harm. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be patient condition because the pump was unable to pass through due to calcification in the vessels. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
10321 | 1220648-2025-27515 | 21845476 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21845476&pc=OZD | 03/20/2025 | 05/28/2025 |  | 04/15/2025 |  |  | 03/20/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026629692 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, there was a suction alarm and requested for support; recommendations were given to decrease the p-level and verify the position via echo. the physician reported that the pump was much too far in the ventricle, so pulled the impella 10 cm out of the ventricle until it was completely in the aorta; due to decrease of rr, recommendation given to start the cpr and implant a new impella. it was noted that during the complete event the impella showed the red failure mode "impella in ventricle", despite the impella being in the aorta. ultimately the patient expired on support. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. based on the clinical details the cause of the positioning issues most likely was use related due the pump was much too far in the ventricle and pull the pump about 10cm out of the ventricle. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings, cautions and precautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿ correct impella catheter position ¿for optimal positioning of the impella catheter, the inlet area of the catheter should be 3.5 cm below the aortic valve annulus and well away from papillary muscle and subannular structures. the outlet area should be well above the aortic valve.¿
10322 | 1220648-2025-27516 | 21845484 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21845484&pc=OZD | 03/20/2025 | 04/15/2025 |  | 04/15/2025 |  |  | 03/20/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025573222 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the access site had drainage/tissue exposure around the blue suture hub. antibiotics were given due to positive blood and tissue cultures. additionally, there was oozing at the access site. bival was withheld and support continued. | [5 empty] | D6 explant date: unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock & impella cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
10323 | 1220648-2025-27517 | 21845522 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21845522&pc=OZD | 03/20/2025 | 05/23/2025 |  | 04/15/2025 |  |  | 03/20/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026625687 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 31 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support.. while on support, pump stop occurred. attempts to restart were unsuccessful. the impella was exchanged for another impella cp. no patient harm was reported. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cp section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation of pump stop has been completed. the cut pump was returned for investigation. the data log was not returned for this case run. no physical damage was noted on the returned product. the pump was soaked in warm water, and some biomaterial was noted on the impeller and purge gap; however, no further testing can be performed due to the cut catheter on the returned pump. the cause of the pump stop issue was most likely biomaterial, based on returned product investigation
10324 | 1220648-2025-27518 | 21845525 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21845525&pc=OZD | 03/20/2025 | 04/15/2025 |  | 04/15/2025 |  |  | 03/20/2025 | 04/15/2025 |  |  |  |  | 04/15/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026629965 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 32 YR |  | Female |  | Hispanic |  | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the shaft of the catheter was nicked inside the pocket, and bleeding at the red plug was occurring. bone wax and tegaderm applied to site that is bleeding until the impella 5.5 was replaced. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10325 | MW5168975 | 21833517 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21833517&pc=OZD | 12/11/2024 | 04/08/2025 |  | 04/14/2025 |  |  |  | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CATH DEVICE |  |  | OZD | [5 empty] | ABIOMED, INC. | [8 empty] | Device was returned to manufacturer | [5 empty] | No | Health Professional |  |  | Voluntary report | Yes | Risk Manager | [6 empty] | 88 YR |  | Female | 67 KG | Non Hispanic |  | Death | Yes | No | Material Frayed | [6 empty] | Heart Failure/Congestive Heart Failure | [12 empty] | Other | Death | [27 empty] | Initial submission | [9 empty] | Impella controller began to alarm, line traced and did find wire fraying. device pulled; pt expired but not attributed to device. pt had signs of heart failure before alarms. no event isolated that caused damage to line. line was taken back to abiomed for research.
10326 | 1220648-2025-27418 | 21836530 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836530&pc=OZD | 03/14/2025 | 04/14/2025 |  | 04/14/2025 |  |  | 03/21/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025575016 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Shock | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the patient was on impella cp support and during support, a hemothorax of unknown cause was confirmed. the bleeding could not be controlled, and the patient died of hemorrhagic shock before being removed. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
10327 | 1220648-2025-27424 | 21836534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836534&pc=OZD | 03/16/2025 | 05/06/2025 |  | 04/14/2025 |  |  | 03/16/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025599973 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the optical sensor had gone out and the staff was educated on how to monitor the flows and motor current. the pump was weaned to p3 and the plan was to continue weaning until explant the following morning. the pump was explanted shortly afterwards. | [5 empty] | The investigation of optical signal issue has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10328 | 1220648-2025-27430 | 21836540 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836540&pc=OZD | 03/17/2025 | 04/14/2025 |  | 04/14/2025 |  |  | 03/17/2025 | 04/14/2025 |  |  | 03/22/2024 |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025492044 |  | 1000482 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  |  |  | Malfunction | No | Yes | Positioning Failure | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the impella 5.5 was dislocated into the aorta by the awakening patient. additionally on support there was high purge pressure and low purge flow 1h after implantation. the 5.5 was removed and replaced with a new impella 5.5. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
10329 | 1220648-2025-27428 | 21836541 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836541&pc=OZD | 04/29/2023 | 04/14/2025 |  | 04/14/2025 |  |  | 03/16/2025 | 04/14/2025 |  |  | 03/31/2025 |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024327966 |  | 0550-0002 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complaints team received an impact study that found that the patient had acute kidney injury while on impella 5.5 support. the study found that there was a possible relationship between the impella device and the impella procedure and the patient was put on dialysis. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10330 | 1220648-2025-27433 | 21836546 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836546&pc=OZD | 03/17/2025 | 04/14/2025 |  | 04/14/2025 |  |  | 03/18/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025578916 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a decision was made to implant an impella cp after a patient was found to have multivessel disease. it was reported that the groin site began bleeding and the nurse tried femstop with no success. the medical doctor opted to remove the device. the patient was stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10331 | 1220648-2025-27421 | 21836547 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836547&pc=OZD | 04/01/2025 | 04/14/2025 |  | 04/14/2025 |  |  | 04/04/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025549019 |  | 1000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient passed away from an infection during impella 5.5 support. the source of the infection is unknown. as such, there is not enough evidence to disassociate the patient's passing with impella support. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿. section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
10332 | 1220648-2025-27438 | 21836551 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836551&pc=OZD | 03/18/2025 | 05/08/2025 |  | 04/14/2025 |  |  | 03/18/2025 | 04/14/2025 |  |  | 01/31/2026 |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025461120 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. there was a brief alarm stating, ¿high motor current, impella stopped¿ without changes in waveforms and immediately the device restarted on own with normal waveforms. after about 15 minutes while still weaning off the alarm occurred with the pump stopping and no changes in waveforms. the pump was restarted at p2 and increased back up to p6 without issues. all alarms resolved and normal waveforms with pump at p6. there was no patient harm. | [5 empty] | The investigation of temporary pump stop has been completed. the impella device was not returned for evaluation. the root cause of the temporary pump stops could not be definitively determined as no product was returned for evaluation. a5 should have been left blank on the initial manufacturer device report number 1220648-2025-27438 as the ethnicity was unknown. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock entering cardiac output (impella cp with smartassist) section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
10333 | 1220648-2025-27426 | 21836554 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836554&pc=OZD | 03/17/2025 | 05/28/2025 |  | 04/14/2025 |  |  | 03/17/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1142852 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The automated impella controller (aic) had a controller error alarm with placement signal issues. this issue occurred during patient use, but resolved and support continued. staff monitored. there was no patient harm. | [5 empty] | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. the console was received for investigation. the data analysis confirmed the reported failure mode. the root cause for the controller error is due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. d9 device returned to manufacturer date added. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10334 | 1220648-2025-27439 | 21836555 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836555&pc=OZD | 03/17/2025 | 05/05/2025 |  | 04/14/2025 |  |  | 03/17/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025611194 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Material Split, Cut or Torn | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the pump was pulled out of the left ventricle and the sterile sleeve tore, and the staff explanted the pump. the patients outcome was stable. | [5 empty] | The investigation of product damage (sleeve damage) and positioning issues has been completed. the impella device was not returned for evaluation. data review: data logs confirmed pump was in aortic area (about 16 hours into the case). pump ran 4 more hours then was stopped manually and disconnected from the console. product was not returned for review. the root cause of positioning issue was most likely patient movement related since it happened during transferring the patient. the root cause of product damage (sleeve damage) was unable to be determined as limited clinical details were provided and no product was returned for review. h5 was erroneously selected on the initial manufacturer device report number 1220648-2025-27439. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10335 | 1220648-2025-27429 | 21836558 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836558&pc=OZD | 03/17/2025 | 11/18/2025 |  | 04/14/2025 |  |  | 03/17/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026621521 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that they encountered high purge pressure and a pump stop. the purge system was changed independently without success. recommendation was given to monitor the motor current and to change the pump. lysis therapy was also discussed and the protocol was forwarded. however, the motor current became high and the pump suddenly stopped. the pump was unable to be restarted and the patient remained stable. | [5 empty] | The investigation has been completed. pump stop: clinical details report that on 18/mar/2025, a pump stop occurred. the pump was unable to be restarted, so it was explanted. data logs were reviewed and there were actually several temporary high motor current pump stops prior to the reported stop. data logs showed the first stop to occur on 17/mar/2025. at this time, the mc spiked to 1451 ma, and there was a slight spike up in purge pressure that continued to trend up for the following 6 hours until a high purge pressure alarm triggered. the pump was able to be restarted, but there continued to be many spikes in the motor current for the rest of the case, sometimes causing temporary pump stops, depending on the duration of the spike. on the 18th, the pump stopped in similar fashion, was able to restart, but stopped again after less than a minute and was disconnected. returned product was investigated, and when connected to a console and attempted to run, high purge pressure and a high motor current pump stop reproduced. when visually inspecting, evidence of bearing wear was seen, as the gap between the motor housing sleeve bearing and impeller was enlarged. the pump was sent out for ct scans, which showed that the motor shaft/ball bearing inner race had slipped away from the ball bearing outer race, and all the balls were floating freely. finally, there was noted to be an uneven wear pattern on the peek retainer of the ball bearing. based on the data log review and returned product showing bearing wear, the root cause of the pump stop was determined to be bearing wear. the reported pump stop occurred on the 4th day of impella support. temporary pump stop: as mentioned above, investigation of data logs found that there was an unreported temporary pump stop on 17/mar/2025. for that reason, and additional failure mode was added. per the investigation above, the cause of the temporary pump stop was determined to be due to bearing wear. the temporary pump stop occurred on the 3rd day of impella support. high purge pressure: clinical details report that a high purge pressure alarm triggered on 17/mar/2025. the purge cassette was replaced but didn't help, and tpa was added to the purge, but it's unknown if this resolved the issue. data logs were reviewed and the high purge pressure alarm was confirmed on the reported date. at the time of the first high motor current temporary pump stop, purge pressure began to rise for 6 hours before triggering a high purge pressure alarm. it remained elevated for 22 hours before stepping down back to ~600 mmhg. fluctuations in the purge pressure for the rest of the case aligned with motor current spikes. returned product was investigated and the only visual abnormality was the noted bearing wear. a window was cut in the catheter just proximal to the motor housing, and no blood ingress was noted. the pump was connected to a console and purged overnight, and high purge pressure reproduced. it resolved when cutting the catheter in the same location as the window. when the pump was sent out for ct scan, the imaging showed the motor shaft/magnet to have slipped up, contacting the bottom of the sleeve bearing. this suggests that when the bearing wear occurred, this blockage of the purge fluid flow path out of the motor housing caused the high purge pressure. based on the high purge pressure aligning with the initial motor current spike in data logs and returned product reproducing the issue and with the motor magnet contacting the sleeve bearing, the high purge pressure was determined to most likely be due to bearing wear. d10 concomitant medical products and therapy dates updated. | As noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ a.2 age should have been left blank on the initial manufacturer device report number 1220648-2025-27429 as it is unknown. a.3 unknown should have been selected on the initial manufacturer device report number 1220648-2025-27429 as it is unknown. e.1 revised facility name as it was entered incorrectly on the initial manufacturer device report number 1220648-2025-27429. g.1 revised mdr reporting contact email as it was entered incorrectly on the initial manufacturer device report number 1220648-2025-27429. h.6 codes 4118, 3221 and 4315 were reported incorrectly on the initial manufacturer device report number 1220648-2025-27429. new codes were added. h.10 the instructions for use were entered incorrectly on the initial manufacturer device report number 1220648-2025-27429 and wer revised. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
10336 | 1220648-2025-27425 | 21836559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836559&pc=OZD | 02/23/2025 | 04/14/2025 |  | 04/14/2025 |  |  | 03/19/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025588668 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that there was an unspecified death. a new impella was inserted after the original one was accidently removed during transportation. after impella cp insertion, ventricular septal perforation (vsp) was found by echocardiogram and surgery was planned, but the patient's overall condition deteriorated and they passed away from severe heart failure. the patient was already in a serious condition and the relationship to impella is unknown. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
10337 | 1220648-2025-27446 | 21836562 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836562&pc=OZD | 03/26/2024 | 04/14/2025 |  | 04/14/2025 |  |  | 03/18/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1125047 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported that the placement signal intermittently disappeared on the automated impella controller (aic) during support for a patient. it was noted that a controller error alarm was present each time the placement signal disappeared. support remained ongoing with no indication that intervention was required. there was no patient harm. upon investigation findings, it was discovered that the event involved an optical bench firmware communication protocol anomaly. | [5 empty] | The investigation for the controller alarm issue has been completed. console logs from the reported day of event are consistent with complaint and showed several instances of loss of opm data and - in case the fault condition lasted more than 2 minutes - associated alarms ¿controller error¿. further analysis of device logs revealed that the rlm was frequently rebooting during prior cases, which resulted in intermittent connection and data transmission to the cloud (impellaconnect). the console was returned for evaluation which revealed that purge disk detect flag was stuck under the disk retainer (which also had a broken mounting post), resulting in console being unable to detect purge disk being removed. we also observed hard-to-remove contamination of optical pump connector (unrelated to loss of optical data). testing of the rlm did not reveal any irregularities, and the unit operated within specification (field issues were not reproduced). neither the purge nor rlm issue were reported to abiomed and are unrelated to the original complaint. the controller error and loss of placement signal is due to an optical bench firmware communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the automated impella controller (aic) software operated as designed. the cause of the aic issue was a software issue.
10338 | 1220648-2025-27431 | 21836564 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836564&pc=PYX | 03/17/2025 | 05/06/2025 |  | 04/14/2025 |  |  | 03/17/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026614335 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Leak/Splash | Adverse Event Without Identified Device or Use Problem | Positioning Problem |  |  |  |  | Hemorrhage/Blood Loss/Bleeding | Sepsis | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was further reported that on day four of impella rp flex support a chest x-ray was completed, and rp was repositioned. the next day, the impella had intermittent purge pressure low alarms. troubleshooting was done. on day 14, the patient was taken to the operating room for rp explant as there was concern for sepsis. all lines were also removed. | The complainant reported that the patient on impella rp support developed bleeding at the rp site. pressure dressing, additional gauze, and angle matching was attempted. new suture placed. |  |  |  |  | The investigation of purge leak ¿ pump, positioning issues, infection/sepsis, and access site bleeding has been completed. unable to run the pump in the lab due to open lines. the root cause of the purge leak - pump was not determined since no leak was reproduced on the returned product and returned pump could not be tested due to observed physical abnormality that was not present in the field. the root cause of the positioning issue was not determined due to insufficient clinical details. the root cause of the access site bleeding and sepsis was not determined due to insufficient clinical details. b1 revised to reflect both product problem and adverse event. b2 revised to reflect intervention. b5 updated with additional information received from the clinical site. h6 updated to reflect additional failure modes reported by the clinical site. instructions for use: impella rp and rp flex impella; rp smart assist system with automated impella controller and rp flex with smart assist system with automated impella controller section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (eg, infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir.¿ section: potential adverse events (united states). ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: use of the repositioning sheath and the 23 fr peel-aways introducer. ¿to prevent contamination and subsequent infections, always use sterile technique on the insertion site. follow institutional protocols for prophylaxis of infection for patients on ventricular assist devices and indwelling lines, as well as protocols for surveillance of such patients. if device-related infection occurs, consider each patient¿s clinical circumstances when deciding whether to continue impella support.¿ section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental mdr will be filed. as noted in the impella ifu: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter. use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿
10339 | 1220648-2025-27432 | 21836565 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836565&pc=OZD | 03/17/2025 | 04/14/2025 |  | 04/14/2025 |  |  | 03/17/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025606056 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Delivered as Unsterile Product | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that they encountered packaging issues. upon opening the 14fr introducer kit, moisture was noted inside of the clear plastic sheath protector. doctor was concerned and another 14fr sheath kit was obtained from another impella kit. prior to opening that introducer set, moisture was noted in that kit as well in the same clear protector. an intra-aortic balloon pump (iabp) was placed in the patient and hospital team was reluctant to use the impellas until documentation and/or replacement catheters were provided. there was no patient harm as a result of delay in placement. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
10340 | 1220648-2025-27423 | 21836567 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836567&pc=OZD | 03/16/2025 | 09/30/2025 |  | 04/14/2025 |  |  | 03/16/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025550400 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant also reported that there were placement signal issues. staff reports that when the white cable was unplugged and reinserted, a red placement signal appeared normal for a short time but then the alarm resumed. support continued. | The complainant reported the patient was on impella cp support and the patient reported losing feeling in her impella leg. the medical doctor inserted reperfusion sheath and symptoms improved. it was further reported that the patient coded and expired. no further information provided. the relationship between the impella and cause of death is unknown. |  |  |  |  | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed optical signal issue and limb ischemia. it was stated that there were placement signal no reliable alarms and that the pump was unplugged from the console and reinserted, the placement signal alarms resolved but a short time but then the alarm resumed. due to lack of product and data logs returned, the root cause of the optical signal issue cannot be determined. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include placement signal issue description. h6 updated to include placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-27423. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states) - ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
10341 | 1220648-2025-27427 | 21836571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836571&pc=OZD | 03/17/2025 | 04/14/2025 |  | 04/14/2025 |  |  | 03/17/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025581896 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced access site bleeding after admission to icu. the pump was removed due to bleeding. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10342 | 1220648-2025-27419 | 21836574 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836574&pc=OZD | 04/01/2025 | 07/16/2025 |  | 04/14/2025 |  |  | 04/01/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Positioning Failure | [6 empty] | Aortic Valve Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Abiomed received a literature article titled "a case of lmpella-lnduced aortic regurgitation that offset the lmpella flow" that stated a patient was on impella 5.5 pump for support. the pump positioning was contributing to/causing heart valve damage in the form of aortic regurgitation (ar). the aortic leaflets were disrupted by the pump positioning. after the removal of the impella, ar was improved, suggesting it was secondary to malcoaptation of the aortic leaflets. the patient was discharged from intenstive care unit and was stable. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. a.1, a.4 and a.5 are unknown. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. h.4 device manufacture date is unknown. literature citation is unknown. | The investigation of heart valve damage and positioning issues has been completed. the impella device was not returned for evaluation. heart valve damage: data logs were not returned as well. therefore, the root cause of the heart valve damage issue was not determined since the clinical information about the damage is insufficient and pump and data logs were not returned. positioning issue: the root cause of the positioning issue was most likely patient condition related patients' anatomy since doctors were able to improve the ar by removing the impella position since impella was interacting with the cardiac morphological characteristics with the angular bend between ascending aorta and left ventricle and the damaged aortic leaflets.
10343 | 1220648-2025-27434 | 21836576 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836576&pc=OZD | 03/18/2025 | 04/14/2025 |  | 04/14/2025 |  |  | 03/18/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025607877 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | Death |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the decision was made to place an impella 5.5 on a patient with multi vessel disease. the impella 5.5 was placed per protocol without issue. during support, the patient experienced significant abdominal pain followed by a lactate of 16 requiring increases in chemical support. computed tomography imaging revealed a thrombus in abdominal aorta that was occluding the superior mesenteric artery (sma) and partially occluding the celiac artery. additionally, there was excessive bleeding from the impella site with coagulopathy/shock liver requiring blood products. the patient was taken to operating room emergently for exploratory laparotomy. necrotic ascending and transverse colon resected followed by sma thrombectomy. additional area of bowel was removed. the belly was left open with wound vac in place. urine output had picked back up since volume resuscitation. the impella pocket was washed out and re-closed. it was also noted that the patient had suction alarms that were volume related. more blood products and volume were administered and suction alarms resolved. additional packed red blood cells and plasma were also being administered. ultimately, the family withdrew care and the patient expired. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
10344 | 1220648-2025-27441 | 21836578 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836578&pc=OZD | 03/17/2025 | 04/30/2025 |  | 04/14/2025 |  |  | 03/17/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025561190 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. on day one of impella 5.5 support, the patient was taken to the operating room for bleeding from the site. the bleeding was found around the pectoralis muscle and tissues. the team intervened and achieved hemostasis. all anticoagulants were held. additionally, the patient developed hemolysis and the impella 5.5 was repositioned. the patient survived the procedure. | [5 empty] | The investigation of access site bleeding and hemolysis has been completed. hemolysis occurred on the 2nd day of support & resolved by repositioning of the pump. hemolysis reoccurred intermittently few more times and resolved by repositioning of the pump. clinical team confirmed hemolysis was not related to impella. data review: data logs showed pump ran without issue during support. there were suction alarms triggered early of the case but resolved quickly. product analysis: pump was returned for analysis. visual inspection found no issues on the pump. no biomaterial was found inside the cannula and pump housing. the root cause of access site bleeding was unable to be determined as limited clinical details were provided and no problem was found on the pump. the root cause of hemolysis was most likely due to pump was not in the proper position as hemolysis was resolved by repositioning of the pump per clinical team and no problem was found on the pump. | The pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿. impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿. section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions, including catheter position, pre-existing medical conditions, and small left ventricular volumes, may also play a role in patient susceptibility to hemolysis.¿. section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿.
10345 | 1220648-2025-27436 | 21836581 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836581&pc=OZD | 03/15/2025 | 04/14/2025 |  | 04/14/2025 |  |  | 03/16/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025562871 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Ventricular Fibrillation | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient on impella 5.5 with a history of cerebrovascular accident (cva) was taken for a head computed tomography and found a new ischemic cva. heparin was held. additionally, the patient experienced ventricular fibrillation arrest that required 33 shocks. care was ultimately withdrawn and the patient expired. medical safety review of the event was completed, there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. although the demise is unlikely for the impella pump given anticipated misuse: patient post-surgery with activated clotting time subtherapeutic for impella use due to the bleed potential. | [5 empty] | The investigation has been completed. no product was returned. the root cause of stroke and ventricular fibrillation cannot be determined since the product was not returned and insufficient clinical details were provided. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10346 | 1220648-2025-27448 | 21836582 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836582&pc=OZD | 03/17/2025 | 04/22/2025 |  | 04/14/2025 |  |  | 03/17/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1129854 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the automated impella controller (aic) had a controller error alarm. there was no patient harm reported. | [5 empty] | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. console logs from the reported day of event are consistent with complaint and show that 15 minutes after case start, console presented with controller error alarm, immediately preceded by error message ¿scu: v error in module: 64¿ that translates to ¿bad voltage: power battery manager (pbm) 2: vpurge¿. approx. 80 seconds earlier, console had been plugged into wall ac. the controller error alarm self-resolved after 4 minutes, but 5 minutes later pump was disconnected and transferred to backup console. during the alarm condition, hemodynamic support was unaffected. during data analysis the measurement of pud voltage delivered from pbm to pud revealed that switching from nominal 16v (aic on batteries) to nominal 24v (aic on ac), produced occasional delays when voltage was approx. 17v. it is most likely that the system recognized ac being plugged in, but voltage was lower than 24v, which triggered ¿pbm voltage out of spec¿ alarm due to pud undervoltage. d4 lot number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27448. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
10347 | 1220648-2025-27444 | 21836586 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836586&pc=PYX | 03/17/2025 | 04/14/2025 |  | 04/14/2025 |  |  | 03/17/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025503354 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, there was noted bleeding and manual pressure was applied. the physician returned to place additional suture and hemostasis was achieved. the site was dressed with gauze, clear dressing, and modified pressure dressing. the impella cp was successfully weaned and explanted. however, it was noted that the site was allowed to bleed very briefly before the physician applied manual pressure and maintained the site and controlled bleeding for 10 minutes. then the nurse also applied pressure for 25 minutes. hemostasis was achieved. the site was cleaned and clear dressing was applied. the procedure was successful. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿
10348 | 1220648-2025-27452 | 21836590 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836590&pc=OZD | 03/04/2024 | 04/14/2025 |  | 04/14/2025 |  |  | 03/18/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1183444 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that a placement signal unreliable alarm appeared following impella 5.5 implant. it was noted that the delivery had been difficult due to the patient's anatomy. the alarm was silenced and support with the pump and console was maintained. upon conclusion of the investigation, a review of the logs showed that the automated impella controller had been swapped immediately following the event. in addition, an optical bench firmware communication protocol anomaly was identified. there was no patient harm. | [5 empty] | The investigation for the reported issue has been completed. console logs from the reported day of event are consistent with the complaint and showed placement signal not reliable alarms and controller error alarms, as well as multiple messages indicating ¿invalid sample due to snr below minimum¿. real time log data showed that signal-to-noise ratio (snr) is pulsatile (varying from 2500 to 0) and correlated with raw optical placement data, and also that placement signal is ¿barcoding¿ when snr is low. inspection of the console revealed that the optical pump connector was heavily contaminated, which would have contributed to lowering the snr. however, this would not explain the pulsatility of snr. analysis on the analog output of the optical bench signal revealed a defect on the ccd, which caused placement signal to be unreliable for certain values of placements signal, by locally lowering the snr. the controller error and loss of placement signal is due to an optical bench firmware communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic software operated as designed. this issue was unrelated to the ccd defect. the cause of the aic issue was a defective optical bench. the cause of the aic issue was optical bench firmware communication protocol anomaly.
10349 | 1220648-2025-27437 | 21836591 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836591&pc=PYX | 03/17/2025 | 05/06/2025 |  | 04/14/2025 |  |  | 03/17/2025 | 04/14/2025 |  |  | 06/30/2025 |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024369366 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was being implanted with an impella rp device for mechanical circulatory support. during impella rp placement, the impella rp was not able to be advanced from the left femoral valve across the pulmonic valve. after several unsuccessful attempts, the impella rp placement was aborted, and veno-arterial extracorporeal membrane oxygenation was placed. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since impella to unable advance across valve due to patient's anatomy. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist: section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
10350 | 1220648-2025-27447 | 21836596 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836596&pc=OZD | 03/19/2025 | 09/24/2025 |  | 04/14/2025 |  |  | 03/19/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026623847 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 was inserted successfully but increased ectopy noted after impella placement. the patient has a history of non-sustained ventricular tachycardia. lidocaine bolus administered and impella repositioned. ectopy resolved and patients remains on support. | The customer reported a placement signal not reliable (psnr) issue. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. no product was returned for evaluation. optical signal issue: data logs were reviewed and first lt log of case showed psnr alarms occurring. at time of psnr alarms, snr dropped to 0, gain and light both decreased, and contrast was observed to be oscillating. oscillating contrast occurred twice during support and all optical parameters resolved in both instances. upon review of the data logs, it is apparent that the console is the potential cause of the issue. please refer to complaint (b)(4) for an investigation into the console. no problem was found with the pump during the case. vf/vt: cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: ¿patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease ¿timing of the arrythmia event in relation to impella support (during or post-implant) ¿electrocardiogram (ekg) recordings of the arrhythmia episodes ¿evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances ¿assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements ¿presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities ¿response to any antiarrhythmic treatments or interventions during the event ¿any documented device-related issues or complications during impella support ¿evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand ¿review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. the cause of the ectopy was unable to be determined due to insufficient clinical details. the failure mode vt/vf/at/af is used as a category for injuries including arrhythmia, asystole, bradycardia, cardiac arrest, heart block, paroxysmal atrial fibrillation, tachycardia, ventricular fibrillation, and ventricular premature complexes. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical signal issue "2917" as part of a retrospective review of optical signal issues in accordance with fda recommendation.
10351 | 1220648-2025-27422 | 21836604 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836604&pc=OZD | 03/18/2025 | 05/15/2025 |  | 04/14/2025 |  |  | 03/18/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1604236 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 23 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that in preparation of an impella placement, an alarm occurred during the automated impella controller (aic) self-check. the aic was turned off and on twice. however, the controller failure alarm occurred. the aic was replaced. there was no patient harm reported. | [5 empty] | Additional information provided in h6 and h10. the investigation into the reported issue has been completed. the console was booted up on march 17, 2025. immediately after boot-up, the log recorded abnormal optical parameters: light: 0.60v, snr: 2.5. the console then displayed ¿controller error (defective optical sensor lamp) ¿ alarm,¿ event # (b)(6). this symptom was likely misreported as ¿alarm occurred during aic self-check.¿ the console was shut down and restarted two more times. each time, the console displayed event # (b)(6). this aligns with the claim that ¿the aic was turned off and on twice,¿ but the ¿controller error¿ was misreported as ¿controller failure.¿ the console was tested after arriving at fs, but the ¿controller error¿ could not be reproduced. the bulb power was 4.42 w, and the 5s parameters were within specifications. the voltage on the pbm at connector j3 was 4.9v, while the average voltage on the optical bench connector p4 was 4.8v as intended. the root cause of the controller error could not be determined due to the inability to reproduce it. it could be due to either lamp malfunction or optical bench malfunction. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
10352 | 1220648-2025-27450 | 21836609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836609&pc=OZD | 03/17/2025 | 05/01/2025 |  | 04/14/2025 |  |  | 03/17/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025608558 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Delivered as Unsterile Product | Device Contaminated at the User Facility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the pump was contaminated as the scrub tech dropped the impella 5.5 on floor during prep. the pump was not used for the procedure. there was no patient harm. | [5 empty] | The investigation of packaging issues- sterility has been completed. the impella device was not returned for evaluation. based on the clinical details the cause of the packaging issues - sterility was use related as the device became contaminated and unusable due to being dropped outside of the sterile field, on the floor, by the scrub tech. h6 medical device problem code 1421 was erroneously entered on the initial manufacturer device report number 1220648-2025-27450. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
10353 | 1220648-2025-27443 | 21836615 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836615&pc=OZD | 03/17/2025 | 04/21/2025 |  | 04/14/2025 |  |  | 03/17/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026621480 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the patient had a hematoma at the access site. hemoglobin/hematocrit/platelet count were noted to have dropped. there was no active bleeding. two units of blood products were given to the patient. the patient was successfully weaned. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10354 | 1220648-2025-27420 | 21836617 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836617&pc=OZD | 04/01/2025 | 04/16/2025 |  | 04/14/2025 |  |  | 04/01/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 0 DA |  | Unknown |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombosis/Thrombus | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Abiomed received information through submitted literature "proceedings of the japanese society of intensive care medicine (2024; vol.31. nosuppl.1,s1051-)" that a patient experienced multiple adverse events during impella support. it was initially noted that a thrombus was discovered via computed tomography scan in the descending aorta around the impella shaft. anticoagulation therapy was strengthened to treat the condition. the patient was then found to have suffered a cerebral hemorrhage. when the original thrombus was found to have shrunk, the decision was made to remove the pump. on a later date, a bare stent was placed in the proximal superior mesenteric artery and the procedure was completed, but the patient's multiple organ failure progressed and the patient passed away. there is not enough evidence to disassociate the patient's passing and impella support. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. e.1 name prefix/title, first name and last name are unknown. h.4 device manufacture date is unknown. literature was not available to attach to this report. | As noted in the impella instructions for use: impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ h.10 added instructions for use as they were inadvertently omitted from the initial manufacturer device report number 1220648-2025-27420.
10355 | 1220648-2025-27442 | 21836619 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836619&pc=OZD | 03/06/2025 | 05/28/2025 |  | 04/14/2025 |  |  | 03/17/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1122979 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on an automated impella controller (aic) for mechanical circulatory support. while on support, the screen froze at the conclusion of the case when the staff selected p-0. the staff had to disconnect the pump in order to get it to stop. while attempting to open diag flows dropped significantly, cardiopulmonary resuscitation was initiated and the patient expired. the patient expired for a reason/cause that is unrelated to the aic event. medical safety review of the event noted that the issue with the aic occurred at the end of the case. the patient had already expired. it appears the patient was supported uninterrupted, and it appears without any delay given the information at hand. | [5 empty] | The investigation of automated impella controller (aic) - display issue has been completed. the reported complaint of the screen froze and need to disconnect pump in order get it to stop was not determined as not able to reproduced. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. medical safety review of the event noted that the issue with the automated impella controller occurred at the end of the case. the patient had already expired.. it appears the patient was supported uninterrupted, and it appears without any delay given the information at hand.
10356 | 1220648-2025-27451 | 21836620 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836620&pc=OZD | 03/17/2025 | 04/14/2025 |  | 04/14/2025 |  |  | 03/18/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025594117 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient developed a hematoma at their insertion site during impella 5.5 support. two units of blood were transfused and a surgical evacuation with drain placement was performed to manage the condition. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10357 | 1220648-2025-27445 | 21836622 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836622&pc=OZD | 03/19/2025 | 04/14/2025 |  | 04/14/2025 |  |  | 03/19/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025591533 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Native Hawaiian Or Other Pacific Islander | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that low pump flow was encountered during impella cp support. it was noted that there was an acute drop in flows at support levels above p-2 coinciding with reports of suction. the device was repositioned multiple times in conjunction with ventricular tachycardia (vt) events. upon further inspection, an echogenic mass was noted on the inlet of the pump. the pump was subsequently replaced as a result of the noted issue. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
10358 | 1220648-2025-27449 | 21836623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836623&pc=OZD | 03/20/2025 | 04/14/2025 |  | 04/14/2025 |  |  | 03/20/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1603745 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Sense | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that intermittent placement signal not reliable alarms occurred. it appeared that the placement signal was currently working, but it was noted that there were several placement signal not reliable alarms that resolve with adjusting the p-level. per the investigation results the placement signal issue is now reportable with the aware date of (b)(6) 2025 when the investigation was approved. | [5 empty] | The investigation has been completed. the controller was returned for investigation. the failure mode was reproduced while running a pump. despite having 5v into the lamp, the lamp was outputting no light causing a placement signal not reliable with oscillating contrast. before the lamp would no longer turn on, the 5s was taken and contrast is low compared to pm proceduce. the lamp was measured to have 5v into it. the oscillating contrast and unreliable placement signal were due to a bad lamp.
10359 | 1220648-2025-27455 | 21836626 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836626&pc=OZD | 03/17/2025 | 04/25/2025 |  | 04/14/2025 |  |  | 03/17/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025599873 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, there were purge blocked alarms. tissue plasma activator was added to the purge and the staff monitored the patient. there was no patient harm reported. | [5 empty] | The investigation of high purge pressure has been completed. the pump was returned for investigation. a data log review was not provided for this case. no physical damage was noted on the returned product. the pump was soaked and tested with a sample purge cassette, and a high alarm for purge pressure was observed after priming. the pump was further cut near the motor housing, and purge fluid was seen coming from the catheter, which proves that there was no blocking in the purge line. furthermore, the outflow case was torn down, and evidence of biomaterial was observed in the purge gap. the cause of the high purge pressure issue was most likely biomaterial based on returned product investigation. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
10360 | 1220648-2025-27454 | 21836628 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836628&pc=OZD | 03/17/2025 | 05/12/2025 |  | 04/14/2025 |  |  | 03/17/2025 | 04/14/2025 |  |  | 08/31/2025 |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024380586 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 40 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Perforation | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, during the ventricular assist device (vad) procedure following a concomitant tricuspid repair: the physician lifted the heart up and saw the tip of the impella 5.5. the impella had punctured the left ventricle (lv) wall. the physician removed the 5.5, repaired the perforated lv and continued with the vad placement. the physician stated that the heart muscle was inflamed and friable. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system and impella 5.5 with smartassist for use during section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist system and impella 5.5 with smartassist for use during cardiogenic shock & impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation of ventricle perforation has been completed. the pump was not returned for investigation. data log review was not required for this case run. as per the clinical, left ventricle (lv) perforation occurred during heart manipulation intraoperatively when the device prolapsed into the heart during the ventricular assist device (vad) procedure. the physician repaired the lv before continuing with the vad placement. the cause of the ventricle perforation issue was most likely use related, based on given clinical details.
10361 | 1220648-2025-27435 | 21836631 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836631&pc=OZD | 03/16/2025 | 04/21/2025 |  | 04/14/2025 |  |  | 03/17/2025 | 04/14/2025 |  |  | 05/31/2025 |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024351825 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was on impella 5.5 support and during support, the patient had bleeding at the insertion site while ambulating. quick clot was applied to the site. re-secured catheter not as tight with three point fixation. support was continued and the patients outcome was stable. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the access site bleeding issue was most likely patient movement, patient bleeding at insertions site after ambulation. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10362 | 1220648-2025-27453 | 21836633 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836633&pc=OZD | 03/17/2025 | 04/14/2025 |  | 04/14/2025 |  |  | 03/17/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025604557 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, upon placing the repositioning sheath, a hematoma developed at the impella cp insertion site. the angle was adjusted to manage the condition, but the hematoma continued to grow. manual pressure was initially applied, and hemostasis was achieved via surgical intervention. following completion of support and removal of the device, the patient developed a cerebral hemorrhage. the cause of the hemorrhage was unknown. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿. section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. impella cp with smartassist system instructions for use for the related event are as follows: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿. section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿.
10363 | 1220648-2025-27440 | 21836638 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836638&pc=PYX | 03/17/2025 | 04/14/2025 |  | 04/14/2025 |  |  | 03/17/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025590540 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella rp device for mechanical circulatory support. while on support, the patient had hemolysis. blood products were given. the hemolysis was never completely resolved before explantation. the procedure outcome was successful with this device. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
10364 | 1220648-2025-27459 | 21836643 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836643&pc=OZD | 03/17/2025 | 04/14/2025 |  | 04/14/2025 |  |  | 03/17/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026618607 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, there was access site bleeding, noted by the oozing. the site was redressed, pressure applied and two units of blood were administered. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation for the reported access site bleeding issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of access site bleeding was most likely anticoagulation management since the bleed started to improve as act was coming in recommended range. as noted in the impella ifu: impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10365 | 1220648-2025-27461 | 21836644 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836644&pc=OZD | 02/15/2025 | 08/01/2025 |  | 04/14/2025 |  |  | 03/17/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025579062 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Renal Impairment | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported through an impact study that the patient encountered the following during impella support: paroxysmal atrial fibrillation, severe vasodilatory shock, acute kidney injury (aki), impella access site hematoma, vascular access site hematoma, acute blood loss anemia, blood loss anemia and hemolysis. the investigator found that the device insertion could trigger arrythmia's and low-level hemolysis can cause aki. additionally, due to the acute blood loss anemia the patient received 3 units of packed red blood cells (prbcs), 3 units of platelets, one unit of cryo, and 500 of anti-inhibitor coagulant complex intraoperatively. | A notification was received via impact study ¿(b)(6) medical center: (b)(4)¿ regarding a patient who experienced arrhythmia and hemolysis during impella 5.5 support, possibly triggered by device insertion. the patient continued on support until the device was successfully weaned. |  |  |  |  | The investigation of vf/vt/af/at, hemolysis, access site bleeding, and renal failure has not been completed. should new information be received, a supplemental manufacturer device report will be filed. b3 date of event was erroneously entered on the initial manufacturer device report number. b5 narrative on the initial manufacturer device report number was incomplete h6 health effect - clinical codes 1888, 4499, 1884, and health effect - impact codes 4643 and 4644 were inadvertently omitted on the initial manufacturer device report number. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
10366 | 1220648-2025-27458 | 21836645 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836645&pc=OZD | 03/17/2025 | 05/12/2025 |  | 04/14/2025 |  |  | 03/17/2025 | 04/14/2025 |  |  | 12/31/2025 |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025436735 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was attempted to be implanted with an impella 5.5 device for mechanical circulatory support. during the impella 5.5 placement, the physician could not advance the 5.5 past the patient anatomy, around the innominate area. after several unsuccessful attempts the physician decided to place an impella cp through the axillary instead. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since impella to unable advance past innominate area due to patient's anatomy.
10367 | 1220648-2025-27457 | 21836648 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836648&pc=OZD | 03/17/2025 | 11/24/2025 |  | 04/14/2025 |  |  | 03/17/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026622696 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Fluid/Blood Leak | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the yellow leur would not disconnect from the purge tubing to yellow leur on the 5.5 purge sidearm, despite a warm water compress soaking on purge. also, the placement signal and left ventricle signal were intermittently dashing out with controller error alarms but without any placement signal alarms. no patient harm has been reported. | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the yellow leur would not disconnect from the purge tubing to yellow leur on the 5.5 purge sidearm, despite warm water compress soaking on purge. also, no placement signal unreliable alarms were noted/identified. there was no patient harm reported. |  |  |  |  | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. priming issue. the clinical details state that they were unable to get the yellow luer to disconnect from the purge tubing during regular changing of the cassette on (b)(6) 2025. on (b)(6) 2025 they opened up a new cassette to replace the broken spike on the stuck luer. there was no product retuned so the root cause of the priming issue cannot be determined. optical signal issue upon review of the data logs, it is apparent that the console is the potential cause of the issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. | After evaluation, the complaint does not meet the criteria for mdr reportability as defined by 21 cfr 803.50. the issue will be tracked and trended internally. | D6 explant date unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 impella 5.5 with smartassist for use during cardiogenic shock. section: cleaning. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
10368 | 1220648-2025-27460 | 21836651 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836651&pc=OZD | 03/17/2025 | 05/05/2025 |  | 04/14/2025 |  |  | 03/17/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026615277 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. during the implant, the pump would not make turn past the arch near the innominate artery. the implant was aborted and the patient was placed on veno-arterial extracorporeal membrane oxygenation (va ecmo). no patient harm was reported. | [5 empty] | The investigation of delivery issues has been completed. clinical reported pump being unable to make turn past the arch near the innominate artery due to patient anatomy and vessel calcification. delivery was aborted the patient was placed on venoarterial extracorporeal membrane oxygenation (va-ecmo). pump was returned and there were no abnormalities found. the root cause of the delivery issue was most likely patient condition related since clinical reported the patient had calcified vessels and the pump was unable to pass the innominate artery. h5 was erroneously selected on the initial manufacturer device report number 1220648-2025-27460. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smartassist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controllerrp flex with smartassist system with automated impella controller section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
10369 | 1220648-2025-27456 | 21836689 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836689&pc=OZD | 03/17/2025 | 04/14/2025 |  | 04/14/2025 |  |  | 03/17/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025572892 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, while pushing in a companion sheath, the impella pump was not stabilized and subsequently perforated the left ventricle. as support continued, the patient's blood pressure began to consistently drop. suction was noted and the pump was pulled back. at that point in time the patient's blood pressure dropped dramatically. an echocardiogram showed a hole in the apex. pericardiocentesis was performed, but the patient became hypovolemic. volume resuscitation was not successful, and cpr was performed for thirty minutes without recovery. the patient passed away. following the case, it was noted that the left ventricle may have been friable due to a recent infarction. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system and impella 5.5 with smartassist for use during section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist system and impella 5.5 with smartassist for use during cardiogenic shock & impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position." ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
10370 | 1220648-2025-27462 | 21836846 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836846&pc=OZD | 03/20/2025 | 04/14/2025 |  | 04/14/2025 |  |  | 03/20/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026624489 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support and during support, the patient had a hematoma at the access site. hematoma was manually decompressed, percloses were tightened, and two units of blood products were given to the patient. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
10371 | 1220648-2025-27463 | 21836847 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836847&pc=OZD | 03/20/2025 | 04/14/2025 |  | 04/14/2025 |  |  | 03/20/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1286543 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that a hospital received a loaner automated impella controller and upon turning it on it alerted with a "controller error" alarm. there was no patient involvement. | [5 empty] | The investigation has been completed. the controller was returned for investigation. console logs from the reported day of event are consistent with complaint and show that upon each boot-up the console reports ¿bad lamp¿ condition, as evidenced by low light parameter in ¿5s¿ measurement results. attempts to self-reboot optical bench or manual attempts to power-cycle the console did not resolve the issue. after console was received in danvers, during inspection a severe contamination of optical pump connector fiber was observed. as the contamination was being removed (more than one cleaning was needed), the value of light parameter would increase, and controller error alarm would no longer reproduce. this is consistent with contamination on the optical fiber facet decreasing amount of light being reflected back into optical bench and measured by voltage produced by ccd detector. after cleaning the connector, despite minor scratches and remaining debris, the ¿5s¿ parameters were within specification.
10372 | 1220648-2025-27464 | 21836865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836865&pc=OZD | 03/13/2025 | 11/11/2025 |  | 04/14/2025 |  |  | 03/17/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025598897 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Study | Health Professional | User facility | Company representation |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Thrombocytopenia | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Impact forwarded to complaints reports postoperative thrombocytopenia with relationship per product insert as "possible" to impella device. impact states following: the patient was borderline thrombocytemic preoperatively and experienced a somewhat pronounced periprocedural drop in thrombocytes, impella may have been a contributing factor. | The complaints team received an impact study for cerebral embolism. the impella was removed using the standard procedure. immediately following removal, the patient developed apnea, bilateral mydriasis and became unresponsive. he was intubated and underwent emergent cranial ct angiography, where a thrombus in the basilar artery was detected. he then underwent successful emergent interventional thrombembolectomy. he was transferred to the icu and extubated. the patient was noted to be recovering but continued to experience neurological sequelae, primarily somnolence. |  |  |  |  | B5 additional information received from the clinical site. h6 updated to reflect additional information received from the clinical site. | The investigation of stroke/cva has been completed. the product was returned and no damages or abnormalities were found. strokes/cvas can occur during or after impella support, and can be either ischemic or hemorrhagic. to make a determination as to the cause of the stroke, the following clinical information must be considered: -patient's medical history, including any previous strokes, cardiovascular disease, or risk factors such as hypertension, diabetes, or atrial fibrillation. -timing of the stroke in relation to impella support (during or post-implant). -detailed description of the symptoms experienced by the patient. -neurological examination findings and any focal deficits observed. -diagnostic imaging results, such as ct scan or mri of the brain, to identify the type and location of the stroke (ischemic or hemorrhagic). -laboratory test results, including coagulation profiles and cardiac markers. -any relevant procedural or device-related factors, such as duration and settings of impella support, presence of embolic protection devices, or any documented procedural complications. -response to any previous anticoagulation or antiplatelet therapies. -evaluation of potential alternative causes of stroke, such as arrhythmias, embolic sources, or underlying vascular pathology. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
10373 | 1220648-2025-27465 | 21836991 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21836991&pc=OZD | 03/17/2025 | 04/17/2025 |  | 04/14/2025 |  |  | 03/17/2025 | 04/14/2025 |  |  |  |  | 04/14/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025612013 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The decision was made to withdraw care and the patient expired. per medical safety officer review use of the device cannot conclusively be associated with the outcome. | The user facility reported that the patient was on impella 5.5 support and during support, mobile thrombus was noted on distal end of pump. patient was placed on heparin protocol and monitored. |  |  |  |  | The investigation of thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b2 date of death should have been left blank on the initial manufacturer device report number 1220648-2025-27465. b5 mso statement was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27465. h10 investigation results were inadvertently omitted on the initial manufacturer device report number 1220648-2025-27465. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10374 | 1220648-2025-27390 | 21824349 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21824349&pc=OZD | 03/26/2025 | 07/01/2025 |  | 04/11/2025 |  |  | 04/03/2025 | 04/11/2025 |  |  |  |  | 04/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025611336 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 89 YR |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient developed bleeding at their puncture site during impella cp support. bleeding was additionally seen for other device insertion sites at the same time and blood transfusions were performed. as the patient's cardiac function had recovered, the decision was made to explant the device early. the patient survived the procedure. | [5 empty] | The root cause of access site bleeding is determined to be patient condition because of the bleeding from other device insertion sites. | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10375 | 1220648-2025-27391 | 21824504 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21824504&pc=OZD | 02/17/2025 | 06/30/2025 |  | 04/11/2025 |  |  | 04/03/2025 | 04/11/2025 |  |  |  |  | 04/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025598058 |  | 1000459 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an automated impella controller (aic) displayed de-air purge incomplete and did not proceed. the aic was replaced. upon investigating the aic, the engineer found that the issue could be related to the pump and possibly the purge cassette. there was no patient harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: cleaning: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of purge leak ¿ pump has been completed. the impella device was not returned for evaluation. purge leak pump: based on the findings from data log review, the root cause of the purge leak - pump was determined to be no problem found.
10376 | 1220648-2025-27392 | 21824579 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21824579&pc=OZD | 03/13/2025 | 04/15/2025 |  | 04/11/2025 |  |  | 03/13/2025 | 04/11/2025 |  |  |  |  | 04/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1762746 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller (aic) had a system self-check failure during setup. the console was replaced with no harm to the patient. | [5 empty] | The automated impella controller device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - system self-check error has been completed. the console logs showed a communication issue with the pump motor driver (pmd) during the initial bootup. due to the communication issue, the aic console rebooted automatically (as designed) to attempt another power on. after rebooting, the console displayed the "system self check failed" screen, and the console was then forcefully shut down. the console was tested on the lab bench, and the reported issue was reproduced. upon bootup, the console rebooted automatically and displayed the "system self check failure" screen. initially, while investigating, found that the bat2 was not charging, while light emitting diode 1 and 3 were flashing, indicating a pf status. additionally, the battery status was verified using the batttest, confirming that the bat2 battery was in a discharged state. further investigating found that the power battery manager (pbm) was defective and most likely failed to provide nominal 20 volt to power pmd circuitry (impellatronic pca) which led to a drop in communication. a new battery was used with the pbm, and while the battery initially charged, it displayed the same pf status after about one minute. both the pbm and batteries were replaced with good ones and the aic ran with pump without any issues. the root cause of the system check failure was determined to be a defective power battery manager board.
10377 | 1220648-2025-27393 | 21824631 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21824631&pc=OZD | 03/16/2025 | 04/29/2025 |  | 04/11/2025 |  |  | 03/16/2025 | 04/11/2025 |  |  |  |  | 04/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025593433 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was noted that when removing the .018 wire from the impella, the wire was hard to remove. when the wire was finally removed, the physician noticed the wire had become unraveled at the tip. the impella started with no complications and it was confirmed on x-ray that no parts or pieces of wore were left in due to the unraveling. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation of delivery issues has been completed. the root cause of delivery issue is not determined because insufficient clinical details were provided and the pump and guidewire were not returned for inspection for evaluation. hence the root cause is not determined. g1 reporting contact email revised on manufacturer device report 1220648-2025-27393 in accordance with updated procedures. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system: section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
10378 | 1220648-2025-27394 | 21824652 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21824652&pc=OZD | 01/13/2025 | 04/11/2025 |  | 04/11/2025 |  |  | 03/13/2025 | 04/11/2025 |  |  |  |  | 04/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025570425 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  |  | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The product was returned due a low placement signal complaint, which was assessed as not reportable. however, analysis of the returned device indicated an upward trend in motor current and evaluation of the pump found biomaterial stuck in the gap. there was no information that indicated harm to the patient (unknown race and ethnicity). | [5 empty] | The impella cp device was returned for analysis and evaluation. no data logs were returned. the pump was evaluated, and mold was seen all along the purge gap and the impeller. it was soaked in warm water with tergazyme for 10 minutes, and then the gap was re-imaged. there was significant biomaterial stuck inside the gap. the pump was started, the motor current was elevated and saturated flows were reproduced. product history review was completed, and the pump passed all post sterile inspection checks. in conclusion, the root cause of the saturated pump flow is due to biomaterial based on the returned product analysis.
10379 | 1220648-2025-27395 | 21824729 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21824729&pc=OZD | 03/15/2025 | 05/19/2025 |  | 04/11/2025 |  |  | 03/15/2025 | 04/11/2025 |  |  |  |  | 04/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025605516 |  | 1000459 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Pumping Stopped | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was on impella 5.5 support when an impella defective alarm occurred, and the pump stopped. the impella was exchanged. the patient expired five days later from multiple organ failure. per abiomed's medical safety officer review, the impella 5.5 did not work/stopped and it cannot be said that the impella 5.5 pump did not contribute to the patient's outcome. | [5 empty] | Per abiomed's medical safety officer review, the impella 5.5 did not work/stopped and it cannot be said that the impella 5.5 pump did not contribute to the patients outcome. the impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of pump stop has been completed. after 5 days on impella , sudden impella failure alarm occurred and pump stopped. it was replaced by a second pump. data logs were not recovered. pump was returned and inspected. no physical abnormalities were noted on the pump however it was unable to connect to any console during investigation. connection to a console resulted in impella defective alarm being triggered. no broken patient cable pins or blood in the red handle were observed on the returned pump. the cause of the pump stop was not determined since data logs were not recovered and the returned pump could not be detected on any console during investigation.
10380 | 1220648-2025-27396 | 21824752 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21824752&pc=OZD | 03/12/2025 | 04/17/2025 |  | 04/11/2025 |  |  | 03/13/2025 | 04/11/2025 |  |  |  |  | 04/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1652655 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller (aic) had a "staticky" display. the aic remained in use and the situation was monitored by the staff. there was no harm to the patient as a result of the event. | [5 empty] | The investigation of automated impella controller (aic) - display issue has been completed. no evidence was found in the logs to confirm the reported failure mode. this console was tested and upon running the pump and cassette for 1 hour, the horizontal flickering screen was reproduced. the input voltage at the 4-in-1 connector x25 is good, measuring 23.8v under ac power. the cables between the 4-in-1 connector and the display are appropriately connected. root cause: the root cause for the flickering screen was not determined. h6 component code 765 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27396. | The automated impella controller device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
10381 | 1220648-2025-27397 | 21824782 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21824782&pc=OZD | 02/04/2025 | 04/11/2025 |  | 04/11/2025 |  |  | 03/17/2025 | 04/11/2025 |  |  |  |  | 04/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025583346 |  | 0048-0024 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Aortic Valve Insufficiency/ Regurgitation | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. after the impella cp was implanted, aortic regurgitation worsened and lactate levels increased. the impella cp was explanted and an intra-aortic balloon pump was placed. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smart assist system: section: warnings & cautions: warnings; section: pre-support evaluation; section: impella cp with smart assist catheter insertion; section: axillary insertion of the impella cp with smart assist catheter; section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: impella cp with smartassist catheter insertion: ¿impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve¿. section: use of echocardiography for positioning of the impella catheter impella catheter inlet to the aorta: ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms: ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿
10382 | 1220648-2025-27398 | 21824796 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21824796&pc=OZD | 03/13/2025 | 10/06/2025 |  | 04/11/2025 |  |  | 03/13/2025 | 04/11/2025 |  |  |  |  | 04/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1268753 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Connection Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 | [18 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support, and during support placement signal issues were encountered. the physician removed the white plug and then put it back in on the automated impella controller (aic). after doing this, the placement signal was present. however, the issue presented again, and the nurse pushed on the white plug portion that connects to the aic. immediately thereafter the placement signal worked. there was no harm reported to the patient as a result of this event. the patient was noted to be in stable condition. | [5 empty] | The investigation of automated impella controller (aic) - cabling issue has been completed. the impella device was not returned for evaluation. based on previous investigation of this console with a different pump, the cause of the aic issue was a defective blue receptacle. | The automated impella controller device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. console logs from the reported day of event are consistent with complaint and show that after pump start, console kept presenting with multiple alarms placement signal not reliable (#130, due to low snr). ltlog and rtlog data showed that optical signal snr was low, with sporadic but brief improvements during the case (which most likely correspond to reported attempts to ¿pushing on the white plug¿). characteristics of optical data - low snr, low contrast, lamp at 100%, maxed-out gain of 2.6 - are consistent with an air gap. console ic4779 arrived in danvers for later investigation 25-42987-2, involving pump 573246, with very similar symptoms (same failure mode). it was tested, and initial measurements of ¿5s¿ parameters using a fiber-coupled test cavity showed parameters within specification. when using a pump (attached via pump connector), the parameters - in particular snr ¿ did not meet specification. inspection of the optical pump connector revealed a piece of debris (presumably a broken piece of pump plug) lodged in the blue receptacle, preventing pump plug from being inserted fully, therefore creating a large air gap. when the piece of debris was removed, optical ¿5s¿ parameters were measured to be within specification (using a pump). the findings and root cause are consistent between the two investigations for ic4779. the cause of the aic issue was a defective blue receptacle. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated to include optical signal coding. d10 concomitant medical products and therapy dates updated.
10383 | 1220648-2025-27399 | 21824854 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21824854&pc=OZD | 03/13/2025 | 04/11/2025 |  | 04/11/2025 |  |  | 03/13/2025 | 04/11/2025 |  |  | 10/31/2025 |  | 04/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024421481 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and developed an access site hematoma. manual pressure was held, and one unit of packed red blood cells was given to the patient. it was further noted that the physician did an assessment. and stated the site was soft. however, it was unknown if there was a hematoma or possibly just related to patient body habitus. after approximately two days of total impella cp mcs, the device was successfully weaned and explanted with a survived outcome. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
10384 | 1220648-2025-27400 | 21824909 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21824909&pc=OZD | 03/13/2025 | 05/08/2025 |  | 04/11/2025 |  |  | 03/13/2025 | 04/11/2025 |  |  |  |  | 04/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026613219 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  |  |  | Death | Yes | Yes | Pumping Stopped | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient (unknown race and ethnicity) noted with post-cardiotomy cardiogenic shock (pccs)/low cardiac output syndrome (lcos) was implanted with an impella 5.5 device for mechanical circulatory support. during support, the pump stopped due to motor current spike. the pump was eventually restarted on p-1 setting and was worked up to p-3. there was no report of patient harm. | Additional information received that the family decided to withdraw care and the patient expired. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The date of explant is unknown the field was left blank accordingly. the impella device was returned, and an evaluation is underway. upon completion of the analysis, a supplemental report will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system. entering cardiac output (impella cp with smart assist), section: direct aortic insertion: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of temporary pump stop has been completed. the pump was returned for investigation. data logs were not provided for the case. no physical damage was noted on the returned product. upon further investigation, a biometer was found wrapped around the impeller. the pump was socked in a bucket of water for 15 minutes and was then able to start, with observed saturated pump flow. the cause of the temporary pump stop issue was not determined since issue was not reproduce on returned product and data logs were not available to review. b1 revised to reflect both adverse event and product problem, based on additional information received from the clinical site. b2 revised to reflect death and date of death based on additional information received from the clinical site. b5 revised with additional information received from the clinical site. d6b explant date added. h1 revised to reflect death. h6 revised based on additional information received from the clinical site.
10385 | 1220648-2025-27401 | 21824990 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21824990&pc=OZD | 03/14/2025 | 04/11/2025 |  | 04/11/2025 |  |  | 03/14/2025 | 04/11/2025 |  |  |  |  | 04/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1606424 |  | 0042-0040-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. there was an automated impella controller (aic) battery failure alarm. the pump was opened for insertion, however, upon turning on two aic's at the facility, the battery failure alarm went off. phone support was given prior on how to turn battery back on as it had been sent for preventative maintenance and returned since last used. on button was not engaged once and so one aic turned back on and started to charge, however, the doctor had proceeded with the intra-aortic balloon pump (iabp) for procedure due to timing and patient needs. the other pump also was not on at the bottom of the console. it was turned back on and now both aic's were charging. the field rep spoke with biomed about proper steps that come on screen of aic and importance to complete prior to return to circulation and not to remove unless steps complete. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10386 | 1220648-2025-27402 | 21825035 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21825035&pc=OZD | 03/14/2025 | 04/11/2025 |  | 04/11/2025 |  |  | 03/17/2025 | 04/11/2025 |  |  |  |  | 04/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025612218 |  | 1000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Unspecified Infection | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that a patient expired and it was noted that the cause of death was infection and intestinal ischemia. the relationship between the cause of death and the impella is unknown. the physician noted it is unlikely the infection is related to the impella insertion as it occurred to early. after escalation to the impella 5.5 and upon returning to the room, the patient developed a fever. the cause of the infection is unknown. information was received indicating that the patient died and the pump was removed. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock & cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
10387 | 1220648-2025-27403 | 21825225 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21825225&pc=OZD | 03/14/2025 | 04/11/2025 |  | 04/11/2025 |  |  | 03/14/2025 | 04/11/2025 |  |  |  |  | 04/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1703855 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) was unable to recognize purge disk during pump priming. priming was retried; however, still the aic was unable to recognize purge disk. consequently, the aic was exchanged and support continued for the patient. there was no harm to the patient as a result of this event. | [5 empty] | Patient-specific information is unavailable at the time of this mdr writing; therefore, respective fields reflect "unknown" or left blank accordingly. the automated impella controller along with data logs were received for analysis and evaluation. the imc logs from the complaint date revealed that the console issued purge disc not detected alarm twice. during the device analysis, the console was examined after arriving, and a broken purge disc flag was identified. product history review was completed, and the no action was required as a result of the review. in conclusion. the purge disc flag was observed to be broken and was the root cause of the reported purge disc not detected alarm issue.
10388 | 1220648-2025-27404 | 21825333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21825333&pc=OZD | 03/15/2025 | 04/11/2025 |  | 04/11/2025 |  |  | 03/15/2025 | 04/11/2025 |  |  |  |  | 04/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577879 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Malposition of Device | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support (mcs) and after start of the support overnight, the patient developed hemolysis. additionally, it was noted the automated impella controller displayed positional issue. the pump was deep; however, the physician chose not to reposition the impella cp pump at such time. following, the patient was transferred to a different hospital, and the intervention for the hemolysis and positioning issues was unclear. the patient continued on mcs for additional six days before the device was successfully weaned and explanted with a survived patient outcome. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
10389 | 1220648-2025-27405 | 21825556 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21825556&pc=OZD | 03/16/2025 | 05/06/2025 |  | 04/11/2025 |  |  | 03/16/2025 | 04/11/2025 |  |  |  |  | 04/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026617619 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Vascular Dissection | Ectopic Heartbeat | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient was taken to operating room for impella 5.5 placement for suspect of ventricular septal defect. during placement it was reported that they had access complications of small vessel with difficulty passing past the anastomosis due to anatomy and small vessel, decision was made at that time to attempt the left side w/ removal of the graft. dissection was noted, the axillary vessel was repaired and moved to the left side. additionally, during support it was noted that the care team had noticed more ectopy. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during: section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. | The investigation of vascular damage - peripheral vessel, vf/vt/af/at, and delivery issues has been completed. the impella device was not returned for evaluation. a dissection of the right axillary artery was noted after a difficult delivery through a small vessel. patient had poor vasculature with history of multiple vascular surgeries and amputations. the root cause of the vascular damage was most likely patient condition related since the patient had small and poor vasculature. as the necessary information was not provided, the root cause of the vt/vf was not determined. the root cause of the delivery issues was most likely patient condition related since the patient had small and poor vasculature.
10390 | 1220648-2025-27406 | 21825676 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21825676&pc=OZD | 03/14/2025 | 05/28/2025 |  | 04/11/2025 |  |  | 03/16/2025 | 04/11/2025 |  |  |  |  | 04/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026618860 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | Defective Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp had high purge pressure and purge system blocked alarm. motor current was in range. purge lysis protocol was administrated successfully. it was further reported that the transparent ring from the dilatator remained on the peel-away sheath. the physician commented that if the sheet didn't have peeled, the ring wouldn't have been noticed. there was no harm to the patient. | [5 empty] | The investigation of high purge pressure and introducer damage ¿ mechanical has been completed. high purge pressure: impella cp was returned cut on catheter and patient cable side was not returned. no biomaterial was present on returned pump. data logs were reviewed and lt logs showed a gradual rise in purge pressure over approximately 10 hours. no motor current spike is present prior to purge pressure rise. purge pressure remained high for only a short while before slowly resolving back to normal range. clinical details noted that purge pressure was increasing with a drop in purge flows, tissue plasminogen activator (tpa) was administered and purge flows and purge pressure were able to resolve. the root cause of the high purge pressure was most likely due to biomaterial build-up as it was able to be resolved with tpa. introducer damage - mechanical: no introducer or patient cable side of pump were returned for evaluation. clinical details noted that the clear plastic connector came disconnected from the dilator and remained connected to introducer after dilator was removed. clinical details noted no excessive force was used when removing dilator from introducer. the root cause of the introducer damage - mechanical could not be definitively determined as no product was returned for analysis. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-27406. h6 component code 4742 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27406. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
10391 | 1220648-2025-27407 | 21825929 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21825929&pc=OZD | 03/16/2025 | 04/15/2025 |  | 04/11/2025 |  |  | 03/16/2025 | 04/11/2025 |  |  |  |  | 04/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026624269 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  |  | Black Or African American | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that on day one of impella cp support motor current spiked and the impella stopped. the cp restarted on the third attempt. plans are to escalate to impella 5.5. additionally, the patient also had acute bleeding from both extracorporeal membrane oxygenation cannula and impella site. manual pressure held, sutures placed, and blood products were administered. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of pump stop and access site bleeding has been completed. the returned product was inspected and there were no physical abnormalities. there was no biomaterial on the impeller or in the cannula. the pump passed electrical testing. the pump was run in the lab and the pump stop did not reproduce. no data logs were provided for this case. the root cause of the temporary pump stop was not determined as no data logs were provided to confirm reason for temporary pump stop. access site bleeding: clinical details mention patient had bleeding from impella access site and ecmo cannula which was resolved with troubleshooting and blood products. the root cause of the access site bleeding - major was most likely patient condition as evidenced by multiple bleeding locations. pump stop instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ h10 instructions for use for pump stop was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27407. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10392 | 1220648-2025-27408 | 21825978 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21825978&pc=OZD | 03/11/2025 | 05/19/2025 |  | 04/11/2025 |  |  | 03/19/2025 | 04/11/2025 |  |  |  |  | 04/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025594272 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facililty reported that a patient was on impella cp support and during transportation, the impella cp was removed due to an incident. the pump fell out due to strong body movement. the pump was replaced. there was no patient harm. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. according to the clinical report, the pump was pulled into the aorta during the patient transfer to the catheterization lab. the cause of the positioning issue was most likely related to use, as the pump position was compromised during the patient's transfer. | The impella device was discarded by the customer and therefore, ¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
10393 | 1220648-2025-27409 | 21826153 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21826153&pc=OZD | 03/18/2025 | 10/01/2025 |  | 04/11/2025 |  |  | 03/18/2025 | 04/11/2025 |  |  |  |  | 04/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025612317 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | Increase in Pressure | Device Sensing Problem |  |  |  |  | Hemolysis | Embolism/Embolus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient on impella 5.5 support had labs that were hemolyzed. the patient had an increase in lactate dehydronase levels and evidence of hemolysis with tea colored urine and elevation in creatinine levels. performance level was initially at p6 and was reduced to p4 when hemolysis was suspected. bedside echocardiogram was performed demonstrating adequate position of the impella. repeat bedside transthoracic echocardiogram performed six hours later re-demonstrated adequate position of the impella in the left ventricle (lv). this was pulled back approximately 1.5 centimeters for a depth of 4.0 centimeters from the aortic valve to the mid portion of the lv inlet. appearance of urine improved. additionally, the impella had high purge pressure and purge system blocked alarms that were resolved with the changing of the purge cassette. there were also false "impella in aorta" alarms. lv position signal was disabled manually. it was further reported that the impella 5.5 was removed for hemolysis and concern for clot at the outlet/motor housing unit. the patient survived the procedure. | The complainant reported that the device exhibited optical sensor issues. |  |  |  |  | The investigation of hemolysis, positioning issues, and high purge pressure has been completed. the impella device was not returned for evaluation. the cause of the hemolysis issue was most likely biomaterial, based on data log review. the cause of the high purge pressure issue was most likely biomaterial, based on data log review. the cause of the positioning issue was not determined since product was not returned and sufficient clinical information was not provided. thrombus: the failure mode of thrombus has been added as an investigated failure mode, as the case information suggests that the device may have interacted with biomaterial. as the necessary information was not provided, the root cause of the thrombus was not determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ impella 5.5 with smartassist: section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the pump was not returned for investigation. optical signal issue: several impella position alarms were noted during the case, with a verified good pump position. the data logs show several motor current spikes during the reported false position alarms, and it was seen that the doctor disabled the placement signal monitoring alarm. the cause of the optical signal issue was most likely biomaterial interaction which cause the false position i aorta alarm. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions. description of the location and characteristics of the thrombus (e.g., size, shape, consistency). relevant laboratory test results, such as coagulation profiles and platelet counts. imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). medications or treatments that the patient was receiving at the time of thrombus formation. surgical or procedural details if applicable (e.g, cardiac catheterization). any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical signal issue "2917" as part of a retrospective review of optical signal issues in accordance with fda recommendation. the results of the optical signal investigation added to h6.
10394 | 1220648-2025-27411 | 21827014 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21827014&pc=OZD | 03/17/2025 | 05/16/2025 |  | 04/11/2025 |  |  | 03/17/2025 | 04/11/2025 |  |  |  |  | 04/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025595784 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemolysis | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The decision was made to withdraw care and the patient expired. per the medical safety officer the use of the impella cp device cannot be dissociated. there is not enough information to exclude as an associated factor. | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the pump position became low, and the patient experienced extremely bloody urine. the next day, continuous renal replacement therapy (crrt) was started. afterwards, the pump was repositioned and pulled back which helped slowly resolve the hemolysis. additionally, the patient suffered a massive stroke. discussions about limited further options were made due to patient suffering severe stroke, crrt, and cardiac requirements needed. |  |  |  |  | The investigation of stoke/cva and hemolysis has been completed. the impella device was not returned for evaluation. based on the clinical details the cause of the hemolysis issue was most likely improper positioning of the device, based on given clinical details. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. b5 medical safety officer statement added b2 death and date of death were inadvertently omitted on the initial manufacturer device report number 1220648-2025-27411. b5 narrative erroneously stated the patient outcome on the initial manufacturer device report number 1220648-2025-27411. h1 type of report was erroneously selected on the initial manufacturer device report number 1220648-2025-27411. h6 health effect - impact code 1802 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27411. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
10395 | 1220648-2025-27412 | 21827226 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21827226&pc=OZD | 11/09/2021 | 04/11/2025 |  | 04/11/2025 |  |  | 11/09/2021 | 04/11/2025 |  |  | 07/31/2023 |  | 04/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022130581 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was being implanted with an impella 5.5 for mechanical circulatory support. for the insertion of the impella 5.5, a graft was added to the axillary artery which was dissected during the suturing of the graft. surgical repair was done and hemostasis was achieved. during insertion of the catheter, it was reported that the catheter would not advance through the subclavian artery. the medical team decided to abort the impella 5.5 implant due to the patient's anatomy. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. there was no evidence of excessive force being applied and the clean procedure was followed and the investigation showed stenosis of innominate from subclavian artery. the root cause of the delivery issues was due to patient anatomy, most likely due to patient condition. he root cause of the vascular damage was determined to be patient condition. impella 5.5 with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings & cautions: warnings ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/ descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with echocardiography guidance.¿ section: warnings & cautions: cautions never advance the guidewire or sheath when resistance is met. determine the cause of resistance using fluoroscopy and take remedial action. this report is being filed as part of a retrospective review of historical records.
10396 | 1220648-2025-27413 | 21827300 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21827300&pc=OZD | 04/29/2022 | 04/11/2025 |  | 04/11/2025 |  |  | 04/29/2022 | 04/11/2025 |  |  | 02/29/2024 |  | 04/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023187469 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | [6 empty] | 66 YR |  | Male |  | Hispanic |  | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in post cardiotomy cardiogenic shock and with low cardiac output syndrome was being implanted with an impella 5.5 for mechanical circulatory support. it was reported that the physician attempted three times and was unable to advance the impella without a guidewire. the impella was removed, and a .018 wire was place into the left ventricle without issue. the pump was then successfully delivered and the patient was able to be weaned off bypass. the patient remained on impella support until successfully weaned. | [5 empty] | The investigation for the reported issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. the cause of the delivery issue was wireless insertion. the impella 5.5 device is not indicated for wireless insertion per instructions for use. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings & cautions: warnings: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/ descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with echocardiography guidance.¿ section: warnings & cautions: cautions: never advance the guidewire or sheath when resistance is met. determine the cause of resistance using fluoroscopy and take remedial action. this report is being filed as part of a retrospective review of historical records.
10397 | 1220648-2025-27414 | 21827439 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21827439&pc=OZD | 03/16/2025 | 04/11/2025 |  | 04/11/2025 |  |  | 03/16/2025 | 04/11/2025 |  |  |  |  | 04/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026625430 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella cp pump was explanted due to concerns of diminished limb perfusion in the right lower extremity. a dopplerable pulse was present in the right foot when the pump and sheath were still in place. the same pulse was subsequently present following explant. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
10398 | 1220648-2025-27415 | 21827623 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21827623&pc=OZD | 04/01/2025 | 06/27/2025 |  | 04/11/2025 |  |  | 04/01/2025 | 04/11/2025 |  |  |  |  | 04/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026617697 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 75 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp placed to support a cardiomyopathy patient. the pump was supporting for just under six hours when it was explanted and replaced due to a sensor failure. there was no reportable patient harm. | [5 empty] | The pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of optical signal issue has been completed. upon review of data logs, clinical details and returned product, it is apparent that the console is the potential cause of the issue. please refer to complaint (b)(4) for an investigation into the console.
10399 | 1220648-2025-27416 | 21827680 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21827680&pc=OZD | 03/16/2025 | 05/30/2025 |  | 04/11/2025 |  |  | 03/16/2025 | 04/11/2025 |  |  |  |  | 04/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025609725 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  |  | Malfunction | No | Yes | Deformation Due to Compressive Stress | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Per the medical safety officer review the impella cp pump 2 was implanted although there was some guidewire difficulty (impeded); however, not contributing to the demise outcome. the patient continued on support for approximately 4 days until his/her demise. the demise outcome is associated with the impella cp pump 1 that stopped under manufacture device report 1220648-2025-27380. | The complainant reported that they encountered a guidewire issue. the impella could not be placed over the guidewire when advanced because the guidewire had kinks. there was no patient harm. |  |  |  |  | The investigation of delivery issues has been completed. the guidewire was returned for investigation. kinks were reported on returned guidewire. no damage was observed at the guidewire tip. data log review was not required for this case run. the cause of the guidewire damage issue was most likely use related during delivery. per the medical safety officer review the impella cp pump 2 was implanted although there was some guidewire difficulty (impeded); however, not contributing to the demise outcome. the patient continued on support for approximately 4 days until his/her demise. the demise outcome is associated with the impella cp pump 1 that stopped under manufacture device report 1220648-2025-27380 b5 revised to include medical safety officer statement. d6b explant date was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27416. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
10400 | 1220648-2025-27417 | 21829747 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21829747&pc=OZD | 02/07/2021 | 04/11/2025 |  | 04/11/2025 |  |  | 04/01/2025 | 04/11/2025 |  |  | 01/31/2022 |  | 04/11/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1465351 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 85 YR |  | Male |  |  | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Mitral Valve Insufficiency/ Regurgitation | Intraoperative Cardiac Valve Injury | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A japan literature piece found an 85-year-old male patient had an impella cp implant. after percutaneous coronary intervention with impella cp indwelling, the patient was admitted to the intensive care unit and showed improvement in hemodynamics, and the impella was removed on the fourth day of hospitalization. after the impella was removed, there was a decrease in blood pressure and an increase in pulmonary artery pressure, and transthoracic and transesophageal echocardiography revealed that the patient had severe mitral regurgitation due to ruptured chordae tendineae. mitral valve replacement and impella 5.0 placement were performed on the same day. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: impella cp with smart assist catheter insertion, section: axillary insertion of the impella cp with smart assist catheter, section: use of impella with transcatheter aortic valves: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: impella cp with smartassist catheter insertion: ¿impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), portable c-arm fluoroscopy, or chest x-ray can help confirm the position of the impella catheter after placement, these methods do not allow visualization of the entire catheter assembly and are inadequate for reliably placing the impella catheter across the aortic valve.¿ section: use of echocardiography for positioning of the impella catheter impella catheter inlet to the aorta. ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿
10401 | 21812674 | 21812674 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21812674&pc=OZD | 02/15/2025 | 04/03/2025 | 04/03/2025 | 04/10/2025 | 04/03/2025 | 04/10/2025 |  | 04/10/2025 |  |  |  |  | 04/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US | [5 empty] | Yes |  |  | Not Available |  |  | No |  |  |  | User Facility report | No | Risk Manager | [6 empty] | 66 YR |  | Male |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Impella 5.5 catheter implanted in right subclavian on [redacted]. catheter sensor failure leading to no systolic and diastolic placement signals starting [redacted] at 1500. patient sent to or for transplant on [redacted].
10402 | 21812713 | 21812713 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21812713&pc=OZD | 09/05/2024 | 04/03/2025 | 04/03/2025 | 04/10/2025 | 04/03/2025 | 04/10/2025 |  | 04/10/2025 |  |  |  |  | 04/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Not Available |  |  | No |  |  |  | User Facility report | No | Risk Manager | [6 empty] | 63 YR |  | Male |  |  |  | Malfunction | No | Yes | Break | Unexpected Shutdown | [5 empty] | Low Blood Pressure/ Hypotension | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Patient (pt) in ct for imaging. during the return transfer to the patient's bed, the impella catheter broke causing the machine to stop. the registered nurse (rn) at the bedside heard a ¿pop¿ and identified the problem immediately and intervened with vasopressors, as needed. the patient had a slight drop in blood pressure, which was corrected with the vasopressors. no other change in condition was noted and the patient tolerated well. cvis [cardiovascular information systems] medical doctor was called to the bedside and the patient was returned to the cvicu [cardiovascular intensive care unit] without issue. cv [cardiovascular] surgeon was notified and took the patient to the or to replace the impella catheter, without incident. presumably the catheter was pulled apart during the move. there was plenty of slack in the line, and the machine was located at the patient's side. the team used the bedsheet to move the patient. the catheter was between the bedsheet and covered by a blanket. a verge report will be completed before shift change. md [redacted] threw the catheter away. there was no malfunction of the machine or suspected manufacturer defect of the catheter itself.
10403 | 1220648-2025-27352 | 21812865 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21812865&pc=OZD | 03/30/2025 | 06/27/2025 |  | 04/10/2025 |  |  | 03/31/2025 | 04/10/2025 |  |  |  |  | 04/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025553876 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient admitted in with cardiogenic shock had an impella cp pump placed to support. the day of implant, there was a purge leak at the connector of the purge filter and purge reservoir. this prompted the team to replace the pump. the patient survived the procedure. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of purge leak ¿ pump has been completed. pump was not returned for investigation. the data log review was not required for this case. as per the clinical notes, the pump was damaged at the purge filter and the pressure reservoir. the image of the damaged purge sidearm was not provided for this case. the cause of the purge leak issue was a leak from the sidearm but the exact location of the leak was unknown without sufficient clinical information or returned product. d6b explant date added.
10404 | 1220648-2025-27353 | 21813028 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21813028&pc=OZD | 04/01/2025 | 04/10/2025 |  | 04/10/2025 |  |  | 04/01/2025 | 04/10/2025 |  |  |  |  | 04/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported via a literature entry "a case of severe heart failure using bicarbonate based impella purge solution as an alternative to heparin" that a patient developed heparin induced thrombocytopenia during impella cp support. argatroban was used as an alternative to heparin, but the impella cp pump developed low purge flow prompting pump replacement. no details were provided on the patients outcome. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. a.1, a.4 and a.5 are unknown.. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. e.1 name prefix/title, first name and last name are unknown. h.4 device manufacture date is unknown. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10405 | 21813052 | 21813052 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21813052&pc=OZD | 03/07/2025 | 03/07/2025 | 04/03/2025 | 04/10/2025 | 04/03/2025 | 04/10/2025 |  | 04/10/2025 |  |  |  |  | 04/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | 0048-0045 |  |  | ABIOMED, INC. |  |  | US |  | 2025601952 |  | 0048-0003 |  | Device was returned to manufacturer |  |  | Not Available |  |  | No | Health Professional |  |  | User Facility report | No | Risk Manager | [6 empty] | 76 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Pump | [6 empty] | Gastrointestinal Hemorrhage | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Patient with an impella cp. previous attempts to wean the impella unsuccessful. was off anticoagulation for gi [gastrointestinal] bleeds. this morning, impella motor current high and suffered pump failure followed by controller failure. catheter pulled back into the aorta by provider bedside.
10406 | 1220648-2025-27354 | 21813384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21813384&pc=OZD | 03/16/2025 | 04/10/2025 |  | 04/10/2025 |  |  | 03/16/2025 | 04/10/2025 |  |  |  |  | 04/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED INC. |  |  | US |  | 2026615861 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention |  |  |  |  | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the patient had limb ischemia. the staff was unable to doppler distal pulses. the pump was explanted and an impella 5.5 was implanted. the patient expired on support of the replacement pump. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10407 | 1220648-2025-27355 | 21813392 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21813392&pc=OZD | 03/13/2025 | 10/17/2025 |  | 04/10/2025 |  |  | 03/14/2025 | 04/10/2025 |  |  |  |  | 04/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026623651 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | AUTOMATED IMPELLA CONTROLLER | [17 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. three days into impella cp support, the patient began to experience swelling in their impella leg. it was noted that a pulse was still present in the extremity. a reperfusion catheter was placed in the leg and the patient's swelling began to subside. support continued with the implanted pump following the intervention. | The complainant also reported that the nurse observed a loss of placement signal for the left ventricle waveforms. placement signal not reliable alarms were noted. the pump was not explanted due to this issue. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. optical signal issue: no device was returned for analysis. the data logs confirm placement signal not reliable alarms. the logs show snr dropped to zero at the instant of placement signal going out of range. the contrast barcoded. the lamp went up to 100. the light and gain also dropped at that instant. the root cause of the optical signal issue was most likely due to a damaged fiber line as evidenced by the data logs pattern observed. limb ischemia: "limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition. any concomitant medication. vascular size or variations thereof. detailed description of the symptoms experienced by the patient. physical examination findings, including pulse assessment and visual examination of the affected limb. diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. laboratory test results, including blood tests for markers of inflammation or clotting disorders. any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension patient's response to previous treatments or interventions for vascular conditions with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include placement signal issue description. h6 updated to include placement signal coding. d10 concomitant medical products and therapy dates updated. g1 reporting contact email updated.
10408 | 1220648-2025-27356 | 21813466 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21813466&pc=OZD | 03/16/2025 | 04/10/2025 |  | 04/10/2025 |  |  | 03/16/2025 | 04/10/2025 |  |  |  |  | 04/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED INC. |  |  | US |  | 2025609768 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was attempted to be implanted with an impella cp device for mechanical circulatory support. the surgeon was unable to get the impella cp to advance past the iliac artery due to heavy calcification. the implant was aborted. there was no patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
10409 | 1220648-2025-27357 | 21813495 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21813495&pc=OZD | 03/14/2025 | 04/15/2025 |  | 04/10/2025 |  |  | 03/14/2025 | 04/10/2025 |  |  |  |  | 04/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026612737 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Defective Device | Material Twisted/Bent | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. there was initial difficulty inserting the companion sheath, however, once tip of sheath was inserted slightly past the hemostatic valve the remainder of the sheath glided through without issue. companion sheath tip was found to be crimped upon removal. there was no patient harm. | [5 empty] | The investigation of introducer damage ¿ mechanical has been completed. the returned companion sheath was inspected and damage was confirmed at the sheath tip. clinical mentioned that the physician had to push harder than usual. the root cause of the introducer damage - mechanical of 7fr companion sheath tip was most likely user related as evidenced by the additional force used and damage on returned product. g1 reporting contact email revised on manufacturer device report 1220648-2025-27357 in accordance with updated procedures. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental mdr will be filed.
10410 | 1220648-2025-27358 | 21813515 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21813515&pc=OZD | 03/16/2025 | 05/06/2025 |  | 04/10/2025 |  |  | 03/16/2025 | 04/10/2025 |  |  |  |  | 04/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED INC. |  |  | US |  | 2025606224 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the complainant reported that upon repositioning the patient, the impella cp was inadvertently pulled out of the left ventricle and into the aorta. the staff stated that they heard a loud pop and noted that the t lock was no longer attached to the repositioning sheath. the pump was explanted in response to the noted issue. there was no patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. | The investigation of positioning issues has been completed. product was not returned for review. based on clinical description, the root cause of positioning issue was most likely patient movement related since the pump pulled out of the left ventricle (lv) while pulling the patient up in bed.
10411 | 1220648-2025-27359 | 21813556 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21813556&pc=OZD | 03/15/2025 | 05/05/2025 |  | 04/10/2025 |  |  | 03/15/2025 | 04/10/2025 |  |  |  |  | 04/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026632716 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. while on support, the impella cp was noted to be entangled in the mitral valve apparatus and was exchanged for an impella 5.5 due to the positioning issue. the patient survived the event. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10412 | 1220648-2025-27360 | 21813636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21813636&pc=OZD | 03/14/2025 | 04/10/2025 |  | 04/10/2025 |  |  | 03/14/2025 | 04/10/2025 |  |  |  |  | 04/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025590659 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during impella cp insertion, the physician pulled the impella out of the left ventricle unexpectedly. the physician reimplanted the same cp. additionally, prior to leaving the procedural area a small hematoma was noted and was evacuated. it was further reported that the patient had dark/red urine and increase in ld to 1074. the impella was repositioned. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) : ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10413 | 1220648-2025-27361 | 21813712 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21813712&pc=PYX | 03/15/2025 | 05/05/2025 |  | 04/10/2025 |  |  | 03/15/2025 | 04/10/2025 |  |  |  |  | 04/10/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025548057 |  | 10000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella rp flex device for mechanical circulatory support. while on support, the impella rp flex flows started to decrease significantly. the impella rp flex was exchanged for a new impella rp flex. upon removal of catheter, a clot was noted at the inflow cage. the patient survived the surgery. | [5 empty] | The investigation of thrombus and low pump flow has been completed. low pump flow: clinical states that the patient was transferred and next morning flows decreased due to which the pump was replaced. at explant a clot was noted. pump was returned and during investigation, a clot was observed at the outlet of the pump with no other physical abnormalities. data logs were investigated and on 14-mar-2025, the logs show rise in placement signal and motor current changes at the same p-level. pump flows are seen to be dropping with numerous suction alarms. this evidence likely supports biomaterial ingestion of which traces were seen on the outlet of the return pump. the root cause of the low pump flow issue was most likely biomaterial as found during product investigation and data log review. as the necessary information was not provided, the root cause of the thrombus was not determined. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
10414 | 1220648-2025-27362 | 21813858 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21813858&pc=PYX | 03/15/2025 | 05/02/2025 |  | 04/10/2025 |  |  | 03/15/2025 | 04/10/2025 |  |  |  |  | 04/10/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025601209 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella rp flex for mechanical circulatory support. the patient was expereincing low flows on durable left ventricular assist device and the decision was made to exchange the impella rp flex for protek duo. | [5 empty] | The investigation of low pump flow has been completed. the root cause of low flow was unable to be determined as limited clinical details were provided and no product was returned for review. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10415 | 1220648-2025-27363 | 21813912 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21813912&pc=PYX | 03/14/2025 | 04/10/2025 |  | 04/10/2025 |  |  | 03/14/2025 | 04/10/2025 |  |  |  |  | 04/10/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2026614359 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  |  | White | Death | Yes | No | Failure to Advance | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that a patient was having watchman procedure in the cath lab when the left atrium was perforated. cpr was initiated and the patient was transferred to the operating room for emergent sternotomy/repair. decision was made for bipella. there were concerns of patient having bilateral upper limb ischemia. vascular had been consulted and discussion amongst the heart team on approach due to vessel sizes. the decision was made to place impella cp via right femoral artery and place rp flex via rij. it was noted that rp flex would not advance into pulmonic valve. the decision was made to remove the rp flex and patient began bleeding where the patient was placed on cardiopulmonary bypass support and emergent sternotomy. the impella cp was already in the left ventricle. ultimately the patient expired in the procedural area. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella rp flex with smartassist system with automated impella controller: section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: potential adverse events (united states): ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings: ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
10416 | 1220648-2025-27364 | 21813948 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21813948&pc=OZD | 03/15/2025 | 04/10/2025 |  | 04/10/2025 |  |  | 03/15/2025 | 04/10/2025 |  |  |  |  | 04/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026626896 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the patient developed bleeding at the insertion site. the pump was removed in response to the condition. the patients outcome is critical. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer, ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10417 | 1220648-2025-27365 | 21814369 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21814369&pc=OZD | 03/14/2025 | 05/08/2025 |  | 04/10/2025 |  |  | 03/14/2025 | 04/10/2025 |  |  |  |  | 04/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1420045 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | No Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the purge disc of the automated impella controller (aic) was not recognizing the cassette. there was no patient involved. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a final report will be filed. a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. there was a repeated inability to get through step 2 of case start as the purge cassette was recognized as inserted but the purge disc was not detected. device analysis: the failure mode was repeated during testing. despite an intact purge retainer and purge flag, the purge disc was not initially detected. however after pushing the disc with more force, the purge disc was able to be completely inserted and then the purge disc was recognized. additionally, it was very difficult to take the purge disc out of the retainer. disassembling the purge system, there were multiple instances of glucose ingress including under the purge retainer that specifically would have provided a sticky counterforce that would have made it difficult to insert the purge disc fully. root cause: the cause of the purge disc not being recognized was a glucose ingress making it more difficult to fully insert the purge disc. g1 reporting contact email revised on manufacturer device report 1220648-2025-27365 in accordance with updated procedures.
10418 | 1220648-2025-27366 | 21814427 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21814427&pc=OZD | 03/15/2025 | 04/21/2025 |  | 04/10/2025 |  |  | 03/15/2025 | 04/10/2025 |  |  |  |  | 04/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026621575 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the impella was placed with no issues. however, once in intensive care unit a hematoma was noted, and manual pressure was held. the physician came to address and asked for 30 minutes then femstop if needed. the patient was dependent on impella with no pulsatility. while holding manual pressure, the hematoma softened. it was noted that heparin was not be started until the next day. the impella cp was successfully weaned and explanted. the patient survived the procedure. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10419 | 1220648-2025-27368 | 21814563 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21814563&pc=OZD | 03/14/2025 | 04/21/2025 |  | 04/10/2025 |  |  | 03/14/2025 | 04/10/2025 |  |  |  |  | 04/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025601606 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 82-year-old female who was being prepped to be implanted with an impella cp, but upon sheath insertion there was too much resistance. the physician decided to abort impella placement. there was no patient harm reported. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since upon sheath insertion there was too much resistance due to patient's anatomy. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the impella cp with smartassist system states the following: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
10420 | 1220648-2025-27369 | 21815104 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21815104&pc=OZD | 02/18/2025 | 05/02/2025 |  | 04/10/2025 |  |  | 03/17/2025 | 04/10/2025 |  |  |  |  | 04/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025594223 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic |  | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, it was noted that there were occasional positioning alarms during ambulation. the patient survived the event. | [5 empty] | The investigation of positioning issues has been completed. based on the clinical details unclear why the impella was deep. no unusual patient conditions stated during the process. the cause of the positioning issues cannot be determined since complaint device not returned for analysis and insufficient clinical data provided. (the cassette was returned for investigation). | The purge cassette was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
10421 | 1220648-2025-27370 | 21815155 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21815155&pc=OZD | 03/15/2025 | 04/10/2025 |  | 04/10/2025 |  |  | 03/15/2025 | 04/10/2025 |  |  |  |  | 04/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025609829 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the 45-year-old male patient was on impella cp support and during support, there was bleeding from the access site due to a failed closure. pressure was applied and blood products were given. the patient was successfully weaned and explanted. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10422 | 1220648-2025-27371 | 21815286 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21815286&pc=OZD | 03/17/2025 | 05/14/2025 |  | 04/10/2025 |  |  | 03/17/2025 | 04/10/2025 |  |  |  |  | 04/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1133148 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that when placing the new impella cp, the customer was prompted to connect grey-to-grey plugs by the automated impella controller (aic). the customer was unable to move past that prompt. an alternate aic was used. there was no patient harm. | [5 empty] | The investigation of pump not detected has been completed. the data analysis confirmed the reported complaint. the console was returned and 10 times the console was put through case start to properly connect with a pump and all ten times was successful without the failure mode reproducing. root cause: the pump detection issue was due to the rarely occurring epm crystal issue. the impellatronic pca does not have permanent damage, but this intermittent pump detection failure occurs infrequently and already has a sw mitigation in place. h6 component code 925 was erroneously entered on the initial manufacturer device report number. | The impella automated controller was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
10423 | 1220648-2025-27372 | 21815443 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21815443&pc=OZD | 03/12/2025 | 04/10/2025 |  | 04/10/2025 |  |  | 03/14/2025 | 04/10/2025 |  |  |  |  | 04/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025599950 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Other | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced dark red/brown urine. the team was concerned for hemolysis. patient was on right ventricular assist device, impella and continuous renal replacement therapy. per the team, they believe that the plasma free hemoglobin could be falsely elevated because patient received a dose of cyanokit. the patient was successfully explanted from impella support. the patient continued on extracorporeal membrane oxygenation support. the patients condition was noted to be critical. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
10424 | 1220648-2025-27373 | 21815617 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21815617&pc=OZD | 03/12/2025 | 10/01/2025 |  | 04/10/2025 |  |  | 03/12/2025 | 04/10/2025 |  |  |  |  | 04/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025592783 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Sepsis | Tachycardia | Ventricular Fibrillation | [10 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella 5.5 as an elective placement and during support, the patient experienced a ventricular tachycardia and fibrillation arrests and developed sepsis. the patient presented and was planned for coronary artery bypass graft surgery. however, upon the initial transesophageal echocardiogram, severe mitral regurgitation was noted, and the decision was made to place an impella 5.5. with coronary artery bypass graft and mitral valve replacement. the patient arrived at the unit hemodynamically unstable. three days later, the patient was placed on continuous renal replacement therapy and continued on impella mechanical circulatory support (mcs). a couple days thereafter, the impella performance level was decreased from p-8 to p-6 with improvement in hemodynamics. approximately seven days later, the patient was intubated after having a ventricular tachycardia arrest. the patient was also now, septic with two positive blood cultures (relationship to the impella is unknown). optic sensor in impella was drifting, and education was performed with staff about not following the aortic placement for pressure monitoring and what to look for regarding potential pump saturation as patient had become coagulopathic. the patient continued on support and was noted to be doing well. however, approximately six days after the first arrhythmic event, the patient had a ventricular fibrillation arrest. the patient was defibrillated and converted with one shock. amiodarone was started and an echocardiogram was completed done at bedside to ensure placement of the impella. after a total of 25 days on support, the patient was successfully weaned from the impella mcs with a survived outcome. | Us complainant reported the patient was on impella 5.5 support and during support, there were placement signal issues. the optic sensor was showing erroneously high ao placement pressure. support continued and staff monitored. |  |  |  |  | The investigation of vf/vt/af/at and infection/sepsis has been completed. vt/vf/at/as": cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease. timing of the arrythmia event in relation to impella support (during or post-implant). electrocardiogram (ekg) recordings of the arrhythmia episodes. evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances. assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements. presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities. response to any antiarrhythmic treatments or interventions during the event. any documented device-related issues or complications during impella support. evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand. review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. infection/sepsis: an infection/sepsis can occur during or after impella support. when diagnosing the cause of an infection in a patient, the following clinical evidence would need to be considered holistically: patient symptoms and medical history obtained by the treating clinician. physical examination findings. results of special investigations: laboratory test results and imaging studies. microbiological culture and sensitivity results. response to previous treatments and therapies. any relevant epidemiological information or exposure history. concomitant devices in use. preexisting patient condition such as underlying infection, remote septic sources such as indwelling lines/catheters or other bioprosthetic material in situ. care of access site. because this evidence has not been provided, the root cause of the infection cannot be determined. | The pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock & cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. optical signal issue: insufficient data logs were provided. the returned product was inspected and there were no visual abnormalities. the 5s parameters were within specification. the pump passed the laser continuity test. the complaint occurrence date was 3/12. the logs were returned for case dates: 02/27 to 03/04. there were no abnormalities seen in the 5s parameters of the returned logs. the root cause of the optical signal issue was unable to be determined as there were no issues with the returned device and insufficient data logs were provided. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical signal issue "2917" as part of a retrospective review of optical signal issues in accordance with fda recommendation.
10425 | 1220648-2025-27374 | 21815795 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21815795&pc=OZD | 03/11/2025 | 06/11/2025 |  | 04/10/2025 |  |  | 03/11/2025 | 04/10/2025 |  |  |  |  | 04/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1761452 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Select Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | Abiomed service depot field service operation representative reported one of the automated impella controller (aic) service loaners on site was having ¿disc insert¿ issues. it was noted that the aic had been running on a case without an issue for a while before displaying an ¿insert disc¿ error. repeated attempts were made to reseat the cassette, and the disc did not resolve the issue. as a result, the user facility opted to exchange the aic, and the issue was resolved. there was no report or indication of patient harm. | [5 empty] | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. the logs from the complaint date revealed that the console issued purge disc not detected alarm several times. the console was examined, and a broken purge disc flag was identified. the purge disc flag was observed to be broken and was the root cause of the reported purge disc not detected alarm issue e3 occupation was erroneously selected on the initial manufacturer device report number 1220648-2025-27374. | Imc logs from the complaint date revealed that the console issued purge disc not detected alarm (event set #402) multiple times. (b)(6). device analysis: the console was examined after arriving in fs. a bent/broken purge disc flag was identified. (b)(6). conclusion: the purge disc flag was observed to be broken and was the root cause of the inability to detect the purge disc. (b)(4) documents the possible root causes of a fracture of the purge flag which is not tied to a manufacturing or design defect, typically caused by fluid ingress into the purge pressure sensor assembly. there is a small amount of glucose residue in (b)(6) purge pressure sensor assembly. repair: before sending this console back to the customer, field service was instructed to replace the purge disc detect flag [3001198] and purge flag spring [3001196], validate sensor accuracy via sections 10-12 of the pm procedure [0042-7309], and perform a full functional check on the console and optical system [0042-7316]. | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
10426 | 1220648-2025-27375 | 21815910 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21815910&pc=OZD | 03/14/2025 | 05/15/2025 |  | 04/10/2025 |  |  | 03/14/2025 | 04/10/2025 |  |  |  |  | 04/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025611136 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 23 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A physician reported a pump stop on an impella cp after a spike in motor current appeared. the pump was not able to be restarted, so the pump was explanted. there was no reported patient harm. | [5 empty] | The investigation of pump stop has been completed. the impella cp was returned for evaluation. upon visual inspection, a large purge gap was present. no data logs were returned for analysis. clinical details noted that the pump stop occurred on day 14 of support. the root cause of the pump stop was due to bearing wear caused by the pump being used beyond intended design life of 8 days per design trace matrix. | The pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
10427 | 1220648-2025-27376 | 21815928 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21815928&pc=OZD | 03/13/2025 | 04/10/2025 |  | 04/10/2025 |  |  | 03/14/2025 | 04/10/2025 |  |  |  |  | 04/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025610985 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that impella cp stopped while supporting a patient. the cp was exchanged. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
10428 | 1220648-2025-27377 | 21816127 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21816127&pc=OZD | 03/14/2025 | 04/14/2025 |  | 04/10/2025 |  |  | 03/14/2025 | 04/10/2025 |  |  |  |  | 04/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1123072 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that when the automated impella controller (aic) was booted up. the aic got very hot, and it smell like burning with a little smoke coming off the vent. no patient involvement. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - a/c power issues has been completed. the console logs were examined as a precaution but did not contain any data relevant to this complaint. the console was tested and during inspection of the console, no evident burning smell or burnt internal components were observed. the i/o fuses were checked but were still intact. conclusion: no problem found while testing and inspecting this console. the console operated as intended g1 reporting contact email revised on manufacturer device report 1220648-2025-27377 in accordance with updated procedures.
10429 | 1220648-2025-27380 | 21816564 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21816564&pc=OZD | 03/16/2025 | 06/02/2025 |  | 04/10/2025 |  |  | 03/16/2025 | 04/10/2025 |  |  |  |  | 04/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025603694 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  |  |  | Death | Yes | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | Per medical safety officer review; the pump stopped during impella cp pump after 12 days start of support requiring replacement and was the cascading event which cannot be dissociated. there is not enough evidence to exclude the impella cp pump 1 as an associated factor to the patient's outcome of death. | The complainant reported the patient was on impella cp support and during support, there was a pump stop. staff attempted to restart the pump to no avail. the pump was explanted and replaced. the patient remained stable and the procedural outcome was the patient survived surgery. |  |  |  |  | The investigation of pump stop has been completed. the impella cp returned for evaluation. upon visual inspection, large purge gap was found. data logs were reviewed and a high motor current pump stop with alarm occurred on 11th day. multiple attempts were made to restart pump without success. the root cause of the pump stop was due to bearing wear as pump was run past indicated design life. | Per the medical safety officer review; the pump stopped during impella cp pump 1 mcs after 12 days start of support requiring replacement and was the cascading event which cannot be dissociated. there is not enough evidence to exclude the impella cp pump 1 as an associated factor to the patient's outcome. based on the medical safety officer (mso) review the following fields have been updated: b1 revised from the initial manufacturer device report number 1220648-2025-27380 to reflect both product problem and adverse event. b2 revised from the initial manufacturer device report number 1220648-2025-27380 to reflect death and date of death b5 revised from the initial manufacturer device report number 1220648-2025-27380 to reflect patient outcome and mso statement h1 type of reportable event revised from the initial manufacturer device report number 1220648-2025-27380 to reflect death h6 health effect - impact code revised from the initial manufacturer device report number 1220648-2025-27380 to reflect death. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk cpi section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
10430 | 1220648-2025-27381 | 21816641 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21816641&pc=OZD | 03/14/2025 | 04/21/2025 |  | 04/10/2025 |  |  | 03/14/2025 | 04/10/2025 |  |  |  |  | 04/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026623932 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during impella cp removal from a patient, a moderate amount of hematoma was noted which was evacuated. additionally, a distal clamp was placed prior to removal to prevent any clot burden from flowing down the right common femoral artery. clot burden was noted and embolectomy was done. the patient survived the procedure. | [5 empty] | The investigation of thrombus and access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: impella cp with smart assist catheter insertion, section: axillary insertion of the impella cp with smart assist catheter, section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
10431 | 1220648-2025-27382 | 21816754 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21816754&pc=OZD | 03/16/2025 | 04/10/2025 |  | 04/10/2025 |  |  | 03/16/2025 | 04/10/2025 |  |  |  |  | 04/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026617681 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the patient experienced limb ischemia. the patient's impella leg was livid and had poor circulation. catecholamine could not be reduced and there was no vascular surgeon on site. the physician did not want to do a crossover by pass because the patient was not able to be transferred. as the patient was not weanable from the impella, it was recommended to remove the pump and implant a new one on the opposite side. care was withdraw and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
10432 | 1220648-2025-27383 | 21816757 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21816757&pc=OZD | 03/15/2025 | 04/11/2025 |  | 04/10/2025 |  |  | 03/15/2025 | 04/10/2025 |  |  |  |  | 04/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025606267 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient had skin tract oozing at their impella cp access site during support. a purse string suture was placed and one unit of packed red blood cells was given to treat the condition. additionally, heparin was temporarily paused and the site was redressed and oozing was resolved. the following day, care was withdrawn and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the root cause of access site bleeding was most likely use issue since hemostasis was achieved by adding additional suture. h.6 investigation findings and investigation conclusions were reported incorrectly on the initial manufacturer device report number 1220648-2025-27383 as the investigation was complete. h.10 investigation results were inadvertently omitted on the initial manufacturer device report number 1220648-2025-27383.
10433 | 1220648-2025-27384 | 21816844 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21816844&pc=OZD | 03/11/2025 | 05/01/2025 |  | 04/10/2025 |  |  | 03/16/2025 | 04/10/2025 |  |  |  |  | 04/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026617571 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | Embolism/Embolus | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an implanted impella 5.5 pump began to experience persistent suction during patient support. as repositioning was unable to resolve the suction, the decision was made to explant the pump. there were no overt signs of cannula obstruction in the device upon removal. it was later reported via abiomed¿s long-term outcome and quality indicator (loqi) impella registry, that the patient also experienced thrombocytopenia, recurring non-sustained ventricular tachycardia (vt), and recurring pleural bleeding during support. the thrombocytopenia and vt were believed to have a probable relationship to the impella device. the vt was treated with amiodarone and electrolyte replacement. the pleural bleeding was believed to have a possible relationship to the device. the patient was taken for a washout/exploration, the site was packed, and hemostasis was performed. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of low pump flow, thrombocytopenia, vf/vt/af/at, and vascular damage - peripheral vessel has been completed. the returned complaint cp device visually inspected. biomaterial observed in purge gap and around impeller. no broken blade , no cannula kink observed. no other issue found. impella washed to dissolve the dried blood around impeller. biomaterial was buildup and there was no evidence of ingestion. low pump flow: the cause of the low pump flow was biomaterial ingestion. thrombocytopenia: the cause of the thrombocytopenia issue was not determined due to insufficient clinical details. vascular damage - peripheral vessel: the cause of the vascular damage - peripheral vessel issue cannot be determined since complaint device not returned for analysis and insufficient clinical data provided. as the necessary information was not provided, the root cause of the vt/vf was not determined. failure mode thrombus was added. as the necessary information was not provided, the root cause of the thrombus was not determined. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-27384. g2 report source; study was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27384. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10434 | 1220648-2025-27386 | 21817697 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21817697&pc=OZD | 03/27/2025 | 06/26/2025 |  | 04/10/2025 |  |  | 03/31/2025 | 04/10/2025 |  |  |  |  | 04/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2025611709 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 80 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Tamponade | Bronchial Hemorrhage | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the 80-year-old male patient had a cp pump support placed during cpr for support due to myocarditis or pericarditis. the pump was placed while undergoing cpr but, the patient expired from pulmonary hemorrhage and cardiac tamponade. it was thought that cpr caused pulmonary hemorrhage and cardiac tamponade, and the death had no relationship to impella pump use. however, there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of vascular damage - peripheral vessel and ventricle perforation has been completed. the impella device was not returned for evaluation. vascular damage - peripheral: clinical details report that pulmonary hemorrhage occurred as a result of cardiopulmonary resuscitation (cpr). product was not returned for investigation and data logs are not relevant to the complication. since clinical details report that the injury occurred due to cpr, the root cause of the vascular damage was determined to most likely be a use issue. ventricle perforation: clinical details report that impella was inserted while cpr was being done. as a result of cpr, a cardiac tamponade occurred. data logs were reviewed, and there were no positioning alarms during the case. product was not returned for investigation. based on the clinical details provided that the cardiac tamponade was a result of cpr, the root cause of the ventricle perforation was determined to most likely be a use issue. b3 date of event was erroneously entered on the initial manufacturer device report number: 1220648-2025-27386. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations: "assess access site for bleeding and hematoma." section: entering cardiac output: use of the repositioning sheath and peel-away introducer: "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site." section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)."
10435 | 1220648-2025-27387 | 21819197 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21819197&pc=OZD | 04/01/2025 | 04/10/2025 |  | 04/10/2025 |  |  | 04/01/2025 | 04/10/2025 |  |  |  |  | 04/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Implant Pain | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Abiomed received a literature article titled "a case of acute thrombotic occlusion in the femoral artery due to impella removal" where it was reported that a patient developed a thrombosed leg while on cp pump support. the patient had coronary cardiac procedure done and then when the cp was explanted the pump limb was seen to be painful and thrombosed. the leg was treated by thrombectomy and stenting, in addition to medical therapy. the patient was discharged alive to the rehab hospital. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ missing details due to limited complaint details provided in literature article: a.1-a5 are unknown. d.4 serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. e.1 name prefix/title, first name and last name, facility name are unknown. h.4 device manufacture date is unknown.
10436 | 1220648-2025-27388 | 21819359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21819359&pc=OZD | 04/01/2025 | 09/22/2025 |  | 04/10/2025 |  |  | 04/01/2025 | 04/10/2025 |  |  |  |  | 04/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | UNKNOWN |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Increase in Pressure | [6 empty] | Thrombocytopenia | No Clinical Signs, Symptoms or Conditions | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | It was reported via a literature entry "a case of severe heart failure using bicarbonate based impella purge solution as an alternative to heparin" that an implanted impella cp pump developed low purge flow prompting pump replacement. there was no noted patient harm. | [5 empty] | Instructions for use: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b1 adverse event was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27388. b2 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27388. h1 type of reportable event was erroneously selected on the initial manufacturer device report number 1220648-2025-27388. h6 health effect - clinical code 4582 was erroneously entered on the initial manufacturer device report number 1220648-2025-27388. h6 health effect - impact code 4614 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27388. h10 instructions for use were inadvertently omitted on the initial manufacturer device report number 1220648-2025-27388. no changes in investigation results previously reported in manufacturer device report number 1220648-2025-27388-1. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the following sections are blank as little details were provided in literature article: d.4 model number, catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. h.4 device manufacture date is unknown. | No product or data logs were returned for evaluation. high purge pressure: clinical details noted that impella cp was used with argatroban as a purge solution and decreased purge flow was observed. pump was replaced and purge was switched to sodium bicarbonate and purge flows and pressures were able to be maintained. the root cause of the high purge pressure could not be definitively determined as no product or log were returned for evaluation. thrombocytopenia: a failure mode of "thrombocytopenia" was added to the investigation based on clinical details that patient was experiencing heparin-induced thrombocytopenia while on impella support. the cause of the thrombocytopenia was not determined due to insufficient clinical details.
10437 | 1220648-2025-27389 | 21822142 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21822142&pc=OZD | 03/29/2025 | 04/10/2025 |  | 04/10/2025 |  |  | 03/29/2025 | 04/10/2025 |  |  |  |  | 04/10/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2025548808 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient with confirmed rca occlusion and lad stenosis that arrested was placed on impella cp and ecmo. it was noted that the impella was placed in the left ventricle without any problems, but because the patient was not sufficiently sedated, it was pulled into the aorta. for safety reasons, the same device was not reinserted, and a new cpsa was placed. there were no problems with the device, and there were no issues with its delivery or placement. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
10438 | 1220648-2025-27313 | 21804493 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21804493&pc=OZD | 03/10/2025 | 04/09/2025 |  | 04/09/2025 |  |  | 03/10/2025 | 04/09/2025 |  |  |  |  | 04/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025592268 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Ischemia | [11 empty] | Other | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and experienced limb ischemia, atrial fibrillation and subsequently passed away. left heart catheterization showed chronic total occlusion to the right coronary artery (rca) graft. an intra-aortic balloon pump (iabp) was placed at the cardiothoracic surgeons request. the patient continued to decline requiring intubation, levophed/vasopressin/amiodarone to be added. cardiothoracic surgery team turned the patient down for surgery due to the patient's instability. the team elected to exchange the iabp for an impella cp to optimize. the patient required two rounds of cardiopulmonary resuscitation between iabp discontinuation and the impella cp placement. temporary pacer was placed. the patient was sent to the unit on support. the same day, it was reported that the patient had no pulses lower extremity bilaterally. two days later, vasopressin was removed. no escalation to an impella 5.5 was determined. however, coronary artery bypass graph (rca) redo is needed when the patient is more stable. later in the evening, it was reported the nurse was unable to obtain pulses. the patient experienced paroxysmal atrial fibrillation with rapid ventricular response. the following morning, the patient was made do not resuscitate and the patient would expire. care was withdrawn. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10439 | 1220648-2025-27314 | 21804569 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21804569&pc=OZD | 03/13/2025 | 08/30/2025 |  | 04/09/2025 |  |  | 03/13/2025 | 04/09/2025 |  |  |  |  | 04/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026625250 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | User facility | Company representation |  |  |  | 61 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | Renal Failure | Thrombosis/Thrombus | Unspecified Vascular Problem | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient encountered both limb ischemia and thrombus. distal limb flow was checked via side port of repositioning sheath after advancement with minimal dye seen. distal reperfusion was flushed and blood did advance into the patient's left leg. bilateral legs were equal in color and temperature despite not being able to find pulses in left lower extremity in catheterization lab. additionally, a clot on the inlet of the impella was observed on the echocardiogram. the doctor was worried that removing the impella while the clot was there would embolize it. the plan was made to give the patient lytics after two units of packed red blood cells in and lab recheck. heparin bolus was given to the patient for the inlet clot. the procedural outcome was the patient survived the surgery. | "we received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured acute kidney injury with a "possible" relationship to the impella device used: ¿on (b)(6) 2025, during index hospitalization, while on impella support, patient developed aki due to cardiogenic shock; possibly related to atn secondary to cardiogenic shock, intravascular hemolysis. baseline cr on admit was 1.04 ((b)(6) 2025); cr 2.50 ((b)(6) 2025). impella setting de-escalated to p2, weaning on the pressors. plan: administer iv furosemide 20 mg x 1, follow-up urine output, and if remains suboptimal we will initiate crrt. her urine output has been minimal even with high-dose diuretics, nephrology recommended starting the patient on crrt. patient became anuric and crrt was initiated. cr improved to 1.31 ((b)(6) 2025),1.30 ((b)(6) 2025). crrt discontinued on (b)(6) 2025 and transitioned to dialysis. cr 3.25 ((b)(6) 2025). transitioned to ihd (b)(6) 2025. right ij tunneled dialysis catheter placed (b)(6) 2025. patient is anuric and dialysis dependent. on (b)(6) 2025, family conference to discuss inability to procure funding for dialysis and cares in the u.s. plan is to obtain a flight to patients¿ country as late in the day as possible. dialysis to be performed just prior to her flight home. patient will follow up at a hospital in her country. continue weekly dialysis per nephrology recommendations. patient discharged from hospital on (b)(6) 2025, cr 4.11 at discharge. patient came back for last dialysis treatment on (b)(6) 2025 due to a flight delay. cr 4.55 ((b)(6) 2025). on (b)(6) 2025, nephrology recommendation: reportedly passing more urine, plan to continue 3 times a week dialysis, when she reaches her home country, she will need to be reassessed for renal recovery. continue avoiding nephrotoxins, dialysis today with uf as tolerated, remains dialysis dependent.¿. the patient not recovered/not resolved." |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves, ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation of thrombus and limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the investigation hemolysis, renal failure, and vascular damaged has not been completed. should any new information be received, a supplemental manufacturer device report will be filed. b5 additional information received from the clinical site h6 health effect - clinical codes added based on the additional information received from the clinical site. instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ g2 report source; study was missing from manufacturer device report 1220648-2025-27314.
10440 | 1220648-2025-27315 | 21804612 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21804612&pc=OZD | 03/12/2025 | 04/25/2025 |  | 04/09/2025 |  |  | 03/12/2025 | 04/09/2025 |  |  |  |  | 04/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025611406 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Nerve Damage | Swelling/ Edema | Limb Pain | [9 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and bridge to transplant. the next day, the patient experienced self-limiting ectopy. seven days later, the patient experienced swelling and pain in the impella arm. there was concern for nerve injury due to impella. two days thereafter, it was noted the swelling significantly decreased. action taken to treat is unknown. however, support continues as the patient waits for heart and liver transplants. | [5 empty] | The device remains implanted, supporting the patient at the time of the mdr writing, and therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of peripheral nerve injury and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6b explant date added.
10441 | 1220648-2025-27316 | 21804718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21804718&pc=OZD | 03/12/2025 | 10/01/2025 |  | 04/09/2025 |  |  | 03/12/2025 | 04/09/2025 |  |  |  |  | 04/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025601986 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ectopic Heartbeat | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The user facility reported a patient was on impella 5.5 support and during support, the patient experienced ectopy. the p level was reduced and the pump was repositioned. support was continued. additionally, the patient had a hematoma at the impella site. computed tomography showed arterial ooze, suspected from the pectoral muscle. a washout was performed at the access site. the patient survived the procedure. | The customer reported that the device exhibited placement signal disruption causing inaccurate positioning alarms. |  |  |  |  | The investigation of vf/vt/af/at and access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. optical signal issue: due to a lack of data logs returned and product, the root cause of the optical signal issue could not be determined. vt/vf: cardiac arrythmia can be reported during impella support. to make a determination as to the cause of the cardiac arrythmia, the following clinical information must be considered: patient's medical history, including any pre-existing cardiac conditions, previous episodes of arrhythmias, or structural heart disease timing of the arrythmia event in relation to impella support (during or post-implant) electrocardiogram (ekg) recordings of the arrhythmia episodes evaluation of any underlying electrical disturbances, such as qt prolongation or electrolyte imbalances assessment of hemodynamic instability during the arrhythmia, including blood pressure and cardiac output measurements presence of other potentially contributory factors, such as medication changes, ischemia, or electrolyte abnormalities response to any antiarrhythmic treatments or interventions during the event any documented device-related issues or complications during impella support evaluation of potential triggers or precipitating factors, such as catheter manipulation, reperfusion injury, or increased myocardial oxygen demand review of any concomitant medications that may influence cardiac electrophysiology with a returned impella, the device can be inspected for any burrs or rough edges that may contact the heart wall. to determine the true root cause of the vt/vf, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the vt/vf was not determined. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical signal issue "2917" as part of a retrospective review of optical signal issues in accordance with fda recommendation.
10442 | 1220648-2025-27317 | 21804872 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21804872&pc=OZD | 03/12/2025 | 04/22/2025 |  | 04/09/2025 |  |  | 03/12/2025 | 04/09/2025 |  |  |  |  | 04/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026617653 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during support, persistent suction alarms were present. support above p2 to p3 was unable to be achieved without triggering a suction alarm regardless of placement and filling pressures. the pump was replaced. no patient harm was reported. | [5 empty] | The investigation of low pump flow has been completed. the pump was placed and run with suction alarms in the catheterization lab and resolved by fluid compensation. suction alarms were triggered again in the unit and the pump was unable to run above p-3. the pump then was replaced. data was not returned for review. the pump was returned for analysis. visual inspection found no issues on the pump. the failure mode could not be reproduced as pump ran without issue during bench test. the root cause of the low flow was most likely a patient condition since there was no issues with the returned product and the clinical details stated that volume was given to the right ventricle and the suction alarms resolved. | The pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
10443 | 1220648-2025-27318 | 21804982 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21804982&pc=OZD | 03/12/2025 | 09/18/2025 |  | 04/09/2025 |  |  | 03/12/2025 | 04/09/2025 |  |  |  |  | 04/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1619563 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | Unintended Application Program Shut Down | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. impella in aorta alarm appeared despite the device having proper positioning in the left ventricle. the physician made the decision to pull the pump and rewire the left ventricle, but the staff was unable to turn the impella off. despite turning support down to p0, the motor current would not completely shut off. the pump was replaced in response to the noted issue. there was no patient harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. the investigation of automated impella controller (aic) - shutdown or restart issue and optical signal issue has been completed. since there was no log evidence indicating that the user attempted to reduce the p-level to p0 before the impella was unplugged, the root cause of the claim ¿unable to turn impella off despite setting it to p0¿ was most likely a user issue, which was misreported as ¿aic - shutdown or restart issue.¿ there was no record of an unexpected aic shutdown or restart on the complaint date. there was no optical-related hardware malfunction on the console. the root cause of the optical signal issue was most likely the patient¿s condition. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d9 device returned to manufacturer date added
10444 | 1220648-2025-27319 | 21805010 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21805010&pc=OZD | 03/13/2025 | 05/09/2025 |  | 04/09/2025 |  |  | 03/13/2025 | 04/09/2025 |  |  |  |  | 04/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026615110 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Positioning Failure | [6 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Us complainant reported the patient was on impella 5.5 support and after the implant, the pump shifted towards septal wall limiting its flow. kink occurred on catheter that did not allow forward maneuver. final position was at 3.4 cm toward mitral valve. the patient's family decided to withdraw care and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of positioning issues has been completed. product was not returned for investigation. data log review was not required for this case run. according to the clinical notes, the doctor was not able to reposition the pump after it became shallow during the advancement of the blue hub. doctor was able to reposition the pump. the cause of the positioning issue was most likely use related since the pump position was compromised during advancing the blue cath-anchor.
10445 | 1220648-2025-27320 | 21805101 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21805101&pc=OZD | 03/10/2025 | 04/09/2025 |  | 04/09/2025 |  |  | 03/10/2025 | 04/09/2025 |  |  |  |  | 04/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026618539 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was initially placed on bipella support (impella cp and impella rp flex) for one day, then left sided support only for three days before being escalated to an impella 5.5 device for continued mechanical circulatory support. approximately four days after the start of the impella 5.5 support, the patient's urine output was pink-tinged. the following day, the urine was still pink/red tinged with small clots. diuretics were discontinued. albumin was given. the patient continued on support. two days later, the patient was reintubated for airway concerns and falling oxygen saturations. continuous veno-venous hemofiltration was started. however, the patient would expire a couple of days later. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
10446 | 1220648-2025-27321 | 21805293 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21805293&pc=OZD | 11/06/2024 | 04/09/2025 |  | 04/09/2025 |  |  | 03/13/2025 | 04/09/2025 |  |  |  |  | 04/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025560824 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 70 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured access site bleeding with a "definite" relationship to the impella device used: ¿participant developed femoral artery access site bleeding and hematoma during impella-supported procedure requiring two units prbcs, iv fluids, and epinephrine to maintain blood pressure. two additional units prbcs received in icu plus albumin. bleeding noted as "abated" (b)(6) 2024. pre-procedure hgb 11.0 g/dl, post-procedure and transfusion hgb 12.6 g/dl.¿. the patient recovered with no sequelae. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10447 | 1220648-2025-27322 | 21805590 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21805590&pc=OZD | 03/11/2025 | 04/09/2025 |  | 04/09/2025 |  |  | 03/11/2025 | 04/09/2025 |  |  |  |  | 04/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025592994 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support and same day post implant experienced bleeding from the impella cp access site. manual pressure was held above the site for approximately 30 minutes; however, the bleed persisted. an existing perclose suture was then cinched down and the bleeding "immediately" stopped. the site was redressed and support continued uninterrupted for 11 days before the device was successfully weaned and explanted at bedside using existing perclose sutures and manual pressure with "excellent" hemostasis. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
10448 | 1220648-2025-27323 | 21805769 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21805769&pc=OZD | 03/12/2025 | 04/09/2025 |  | 04/09/2025 |  |  | 03/12/2025 | 04/09/2025 |  |  |  |  | 04/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1731823 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Select Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automated impella controller (aic) would not be operational with purge cassettes. further details are unknown. however, it was indicated the complaint is not related to a clinical procedure. thus, no patient involvement. | [5 empty] | The automated impella controller (aic) data logs and console were returned for evaluation and analysis. data analysis of the console logs was consistent with the reported complaint. multiple instances of purge fluid were not detected. errors were observed in the imc logs. rtlog data showed that just a couple seconds before one of the purge fluids not detected errors, the purge pressure peaked at less than 15mmhg, which is very low. device analysis, the console was tested after arriving in field service. no purge cassette issues were able to be reproduced. the purge disc retainer/flag were examined but no issues were found. console was able to detect the purge cassette and purge fluid without any issues. product history review as completed and not action was required as a result of this review. in conclusion, the root cause of the reported purge cassette issue could not be determined due to the inability to reproduce. purge pressure was very low which could be indicative of a kink in the purge line or a possible leak but no there was no conclusive evidence. console operated as intended while testing.
10449 | 1220648-2025-27324 | 21805876 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21805876&pc=OZD | 03/12/2025 | 06/24/2025 |  | 04/09/2025 |  |  | 03/12/2025 | 04/09/2025 |  |  |  |  | 04/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026624386 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Detachment of Device or Device Component | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient began to complain of lower extremity pain and nurse was unable to find doppler pulses anywhere on the leg. additionally, after impella cp was placed the cath lab staff locked device. in icu then noted touhy bourst valve was not locked, nurse proceeded to tighten it. upon arrival to the higher care center, it was noted that touhy bourst valve again was not locked. upon tightening nurse noted connecting point of touhy bourst valve would continue to turn 180 degree and attempted to rotate back end of sleeve to loosen tension by red plug and was unable to rotate sterile sleeve. it was further reported that the patient¿s urine had become red tinged upon arrival to the higher care center. the impella was explanted due to complications. | [5 empty] | The investigation of product damage (sleeve damage), limb ischemia, and hemolysis has been completed. product damage - sleeve: the returned product was investigated and the touhy-borst valve was returned loose. however, it was able to be tightened and the catheter was locked in place while the valve was tight. the same was done to the valve on the proximal end of the sterile sleeve, and again, there were no abnormalities noted. it is unclear what was causing the team in the field to struggle to tighten the valves, however, based on returned product functioning properly, the root cause of the product damage was determined to most likely be a use issue. limb ischemia: as all the necessary information was not provided, the root cause of the limb ischemia was not determined. hemolysis: clinical details report that when being assessed after hospital transfer, the patient was exhibiting hemolysis symptoms by means of red urine. pump position was checked, and the pigtail was contacting a papillary and the pump inlet was at the mitral valve. blood products were not administered, and hemolysis symptoms resolved after the pump was explanted. data log review showed that in the last 30 minutes of support, motor current modulation dropped and pump flows decreased while the pump was at p-7. there was a 30-second position in aorta alarm early in the case, however, it's does not appear to be related to the complication. additionally, the placement signal (ps) was trending down while running at p-7, potentially shifting closer to the ventricle, supporting the report that the pump was contacting the mitral apparatus. however, it was not so deep to trigger a position in ventricle alarm. returned product was investigated and there were no visual abnormalities noted. based on the report of the pump contacting the mitral apparatus, trends of pump metrics in data logs, no visual abnormalities on returned product, and symptoms resolving after the pump was explanted, the most likely root cause of hemolysis was improper positioning. d9 date device returned to manufacturer added. h6 component codes 3038 and 4723 were inadvertently omitted on the initial manufacturer device report number 1220648-2025-27324. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
10450 | 1220648-2025-27325 | 21805904 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21805904&pc=OZD | 03/22/2025 | 04/09/2025 |  | 04/09/2025 |  |  | 03/25/2025 | 04/09/2025 |  |  |  |  | 04/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025581017 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 76 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on impella cp support and during support, the patient had a hematoma. the patient's dynamics were stable so the pump was removed. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10451 | 1220648-2025-27326 | 21806035 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21806035&pc=OZD | 03/12/2025 | 05/07/2025 |  | 04/09/2025 |  |  | 03/12/2025 | 04/09/2025 |  |  |  |  | 04/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1719481 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | User facility | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported by biomed at the user facility that the automated impella controller (aic) arrived with water damage. the water damage was of unknown origin and was on/in the box and console. there was no patient involvement. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - damage before therapy has been completed. the console was powered up on the complaint date, the date/time was set then the console was turned off. there was nothing in the logs confirming nor denying the water in the shipping box. the console returned for testing and performed as expected during testing as well as in the field. the cause of the water logged console was most likely from shipping. d9 date device returned to manufacturer added.
10452 | 1220648-2025-27327 | 21806102 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21806102&pc=OZD | 03/14/2025 | 06/10/2025 |  | 04/09/2025 |  |  | 03/14/2025 | 04/09/2025 |  |  |  |  | 04/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1364511 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that during regular inspection of the automated impella controller, it was confirmed that the optical bench check signal was outside the specifications. there was no patient involvement. | [5 empty] | After evaluation, the complaint does not meet the criteria for mdr reportability as defined by 21 cfr 803.50. the issue will be tracked and trended internally. | The console device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
10453 | 1220648-2025-27328 | 21806270 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21806270&pc=OZD | 03/09/2025 | 04/09/2025 |  | 04/09/2025 |  |  | 03/14/2025 | 04/09/2025 |  |  |  |  | 04/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025594993 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support and during support, there was a pump stop. the pump was used beyond the expected period of use. attempts were made to restart the pump to no avail. the pump was explanted. there was no harm to the patient. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention & impella rp with smartassist system section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
10454 | 1220648-2025-27329 | 21806453 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21806453&pc=OZD | 03/14/2025 | 04/30/2025 |  | 04/09/2025 |  |  | 03/14/2025 | 04/09/2025 |  |  |  |  | 04/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026613503 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  |  |  | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support, there was a purge leak from the purge sidearm. purge flow had been fluctuating on this device since insertion. purge pressure was in the 500's consistently. support was continued. weaning was successful. the device was explanted and the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of purge leak - pump has been completed. the returned product was severed at the catheter and the cannula was not returned. the rest of the pump was inspected and there were no physical abnormalities. there was no observed damage to the purge sidearm luer connection or filter. the severed pump was purged and the fluid ran through the sidearm without leaks and the fluid exited the purge tubing at the severed point. the reproduction test was carried out for 24 hours and no leak was observed. purge flow values were stable at 30ml/hr due to severed pump/. there was no leak observed. the data logs from the case show a slight downwards trend in purge flow values before a bag change on 3/14 at 5:20am. leak location unable to be identified from data logs. the root cause of the purge leak - pump was not determined as the leak was not reproduced and limited information related to failure was provided. | The pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: cleaning. ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
10455 | 1220648-2025-27335 | 21806558 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21806558&pc=OZD | 03/25/2025 | 04/09/2025 |  | 04/09/2025 |  |  | 03/25/2025 | 04/09/2025 |  |  |  |  | 04/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2025594828 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a 77-year-old male patient developed hemolysis during impella cp support. when the condition did not improve, the decision was made to remove the pump. an intra-aortic balloon pump was subsequently placed for ongoing support. the patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
10456 | 1220648-2025-27330 | 21806588 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21806588&pc=OZD |  | 04/09/2025 |  | 04/09/2025 |  |  | 03/11/2025 | 04/09/2025 |  |  | 11/30/2024 |  | 04/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023281005 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Literature | Health Professional | User facility | Company representation |  |  | 75 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | Life Threatening | [22 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | Abiomed japan forwarded a publication titled "echocardiography-guided percutaneous ventricular assist device for postinfarct ventricular septal rupture," which noted a patient in cardiogenic shock from an acute myocardial infarction was implanted with an impella 5.5 and placed on extracorporeal membrane oxygenation (ecmo) support. in summary, the patient experienced "uncontrolled bleeding in setting of systemic anticoagulation, leading to hemorrhagic shock." care was withdrawn, and the patient expired. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. the actual or best estimate date of the event is unknown/could not be determined. consequently, section b5 was left blank. publication reference: avinash d¿souza, do et. al (2021). "echocardiography-guided percutaneous ventricular assist device for postinfarct ventricular septal rupture." cardiovascular imaging case reports. december 2024 volume 8 number 12.
10457 | 1220648-2025-27331 | 21806682 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21806682&pc=PYX | 03/14/2025 | 05/19/2025 |  | 04/09/2025 |  |  | 03/14/2025 | 04/09/2025 |  |  | 01/31/2026 |  | 04/09/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025495229 |  | 0046-0011 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  | White | Malfunction | No | Yes | Defective Device | Material Split, Cut or Torn | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that it was not possible to push the impella rp over the guidewire into the pulmonary artery during the impella rp insertion. after removal of the impella rp, it was noticed that the sheath was broken on the tip, and a bend in the wire on fluoroscopy was seen. the sheath, guidewire, and impella rp were exchanged. a new impella rp was successfully implanted. there was no patient harm. | [5 empty] | The pump and introducer were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of introducer damage ¿ mechanical and product damage (guidewire damage - peripheral vessels) has been completed. introducer damage - mechanical: the pump and introducer were returned and inspected. the introducer was noted to be kinked and a piece of it was chipped off at the bottom. the chipped piece was found lodged in the inflow cage of the pump. the cause of the introducer damage - mechanical was most likely a sheath tip split as observed on the returned introducer. however, the cause of the split is not known due to insufficient clinical details. product damage (guidewire damage - peripheral vessels): the guidewire was not returned. the cause of the product damage (guidewire damage - peripheral vessels) was not determined since product was not returned and insufficient clinical details.
10458 | 1220648-2025-27332 | 21806701 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21806701&pc=OZD | 03/12/2025 | 04/30/2025 |  | 04/09/2025 |  |  | 03/12/2025 | 04/09/2025 |  |  |  |  | 04/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026622619 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 46 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device as elective placement, and during support, the impella catheter became damage from patient movement. the purge pressure was low and purge seen leaking from red impella plug. as the patient required a long runway for heart recovery, the decision made to exchange for right axillary impella 5.5 device implant, which was successfully completed. the patient was reported to be in stable condition following the event. | [5 empty] | The investigation of purge leak ¿ pump has been completed. product was returned and evaluated. the purge leak through the red handle was reproduced with the source of the leak coming from the catheter side. the catheter was opened from the red pump plug to the motor housing to reveal the purge tubing. at the catheter's 20 cm mark, the catheter appeared to be squashed and the coil jacket just inside the catheter was flattened and broken. the purge tubing at this location was damaged. the purge fluid leaking from this location built up inside the catheter and leaked up towards the red pump plug, causing the visible leaking. data log review showed stable purge pressure and flow until 17-mar-2025 when purge pressure low alarms were triggered. the purge pressure at this time initial drops to 350 mmhg for 3 hours and then drops below 200 mmhg for the last hour of the case. the purge flow begins rising to 30 ml/hr at the time of the first pressure drop. motor current was stable. the root cause of the pump purge leak was most likely damaged purge tubing since the returned product showed evidence of damaged catheter and purge tubing that reproduced the leak from the red pump plug and clinical details mention the patient was not cautious with the catheter. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27332 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was returned, and an evaluation is underway. upon completion of the analysis, a supplemental report will be filed.
10459 | 1220648-2025-27336 | 21807100 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21807100&pc=OZD | 03/15/2025 | 04/09/2025 |  | 04/09/2025 |  |  | 03/15/2025 | 04/09/2025 |  |  |  |  | 04/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025568028 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, the patient developed a hematoma at the access site. sandbags were applied to the site and eight units of packed red blood cells were infused to counteract the patient's unstable hemoglobin levels. support continued with the implanted pump. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10460 | 1220648-2025-27333 | 21807696 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21807696&pc=OZD | 03/30/2025 | 05/05/2025 |  | 04/09/2025 |  |  | 03/30/2025 | 04/09/2025 |  |  |  |  | 04/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026613679 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 22 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was reported that a patient on impella cp was escalated to axillary impella 5.5 and extracorporeal membrane oxygenation decannulation. it was noted of point-of-care ultrasound done with concerns for possible left ventricle thrombus. due to pulmonary hemorrhage and bronchial blocker insitu, unable to anticoagulated. sedated and paralyzed for lung compliance and therefore unable to assess neurological status. formal transthoracic echocardiogram this am, shows no thrombus and new left ventricular ejection fraction recovery of 53%. the pump was explanted the following day and the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states), ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The failure mode of ¿thrombus¿ is to be used when a thrombus is noted but cannot be connected to any other adverse events or decrease in pump performance. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: patient's medical history, including any previous thrombotic events or conditions description of the location and characteristics of the thrombus (e.g., size, shape, consistency) relevant laboratory test results, such as coagulation profiles and platelet counts imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states) medications or treatments that the patient was receiving at the time of thrombus formation surgical or procedural details if applicable (e.g, cardiac catheterization) any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events) presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary clinical information was not provided, the root cause of the thrombus was not determined.
10461 | 1220648-2025-27337 | 21807861 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21807861&pc=OZD | 03/10/2025 | 04/22/2025 |  | 04/09/2025 |  |  | 03/10/2025 | 04/09/2025 |  |  |  |  | 04/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026621449 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Hemorrhagic Stroke | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock underwent percutaneous coronary intervention and was initially implanted with an impella cp device for mechanical circulatory support (mcs). the patient was transferred to a receiving facility for escalation in care to an impella 5.5 device for continued mcs. the same day, post implant of the impella 5.5, it was indicated that the patient experienced bleeding, and anticoagulation was withheld (notes reflected "will start systemic anticoagulation tomorrow pending bleeding."). the following day heparin still was withheld. the patient had clear to yellow urine since being placed on the impella 5.5, and remained on three pressors with plans to wean off sedation in hopes to extubate. the patient continued on impella 5.5 mcs. three days later, it was reported that the patient had been going in and out of ventricular tachycardia/torsades de pointes into the following day. the patient was defibrillated. two days later, the patient was placed on venoarterial extracorporeal membrane oxygenation (va ecmo) and temporary pacing wires. the patient received four units of packed red blood cells, three units of fresh frozen plasma, and albumin were noted in the last 24 hours. after approximately three days on va ecmo, the patient was decannulated and was noted to be doing well with a plan to eventually start weaning the impella 5.5 device. the patient continued on impella mcs, and approximately 14 days later, the patient experienced a sudden mental status change. a head computed tomography scan confirmed a "large" cystic lesion with hemorrhage in the right frontal region with large volume intraventricular hemorrhage in lateral ventricles. neurological exam was consistent with stated findings. there were notable electrocardiogram changes the morning of this day, suddenly, widening qrs, and the patient would expire this day. care was withdrawn. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of access site bleeding and stroke/cva has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
10462 | 1220648-2025-27338 | 21808252 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21808252&pc=OZD | 03/15/2025 | 05/15/2025 |  | 04/09/2025 |  |  | 03/15/2025 | 04/09/2025 |  |  |  |  | 04/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026612491 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. while on support, following advancement into the aorta, the impella 5.5 pump began to display strange signals with fluctuations in the placement signal. the pump was subsequently replaced as a result of the noted issue. there was no patient harm. | [5 empty] | The investigation of optical signal issue has been completed. the pump was returned for investigation. data logs were not provided for the investigation. no physical damage was observed on the returned product. the pump passed the laser continuity test, water bucket test, and pressure test for optical system. clinical assessment suggests that the pump was explanted due to an abnormal placement signal waveform. the cause of the optical signal issue was not determined since the returned product did not have any issues, and sufficient clinical information was not provided. d9 date device returned to manufacturer added. | D6a / d6b implant date and explant date unknown. the impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10463 | 1220648-2025-27334 | 21808292 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21808292&pc=OZD | 03/19/2025 | 06/16/2025 |  | 04/09/2025 |  |  | 03/20/2025 | 04/09/2025 |  |  |  |  | 04/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5-JP |  |  | ABIOMED, INC. |  |  | US |  | 2025574192 |  | 1000459 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 60 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that immediately after insertion of the impella 5.5 on a 60 year old male patient, they experienced a pump flow issue. the pump was replaced with a new 5.5. the patient survived the procedure. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The investigation of low pump flow has been completed. the data logs were not returned. product damage (cannula kink) was added since a cannula kink was observed on the returned product. low flow - reported low flow immediately after insertion being unable to run pump above p4 without decreased flow rates. blood volume was given was given and pump was repositioned but it does not resolve the low flow. pump was replaced which resolved the issue. data logs were not returned. impella 5.5 was returned and a cannula kink and catheter kink was observed. the root cause of the low pump flow was most likely a kinked cannula since a cannula kink was found on the returned pump. cannula kink - the root cause of the product damage was a cannula kink.
10464 | 1220648-2025-27339 | 21808660 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21808660&pc=OZD | 03/15/2025 | 10/01/2025 |  | 04/09/2025 |  |  | 03/15/2025 | 04/09/2025 |  |  |  |  | 04/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026613465 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. while on support, the physician was able to feel a deep hematoma on the patient¿s felt flank. the patient's hemoglobin dropped from 12 to 8 overnight and the patient felt pain along the impella access site. two units of packed red blood cells were provided to the patient. the pump was removed. | Us complainant reported the patient was on impella cp support and during support, there were placement signal issues. the position was confirmed to be correct. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smart assist system instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. the pump was not returned for investigation. data logs were not provided for this case run. the cause of the optical issue was not determined since product was not returned and sufficient clinical information was not provided including data log. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical signal issue "2917" as part of a retrospective review of optical signal issues in accordance with fda recommendation.
10465 | 1220648-2025-27340 | 21808701 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21808701&pc=OZD | 03/24/2025 | 04/09/2025 |  | 04/09/2025 |  |  | 03/24/2025 | 04/09/2025 |  |  |  |  | 04/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 0 DA |  | Unknown |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A literature report titled "five cases of preemptive impella insertion before weaning from cardiopulmonary bypass in high-risk patients with cardiogenic shock after open heart surgery" was received by abiomed which stated that "a patient in japan underwent total arch replacement, cabg, and insertion of an impella 5.5 (through the ascending aorta) for distal aortic arch aneurysm and asymptomatic myocardial ischemia. after surgery, the impella flow rate gradually decreased, and re-thoracotomy was performed for cardiac tamponade the day after surgery. however, the hemodynamics remained unstable, and the patient expired on the third day after surgery." there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. a.1, a.4 and a.5 are unknown. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. e.1 name prefix/title, first name and last name are unknown. h.4 device manufacture date is unknown.
10466 | 1220648-2025-27341 | 21808843 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21808843&pc=OZD | 03/13/2025 | 04/11/2025 |  | 04/09/2025 |  |  | 03/14/2025 | 04/09/2025 |  |  |  |  | 04/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1760216 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported that the batteries were deeply discharged and the automated impella controller permanently displayed a battery error. as soon as it was unplugged from the socket, it goes out. they tried overnight charging, however it had no effect. there was no patient involvement. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - battery charging/other issues has been completed. the data analysis displayed a battery failure. the ltpower logs showed that both batteries in the field were around 2.5v per cell by that point. the console was returned and batttest confirmed batteries were still depleted and would not charge. known good batteries were put in, discharged through running a pump on battery power and the console was able to charge those batteries. root cause: the cause of the depleted batteries was most likely due to not routinely charging the batteries while console was left in storage.
10467 | 1220648-2025-27342 | 21808886 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21808886&pc=OZD | 03/27/2025 | 04/09/2025 |  | 04/09/2025 |  |  | 03/29/2025 | 04/09/2025 |  |  |  |  | 04/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025611426 |  | 1000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  | Asian | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported that during insertion of the impella 5.5 the catheter shaft kinked and the insertion was aborted. the patients outcome is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
10468 | 1220648-2025-27343 | 21809361 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21809361&pc=OZD | 03/12/2025 | 04/09/2025 |  | 04/09/2025 |  |  | 03/12/2025 | 04/09/2025 |  |  |  |  | 04/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026624274 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | High Blood Pressure/ Hypertension | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient (unknown race) with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and two days after start of the support, the patient experienced an acute right groin hematoma and bleeding with hypotension. heparin was discontinued, two units of packed red blood cells (prbcs) were given, manual pressure applied and femstop was placed. the patient was hemodynamically stable after administration of the prbcs. a couple of hours later, the device was successfully weaned and explanted (noted not as a result of the complication) with the patient's outcome as survived. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events. ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
10469 | 1220648-2025-27344 | 21809814 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21809814&pc=OZD | 03/13/2025 | 04/09/2025 |  | 04/09/2025 |  |  | 03/13/2025 | 04/09/2025 |  |  |  |  | 04/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025610580 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported the patient was on impella 5.5 support and during support, the patient had a left ventricle thrombus. total parenteral nutrition was administered in the purge and heparin was given systemically. the pump was explanted and upon explant, a thrombectomy was performed with no thrombus retrieved. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
10470 | 1220648-2025-27345 | 21809821 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21809821&pc=OZD | 03/14/2025 | 04/09/2025 |  | 04/09/2025 |  |  | 03/14/2025 | 04/09/2025 |  |  |  |  | 04/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025555526 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced a large hematoma on the left groin that required a fem stop and two units of red blood cell's. the pump was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10471 | 1220648-2025-27346 | 21809839 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21809839&pc=OZD | 03/14/2025 | 04/21/2025 |  | 04/09/2025 |  |  | 03/14/2025 | 04/09/2025 |  |  |  |  | 04/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026632984 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that during an attempted insertion of an impella cp on 76 year old male, the procedure was aborted due to the patient¿s anatomy. a balloon pump was placed instead and the cardiac procedure was completed successfully. the patient survived the procedure. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, impella could not be placed due to vascular anatomy. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27346 as it is unknown if the initial reporter also sent the report to the food and drug administration. h5 was erroneously selected on the initial manufacturer device report number 1220648-2025-27346. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
10472 | 1220648-2025-27347 | 21809841 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21809841&pc=OZD | 03/14/2025 | 04/14/2025 |  | 04/09/2025 |  |  | 03/14/2025 | 04/09/2025 |  |  |  |  | 04/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026613455 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that the 68 year old female patient had limb ischemia. distal flow was assessed due to concerns for small vessel versus vasoconstriction from drips. the patient's left leg was found to have no flow and so it was accessed with micro puncture and a 6 french sheath distally. the right arterial line was exchanged for the sheath and contralateral flow was established with a good result. the patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. b.7 added information that was inadvertently omitted from the initial manufacturer device report number 1220648-2025-27347. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-27347 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.4 device manufacture date was inadvertently omitted from the initial manufacturer device report number 1220648-2025-27347 and was added. h.6 component code 925 was reported incorrectly on the initial manufacturer device report number 1220648-2025-27347. a new code has been added. h.10 investigation results were inadvertenly omitted from the initial manufacturer device report number 1220648-2025-27347 and were added.
10473 | 1220648-2025-27348 | 21809859 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21809859&pc=OZD | 03/14/2025 | 04/25/2025 |  | 04/09/2025 |  |  | 03/14/2025 | 04/09/2025 |  |  |  |  | 04/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | Thrombosis/Thrombus | [11 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | It was reported to abiomed via a literature article title "a case of fatal acute saddle embolism of the terminal aorta after long-term support using impella cp" that a 69-year-old male patient had a cp placed due to subacute myocardial infarction. upon admission, echocardiography revealed ventricular septal rupture (vsr). the patient was promptly supported via veno arterial (va) extracorporeal membrane oxygenation (ecmo) and impella cp before surgical vsr repair. after surgery, it was decided that the patient would be transferred to the intensive care unit under new va-ecmo support. subsequently, removal of the impella cp and reinsertion of the arterial cannulation were performed at the ipsilateral site, and no pulsatile bleeding from the arterial cannulation site was observed. an emergency aortography revealed an imaging defect in the terminal aorta. because of the possibility of acute thrombotic occlusion, a fogarty procedure was performed from the bilateral common femoral arteries (cfa). however, no thrombus was retrieved. a contrast-enhanced computed tomography (ct) examination revealed total occlusion of the bilateral common iliac arteries extending to the abdominal aorta. uncontrolled and rapid progression of acidemia led to sudden cardiac arrest. it is unknown the patient outcome. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ missing information unable to be obtained as follow up could not be performed due to information received via literature article: a1-a5 patient demographic information. d4-catalog number, serial number, udi number, lot number, exp date. e1-initial reporter name and email/phone number. h4-mfr date. | The investigation of thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. citation: kishimoto, s., hiraoka, a., chikazawa, g., & yoshitaka, h. (2025). a case of fatal acute saddle embolism of the terminal aorta after long-term support using impella cp. journal of artificial organs. https://doi.org/10.1007/s10047-025-01499-7 e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27348 as it is unknown if the initial reporter also sent the report to the food and drug administration. h10 citation of literature article was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27348.
10474 | 1220648-2025-27349 | 21809941 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21809941&pc=OZD | 03/14/2025 | 04/16/2025 |  | 04/09/2025 |  |  | 03/14/2025 | 04/09/2025 |  |  |  |  | 04/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025592716 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that upon insertion of impella cp on a 41 year old female, catheter was advanced too far, and the patient went into ventricular tachycardia (vt) but remained consciousness. after a few moments, patient lost consciousness and was defibrillated back to normal sinus rhythm. va ecmo was placed and impella weaned to p2. the impella was repositioned under tee guidance. impella was still pointing toward mitral structure, and was unable to be torque in opposite direction. several days later the patient was escalated to impella 5.5. the patient survived the procedure. | [5 empty] | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27349 as it is unknown if the initial reporter also sent the report to the food and drug administration. h6 health effect - impact code 4647 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27349. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10475 | 1220648-2025-27350 | 21809949 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21809949&pc=OZD | 03/15/2025 | 04/21/2025 |  | 04/09/2025 |  |  | 03/15/2025 | 04/09/2025 |  |  |  |  | 04/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026621002 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that the 83-year-old male patient was on impella cp support and during support, the patient had a hematoma at the access site. the patient was taken to the operating room to have the hematoma removed. the patient was eventually successfully weaned and explanted. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27350 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10476 | 1220648-2025-27351 | 21810156 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21810156&pc=OZD | 03/25/2025 | 06/25/2025 |  | 04/09/2025 |  |  | 03/27/2025 | 04/09/2025 |  |  |  |  | 04/09/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2025604828 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Life Threatening | Other | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported the patient was on impella cp support and during support, the impella flow suddenly dropped to only 0.1 l/min. since vital signs were stable, the impella was removed. when the impella cannula was inspected, a blood clot was found in it. the patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: impella cp with smart assist catheter insertion, section: axillary insertion of the impella cp with smart assist catheter, section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation of low pump flow and thrombus has been completed. low flow occurred during the procedure. pump was removed since the patient hemodynamics was stable. biomaterial found post explant. blood volume not given. no kinks noted. data review: data confirmed the failure mode. logs showed pump started & run without issue. about 23 hours into the case, flow suddenly dropped to 0.5 l/min with flows then continuing to trend downward. there are multiple mc spikes that triggered auto suction response and suction alarms. this event remained to the rest of the case. product was not returned for review. based on clinical description and data review, the root cause of low flow was most likely due to biomaterial since biomaterial was stated to be found post explant and data shows indications of biomaterial. as the necessary information was not provided, the root cause of the thrombus was not determined b2 outcome was erroneously selected on the initial manufacturer device report number 1220648-2025-27351. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-27351.
10477 | 1220648-2025-27290 | 21795528 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21795528&pc=OZD | 03/07/2025 | 05/13/2025 |  | 04/08/2025 |  |  | 03/26/2025 | 04/08/2025 |  |  |  |  | 04/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025575936 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Death | Required Intervention |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was supported by ecpella for mechanical circulatory support. extracorporeal membrane oxygenation (ecmo) was removed and the next day, the patient experienced ventricular fibrillation. ecmo was reinserted, but the patient could not be saved and subsequently the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
10478 | 1220648-2025-27291 | 21795641 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21795641&pc=OZD | 03/10/2025 | 05/01/2025 |  | 04/08/2025 |  |  | 03/10/2025 | 04/08/2025 |  |  |  |  | 04/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026623681 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and experienced access site bleeding and limb ischemia requiring intervention. the patient underwent percutaneous coronary intervention (pci) and was discharged home on lifevest. the patient presented to the emergency department after being shocked twice by the lifevest and was taken to the catheterization lab for a right/left heart catheterization. the left main was being intervene when the patient coded. the decision was made to place an impella cp device. return of spontaneous circulation was scheduled. pci was completed. oozing around the repositioning sheath was noted by the physician and manual pressure was applied. the oozing continued and a hematoma was noticed on the impella site. the patient was sent to the unit on support and later that day, it was noted there was loss of distal flows. the physician decided to place a 14 french peel away sheath and place a femoro-femoral bypass. following, it was noted the hematoma resolved. however, there was minimal oozing and elevated activated clotting time of 237 and activated partial thromboplastin time of 177. the purge fluid was changed to bicarbonate. the impella cp was successfully weaned and explanted after a total of approximately three days on support. the patient was noted to be in stable condition. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding and limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. the root cause of hematoma is determined to be anticoagulation management because the activated clotting time (act) were more than therapeutic range and hematoma resolved by purge fluid changed to bicarbonate.
10479 | 1220648-2025-27292 | 21795719 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21795719&pc=OZD | 03/10/2025 | 04/08/2025 |  | 04/08/2025 |  |  | 03/10/2025 | 04/08/2025 |  |  |  |  | 04/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026618396 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and the following day, a newly formed thrombus was discovered. action noted as continue to collaborate with heart failure for future plan of care. two days later, the patient was transferred to a facility for further evaluation on risk benefit for left ventricular assist device (lvad) placement. due to thrombus, there were no plans to manipulate the impella cp. it was further noted, the patient was not an escalation candidate at said time, and the plan was for expedited transplant lvad work up. the device was successfully weaned and explanted with a survived outcome. the patient was reported to be stable following the event. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
10480 | 1220648-2025-27293 | 21796256 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21796256&pc=OZD | 03/12/2025 | 05/09/2025 |  | 04/08/2025 |  |  | 03/12/2025 | 04/08/2025 |  |  |  |  | 04/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025580304 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during direct insertion of an impella 5.5 pump, the device was advanced too quickly and perforated through the septum. a porcine patch was placed to repair the perforation and the pump was inserted successfully. no additional intervention required and the patient remained stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings and cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿. ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings and cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿. section: warnings and cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿. | The investigation of ventricle perforation has been completed. the pump was returned for investigation. there was no evidence of physical damage on the returned product. no further returned product investigation was required for this case. a review of the pump data log was not necessary for this case run. according to the clinical details, the left ventricle (lv) was perforated while the physician quickly inserted the pump through the aorta to the left ventricle (lv). the lv repair was successful. the cause of the ventricle perforation issue was most likely use related based on given clinical details. d9 date device returned to manufacturer added.
10481 | 1220648-2025-27294 | 21797154 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21797154&pc=OZD | 03/12/2025 | 04/29/2025 |  | 04/08/2025 |  |  | 03/12/2025 | 04/08/2025 |  |  |  |  | 04/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025604010 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 31 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the pump was leaking from the purge sidearm. fluid is clear in color and dripping at a very slow rate. purge flow/pressure was within normal limits and no alarms were present. no visible crack was in the cassette. all connections were checked and tightened. no patient harm was reported. | [5 empty] | The investigation of purge leak ¿ pump has been completed. after 6 days on support, leak form purge sidearm was noted. all connections were checked and tightened. the pump was returned. data log review showed no abnormalities. the returned pump covered in dried dextrose and thus was soaked in warm water and run for 1 hour with leak unable to reproduce. device was unable to get up to nominal purge pressure as the pump was cut off in the field before product was returned for investigation. investigator removed purge sidearm cover and small droplets were noted on the inside cover. purge reservoir was back loaded with syringe and purge tubing was blocked where the pump should have been. despite pressurizing the purge system to levels beyond anticipated use case, the leak was unable to be recreated. the cause of the purge leak - pump was not determined since issue could not be recreated during investigation. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27294 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella 5.5 impella 5.5 with smartassist for use during cardiogenic shock. section: cleaning, ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
10482 | 1220648-2025-27295 | 21797299 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21797299&pc=PYX | 03/11/2025 | 04/08/2025 |  | 04/08/2025 |  |  | 03/11/2025 | 04/08/2025 |  |  | 02/28/2026 |  | 04/08/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025478760 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella rp flex device for mechanical circulatory support and the following day had a dialysis catheter was placed with a sudden change in the placement signal/motor current as well as purge flows/pressure. approximately four days later it was note there was a possible clot entrainment leading to pump stop. information indicates the impella rp flex was explanted with a survived patient outcome. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: rp flex with smart assist system. section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
10483 | 1220648-2025-27296 | 21797544 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21797544&pc=OZD | 03/10/2025 | 04/08/2025 |  | 04/08/2025 |  |  | 03/10/2025 | 04/08/2025 |  |  |  |  | 04/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025599767 |  | 0048-0045 |  | No |  |  | No |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and experienced limb ischemia and would subsequently expire. the patient presented to the emergency department and due to extensive past cardiac history and patient found with new low ejection fraction, the patient was placed on an impella cp device. the impella cp was primed and backloaded over 0.18 wire into the left ventricle and started on auto flowing 3.7 liters per minute once the wire was removed. the position was optimized under fluoroscopy and transesophageal echocardiogram. post implant, the impella cp device was repositioned once the peel away sheath was removed, and the repo sheath was sutured into place. distal flow angiogram was taken with limited flow past the impella sheath, put pulses were dopplerable. feet were blue when the patient arrived in the room. previous below bilateral lower extremities ultrasound was done in the emergency room before going to the catheterization lab. however, results had not return. touhy borst was locked and tightened and angle matched dressing at 90 centimeters. the patient was taken to the unit for management. the following day the patient was escalated to an impella 5.5 and would expire approximately 17 days later. | [5 empty] | Medical safety review of the event noted limb ischemia coming in is unclear. the impella device may have exacerbated the ischemia. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome (demise). the impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10484 | 1220648-2025-27297 | 21797568 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21797568&pc=OZD | 01/19/2025 | 05/12/2025 |  | 04/08/2025 |  |  | 01/19/2025 | 04/08/2025 |  |  |  |  | 04/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025568679 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that they encountered a delivery issue. the patient's femoral arteries were small. the peel away sheath was partially inserted hoping the impella cp would be able to navigate into the aorta. however it would not fit and the procedure was aborted. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high-risk pci. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical detail the root cause of delivery issue is determined to be patient condition because the cp would not fit into the aorta and the insertion was aborted.
10485 | 1220648-2025-27298 | 21797651 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21797651&pc=OZD | 03/12/2025 | 08/14/2025 |  | 04/08/2025 |  |  | 03/12/2025 | 04/08/2025 |  |  |  |  | 04/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1438016 |  | 0042-0010-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Connection Problem | [6 empty] | Cardiac Arrest | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella support and during support, the automated impella controller (aic) was not charging when plugged into power. the aic plug was noted to have a visible fractured wire. aic to aic changeout was preformed without issue. the patient experienced pulseless electrical activity arrest and cardiopulmonary resuscitation was preformed. however, return of spontaneous circulation was unable to be achieved. the patient expired and the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | H.6 code 4647 was added as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-27298.
10486 | 1220648-2025-27299 | 21798538 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21798538&pc=OZD | 03/12/2025 | 05/02/2025 |  | 04/08/2025 |  |  | 03/12/2025 | 04/08/2025 |  |  |  |  | 04/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1232827 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported the staff noted that the left ventricle placement signal was not displayed on the automated impella controller (aic). the staff contacted the field service manager and was given the smart assist metrics code and receiving 3 different metric codes to enable metrics on the console, however an incorrect code entered message would appear. this console metrics were disabled for both cp and 5.5 and despite multiple attempts, remain unable to be displayed. there was no harm to the patient. | [5 empty] | The investigation of automated impella controller (aic) - display issue has been completed. data analysis: the console was ran for 80 minutes and during this time there was consistent optical signals with no abnormalities device analysis: the console was returned for testing and while running a pump, the console had left ventricle (lv) disabled as the default choice. the lv could be enabled in the menu. root cause: the lv placement signal not displaying was most likely a use issue of having the lv disabled in the menu. d9 date device returned to manufacturer added. h5 was erroneously entered on the initial manufacturer device report number 1220648-2025-27299. h8 was erroneously entered on the initial manufacturer device report number 1220648-2025-27299. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10487 | 1220648-2025-27300 | 21798608 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21798608&pc=OZD | 03/12/2025 | 04/08/2025 |  | 04/08/2025 |  |  | 03/12/2025 | 04/08/2025 |  |  |  |  | 04/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025569549 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 24 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that the patient¿s left ventricle appeared globally hypokinetic on echocardiogram. decision was made to place impella cp for unloading. impella placed without complication. during access, it was noted that the patient had extremely narrow femoral artery distal to initial access point. reperfusion sheath placed. in intensive care unit there was concern for the patient having cold leg overnight, but both legs were warm in the morning, but not able to doppler pulses. the patient remained on support for three additional days until successful explanation. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
10488 | 1220648-2025-27301 | 21798617 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21798617&pc=OZD | 03/11/2025 | 04/08/2025 |  | 04/08/2025 |  |  | 03/11/2025 | 04/08/2025 |  |  | 11/30/2025 |  | 04/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024428628 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Sepsis | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced continuous bleeding and sepsis during support. three units of packed red blood cells, as well as fresh frozen plasma and cryo were administered to counteract the bleed. bioglue was placed on the graft as well as surgiflo and a pressure dressing to manage the condition. due to the noted bleed and sepsis, the decision was ultimately made to remove the pump. the site was washed out and the pump was explanted successfully. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock. section: table 3.2 impella catheter components. ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
10489 | 1220648-2025-27302 | 21798672 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21798672&pc=OZD | 03/12/2025 | 04/08/2025 |  | 04/08/2025 |  |  | 03/12/2025 | 04/08/2025 |  |  |  |  | 04/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571473 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Difficult to Remove | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. on day 2 of impella cp support the cp was repositioned under ultrasound guidance but still the impella was not able to go over p5 without suction. it was further reported that during impella removal two perclose were used but there was still bleeding around the 8fr sheath. two times angio-seals were attempted to be deployed, but were unsuccessful. decision was made to hold manual pressure and hemostasis was achieved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: warnings & cautions: warnings ¿avoid manual compression of the inlet and outlet areas of the cannula assembly.¿ section: warnings, cautions, and precautions ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿
10490 | 1220648-2025-27303 | 21800143 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21800143&pc=OZD | 03/11/2025 | 05/08/2025 |  | 04/08/2025 |  |  | 03/12/2025 | 04/08/2025 |  |  |  |  | 04/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025587177 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient developed a mediastinal hematoma during impella cp support. the hematoma was documented to have been caused by extracorporeal membrane oxygenation. as heparin was being used at the time of the noted condition, a relationship between pump support and the condition is unknown. the impella was removed and the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of vascular damage - peripheral vessel has been completed. the root cause of vascular damage peripheral vessel cannot be determined since the product was not returned and insufficient clinical details were provided.
10491 | 1220648-2025-27304 | 21800221 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21800221&pc=OZD | 03/13/2025 | 04/25/2025 |  | 04/08/2025 |  |  | 03/13/2025 | 04/08/2025 |  |  |  |  | 04/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025587444 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella 5.5 pump to support a patient admitted in with acute myocardial infarction/cardiogenic shock. the pump was supporting for two weeks with success when the distal end of sterile sleeve ripped. it was cleaned with chlorhexidine gluconate and covered with sterile tegaderm. no excessive force was applied and it is unknown how the plastic separated. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of product damage (sleeve damage) has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
10492 | 1220648-2025-27305 | 21800649 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21800649&pc=OZD | 03/12/2025 | 04/08/2025 |  | 04/08/2025 |  |  | 03/12/2025 | 04/08/2025 |  |  |  |  | 04/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025573285 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was attempted to be implanted with an impella device for mechanical circulatory support. it was noted by the staff that the pump was unable to cross into the aorta due to right subclavian artery stenosis noted via angiogram. despite multiple attempts, the implant was not successful and the pump was removed. an impella cp pump was subsequently implanted for planned support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings: ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when. manipulating the heart during cardiac surgery, evaluate the position of the pump using. imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
10493 | 1220648-2025-27306 | 21801100 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21801100&pc=OZD | 03/13/2025 | 04/30/2025 |  | 04/08/2025 |  |  | 03/13/2025 | 04/08/2025 |  |  |  |  | 04/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025565412 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a patient on impella cp support. it was reported that the patient had dark urine output but improved with device repositioning. the patient remained stable and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of hemolysis has been completed. clinical states that the cp pump was implanted and soon dark red uo observed after a few hours. the pump was repositioned after 4 hours and the color stated improving. pump was returned for investigation. pump was clear of biomaterial and any physical damage of impellers. cannula kink was observed in the middle area of the cannula. pump ran as specification for 30 minutes in a bucket of water. data logs around the time of hemolysis are not available. therefore, as per the clinical information and product investigation, the root cause of the hemolysis issue was most likely improper positioning of the device due to improper technique which was resolved by repositioning the device. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿.
10494 | 1220648-2025-27307 | 21801704 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21801704&pc=OZD | 03/12/2025 | 05/21/2025 |  | 04/08/2025 |  |  | 03/12/2025 | 04/08/2025 |  |  |  |  | 04/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1319817 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Defective Alarm | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the audible alarm of an automated impella controller (aic) sounded like a rotary phone. the aic was swapped with no harm to the patient. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - buzzer/ alarm inaudible has been completed. there is no evidence available in the logs to confirm the reported failure mode, which was related to the alarm's audio issue. this console was tested and abnormal audio was intermittently reproduced alongside the alarm audio. no loose components were found in the speaker assembly. upon testing the known good speaker, the reported failure mode could not be reproduced. the root cause of the abnormal audio alarm was most likely due to a speaker malfunction. d2 common device name revised on manufacturer device report 1220648-2025-27307 in accordance with updated procedures. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27307 as it is unknown if the initial reporter also sent the report to the food and drug administration.
10495 | 1220648-2025-27308 | 21801786 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21801786&pc=OZD | 03/12/2025 | 04/14/2025 |  | 04/08/2025 |  |  | 03/12/2025 | 04/08/2025 |  |  |  |  | 04/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025573246 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | Device Sensing Problem | Positioning Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. an "impella in aorta" alarm appeared despite the device having proper positioning in the left ventricle. the physician made the decision to pull the pump and rewire the left ventricle, but the staff was unable to turn the impella off. despite turning support down to p0, the motor current would not completely shut off. the pump was replaced in response to the noted issue. there was no patient harm. | [5 empty] | The investigation was completed. the returned product was inspected and there were no visual abnormalities. the 5s parameters on the returned product were within specification. the logs showed the impella position in aorta alarm with stable 5s parameters. clinical details mentioned that the pump position was in the left ventricle. there were no motor current spikes in the logs prior to alarm. the root cause of the optical signal issue was most likely patient condition related as evidenced by false position in aorta alarm despite good positioning and no motor current spikes or ingestions. the data logs and logs for the case were analyzed. the logs revealed that the user did not turn the pump down to p0. the pump was run in the lab and was able to turned down to p0. the root cause of the positioning issue was most likely use-related as there was no evidence in evidenced by the failure to complete the procedure of turning pump to p0 and as evidenced by the issue not reproducing on the returned product. the failure mode pump not detected was changed to positioning issues due investigation findings. b.7 information was added that was inadvertently omitted from the initial manufacturer device report number 1220648-2025-27308. e.1 us phone number was added as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-27308. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-27308 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 investigation conclusions code 4315 and 3221 was erroneously entered on the initial manufacturer device report number 1220648-2025-27308. codes 3233 and 11 should of been reported on the initial manufacturer device report number 1220648-2025-27308 as the device was returned for evaluation. code 1581 was changed to 3009 due to investigation results. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
10496 | 1220648-2025-27309 | 21801794 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21801794&pc=OZD | 03/12/2025 | 04/08/2025 |  | 04/08/2025 |  |  | 03/12/2025 | 04/08/2025 |  |  |  |  | 04/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026613720 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the there were placement signal/motor current issues with the pump. the pump position was assessed as correct, but the pump was replaced. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10497 | 1220648-2025-27310 | 21801922 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21801922&pc=OZD | 03/12/2025 | 11/04/2025 |  | 04/08/2025 |  |  | 03/12/2025 | 04/08/2025 |  |  |  |  | 04/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 EU |  |  | ABIOMED, INC. |  |  | US |  | 2025592663 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Detachment of Device or Device Component | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that the patient was on impella 5.5 support and during support, there were placement signal issues. the pump functioned as normal. it is suspected that cable damage is the cause (there was a small irregularity felt at the point where it emerged from the bed, at the level of the guardrails; likely, the cable was pinched between them). the pump was eventually explanted. no updates were provided on the patient outcomes. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. the following sections had missing information: a1-a5 patient identifier, age, weight, sex, and ethnicity/race. e1-initial reporter name. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation was completed for the returned impella 5.5 pump. the unit was received with insufficient material for full evaluation, and data logs were not returned. according to the clinical team, loss of placement signal occurred after three days of support, possibly due to cable damage from being pinched between the patient¿s bed and guardrail. inspection of the returned pump confirmed a kink in the patient cable near the red pump plug side. laser continuity testing passed up to the cut point in the catheter. the observed kink supports the physician¿s suspicion that the cable was damaged by being pinched, leading to loss of signal. although the exact root cause of the optical signal issue could not be definitively determined due to limited returned product and missing data logs, the most probable cause is cable damage resulting from the patient cable being pinched between the bed and guardrail. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated. d10 concomitant medical products and therapy dates updated.
10498 | 1220648-2025-27311 | 21802461 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21802461&pc=OZD | 03/13/2025 | 05/08/2025 |  | 04/08/2025 |  |  | 03/13/2025 | 04/08/2025 |  |  |  |  | 04/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1456617 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The automated impella controller (aic) had a controller error. the issue was discovered during use. no further details were provided. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. console logs are consistent with complaint and show that the pump was transferred from the 1st console (b)(6) 2025. there were no issues with placement signal until (b)(6) 2025 when console presented with controller error alarms. after that, pump was transferred to backup console where alarms did not repeat. real time (rt) logs confirmed that prior to and during alarms, all signals received from optical bench (placement, snr, contrast, lamp, gain) were represented as -16384 values due to the crc error. the controller error and loss of placement signal was due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27311 as it is unknown if the initial reporter also sent the report to the food and drug administration.
10499 | 1220648-2025-27312 | 21802470 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21802470&pc=OZD | 03/14/2025 | 05/06/2025 |  | 04/08/2025 |  |  | 03/14/2025 | 04/08/2025 |  |  |  |  | 04/08/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025570542 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp insertion was aborted due to the impella not able to traverse the proximal iliac artery. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since impella was unable to traverse proximal iliac artery. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27312 as it is unknown if the initial reporter also sent the report to the food and drug administration.
10500 | 1220648-2025-27255 | 21784433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21784433&pc=OZD | 03/10/2025 | 04/24/2025 |  | 04/07/2025 |  |  | 03/10/2025 | 04/07/2025 |  |  |  |  | 04/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026622077 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 39 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | Failure to Read Input Signal | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support (mcs), and an "impella defective" alarmed triggered prompting explant of the device. the patient was admitted with shortness of breath and intermittent runs of ventricular tachycardia on the automated implantable cardioverter defibrillator. the impella cp was implanted without complications. the patient was subsequently placed on extracorporeal membrane oxygenation (ecmo). approximately two days after the start of the impella mcs, the patient was transferred to a tertiary center and during the automated impella controller (aic) - aic to aic transfer, the nurse plugged in the drive line to the receiving facility's aic and an ¿impella defective¿ error triggered. trouble shooting occurred thereafter which involved plugging the drive line back into original aic, turning the controller on and off still with ¿impella defective¿ error indicating there is a problem with the impella catheter electronics. the team was notified that the impella cp pump needed to be pulled across the valve and into the aorta. catheterization lab was consulted for replacement pump insertion. the catheterization lab team arrived and pulled the impella catheter into the aorta. it was noted that antegrade 4f sheath of the right femoral artery was connected (male to male) to the wire access port. the interventional cardiologist pulled a 7-inch string of clot from the access port and opted not to access the port due to clot concern. the impella cp pump was removed, and the right femoral artery was closed with an intact perclose. the patient was taken to the unit and indicated to be in stable condition. | [5 empty] | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The failure mode 'pump stop' was changed to 'pump not detected' as investigation revealed that 'impella defective' alarm was observed when console swap was attempted. no logs were provided from attempted console swap. the data logs from the first console before attempted swap do not show any impella defective alarms the returned product was inspected. a visual abnormality on one side of the patient cable housing was observed where there was an unusual gap. the impella defective alarm reproduced upon connection to a lab console. the red pump plug was opened to test electrical communication lines between the patient cable plug pins and the pcb/eeprom. an open line was found between pin 5 and the corresponding yellow cable connected to the pcb, responsible for communication between the console and pump. further testing revealed that heating up the patient cable plug with a heat gun reestablished the electrical connection when also applying downward force on the affected pin. the root cause of the pump not detected issue was determined to be due to an open line at the joint between the patient cable pin 5 and the yellow line in the patient cable as was found with testing of returned product.
10501 | 1220648-2025-27256 | 21784474 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21784474&pc=OZD | 03/10/2025 | 04/18/2025 |  | 04/07/2025 |  |  | 03/10/2025 | 04/07/2025 |  |  |  |  | 04/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025586067 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient, non-surgical candidate, undergoing a high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. however, after initiation of support, flows were unable to go above 1.8 liters per minute. the motor current noted to have a spike and then was noted to be flat with the pump appearing to be in proper position. the decision was made to explant and replace with a new impella cp device. the patient remained hemodynamically stable with no harm as a result of the event. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of low pump flow has been completed. low flow occurred when pump started with motor current (mc) spike and flat. pump was removed after 6 minutes of support. data review: data logs confirmed the failure mode and clinical narrative. logs showed after pump started, mc spiked followed low flow that triggered auto suction response & suction alarms. this event remained to the rest of the case. product was not returned for review. based on clinical description and data review, the root cause of low flow was most likely biomaterial ingestion.
10502 | 1220648-2025-27257 | 21784497 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21784497&pc=OZD | 03/10/2025 | 04/28/2025 |  | 04/07/2025 |  |  | 03/10/2025 | 04/07/2025 |  |  |  |  | 04/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026621989 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, approximately two days after start of the support, suction alarm with spike in motor current occurred and there was a drop in flows to 0.8 liters per minute. the performance level was decreased to p-2. however, suction alarms remained persistent. stat echocardiogram was completed and confirmed an apical thrombus and visible clot on the device. the decision was made to explant the device and support the patient on pressors. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation of low pump flow and thrombus has been completed. returned complaint device visually inspected. imeplla washed with water to remove the dried dextrose. biomaterial observed around impeller and purge gap. no cannula kink or broken blade seen. data log reviewed: mc spike and suction alarm observed around same time issue reported followed by impella position in aorta alarm cause drop in flow ~0.8lpm at p-2. these alarms occurred more often to end of the case and impella explanted after around 40 mins. suction alarms and mc spike occurred before the aorta position alarm, clinical confirmed apical thrombus and visible clot on device and device evaluation confirmed biomaterial ingestion. the cause of the low pump flow issue was biomaterial ingestion. as the necessary information was not provided, the root cause of the thrombus was not determined.
10503 | 1220648-2025-27258 | 21784562 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21784562&pc=OZD | 03/10/2025 | 05/15/2025 |  | 04/07/2025 |  |  | 03/10/2025 | 04/07/2025 |  |  |  |  | 04/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025570772 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemolysis | Hematuria | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient (unknown race and ethnicity) noted with cardiomyopathy was implanted with an impella cp device for mechanical circulatory support (mc). during support, approximately seven days after start of impella mcs, suction alarms were noted indicating low pump flow. the patient was noted to have hematuria and hemolysis. a thrombus was also noted while performing transesophageal echocardiogram procedure. the pump was subsequently explanted and replaced. the patient continued on impella mcs with subsequent plan to place a left ventricular assist device. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the analysis, a supplemental report will be filed. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | Additional information provided in h3, h6 and h10. the investigation into the reported issues has been completed. hemolysis: after seven days on support, positive macro hematuria and hemolytic blood was noted. the pump was returned and inspected. after soaking in warm water and tergazyme, biomaterial was found wrapped around the impellor blade and attached to the inner cannula wall. the cause of the hemolysis was most likely biomaterial due to the biomaterial found wrapped around the impellor blades and attached to the inner wall of the cannula. low pump flow: after 7 days on support, it was noted that pump flows at p-5 were at 2.2 l/min, lower than the expected range of 2.3 - 2.7 l/min for cp. data logs were not recovered. the pump was returned and inspected. after soaking in warm water and tergazyme, biomaterial was found wrapped around the impellor blade and attached to the inner cannula wall. the cause of the low pump flow was most likely biomaterial due to the biomaterial found wrapped around the impellor blades and attached to the inner wall of the cannula. thrombus: after 7 days on support, the pump was explanted due to low pump flows and hemolytic blood. after pump was removed, md stated they saw thrombus on explanted impella. the pump was returned and inspected. after soaking in warm water and tergazyme, biomaterial was found wrapped around the impellor blade and attached to the inner cannula wall. thrombus can be observed in the body or on the pump upon explant. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, adequate clinical information mentioned would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined.
10504 | 1220648-2025-27259 | 21784599 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21784599&pc=OZD | 03/12/2025 | 04/22/2025 |  | 04/07/2025 |  |  | 03/12/2025 | 04/07/2025 |  |  |  |  | 04/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025610386 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [10 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was noted that the patient experienced hemolysis. a drop in blood pressure was observed coinciding with the hemolysis and a gastrointestinal bleed prompting administration of levophed. heparin was discontinued and the pump was repositioned. despite the troubleshooting, plans were made to explant the pump later that morning. the patients condition was noted to be stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. the root cause of hemolysis was most likely due to pump was not in the proper position since clinical team confirmed pump was in improper position and hemolysis was resolved post repositioning of the pump.
10505 | 1220648-2025-27260 | 21784692 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21784692&pc=OZD | 03/13/2025 | 05/21/2025 |  | 04/07/2025 |  |  | 03/13/2025 | 04/07/2025 |  |  |  |  | 04/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025514868 |  | 1000482 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 54 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | Positioning Problem | [5 empty] | Perforation of Vessels | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient experienced a left ventricle perforation during impella 5.5 implant. the perforation was repaired and the pump was placed for patient support. ten days into support, the patient passed away. there is not enough evidence to disassociate the patient's passing with the noted complication. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | The investigation of ventricle perforation has been completed. the impella device was not returned for evaluation. added failure mode (fm) positioning issues since the pump was reported to be placed too deep. no product or logs were returned for evaluation. clinical details noted that the physician did not use the insertion template prior to attaching the graft direct to the aorta and impella caused a ventricle perforation. the root cause of the ventricle perforation was most likely a use related issue as the inserting physician did not measure and place the graft in the proper position, resulting in the pump being placed too deep in the ventricle upon insertion.
10506 | 1220648-2025-27261 | 21784858 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21784858&pc=OZD | 03/10/2025 | 04/07/2025 |  | 04/07/2025 |  |  | 03/10/2025 | 04/07/2025 |  |  |  |  | 04/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026623929 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | High Blood Pressure/ Hypertension | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with undergoing high-risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support and experienced access site bleeding requiring intervention and removal of the device. the left femoral arterial access was obtained in routine fashion and upsizing to impella14fr x 13cm sheath was performed. impella was inserted, and flows were initiated. pci of left anterior descending artery (lad) was attempted but due to lesion at ostium of lad attempts to cross were unsuccessful. the decision was then made to halt pci and send patient for surgical evaluation. decision was made to leave impella cp device in and transport patient back to unit for care while awaiting surgical decision. the sheath exchange was performed in routine fashion. once repositioning sheath was introduced patient began to bleed continuously through access site with no resolve. 10mg of labetalol was given to treat hypertension and manual pressure was held. despite pressure lowering, manual pressure, and forward tension on sheath patient still exhibited profuse oozing from insertion site. at this point, the physician made the decision to explant impella cp device. dry closure was performed using a 10mm balloon at the iliac, and post closure was performed with perclose x2. post closure angiogram shown patent artery with no bleeding. the patient was noted to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
10507 | 1220648-2025-27262 | 21784884 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21784884&pc=OZD | 12/28/2023 | 04/07/2025 |  | 04/07/2025 |  |  | 03/14/2025 | 04/07/2025 |  |  | 10/31/2025 |  | 04/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024416018 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Cardiogenic Shock | Heart Failure/Congestive Heart Failure | [8 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was documented via impact study that a patient experienced multiple adverse events during impella 5.5 support. it was initially documented that the patient experienced cardiogenic/vasogenic shock with a possible relationship to the impella device/procedure. the condition was further believed to have a definite relationship to the associated surgical procedure. on three separate occasions, the patient also experienced bleeding that was believed to have possible to probable relationships to the impella device, impella procedure, or surgical procedure. seventeen days into support, the patient was diagnosed with right heart failure. the specific therapy for the right ventricle was added at the time of the event, but the condition was believed to have a possible relationship to the impella device/procedure and a probable relationship to the surgical procedure. roughly twenty nine days into support, the patient experienced hemolysis. while not believed to be related to the impella procedure, the hemolysis was documented to have a probable relationship to the impella device and a possible relationship to the surgical procedure. finally, the patient was diagnosed with acute kidney failure. the condition was believed to have a possible relationship to the impella device/procedure and surgical procedure. the patient was started on continuous renal replacement therapy after the condition worsened following hemolysis. the patient survived the procedure. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
10508 | 1220648-2025-27263 | 21785010 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21785010&pc=OZD | 10/03/2024 | 05/01/2025 |  | 04/07/2025 |  |  | 03/14/2025 | 04/07/2025 |  |  |  |  | 04/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025548426 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | Followup | [6 empty] | The impact study reported that an impella 5.5 was inserted, and the patient developed bleeding requiring reoperation. the relationship to the impella and to the impella procedure had a relationship of "probable". the patient survived the procedure. | The impact study reported that the patient on impella 5.5 support experienced ventricular arrhythmia with a ¿possible¿ relationship to the impella device. amiodarone was administered. the patient recovered and arrhythmia was resolved. |  |  |  |  | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. | The investigation of access site bleeding has been completed. the root cause of the access site bleeding issue was not determined since product was not returned and insufficient clinical details provided. b5 additional information received from the clinical site. h6 revised to reflect additional information received from the clinical site. | The investigation of vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. h10 investigation results for vt/vf were inadvertently omitted from manufacturer device report 1220648-2025-27263-1. | G.2 report source study was added as it was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27263.
10509 | 1220648-2025-27264 | 21785208 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21785208&pc=OZD | 03/14/2025 | 04/07/2025 |  | 04/07/2025 |  |  | 03/14/2025 | 04/07/2025 |  |  |  |  | 04/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026612902 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that complications were encountered during attempted delivery of an impella cp pump for a patient on support for a high-risk cardiac procedure. after successfully placing a 14 french x 13 centimeter introducer, the pump was inserted but would not advance through the patient's heavily calcified and tortuous vasculature. the short sheath was swapped for a long sheath, but the sheath could not be fully inserted due to the patient's anatomy. a second attempt with the short sheath was performed, but once again, the pump was unable to advance through the patient's vasculature. the implant was subsequently aborted. the patient survived the procedure. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.¿
10510 | 1220648-2025-27265 | 21785517 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21785517&pc=OZD | 03/14/2025 | 04/21/2025 |  | 04/07/2025 |  |  | 03/14/2025 | 04/07/2025 |  |  |  |  | 04/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025545650 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that the groin site experienced some oozing following impella cp implant of a patient. after two days of oozing, the staff adjusted the angle of entry and a mattress suture was placed to resolve the condition. support was continued uninterrupted. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of access site bleeding is determined to be use issue because the bleeding was controlled by angle match and placing mattress sutures at the bleeding site.
10511 | 1220648-2025-27266 | 21785636 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21785636&pc=PYX | 03/14/2025 | 05/01/2025 |  | 04/07/2025 |  |  | 03/14/2025 | 04/07/2025 |  |  |  |  | 04/07/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025574576 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with an impella rp flex due to acute right heart failure. the impella rp flex was placed successfully. during support, it was noted that the sterile sleeve was torn. additionally, the patient was reported being heparin-induced thrombocytopenia (hit) positive. they started bicarbonate in the purge and no systemic anticoagulants. the patient as eventually weaned and explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of product damage (sleeve damage)and thrombocytopenia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
10512 | 1220648-2025-27267 | 21785790 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21785790&pc=OZD | 03/14/2025 | 04/07/2025 |  | 04/07/2025 |  |  | 03/14/2025 | 04/07/2025 |  |  |  |  | 04/07/2025 |  | EMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026622644 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that during an attempted delivery of a impella 5.5 pump on a patient, the pump met resistance in the axillary artery. the implant was aborted with no reported harm to the patient. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings, ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions, ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist: section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings, ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
10513 | 1220648-2025-27268 | 21786932 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21786932&pc=PYX | 03/08/2025 | 10/30/2025 |  | 04/07/2025 |  |  | 03/08/2025 | 04/07/2025 |  |  |  |  | 04/07/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025578270 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | Device Sensing Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was placed on bipella support for mechanical circulatory support. during support, the patient developed oozing at the rp insertion site. pressure was held and bleeding was resolved. additionally, the white connector cable had several twists with a significant kink but was correctly attached to the automated impella controller, but placement signal not reliable alarm occurred. flow calculation was disabled and there was rotation of the impellar icon. in addition, the motor current had not changed. echo performed later in the evening showed that the rp was working well. the patient continued on support. additional information noted care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella rp flex with smartassist system with automated impella controller: section: warnings, cautions, and precautions. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. no product or data logs were returned for investigation. based on the clinical details, it was seen that the patient was on cp, rp, and crrt support. patient had seizures during trial weans of rp, bleeding was observed from the chest tubes, and blood transfusions were made throughout the case. there were psnr alarms during the case, and during a call with the nurse on (b)(6) 2025 it was seen that the white connector cable had several twists with a significant kink. echo and x-ray performed after the call on the same day verified that the rp was functioning well as there were no placement changes. patient was diuresed throughout the same day and their cvp had dropped to 6-7mmhg with the hemo calculations suggesting euvolemia due to significantly decompressed rv and easily compressible ivc. according to the patient updates on (b)(6) 2025, both right and left impella devices were in great position and decompressing, with the patient's x-ray results getting worse. there was no pericardial effusion, and more blood products were given for their drop in hgb. on the day of rp explant, the patient was still on crrt support, with rp placement signals not functioning. placement signal issue: the root cause of the placement signal issue was not determined due to inconclusive evidence from the clinical details, unreturned product, and lack of data logs for analysis. product damage (cannula kink): the root cause of the product damage (cannula kink) was most likely patient cable kink based on the clinical details provided. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6 updated.
10514 | 1220648-2025-27269 | 21786939 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21786939&pc=OZD | 03/10/2025 | 04/07/2025 |  | 04/07/2025 |  |  | 03/10/2025 | 04/07/2025 |  |  | 06/30/2025 |  | 04/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024430738 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | Medtronic received information post implant of this 33mm mitral bioprosthetic valve, intermittent atrial fibrillation and accelerated junctional rhythm (ajr) were reported. it was reported that the patient was treated with antiarrhythmic medication. no additional adverse patient effects were reported. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
10515 | 1220648-2025-27270 | 21786972 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21786972&pc=OZD | 03/09/2025 | 04/29/2025 |  | 04/07/2025 |  |  | 03/09/2025 | 04/07/2025 |  |  |  |  | 04/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2026623476 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Positioning Problem | [5 empty] | Bradycardia | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [10 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient experienced ventricular tachycardia (vt). the patient was converted without defibrillation and an episode of bradycardia warranting administration of atropine. support was continued. additionally, the patient had access site bleeding. pressure was applied and heparin was stopped temporarily until in therapeutic range. angle-match dressing was reinforced. support was continued. furthermore, the impella was slightly repositioned under echocardiogram helping with suction alarms, but the patient continuously required more support. the patient expired on support. it is unknown if the use of the impella was related to the patients death. thus, there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding, low pump flow positioning issues, and vf/vt/af/at has been completed. the impella device was not returned for evaluation. data log reviewed: no motor current (mc) spike outside of p-level changes observed. many suction alarms occurred on the day low flow issue reported. impella position in aorta alarms occurred around end of the impella support. clinical stated the impella was slightly repositioned under echocardiogram on (b)(6) helping with suction alarms but patient continuously required more support. unclear evidence obtained from the account data why the position issue happened and what other patient support performed to attempt improve low flow issue. no indications from data logs obtained. the cause of the access site bleeding most likely was anticoagulation management since the heparin stopped temporarily until in therapeutic range to solve the issue. the cause of the low pump flow issue cannot be determined since complaint device not returned for analysis and insufficient clinical data provided. the cause of the positioning issues cannot be determined since complaint device not returned for analysis and insufficient clinical data provided. as the necessary information was not provided, the root cause of the vt/vf was not determined.
10516 | 1220648-2025-27271 | 21786986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21786986&pc=OZD | 03/10/2025 | 04/28/2025 |  | 04/07/2025 |  |  | 03/10/2025 | 04/07/2025 |  |  |  |  | 04/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026613611 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the insertion of impella 5.5 device to a patient in acute myocardial infarction/cardiogenic shock was unsuccessful as the device was unable to be advanced across the axillary graft anastomosis. a buddy wire technique was then attempted and unsuccessful as well. the decision was then made to place an impella cp device via the axillary insertion site, which was successfully completed. there was no report of adverse effect to the patient as a result of the event. | [5 empty] | The investigation of delivery issues has been completed. clinical reported having delivery issues advancing the 5.5 across the axillary graft anastomosis. delivery was aborted and downsized to a cp which was implanted successfully via the axillary insertion site. returned usb was empty. pump was returned and the cannula was deformed near the inflow cage. the root cause of the delivery issue was most likely patient condition related since it was reported the 5.5 was unable to be advanced across the axillary graft anastomosis and the cp was placed successfully afterwards. instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella 5.5 with smartassist for use during with automated impella controller. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ d6a and d6b was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27271. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27271 as it is unknown if the initial reporter also sent the report to the food and drug administration. h10 instructions for use was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27271. | The impella device was returned, and an evaluation is underway. upon completion of the analysis, a supplemental report will be filed.
10517 | 1220648-2025-27272 | 21787034 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21787034&pc=OZD | 03/10/2025 | 05/14/2025 |  | 04/07/2025 |  |  | 03/10/2025 | 04/07/2025 |  |  |  |  | 04/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025583685 |  | 1000459 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was initially implanted with an impella cp device for approximately eight days for mechanical circulatory support until escalation to an impella 5.5 with extracorporeal membrane oxygenation (ecmo) placement for continued therapy. during the impella 5.5 device support, approximately 41 days after starting, high purge pressure was encountered. the pump flow rate was low compared to the performance level, and at p-4 the flow was around 0.6-0.9l. it was noted that sometimes it goes up to 2l, but usually did not exceed 1l. further noted was because the device had been used for longer than the specified number of days, hospital staff monitored the motor's current consumption over time, but there was no increase in the amount of current. there was a 30-second blackout with the automated impella controller, but the patient, flow rate was ecmo dependent, and the impella was for venting purposes with no effect on the patient's vital signs. the patient continued on impella and ecmo therapies for an additional 10 days before expiring noted to be due to multi-organ failure. | [5 empty] | The investigation into the reported issue has been completed. low pump flow: data logs were not recovered. the pump was returned and inspected. investigation reveled that biomaterial had been wrapped around the impellor blade and accumulated on the pump motor housing. the cause of the low pump flow was biomaterial due to biomaterial found on the impeller blade and around the motor housing. after 52 days on support, the patients condition worsened which led to multiple organ failure and death. md noted impella was not involved in the death and the pump continued to function without any problems until the end. unrelated death fm was removed under investigation since the adverse event had no relation with impella per voice of customer. | Medical safety review of the event for the demise outcome noted the impella 5.5 device, there was no change in the therapy. the flow rate was dependent on ecmo. this is one of two devices related to the patient's implant experience. refer to medical device report for the automated impella controller serial number (b)(6) with date of event 10-mar-2025 and with same patient identifier as this report. the impella device was returned, and an evaluation is underway. upon completion of the analysis, a supplemental report will be filed.
10518 | 1220648-2025-27273 | 21787044 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21787044&pc=OZD | 03/12/2025 | 04/30/2025 |  | 04/07/2025 |  |  | 03/12/2025 | 04/07/2025 |  |  |  |  | 04/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025589260 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  |  | White | Injury | Yes | Yes | Increase in Pressure | Device Sensing Problem | Infusion or Flow Problem |  |  |  |  | Thrombosis/Thrombus | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. immediately following implant, a pump in aorta alarm appeared and pump saturation was noted. positioning was verified to be correct, but the displayed flow rate was not appropriate for the set p-level. it was additionally documented that the purge flow was low (2ml/hr). the pump was removed and clotting was seen in the inlet and graft. a new impella 5.5 pump was placed for ongoing support. | [5 empty] | The investigation of optical signal issue, saturated flow, thrombus, and high purge pressure has been completed. the root cause of the optical signal issue, saturated flows, and high purge pressure was most likely due to biomaterial ingestion based on returned data log analysis. as the necessary information was not provided, the root cause of the thrombus was not determined e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27273 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
10519 | 1220648-2025-27274 | 21787056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21787056&pc=OZD | 03/10/2025 | 04/17/2025 |  | 04/07/2025 |  |  | 03/10/2025 | 04/07/2025 |  |  |  |  | 04/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025592893 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  |  | Native Hawaiian Or Other Pacific Islander | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | A |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that during support, the impella cp dropped from p-8 to p-6 in the intensive care unit due to brief suction and robust hemodynamics with preserved ejection fraction and administered intravenous bolus. the impella cp was successfully weaned and explanted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of low pump flow has been completed. the pump was not returned for investigation. the data log shows several suction alarms and low pump flow events during the 8 hours post-implant. there were also some placement signal trends observed during the suction and low pump flow events. flow was resolved later during the case run. the cause of the low pump flow issue was most likely patient condition related, based on data log review.
10520 | 1220648-2025-27276 | 21787092 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21787092&pc=OZD | 03/12/2025 | 04/07/2025 |  | 04/07/2025 |  |  | 03/12/2025 | 04/07/2025 |  |  |  |  | 04/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026625061 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  |  | White | Death | Yes | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | Cardiac Arrest | [12 empty] | Death | Required Intervention |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was attempted to be implanted with an impella device for mechanical circulatory support. it was noted that it was not possible to cross the aortic valve due to patient's severe stenosis and presence of an aortic valve replacement scaffolding. the patient went into ventricular tachycardia requiring defibrillation and then subsequent rounds of cpr and code drugs. the patient remained in pulseless electrical activity with no pulse despite 15 minutes of cpr. the physician was successful at getting guidewire and catheter into left ventricle, but the impella catheter would never cross the severely stenotic aortic valve and the scaffolding of the avr he had previously. pt was pronounced expired on the cath lab table. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves, ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
10521 | 1220648-2025-27275 | 21787094 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21787094&pc=OZD | 03/12/2025 | 04/30/2025 |  | 04/07/2025 |  |  | 03/12/2025 | 04/07/2025 |  |  | 08/31/2025 |  | 04/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024388672 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was attempted to be implanted with an impella 5.5 device for mechanical circulatory support. the pump was unable to advance beyond the innominate artery due to stenosis at the ostial takeoff. after two failed attempts, the implant was aborted and an impella cp was placed for planned support. there was no patient harm. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, the impella was unable to pass through due to stenosis of the vessels. d6b was erroneously entered on the initial manufacturer device report number 1220648-2025-27275. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27275 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controllerrp flex with smartassist system with automated impella controller section: warnings & cautions: warnings. ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings. section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
10522 | 1220648-2025-27277 | 21787095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21787095&pc=OZD | 03/13/2025 | 04/08/2025 |  | 04/07/2025 |  |  | 03/13/2025 | 04/07/2025 |  |  |  |  | 04/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026615200 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Pump | Failure to Read Input Signal | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 71-year-old male patient who had transferred on the 5.5 support from the community to the (b)(6) center. upon transfer to the 2nd console (aic), the pump failed and could not be restarted. the team took the patient back to the or for a new pump in the same conduit. the patient survived the procedure. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock states the following: section: direct aortic insertion. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation of pump stop has been completed. the failure mode 'pump stop' was changed to 'pump not detected' as that was the complaint from the customer. the returned product was inspected and dried blood was observed in the red pump plug. a hole in the catheter was observed at the 35 centimeter mark which led to the blood in the red pump plug. the pump was plugged into the console and the impella defective alarm did not reproduce. the pump passed all electrical testing. the data logs for the first console did not show any impella defective alarms. the root cause of the pump not detected issue was blood ingress due to a hole in the catheter resulting from improper use as evidenced by returned product analysis. b.7 added information that was inadvertently omitted from the initial manufacturer device report number 1220648-2025-27277. e.1 added us phone number as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-27277. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-27277 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 code 3221 and 4315 were reported incorrectly on the initial manufacturer device report number 1220648-2025-27277 as the investigation resulted were available at that time. medical device problem code was revised due to investigation results. h.10 additional manufacturer narrative was incorrect on the initial manufacturer device report number 1220648-2025-27277 as the investigation has been completed at that time.
10523 | 1220648-2025-27278 | 21787108 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21787108&pc=OZD | 03/13/2025 | 05/15/2025 |  | 04/07/2025 |  |  | 03/13/2025 | 04/07/2025 |  |  |  |  | 04/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026614732 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 66 year old male patient that had a 5.5 pump being placed to support a surgical cardiac procedure. during delivery, there were guidewire issues that prompted the team to replace the .018 wire that had kinked at the delivery. the pump was successfully placed and eventually weaned. the patient's outcome from the procedure is unknown. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter." ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. | Reported having issues inserting the guidewire into the pigtail and it came out with kinking. there was resistance upon delivery of the guidewire and the patient was obese. a new wire was used and the pump was implanted. the guidewire was returned and kinking was confirmed. delivery issue - the root cause of the delivery issue was not determined due to insufficient clinical details provided. product damage - the root cause of the product damage was not determined due to insufficient clinical details provided. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27278 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact address and country was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27278.
10524 | 1220648-2025-27279 | 21787157 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21787157&pc=OZD | 03/13/2025 | 04/22/2025 |  | 04/07/2025 |  |  | 03/13/2025 | 04/07/2025 |  |  |  |  | 04/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025609432 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | TRANSCATHETER AORTIC VALVE REPLACEMENT | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that after placement of the impella cp in a 58 year old female they encountered a kink just distal to the outlet. the physician decided to remove and exchange the pump due to patient safety concerns. there was no patient harm as a result. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The investigation of delivery issues has been completed. the kinked cannula on transcatheter aortic valve replacement (tavr) patient. pump was replaced. visual inspection found a kink on can about 15 mm from of cage and can bonding site. no other issues were found on the rest of the pump. the root cause of cannula kink was most likely interaction between tvar and impella during placement. d10 tavr added. instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella cp with smartassist system with smart assist system with automated impella controller. section: warnings & cautions: warnings. section: pre-support evaluation. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention with automated impella controller. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system with automated impella controller. section: warnings & cautions: warnings. section: pre-support evaluation. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27279 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27279. h10 instructions for use was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27279.
10525 | 1220648-2025-27280 | 21787172 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21787172&pc=OZD | 03/13/2025 | 04/30/2025 |  | 04/07/2025 |  |  | 03/13/2025 | 04/07/2025 |  |  |  |  | 04/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026627688 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that when impella cp was removed on a 79 year old male patient, closure of the access site was attempted but unsuccessful. the surgeon was called to repair the femoral artery. the patient survived the procedure. | [5 empty] | The investigation of vascular damage - peripheral vessel has been completed. clinical states that at the time of pump explant of 2 hour case, a femoral damage was observed. at implant, imaging was done and no excessive force was applied. no other supporting information was provided. pump was not returned. therefore, the root cause of the vascular damage issue was not determined since insufficient clinical information about the time and the reason of the damage was provided. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27280 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10526 | 1220648-2025-27281 | 21787181 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21787181&pc=OZD | 03/14/2025 | 04/21/2025 |  | 04/07/2025 |  |  | 03/14/2025 | 04/07/2025 |  |  |  |  | 04/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026618453 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that the impella 5.5 insertion on a 84 year old male was aborted due to the patient vasculature. impella motor would not make the turn from the innominate into the aorta. there was no patient harm as a result. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since the pump could not be implanted due to patient's anatomy. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27281 as it is unknown if the initial reporter also sent the report to the food and drug administration.
10527 | 1220648-2025-27282 | 21787187 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21787187&pc=OZD | 03/14/2025 | 04/16/2025 |  | 04/07/2025 |  |  | 03/14/2025 | 04/07/2025 |  |  |  |  | 04/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025607687 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Tachycardia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that the 86 year old female patient encountered limb ischemia and ventricular fibrillation (vf). a final groin shot was taken to check for distal limb flow and flow appeared to be sluggish. the patient was then transferred to the unit. the patient's distal flow in the right leg was checked and bedside team was unable to find a pulse so an ultrasound of the leg was ordered. at this time, the doctor did not feel a fem fem by pass was a good option as the pump was only placed in relation to the patient coding post cath. while prepping the patient to go back to the unit, the patient bradied down to the fourties, atropine was administered and rhythm was converted. however patient blood pressure was 60s systolic so levophed was started, dobutamine dosing was increased, and dopamine was started in the cath lab. however, the patient converted to v-fib and cpr was initiated. the patient ultimately expired in the lab. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of limb ischemia and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27282 as it is unknown if the initial reporter also sent the report to the food and drug administration. h6 component code 3038 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27282. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smartassist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
10528 | 1220648-2025-27283 | 21787199 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21787199&pc=OZD | 03/14/2025 | 04/11/2025 |  | 04/07/2025 |  |  | 03/14/2025 | 04/07/2025 |  |  |  |  | 04/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026612809 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that following impella cp implant on a 64 year old female patient, the patient experienced oozing from their access site. surgi-seal was applied and the dressing was angle matched to manage the condition. it was additionally reported that the patient lost pulse in their right foot. the condition was monitored with plans to explant the pump when possible. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. as the necessary clinical information was not provided and the device was not returned, the root cause of the limb ischemia was not determined. b.7 information was added that was inadvertently omitted from the initial manufacturer device report number 1220648-2025-27283. e.1 added fax number and us phone number as they were inadvertently omitted from the initial manufacturer device report number 1220648-2025-27283. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27283 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 code 925 was entered incorrectly on the initial manufacturer device report number 1220648-2025-27283 and a new code was added.
10529 | 1220648-2025-27284 | 21791924 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21791924&pc=OZD | 03/10/2025 | 06/03/2025 |  | 04/07/2025 |  |  | 03/10/2025 | 04/07/2025 |  |  |  |  | 04/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1383039 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  | Non Hispanic | American Indian Or Alaskan Native | Death | Yes | Yes | Display or Visual Feedback Problem | Unexpected Shutdown | [5 empty] | Cardiac Arrest | [12 empty] | Death | Required Intervention | [23 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was initially implanted with an impella cp device for approximately eight days for mechanical circulatory support until escalated to an impella 5.5 with extracorporeal membrane oxygenation (ecmo) placement for continued therapy. approximately 41 days after the impella 5.5 mcs started, the automated impella controller suddenly went black without any warning or alarm and restarted on its own. per the report, it took approximately 30 seconds for the aic to restart. after rebooting the aic worked without any problems. however, as a precaution, the aic was subsequently replaced. the patient continued impella and ecmo therapies for an additional 10 days before expiring noted to be due to multi-organ failure. | [5 empty] | Medical safety review of the event noted there was interruption with the automated impella controller (aic), and the aic was exchanged as precaution. there is not enough evidence to exclude the aic as an associated factor in the patient's outcome. there was interruption (possibly brief); impact to the event outcome is unknown. this is one of two devices related to the patient's implant experience. refer to manufacturing report number 1220648-2025-27272 for the impella 5.5 report. the automated impella controller device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - display issue has been completed. the console logs from the reported day of event are consistent with complaint and show that around 13:03 on (b)(6) 2025 console performed full reset (reboot), preceded by unresponsive software process (calculator). prior to reboot, the logs indicate system¿s inability to perform soft-reset (to restart unresponsive process), as well as malfunction of other software elements dependent on system resources. although the log is incomplete and does not include entries before 20:32 on (b)(6) 2025, it clearly shows intermittent, and later permanent, loss of data streaming (usb data transmission from aic to rlm) as evidenced by multiple failed handshake attempts of the icb. after the reboot, pump (b)(6) resumed operation. analysis of rlm journal could not be performed as the unit was unable to fully boot-up and kept rebooting. this correlates with aic¿s inability to communicate with rlm. there was no issue found during the case. the display functioned/operated to specification. console¿s behavior in the field is consistent with unavailability of aic system resources caused by rlm¿s prolonged erratic behavior. the aic software behavior was investigated separately. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27284 as it is unknown if the initial reporter also sent the report to the food and drug administration. h6 component code 765 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27284.
10530 | 1220648-2025-27286 | 21792612 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21792612&pc=OZD | 03/09/2025 | 05/01/2025 |  | 04/07/2025 |  |  | 03/09/2025 | 04/07/2025 |  |  |  |  | 04/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025555528 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Asystole | [11 empty] | Death | Other |  |  |  |  | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, the patient experienced limb ischemia as the patient only had a right posterior tib pulse in the impella leg. the legs were cool and the right foot started to become dusky. the limb ischemia improved with manipulation of the angle of insertion but it was still present. the patient was noted to have worsening multiple organ dysfunction syndrome. maximal impella support eventually progressed and could not increase beyond p5. the patient was deemed not to be a candidate for additional mechanical circulatory support in patient discussions and continued support was no longer wanted by the patient. the patient became non-pulsatile and progressed to asystole and expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27286 as it is unknown if the initial reporter also sent the report to the food and drug administration. h6 component code 3038 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27286.
10531 | 1220648-2025-27285 | 21792832 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21792832&pc=OZD | 02/21/2025 | 10/27/2025 |  | 04/07/2025 |  |  | 03/10/2025 | 04/07/2025 |  |  |  |  | 04/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025603543 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Tachycardia | Thrombocytopenia | [9 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient who experienced recurrent thrombocytopenia, non-sustained ventricular tachycardia and worsening renal/hepatic failure during impella 5.5 support. additionally, the patient experienced some bleeding complications as a large hemothorax was later found in the right pleural space. additionally, the patient was on extracorporeal life support. platelet count presented several times above the protocol threshold. the patient recovered with sequelae. as noted, the patient developed recurring runs of non-sustained ventricular tachycardia as well as frequent recurring premature ventricular contractions. one particular episode of ventricular tachycardia was attributed to mechanical irritation from an attempted endoscopy. the patient outcome secondary to this event is unknown. with regard to the worsening renal failure/hepatic failure, it was noted that the patient was admitted to this facility with acute kidney injury on chronic kidney disease; however, the acute kidney injury deteriorated to renal failure likely secondary to poor end organ perfusion. the patient was started on dialysis. the patient outcome secondary to this event is unknown. | [5 empty] | The investigation of vf/vt/af/at, renal failure, thrombocytopenia, and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-27285. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27285 as it is unknown if the initial reporter also sent the report to the food and drug administration. g2 report source; study was missing from manufacturer device report 1220648-2025-27285. h6 health effect - clinical code 2095 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27285. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10532 | 1220648-2025-27287 | 21793580 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21793580&pc=OZD | 03/10/2025 | 04/15/2025 |  | 04/07/2025 |  |  | 03/10/2025 | 04/07/2025 |  |  |  |  | 04/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026628477 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that them pump presented high purge pressure. the high purge pressure alarm was discussed with the nurse and a kink was found in line near the purge cassette door. however, a few days later the family decided to withdraw care and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen." | The investigation of high purge pressure has been completed. pump was not returned for investigation. the data log shows a quick rise in purge pressure with a purge system blocked alarm, which resolved in 10 minutes. according to the clinical details, a purge line kink was noted near the purge cassette door. the cause of the high purge pressure issue was most likely kinked purge line near purge cassette door, based on given clinical. b2 date of death was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27287. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27287 as it is unknown if the initial reporter also sent the report to the food and drug administration.
10533 | 1220648-2025-27288 | 21793660 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21793660&pc=OZD | 03/12/2025 | 04/28/2025 |  | 04/07/2025 |  |  | 03/12/2025 | 04/07/2025 |  |  |  |  | 04/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025611454 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Other | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient was noted to have new neurological deficits consistent with signs and symptoms of stroke. the physician did not necessarily believe it was directly connected to the impella versus normal surgical complication risk. no known actions were taken. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of stroke/cva has been completed. the clinical reported the patient having new neurological deficits a day after implant. pump was returned and biomaterial was observed in the outflow cage. strokes/cvas can occur during or after impella support, and can be either ischemic or hemorrhagic. to make a determination as to the cause of the stroke, the following clinical information must be considered: -patient's medical history, including any previous strokes, cardiovascular disease, or risk factors such as hypertension, diabetes, or atrial fibrillation -timing of the stroke in relation to impella support (during or post-implant) -detailed description of the symptoms experienced by the patient -neurological examination findings and any focal deficits observed -diagnostic imaging results, such as ct scan or mri of the brain, to identify the type and location of the stroke (ischemic or hemorrhagic) -laboratory test results, including coagulation profiles and cardiac markers -any relevant procedural or device-related factors, such as duration and settings of impella support, presence of embolic protection devices, or any documented procedural complications -response to any previous anticoagulation or antiplatelet therapies -evaluation of potential alternative causes of stroke, such as arrhythmias, embolic sources, or underlying vascular pathology as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27288 as it is unknown if the initial reporter also sent the report to the food and drug administration. h5 was erroneously entered on manufacturer device report 1220648-2025-27288.
10534 | 1220648-2025-27289 | 21793671 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21793671&pc=OZD | 03/07/2025 | 04/07/2025 |  | 04/07/2025 |  |  | 03/08/2025 | 04/07/2025 |  |  |  |  | 04/07/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026623861 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the complainant reported that there was access site bleeding. the patient had bleeding from the groin overnight and one unit of packed red blood cells (prbc) was given. the groin site was assessed with intensivist and it was re-sutured. as a result, the bleeding stopped. the patient outcome was noted as stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10535 | 1220648-2025-27229 | 21772252 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21772252&pc=OZD | 03/10/2025 | 04/30/2025 |  | 04/04/2025 |  |  | 03/11/2025 | 04/04/2025 |  |  |  |  | 04/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026620882 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp for mechanical circulatory support. the complainant reported that low pump flows were encountered. despite successful repositioning of the pump, the low flows persisted and the decision was made to remove the pump. there was no patient harm. | [5 empty] | The investigation of low pump flow has been completed. low pump flow: clinical states that pump was repositioned since it was facing mitral apparatus, patient was already put on extracorporeal membrane oxygenation (ecmo) and later intermittent suction alarms were observed where pump flows dropped to 0l/min. pump was repositioned but no success to get better flow. pump was not returned for investigation. the root cause of the low pump flow issue was most likely biomaterial since a spike in motor current corresponding with a lost in motor current pulsatility and a drop in pump flow was observed in the data logs. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10536 | 1220648-2025-27230 | 21772593 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21772593&pc=OZD | 03/11/2025 | 04/22/2025 |  | 04/04/2025 |  |  | 03/11/2025 | 04/04/2025 |  |  |  |  | 04/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1167410 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller had an error. the error message was ¿impella controller error, please replace controller¿. no patient was involved. | [5 empty] | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. the console logs from the reported day of event showed that upon boot-up the console presented with controller error alarm. low values of light parameter indicated there was no light reaching the ccd detector in the optical bench. issue was not reproduced during testing, but log data is consistent with a failure of either lamp assembly or optical bench electronics. the cause of the aic issue was a defective optical bench. d.2 revised common device name as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-27230. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
10537 | 1220648-2025-27231 | 21772635 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21772635&pc=OZD | 03/11/2025 | 04/04/2025 |  | 04/04/2025 |  |  | 03/11/2025 | 04/04/2025 |  |  |  |  | 04/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026619180 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient had an attempted impella 5.5 implant and the pump was unable to cross due to vascular stenosis. the implant was aborted due to patient's anatomy. there was no reported patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.
10538 | 1220648-2025-27232 | 21772680 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21772680&pc=OZD | 03/11/2025 | 05/29/2025 |  | 04/04/2025 |  |  | 03/11/2025 | 04/04/2025 |  |  |  |  | 04/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026619297 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | Material Split, Cut or Torn | [5 empty] | Vascular Dissection | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient undergoing high risk percutaneous coronary intervention (hrpci) was implanted with an impella cp device for mechanical circulatory support. it was reported that the team made the decision to also place a 7 french companion sheath for the hrpci. the team observed a vascular damage/perforation of the iliac artery from the wireless insertion of the 7 french sheath. the vessel damage prompted the pump to be explanted and covered with two stents to repair the artery. the patient later expired. it is unknown if the use of the impella was related to patient death. thus, there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | Addition findings during the investigation that there was introducer damage issue. there was a sheath tip split which happened due to the improper implant technique. h10 additional investigation findings were inadvertently omitted on the initial manufacturer device report number 1220648-2025-27232-1. h6 medical device problem code 4008 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27232-1. | The investigation of vascular damage - peripheral vessel has been completed. vascular damage - peripheral vessel: clinical states that, at the time of implant, the companion sheath was inserted without a guidewire resulting in artery dissection. companion sheath, 14f introducer, and pump were returned. the companion sheath had a damaged tip and the partial tip missing. these supports the clinical that the companion sheath was used without guidewire. therefore, the root cause of the vascular damage issue was use related to improper technique as evidenced in the product investigation. | The pump and introducer were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
10539 | 1220648-2025-27234 | 21772755 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21772755&pc=OZD | 03/11/2025 | 04/08/2025 |  | 04/04/2025 |  |  | 03/11/2025 | 04/04/2025 |  |  |  |  | 04/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025580360 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Pump | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  | MW5168064 | Yes | Initial submission | Followup | [8 empty] | The complainant reported that upon starting support, an "impella stopped, motor current high" alarm appeared. multiple attempts to start the pump at p-2 were made with only one successful result with motor current in the 900's. support was discontinued and the pump was replaced. the patient remained stable. | [5 empty] | E.4 added yes to reflect that the initial reporter also sent report to the food and drug adminstration h.9 med watch report was added. regulatory attachments - med watch report was added. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10540 | 1220648-2025-27233 | 21772759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21772759&pc=OZD | 03/11/2025 | 05/15/2025 |  | 04/04/2025 |  |  | 03/11/2025 | 04/04/2025 |  |  |  |  | 04/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025588393 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 31 YR |  | Male |  |  | White | Malfunction | No | Yes | Restricted Flow rate | Pumping Stopped | Infusion or Flow Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp was exchanged due to low pump flow. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of low pump flow has been completed. the pump was returned for investigation. no physical abnormalities were reported on the returned product. further investigation revealed large purge gap, which was due to bearing wear of the device. the pump was further tested in a bucket of water, and it was able to run with observed saturated pump flow. the data log shows no low pump flow at the reported time of the suction complaint. a temporary pump stop was seen, followed by the saturated pump flow at end of case run. there were several suction alarms and low pump flow during the case run without any noted positioning issues or motor current spikes. the failure modes of temporary pump stop and saturated pump flow were added as investigated failure modes based on the returned product investigation. the cause of the low pump flow issue was not determined since sufficient clinical information was not provided for the investigation. the cause of the temporary pump stop was bearing wear, which was due to pump's use beyond the design life of the product. the cause of the saturated pump flow was bearing wear, which was due to pump's use beyond the design life of the product. instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: entering cardiac output : ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention, section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
10541 | 1220648-2025-27235 | 21772869 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21772869&pc=OZD | 03/11/2025 | 04/28/2025 |  | 04/04/2025 |  |  | 03/11/2025 | 04/04/2025 |  |  |  |  | 04/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025561065 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a 14 french and an impella cp sheath placed to the femoral to support an acute myocardial infarction/cardiogenic shock patient. the patient placed perclose and performed the percutaneous coronary intervention to the coronary arteries and then had a groin complication. the team accidentally cut the perclose suture and had a groin bleed that required the medical doctor to scrub back in and place a hemastat to the vessel and pull on the suture. the bleeding did not resolve and femstop was placed. the patient remained stable. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the access site bleeding was most likely use issue since team accidentally cut the preclose suture and had a groin bleed. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10542 | 1220648-2025-27237 | 21773226 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21773226&pc=OZD | 03/11/2025 | 04/28/2025 |  | 04/04/2025 |  |  | 03/11/2025 | 04/04/2025 |  |  |  |  | 04/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026612834 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Sepsis | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that despite adequate fluid status as well as several repositioning attempts, the impella 5.5 has continuous suction above p4. the patient was taken in to optimize position with fluoro and transthoracic echocardiogram guidance. the inlet was advanced and rotated to be in mid-cavity. this did not allow for improved flow. team will give some volume to see if that improves ability for higher p-level. the patient remains on support. | It was further reported that the patient became septic overnight and had a spike in white blood cell count. levo and antibiotics were given and condition seem to have improved. the next day, the patient remained septic but slowly improving |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of infection/sepsis and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6b added b1 should have reflected both adverse event and product problem on the initial manufacturer device report number 1220648-2025-27237. b5 additional information that was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27237. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27237 as it is unknown if the initial reporter also sent the report to the food and drug administration. h1 type of report was erroneously selected on the initial manufacturer device report number 1220648-2025-27237. h6 health effect - clinical code 2067 and health effect - impact codes 4644 and 4641 were inadvertently omitted on the initial manufacturer device report number 1220648-2025-27237.
10543 | 1220648-2025-27238 | 21773338 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21773338&pc=OZD | 03/12/2025 | 04/28/2025 |  | 04/04/2025 |  |  | 03/12/2025 | 04/04/2025 |  |  | 08/31/2025 |  | 04/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025443697 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient had an impella cp pump placed for support of acute myocardial infarction. the pump was placed to allow for percutaneous coronary intervention. the patient lost pulse bilaterally but, the pump remained on for support until the family made the decision to withdraw care after 27 hours. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g4 pma/510(k) number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27238. h6 component code 3038 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27238. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10544 | 1220648-2025-27239 | 21773378 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21773378&pc=OZD | 03/13/2025 | 04/04/2025 |  | 04/04/2025 |  |  | 03/13/2025 | 04/04/2025 |  |  |  |  | 04/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025592201 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Pericardial Effusion | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a cp to support a 84 year old male patient undergoing cardiac procedure for coronary artery disease. the team had a groin site bleed that prompted treatment. the site was re-sutured due to heavy bleeding and angle matching done. the pump remains on for support and the patient was transferred to the tertiary center. additionally, patient was found to have a large pericardial effusion and was opened emergently. after multiple attempts to repair lv, family was contacted and patient was sent to unit in preparation for passing. impella had been placed in the aorta in surgical mode before case was started. surgeon uncertain of what caused perforation, as the impella was shallowed out to aortic valve before perforation. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10545 | 1220648-2025-27240 | 21773441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21773441&pc=OZD | 03/13/2025 | 05/08/2025 |  | 04/04/2025 |  |  | 03/13/2025 | 04/04/2025 |  |  |  |  | 04/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025581108 |  | 1000459 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 50 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was escalated to an impella 5.5, the purge pressure increased after the pump was driven. the cassette was replaced with no improvement so the pump was replaced in consideration of future events. the patient survived the procedure. | [5 empty] | The pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | The investigation of high purge pressure has been completed. the pump was sent for investigation. no physical abnormalities were reported on the returned product. the purge gap and impeller were free of biomaterial. the pump was primed successfully, and no issues were observed with purge pressure. additionally, the pump was run as per specifications in a bucket of water. the data log was reviewed, showing a 15-minute gradual rise in blood pressure without reporting any motor spikes. the pump was replaced after 30 minutes of operation. the cause of the high purge pressure issue was not determined since the issue could not be reproduced on the returned product, and the data log analysis was not sufficient to derive the root cause without sufficient clinical information.
10546 | 1220648-2025-27241 | 21773629 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21773629&pc=OZD | 03/13/2025 | 04/04/2025 |  | 04/04/2025 |  |  | 03/13/2025 | 04/04/2025 |  |  | 12/31/2025 |  | 04/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025449318 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that following impella cp implant on a patient, the patient developed a small hematoma prior to the peel away sheath being removed. an aggressive massage by the provider exacerbated the bleed which prompted a blood transfusion. the patient was transferred and the decision was made to explant the pump due to the bleed. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10547 | 1220648-2025-27242 | 21773672 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21773672&pc=OZD | 03/13/2025 | 05/20/2025 |  | 04/04/2025 |  |  | 03/13/2025 | 04/04/2025 |  |  |  |  | 04/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025509232 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that during impella cp support of a 71 year old male patient, they encountered positioning issues. after placement of the impella, it was pushed into the left ventricle, the cannula seemed to be bent and there were low flows. repositioning of the pump was unsuccessful and suction alarms persisted. ultimately, the physician decided to exchange the device. the patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of positioning issues, product damage (cannula kink), and low pump flow has been completed. product and data was not returned for review. based on clinical description, the root cause of positioning issue was most likely use related since the pump was pushed into left ventricle (lv) after exchanging the 14 fr. sheath with the repositioning sheath. the root cause of low pump flow was most likely due to a kinked cannula since the clinical stated the low flow occurred after kinked cannula was found in lv. the root cause of product damage (cannula kink) was most likely use related since the impella was accidentally pushed deeper into the lv after exchanging the 14 fr. with the repositioning sheath. g1 contact address and country was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27242.
10548 | 1220648-2025-27243 | 21773818 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21773818&pc=OZD | 03/13/2025 | 04/04/2025 |  | 04/04/2025 |  |  | 03/13/2025 | 04/04/2025 |  |  |  |  | 04/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025571116 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 23 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Thrombosis/Thrombus | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient had an impella 5.5 pump placed for support, along with extracorporeal membrane oxygenation, for cardiomyopathy. it was reported that during support, there was a clot on the pump outlet and in the right leg and mediastinum that prompted a surgical washout. additionally, the patient begain having high purge pressures and blocked alarms. the purge was changed from bicarbonate to tissue plasminogen activator and the purge pressures improved. the patient remained on support in critical care in a tenuous state. after some time, the family decided to withdraw care and the patient expired. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist. section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smart assist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.". in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿
10549 | 1220648-2025-27244 | 21773871 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21773871&pc=OZD | 03/13/2025 | 04/04/2025 |  | 04/04/2025 |  |  | 03/13/2025 | 04/04/2025 |  |  |  |  | 04/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025585496 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Difficult to Insert | Defective Device | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that during an attempted delivery of a 82 year old male patient, the impella cp pump was unable to traverse through an inserted 14x25 peel away sheath. it was noted that the patient's artery was tortuous with some calcification. upon removal of the sheath, the component was found to be bent / kinked. a 16fr gore sheath was subsequently inserted and the impella cp implant was successful. there was no patient harm. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
10550 | 1220648-2025-27245 | 21773889 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21773889&pc=OZD | 03/12/2025 | 04/30/2025 |  | 04/04/2025 |  |  | 03/12/2025 | 04/04/2025 |  |  |  |  | 04/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026618466 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was attempted to be implanted with an impella 5.5 device for mechanical circulatory support. the pump was unable to advance through the patient's anatomy during attempted delivery. the pump was removed and an impella cp pump was placed for patient support. there was no patient harm. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the impella 5.5 unable to pass around the anatomy and the case was aborted and a cp was used. the cause of the delivery issue was most likely patient condition related as the pump was unable to pass through patient's anatomy and resistance was met at axillary artery. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
10551 | 1220648-2025-27246 | 21774337 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21774337&pc=OZD | 03/09/2025 | 04/14/2025 |  | 04/04/2025 |  |  | 03/09/2025 | 04/04/2025 |  |  |  |  | 04/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025552564 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient developed a gastrointestinal bleed during impella 5.5 support. heparin was being administered through the purge at the time of the bleed. one unit of packed red blood cells was given to counteract the bleed. the bleed persisted for roughly two weeks following the initial intervention. care was eventually withdrawn after thirty days of support and the patient passed away. there is not enough evidence to disassociate the patient's passing with the noted condition during support. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. h1 type of report was erroneously selected on the initial manufacturer device report number 1220648-2025-27246. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10552 | 1220648-2025-27247 | 21774442 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21774442&pc=OZD | 03/11/2025 | 04/04/2025 |  | 04/04/2025 |  |  | 03/11/2025 | 04/04/2025 |  |  |  |  | 04/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025599731 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an implanted impella cp pump experienced positioning complications during patient support due to the presence of a recent transcatheter aortic valve replacement. the pump was removed and a new pump was placed under fluoroscopy and an echocardiogram for optimized positioning. there was no patient harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10553 | 1220648-2025-27248 | 21774855 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21774855&pc=OZD | 03/07/2025 | 04/04/2025 |  | 04/04/2025 |  |  | 03/07/2025 | 04/04/2025 |  |  |  |  | 04/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025558284 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | Yes | Pumping Stopped | [6 empty] | Cardiac Arrest | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient undergoing a high-risk percutaneous coronary intervention (pci) was implanted with an impella cp device for mechanical circulatory support, and experienced a pump stop unexpectedly during the support. it was noted that the automated impella controller screen returned to the home screen and showed a prompt stating "press ok to resume displayed performance level" at the time of the failure. the pump restarted on its own. however, it was explanted following pci completion. the impella cp was uneventfully removed. the patient noted to be hemodynamically stable, and hemostasis was achieved with two perclose. however, following the explant, the patient expired. no details surrounding the demise outcome was provided. the impact of the temporary pump stop is unclear. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. this is one of two devices associated with the event. refer to manufacturing report number 1220648-2025-27153 for the automated impella controller device report. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: entering cardiac output: ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
10554 | 1220648-2025-27249 | 21775890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21775890&pc=OZD | 06/24/2024 | 04/04/2025 |  | 04/04/2025 |  |  | 03/11/2025 | 04/04/2025 |  |  |  |  | 04/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | UNKNOWN |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Cardiac Arrest | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Abiomed japan forwarded a publication from account university of washington entitled: successful treatment of immune checkpoint inhibitor-associated fulminant myocarditis with abatacept and ruxolitinib: a case report. the publication speaks to "he required inotropes and a percutaneous microaxial flow pump placement for temporary mechanical circulatory support and then had sustained ventricular tachycardia resulting in cardiac arrest requiring several rounds of cardiopulmonary resuscitation and pericardial tamponade requiring urgent pericardiocentesis. after cardiac arrest, he was admitted to the intensive care unit, intubated..." no additional information provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
10555 | 1220648-2025-27250 | 21776353 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21776353&pc=OZD | 03/11/2025 | 04/04/2025 |  | 04/04/2025 |  |  | 03/11/2025 | 04/04/2025 |  |  | 04/30/2021 |  | 04/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 1425034 |  | 0048-0034-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Literature | Health Professional | User facility | Company representation |  | 71 YR |  | Male |  |  | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | Ventricular Fibrillation | Thrombosis/Thrombus | Intraoperative Cardiac Injury | [9 empty] | Death | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | A |  |  |  | Initial submission |  |  |  |  | 03/11/2025 |  |  |  |  | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. the patient was supported for 92 hours before the patient was explanted and supported by extracorporeal membrane oxygenation. in 2025, a published article was identified. the authors noted the following: "we performed the bentall procedure, mitral valve repair, tricuspid annuloplasty, and coronary artery bypass graft on a 73-year-old man admitted with congestive heart failure caused by annuloaortic ectasia, along with severe aortic and mitral regurgitation. va-ecmo and impella were required, since the cardiopulmonary bypass weaning was difficult. impella was removed on postoperative day 4. on postoperative days 5 and 6, laboratory data showed worsening renal dysfunction, lactate levels, and acidosis. contrast-enhanced computed tomography showed thrombosis in the celiac and superior mesenteric arteries. aortic thrombectomy was performed. hyperkalemia, caused by a reperfusion injury, resulted in ventricular fibrillation. continuous hemodiafiltration improved the hyperkalemia. however, irreversible acidosis progressed, and the va-ecmo flow rate could not be sustained. on postoperative day 7, the patient died." no further information was provided. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. literature citation: yamana, f., domae, k., kawasumi, r., sakamoto, t., hata, m., shirakawa, y., masai, t., & sawa, y. (2023). aortic thrombosis with visceral malperfusion during circulatory support with a combination of impella and extracorporeal membrane oxygenation for postcardiotomy cardiogenic shock. journal of artificial organs, 26(4), 330¿334. https://doi.org/10.1007/s10047-023-01382-3. as noted in the impella instructions for use: impella cp with smart assist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation, section: impella cp with smart assist catheter insertion, section: axillary insertion of the impella cp with smart assist catheter, section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
10556 | 1220648-2025-27251 | 21777039 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21777039&pc=OZD | 03/10/2025 | 04/04/2025 |  | 04/04/2025 |  |  | 03/11/2025 | 04/04/2025 |  |  |  |  | 04/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1147985 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that they experienced an automated impella controller (aic) shutdown/restart issue. a brief support interruption was encountered with a short pump stop. the console was reset and there was a dark screen. the system then came back on and the pump resumed operating. as result, the aic console was replaced and the case resumed on an alternate aic. there was no patient harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10557 | 1220648-2025-27252 | 21777316 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21777316&pc=OZD | 03/12/2025 | 04/30/2025 |  | 04/04/2025 |  |  | 03/12/2025 | 04/04/2025 |  |  |  |  | 04/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026617650 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the flow of an impella cp pump began to decrease. the device was repositioned, but the issue persisted. the pump was subsequently replaced in response to the failure. there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of low pump flow has been completed. the returned complaint device visually inspected. damaged impeller blades found. biomaterial observed around impeller. the cause of the low pump flow issue was most likely damaged impeller blades due to transcatheter aortic valve replacements (tavr) interaction. biomaterial mostly accumulated during low pump flow issue. per 0048-9007, "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades." g4 pma/510(k)number was erroneously entered on the initial manufacturer device report number 1220648-2025-27252.
10558 | 1220648-2025-27253 | 21777720 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21777720&pc=OZD | 03/12/2025 | 04/04/2025 |  | 04/04/2025 |  |  | 03/12/2025 | 04/04/2025 |  |  |  |  | 04/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025609840 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. due to anatomical issues, the pump kept dislodging from the left ventricle and the decision was made to explant the pump. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10559 | 1220648-2025-27254 | 21778015 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21778015&pc=OZD | 03/13/2025 | 04/04/2025 |  | 04/04/2025 |  |  | 03/13/2025 | 04/04/2025 |  |  |  |  | 04/04/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025594119 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the 64 year old female patient on impella cp support had no pulse present in the right lower extremity. reperfusion sheath was placed but perfusion did not improve. there were no options for escalation of care. the patient was successfully weaned. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10560 | 1220648-2025-27199 | 21760641 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21760641&pc=PYX | 03/09/2025 | 04/03/2025 |  | 04/03/2025 |  |  | 03/09/2025 | 04/03/2025 |  |  |  |  | 04/03/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025540753 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the medical doctor elected to cardiovert non emergently the patient mid-procedure without transesophageal echocardiogram. immediately after, flows were decreased and pulmonary artery pressures increased with a flat motor current. clot ingestion was suspected. the pump was removed due to suspected clot ingestion. high pulmonary artery placement signal with low flows on p-8 were noted. the patient was subtherapeutic on heparin. the patient had an intra-aortic balloon pump placed. the patients outcome was stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10561 | 1220648-2025-27200 | 21760665 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21760665&pc=OZD | 03/10/2025 | 04/03/2025 |  | 04/03/2025 |  |  | 03/10/2025 | 04/03/2025 |  |  |  |  | 04/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025601622 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the same day of impella cp insertion on a patient, the groin site had a slight ooze. the partial thromboplastin time was over 200, so systemic heparin was held. the next day, the patient was complaining of numbness/tingling in leg, unable to doppler distal pulse. the patient was brought back to the catheterization lab, and the impella cp was removed. hemostasis was achieved. the groin was shot from contralateral groin showing flow restored down leg post removal and closure. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10562 | 1220648-2025-27201 | 21760703 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21760703&pc=OZD | 03/10/2025 | 04/25/2025 |  | 04/03/2025 |  |  | 03/10/2025 | 04/03/2025 |  |  |  |  | 04/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026613881 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella cp was not able to be inserted in a patient due to the patient's anatomy. there was resistance in the femoral artery. the procedure was aborted. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, impella could not be inserted because of kinking in the femoral artery.
10563 | 1220648-2025-27202 | 21760746 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21760746&pc=OZD | 03/10/2025 | 04/14/2025 |  | 04/03/2025 |  |  | 03/10/2025 | 04/03/2025 |  |  |  |  | 04/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1561819 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient had an automated impella controller (aic) being powered up for use with an impella cp. an error occurred after starting up the aic. the aic could not be used which prompted aic replacement. there was no patient harm reported. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - system self-check error has been completed. the console logs showed a communication issue with the pbm (power battery manager) during the initial bootup. due to the communication issue, the aic console rebooted automatically (as designed) to attempt another power on. after rebooting, the console displayed the "system self check failed" screen, and the console was then forcefully shut down. visual inspection confirmed the pbm contamination, which has impacted a neighboring rs232 communication channel. root cause: the root cause of the ¿system self check error¿ was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps.
10564 | 1220648-2025-27203 | 21760834 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21760834&pc=OZD | 03/07/2025 | 05/21/2025 |  | 04/03/2025 |  |  | 03/11/2025 | 04/03/2025 |  |  |  |  | 04/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025605432 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 53 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the impella cp was dislodged into the aorta. after inserting the impella, the touhy-borst was not tightened and the impella cp was dislodged from the left ventricle when the patient's position was changed by the nurse. the impell a cp was removed. the patient was managed with extracorporeal membrane oxygenation alone, but the aortic valve was closed and a haze was visible in the left ventricle on ultrasound. the next day, a new impella cp pump was placed. there was no resulting patient harm.00 | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. clinical stated impella cp was dislodged into the aorta because the touhy-borst was forgotten to be tightened after insertion and the pump came out of position from the left ventricle when the patient's position was changed. the cause of the positioning issues most likely was a touhy-borst issue due to improper technique since clinical stated the toughy was not tightened causing the pump to be dislodged from the left ventricle when the patient's position was changed.
10565 | 1220648-2025-27204 | 21761436 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21761436&pc=OZD | 03/08/2025 | 05/13/2025 |  | 04/03/2025 |  |  | 03/08/2025 | 04/03/2025 |  |  |  |  | 04/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026612673 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | Yes | Failure to Pump | [6 empty] | Cardiac Arrest | [12 empty] | Life Threatening | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following a difficult impella 5.5 device implant, complications were encountered when attempting to start the pump. the staff noted that a spike in motor current occurred during each attempt followed by an immediate stop. the staff was advised to exchange the pump; however, this option was not available. the impella 5.5 device implant ultimately failed. the patient's outcome was subsequently updated to expired. no further details surrounding the demise was provided. the impact of the pump's failure to start to the event is unclear. | [5 empty] | The investigation of pump did not start has been completed. pump did not start: upon visual inspection, a severe kinking of the catheter present just below motor housing. electrical testing was performed and open lines were present in all three motor cable lines. broken motor cable lines were present at kinking in catheter below motor housing. two motor cable lines were found to be completely broken, and third motor cable line showed excessive necking. data logs were reviewed and real time (rt) log at beginning of case shows pump was attempted to be started multiple times with immediate motor current spikes to 30000 with "impella stopped (rpm deviations)" alarms being triggered. clinical details noted that the clinical team had difficulty inserting the pump due to patients anatomy and attempted insertion multiple times before insertion was successful. the root cause of the pump not starting was due to an electrical open of the motor cable lines as a result of excessive force when inserting the pump based on review of returned product and data log analysis. b2 life-threatening was erroneously selected on the initial manufacturer device report number 1220648-2025-27204. g1 manufacturing site name, address, country and fax number were erroneously entered on the initial manufacturer device report number 1220648-2025-27204. | The impella device was returned, and an evaluation is underway. upon completion of the analysis, a supplemental report will be filed. "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)."
10566 | 1220648-2025-27205 | 21761734 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21761734&pc=OZD | 03/10/2025 | 05/01/2025 |  | 04/03/2025 |  |  | 03/10/2025 | 04/03/2025 |  |  |  |  | 04/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026622553 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an impella 5.5 was attempted to be placed in a patient. after a lengthy attempt utilizing multiple troubleshooting techniques, the axillary site was aborted. attention was then turned to direct aortic insertion. unfortunately, there was not enough length on the aorta for graft placement. there were no more 10 centimeter (cm) grafts in-house, so an eight (cm) hemashield platinum was sewn onto the innominate artery. the sheath was placed. attempts to pass the wire were unsuccessful as the wire and catheter kept passing into the descending aorta. at this time, the decision was made to place the impella 5.5 wirelessly. however, they could not pass the impella 5.5 across the aortic valve even with manual manipulation of the aortic annulus, epinephrine and full ejection of the heart. the patient was still on bypass at this time with minimal room on her aorta. it was discussed that aortotomy was likely the only option left. however, there was no room for cross clamp beneath the aortic cannula without clamping the impella motor/cannula. flows were stopped from cpb, the descending aorta was clamped, a transverse aortotomy was made, the pump was pulled up through and out of the aortotomy, a cross clamp was beneath aortic cannula across the catheter of the impella, descending clamp was removed and flows restarted from cpb. the impella was then placed across the aortic valve under direct visualization. upon inspection of the aortic valve, it was discovered the patient had aortic stenosis was fusion of leaflets. the aortotomy was closed and the impella position was confirmed with transesophageal echocardiogram. the patient was then weaned from bypass while increasing impella flows. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be patient condition because attempt was made to place impella via right axillary artery and the insertion was aborted due to resistance in innominate artery.
10567 | 1220648-2025-27206 | 21761759 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21761759&pc=OZD | 03/10/2025 | 04/03/2025 |  | 04/03/2025 |  |  | 03/10/2025 | 04/03/2025 |  |  |  |  | 04/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025560046 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Positioning Failure | Restricted Flow rate | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that the 85-year-old male patient was on impella cp support for support during a cardiac procedure. it was noted that after the implant, the patient had a large hematoma at the access site. at the same time the device was unable to be repositioned out of the mitral apparatus and was impeding flows with the calcification on both valves. the pump was explanted. it was later reported that care was withdrawn and the patient expired. there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations: "assess access site for bleeding and hematoma." section: entering cardiac output: use of the repositioning sheath and peel-away introducer: "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site." section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)."
10568 | 1220648-2025-27207 | 21762274 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21762274&pc=OZD | 03/06/2025 | 04/03/2025 |  | 04/03/2025 |  |  | 03/06/2025 | 04/03/2025 |  |  |  |  | 04/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025591434 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient on impella cp had encountered limb ischemia. when the bedside nurse rechecked, the distal pulse was unable to be dopplered and so the vascular team was consulted. the decision was made to intubate the patient and place a distal perfusion catheter. the vascular surgeon successfully placed a femorofemoral distal limb perfusion sheath with positive flow to the impella side, noted on vascular ultrasound. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10569 | 1220648-2025-27208 | 21762298 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21762298&pc=OZD | 03/10/2025 | 05/05/2025 |  | 04/03/2025 |  |  | 03/10/2025 | 04/03/2025 |  |  |  |  | 04/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2025605737 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 68 YR |  | Male |  |  |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Embolism/Embolus | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that an implanted impella cp pump in a 68 year old male patient began to experience low flows during support. biomaterial was observed and the decision was made to escalate to impella 5.5 support. the patient survived the procedure. | [5 empty] | The investigation of low pump flow and thrombus has been completed. the cut pump was sent for investigation. data logs were not provided for this case run. the returned pump had biomaterial inside the cannula and on the impeller. no further investigation can be performed. the cause of the low pump flow issue was not determined since sufficient product and clinical information was not returned. as the necessary information was not provided, the root cause of the thrombus was not determined. h6 code 4582 was reported incorrectly on manufacturer device report 1220648-2025-27208. new code was added to health effect - clinical code. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
10570 | 1220648-2025-27209 | 21762458 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21762458&pc=OZD | 03/11/2025 | 05/21/2025 |  | 04/03/2025 |  |  | 03/11/2025 | 04/03/2025 |  |  | 11/04/2020 |  | 04/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1200534 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | Application Program Problem | [5 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient had been escalated from an impella cp to an impella 5.5 pump. the pump was successfully placed after some issues in gaining the left ventricle access. during support, there were multiple controller alarms which made the pump stop. the automated impella controller was replaced and there were no issues. subsequently, the patient expired. | [5 empty] | The investigation of automated impella controller (aic) - controller failure (red alarm) and aic - pump stop has been completed. aic - controller error (yellow alarm) was added as a failure mode. console logs from the reported day of the event are consistent with the complaint. the case started on (b)(6) 2025 and on (b)(6) 2025 multiple epm reset messages was seen after which multiple "controller error (power battery manager (pbm) voltage out-of-spec) was triggered along with them multiple controller error (loss of (epm). subsequently, impella stopped (mechanical/electrical issue) alarm was triggered on which the purge cassette was removed as well as shutdown was requested, however "shutdown requested by user but pump is connected" error was displayed. additionally, further errors were logged, including "placement signal not reliable. pump motor driver (pmd) has detected a pump speed deviation - alarm issued and "controller error (opm comm crc invalid) [optical pump connected]." after multiple requests for shutdown, a hard shutdown was ultimately initiated at 09:38:22 on 11th march 2025. device analysis: the console was returned to field service for further investigation. a thorough visual inspection of the internal components was conducted, but no physical problems were found. the power supply from the pbm to the impellatronic pca and the 4-in-1 unit was stable. while test the aic the grey push button stopped working, and while investigating found that the 4in1 was messing up the communication with all subsystems within the aic, further looking into the real time logs, was able to see the some of the similar error messages as displayed during the case. replacing the 4in1, aic worked as required and no issues were observed. root cause: the root cause of the controller errors "loss of comm (epm) and pbm voltage out-of-spec" was most likely due to the defective 4in1 board. the root cause of the impella stopped "mechanical/electrical issue and pmd has detected a pump speed deviation" was also most likely due to the defective 4in1 board. there were no controller failure alarm was discovered during the case. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
10571 | 1220648-2025-27210 | 21762691 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21762691&pc=OZD | 03/11/2025 | 04/03/2025 |  | 04/03/2025 |  |  | 03/11/2025 | 04/03/2025 |  |  |  |  | 04/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1233223 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported by biomedical that an console (aic) experienced a battery failure. there was no noted patient involvement. | [5 empty] | The investigation for the aic battery alarm has been completed. console logs were consistent with what was reported in the complaint. multiple ¿battery failure (transport connection blown/missing) alarms [events set #360] were observed in the imc logs. console was returned for evaluation. the reported battery failure issue was able to be reproduced. results of running batttest revealed that cell 3 of battery 2 had a voltage of 0mv. isolative testing was done by swapping battery 2 with a known good one which resolved the alarm. the root cause of the battery failure was determined to be a cell imbalance in battery 2.
10572 | 1220648-2025-27211 | 21762841 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21762841&pc=OZD | 03/06/2025 | 05/20/2025 |  | 04/03/2025 |  |  | 03/06/2025 | 04/03/2025 |  |  |  |  | 04/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025610318 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hematoma | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and later this day, venoarterial extracorporeal membrane oxygenation was placed (va ecmo). additionally, a hematoma was noticed around impella site. manual pressure was held for one hour and hematoma was resolved. the patient required two units of blood. positive pedal pulses remained. va ecmo and impella were consistently having suction events. impella unable to run above p-1 performance level. a gastrointestinal bleed was noted and the patient received eight more units of blood. transfer to a larger facility was recommended and completed the next day. upon arrival, the patient was still critically ill. flows on va ecmo 3.5, highest of 5.0. impella p-1 to p-2, flows 0.7 highest of 2.4 with ecmo circuit change. there was a concern about retroperitoneal bleed. continuous renal replacement therapy was started. the patient received multiple pushes of epinephrine en route for bradycardia. impella insertion site was dry with known bruising, soft to palpation. distal perfusion was poor with patient¿s current state with no plans to escalate to impella 5.5 as left ventricle was well vented with the impella cp. the overall plan was to get the patient hemodynamically stable overnight, evaluate the bleeding. the following day, it was noted that the patient expired. there is not enough information to exclude the impella cp as an associated factor with the event outcome. however, to note, the nurse noted the patient came in asystole, chemically coded for four hours and massive transfusion protocol followed. | [5 empty] | The investigation of access site bleeding and low pump flow has been completed. the impella device was not returned for evaluation. bleeding reported around the access site, blood given and manual pressure held. the root cause of access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided. low pump flow: shortly after arrival to the unit patient had suction alarms. they had to be put on extracorporeal membrane oxygenation (ecmo). they also had access site bleeding requiring 2 units of blood. the next day both cp and ecmo with reports of suction. data logs confirm suction alarms and show the placement signal (ps) trending down. no motor current (mc) spikes outside of p-level changes. the cause of the low pump flow was most likely patient condition related to the patient's deteriorating condition, bleeding issues requiring blood to be given, and ecmo also having suction. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10573 | 1220648-2025-27212 | 21762886 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21762886&pc=OZD | 03/11/2025 | 04/03/2025 |  | 04/03/2025 |  |  | 03/11/2025 | 04/03/2025 |  |  |  |  | 04/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1154176 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported that the console (aic) had a "self-check" error which prompted the aic to be replaced. no harm or injury noted to the patient. | [5 empty] | The investigation for the console (aic) self check error has been completed. data logs were reviewed finding the console rebooted automatically due to powermod_1 communication issue. during the reboot, the console displayed a 'system self check failed' screen because of the persistent powermod_1 communication issue. this confirms the reported failure mode. this console was returned or evaluation. confirmed the pbm corrosion on both the top and bottom sides. the root cause for the system self check failure was pbm corrosion.
10574 | 1220648-2025-27213 | 21762888 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21762888&pc=OZD | 10/26/2024 | 09/19/2025 |  | 04/03/2025 |  |  | 10/26/2024 | 04/03/2025 |  |  | 02/28/2025 |  | 04/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024337612 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | During patient support, the pump generated false 'impella in aorta' alarms. the physician confirmed proper positioning via echocardiography. | The complainant reported a patient was on impella cp support. while on support, the patient went into ventricular tachycardia overnight with multiple defibrillations. the impella cp remained in place for left ventricle unloading and the procedural outcome was the patient survived the surgery. |  |  |  |  | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. no product was returned for evaluation. insufficient data logs from time of "impella in aorta" alarm were provided. the root cause of the optical signal issue cannot be definitively determined as insufficient data logs were provided. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated with additional information from the initial manufacturer device report number 1220648-2025-27213. g1 reporting contact email revised on manufacturer device report in accordance with updated procedures from the initial manufacturer device report number 1220648-2025-27213. h6 medical device problem code revised to include 2917 device sensing problem from the initial manufacturer device report number 1220648-2025-27213. | The investigation has been completed. no product was returned. as the necessary information was not provided, the root cause of the ventricular tachycardia was not determined.
10575 | 1220648-2025-27214 | 21763359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21763359&pc=OZD | 02/13/2025 | 04/03/2025 |  | 04/03/2025 |  |  | 02/13/2025 | 04/03/2025 |  |  |  |  | 04/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025589966 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Material Split, Cut or Torn | [6 empty] | Perforation | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had placed an impella 5.5 to support a patient admitted in with cardiomyopathy. the pump was placed and heart transplant decisions were being made. the pump supported and on day two, there was an access site/muscle bleed that prompted surgical intervention and red blood cells were given. the pump additionally was damaged at the sterile sleeve and was remedied by use of tegaderm. the pump remained on for over 42 days. however, the patient was put on comfort care and expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation has been completed. no product was returned. the root cause of the vascular damage issue was not determined due to insufficient clinical information and since the product was not returned. the root cause of the product damage (sterile sleeve) issue was not determined since the product was not returned and due to insufficient clinical information. as noted in the impella instructions for use: impella 5.5 with smartassist for use during section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
10576 | 1220648-2025-27215 | 21763373 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21763373&pc=OZD | 03/11/2025 | 04/03/2025 |  | 04/03/2025 |  |  | 03/11/2025 | 04/03/2025 |  |  |  |  | 04/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1107436 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that during cath lab training, the impella automated impella controller presented controller error alarms. no patient was involved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10577 | 1220648-2025-27216 | 21763589 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21763589&pc=OZD | 03/11/2025 | 04/28/2025 |  | 04/03/2025 |  |  | 03/11/2025 | 04/03/2025 |  |  |  |  | 04/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025594264 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  |  | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support, there was suctioning with p level dropped to p-2. albumin and blood products were given. the doctor believed the pump was possibly flipped when placed at an outside facility as the patient was transferred in the late evening. the staff was told the pump was repositioned successfully. the pump was explanted and the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of low pump flow has been completed. clinical states that pump was placed at a different facility with suction alarm and questionable positioning but were deemed okay. after transferring, low flows were still persistent of 0.6 - 0.8 l/min at p-2 but echo confirmed position was correct. blood was given but no change was observed in flows. pump was not returned. data logs were not returned as well. therefore, the root cause of the low flow issue was not determined since product and data logs were not returned for investigation. clinical information provided is not sufficient to investigate it further.
10578 | 1220648-2025-27217 | 21763666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21763666&pc=OZD | 03/11/2025 | 05/15/2025 |  | 04/03/2025 |  |  | 03/11/2025 | 04/03/2025 |  |  |  |  | 04/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025600602 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 31 YR |  | Male |  |  | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. there was low flow and suction alarm. the impella was removed. no patient harm was reported. | [5 empty] | The investigation of low pump flow has been completed the data logs confirmed low flow after pump started and remained to the rest of the case. no motor current spikes were observed without any p-level changes. biomaterial was observed on the pump, removed and sent in for analysis. visual inspection confirmed that was biomaterial. no other issue was found on the rest of the pump. the root cause of low flow was most likely due to biomaterial ingestion. g1 reporting contact address and country were inadvertently omitted on the initial manufacturer device report number 1220648-2025-27217. g1 manufacturing site name, address, country, and fax number were erroneously entered on the initial manufacturer device report number 1220648-2025-27217. h6 type of investigation code 4114 was erroneously entered on the initial manufacturer device report number 1220648-2025-27217. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
10579 | 1220648-2025-27218 | 21763886 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21763886&pc=OZD | 03/11/2025 | 04/29/2025 |  | 04/03/2025 |  |  | 03/11/2025 | 04/03/2025 |  |  |  |  | 04/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026613053 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 18 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Unspecified Vascular Problem | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that multiple complications were encountered during impella 5.5 support. during implant, it was noted that the patient's artery was friable and required a repair stitch. the combination of the stitch with the patient's existing poor peripheral flow resulted in the impella catheter being occlusive. pulses were lost in the radial and ulner arteries. the pump was removed and a right forearm fasciotomy was performed to restore flow down the arm. an impella cp pump was placed in the femoral artery in conjunction with a distal perfusion catheter for ongoing support. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of vascular damage - peripheral vessel and limb ischemia has been completed. impella 5.5 was returned cut at catheter. no red handle or patient cable side of the pump were returned for evaluation. vascular damage - peripheral vessel: clinical details noted that vascular complications occurred when sewing graft onto artery and artery was friable. it was not noted if friability of artery was patient condition related or as a result of the graft anastomosis. a repair stich was added on the 10mm graft. the root cause of the vascular damage count not be definitively determined as limited clinical details were provided. limb ischemia - reversible: clinical details noted that the patient lost radial pulse after pump was inserted and also was experiencing poor peripheral flows due to shock state. pulse was returned after pump was removed. limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: ¿ patients pre-morbid condition and peri-procedural condition ¿ any concomitant medication ¿ vascular size or variations thereof ¿ detailed description of the symptoms experienced by the patient ¿ physical examination findings, including pulse assessment and visual examination of the affected limb ¿ diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography ¿ laboratory test results, including blood tests for markers of inflammation or clotting disorders ¿ any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension ¿ patient's response to previous treatments or interventions for vascular conditions with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. h6 component code 3038 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27218. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
10580 | 1220648-2025-27219 | 21766510 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21766510&pc=OZD | 03/06/2025 | 04/03/2025 |  | 04/03/2025 |  |  | 03/11/2025 | 04/03/2025 |  |  | 04/30/2025 |  | 04/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024330676 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 57 YR |  | Male |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Sepsis | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Information received via long-term outcome and quality indicator (loqi) impella registry regarding a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. he was admitted while awaiting advanced therapy decisions, such as left ventricular assist device (lvad) or not. it was reported that the patient endured hemolysis with "possibly related" to the use of the pump. per loqi: hemoglobin dropped on (b)(6) 2025 from 7.6 from 8.6. plasma hgb changed. patient was diagnosed with hemolytic anemia on (b)(6) to impella. 2 units of rbcs total were given during admission. patient continued to show hemolysis until death per notes and pi. plama free hgb nadir was 7.5 (on (b)(6) 2025) peak 67.1 (on (b)(6) 25); baseline unknown for ldh and plasma hgb. peak ldh 759 (on (b)(6) 2025) total bil peak 5.4 (on (b)(6) 2025). additionally, the patient endured sepsis with "possibly related" to the pump use. loqi stated: patient had b/c on (b)(6) 2025 +mrse. bc on (b)(6) 2025 gpc clusters. patient's cause of infection, per id, likely related to impella. patient treated with antibiotics for bacteremia. | [5 empty] | The impella pump was not returned for investigation. should new information be received, a supplemental manufacturer device report will be filed. b2 date of death is unknown.
10581 | 1220648-2025-27220 | 21766560 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21766560&pc=OZD | 03/11/2025 | 04/24/2025 |  | 04/03/2025 |  |  | 03/11/2025 | 04/03/2025 |  |  |  |  | 04/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1177189 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an automated impella controller experienced a system self-check failure. there was no patient involved. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - system self-check error has been completed. data analysis: the console logs showed a communication issue with the pbm (power battery manager) during the initial bootup. due to the communication issue, the aic console rebooted automatically (as designed) to attempt another power on. after rebooting, the console displayed the "system self check failed" screen, and the console was then forcefully shut down. device analysis: the reported failure mode was reproduced during testing. upon bootup, the console rebooted automatically and showed the "system self check failure" screen. visual inspection confirmed the pbm contamination which has impacted a neighboring rs232 communication channel. root cause: the root cause of the ¿system self check failure¿ was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps. d2 common device name revised on manufacturer device report 1220648-2025-27220 in accordance with updated procedures. g1 manufacturing site facility name was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27220. g2 report source selections were inadvertently omitted on the initial manufacturer device report number 1220648-2025-27220.
10582 | 1220648-2025-27221 | 21766590 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21766590&pc=OZD | 03/11/2025 | 04/03/2025 |  | 04/03/2025 |  |  | 03/11/2025 | 04/03/2025 |  |  |  |  | 04/03/2025 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1197229 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a console observed to have a battery failure. the console was removed from use. there was no patient involved. | [5 empty] | The investigation into the reported issue has been completed. console logs were consistent with what was reported in the complaint. the reported battery failure issue was able to be reproduced. results of running batttest revealed that cell 3 of battery 1 had a voltage that was significantly less than the voltages of the rest of the cell stacks in the battery. isolative testing was done by swapping battery 1 with a known good one which resolved the alarm. the root cause of the battery failure was determined to be a cell imbalance in battery 1.
10583 | 1220648-2025-27222 | 21766629 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21766629&pc=OZD | 03/11/2025 | 04/03/2025 |  | 04/03/2025 |  |  | 03/11/2025 | 04/03/2025 |  |  |  |  | 04/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025606274 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient undergoing aortic valvuloplasty was implanted with an impella cp device for mechanical circulatory support. it was reported that high purge pressure/low purge flows were encountered during impella cp support. the purge cassette was replaced during troubleshooting, but the issue persisted. it was noted that the purge was running without sodium bicarbonate at the time of the noted issue. a tissue plasminogen activator was ordered to help manage the condition. support with the implanted pump continued uninterrupted. two days later, care was withdrawn and the patient passed away. there is not enough evidence to disassociate the use of the impella with the patient's passing. | [5 empty] | The impella pump was not returned for investigation. should new information be received, a supplemental manufacturer device report will be filed.
10584 | 1220648-2025-27223 | 21766655 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21766655&pc=OZD | 03/11/2025 | 04/03/2025 |  | 04/03/2025 |  |  | 03/11/2025 | 04/03/2025 |  |  |  |  | 04/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026624566 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that the impella cp was removed dure to suspected clot ingestion. the patient was taken to the operating room to remove the pump and close the access site on the left femoral artery. the pump was not replaced, the patient is also on veno-veno-arterial extracorporeal membrane oxygenation with a positive pressure ventilation. a clot was visible in the outlet cage once removed from the body. the pump did stop one time but it was restarted immediately and did not stop again. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
10585 | 1220648-2025-27224 | 21766780 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21766780&pc=OZD | 03/11/2025 | 05/14/2025 |  | 04/03/2025 |  |  | 03/11/2025 | 04/03/2025 |  |  |  |  | 04/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025481990 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an implanted impella cp pump began to experience low pump flow coinciding with suction alarms during patient support. biomaterial was noted during troubleshooting, but the left ventricle remained clear of any notable thrombus. repositioning was unable to resolve the issue and the decision was made to replace the pump. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of low pump flow and thrombus has been completed. data logs were not recovered. pump was returned and inspection revealed biomaterial wrapped around the impellor blade. the cause of the low pump flow was most likely biomaterial due to biomaterial found wrapped around the returned impellor blade. as the necessary information was not provided, the root cause of the thrombus was not determined. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
10586 | 1220648-2025-27225 | 21766786 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21766786&pc=OZD | 03/11/2025 | 04/03/2025 |  | 04/03/2025 |  |  | 03/11/2025 | 04/03/2025 |  |  |  |  | 04/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025554988 |  | 0048-0045 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. impella inserted via left femoral artery prior to percutaneous coronary intervention (pci) with flows on auto of 3.8l throughout case. patient was given bolus twice at beginning of case due to small size of left ventricle and only required small dose of neo during prolonged hypotension. the patient tolerated impella wean at the conclusion of pci and removal of the sheath was slow and with caution due to patient's inability to receive blood products (jehovah witness). hemostasis was achieved with perclose x2 and light manual pressure for 10 minutes. the patient was feeling well post pci. on 11-mar-2025, a notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry indicating that the patient experienced ventricular arrhythmia with a "remote (unlikely)" relationship to the impella device used. the patient expired on (b)(6) 2024. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10587 | 1220648-2025-27226 | 21766794 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21766794&pc=OZD | 03/12/2025 | 05/15/2025 |  | 04/03/2025 |  |  | 03/12/2025 | 04/03/2025 |  |  |  |  | 04/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2026620372 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 85 YR |  | Female |  |  |  | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 85 year old female patient that had an impella cp pump implanted that experienced persistent suction at p2 after the patient required cpr following guideline catheter intubation of the left main. the pump was removed in response to the noted issue. the patient achieved return of systemic circulation. | [5 empty] | The investigation of low pump flow has been completed. returned complaint cp pump visually inspected. no biomaterial observed. no cannula kink or broken blade found. dried dextrose washed for better investigation first with no observation. impella cp connected to the console to reproduce reported low pump flow issue and flow was with specified limit at p-9. no low flow can be reproduced and no other issue observed. data logs not available to review. clinical states suction alarms. no biomaterial observed on explant. blood volume not given. pump not repositioned. low pump flow: the cause of the low pump flow cannot be determined due to lack of clinical details, no data logs available, and low flow not reproduced in investigation lab. g1 reporting contact address and country were inadvertently omitted on the initial manufacturer device report number. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
10588 | 1220648-2025-27227 | 21766844 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21766844&pc=OZD | 03/11/2025 | 04/03/2025 |  | 04/03/2025 |  |  | 03/11/2025 | 04/03/2025 |  |  |  |  | 04/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026633330 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the team did not peel away the sheath due to underlying anemia and red blood cell (rbc) infusion. the pump/sheath limb became ischemic as pulses were lost and limb color changed. the team explanted and escalated to the 5.5 pump. also, they performed thrombectomy to retrieve the thrombus from the vessel. the patient was reported to be in stable condition. | [5 empty] | The impella pump was not returned for investigation. should new information be received, a supplemental manufacturer device report will be filed. instructions for use as follows: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10589 | 1220648-2025-27228 | 21767094 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21767094&pc=OZD | 03/12/2025 | 05/08/2025 |  | 04/03/2025 |  |  | 03/12/2025 | 04/03/2025 |  |  |  |  | 04/03/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1887775 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the automated impella controller (aic) displayed a controller error until the console was plugged in. the aic was swapped in response to the error message. there was no patient harm. | [5 empty] | The investigation of automated impella controller (aic) ¿ controller error (yellow alarm) has been completed. the controller error is due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed. h6 code 925 was reported incorrectly on manufacturer device report 1220648-2025-27228. new code was added to component code. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10590 | 1220648-2025-27170 | 21750684 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21750684&pc=OZD | 03/09/2025 | 04/02/2025 |  | 04/02/2025 |  |  | 03/10/2025 | 04/02/2025 |  |  |  |  | 04/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025600819 |  | 0048-0044 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on impella cp support and during support, there was a cassette failure that led to a pump failure error (pump stop). the pump was restarted and the cassette was replaced. once the cassette was changed, the error resolved. no harm or significant loss of support occurred. the patient continued on support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock. section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
10591 | 1220648-2025-27171 | 21750716 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21750716&pc=OZD | 03/10/2025 | 04/25/2025 |  | 04/02/2025 |  |  | 03/10/2025 | 04/02/2025 |  |  |  |  | 04/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1678477 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Unintended Application Program Shut Down | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that after support was completed on a patient, the automated impella controller would not shutdown. the impella was disconnected from the console and required a hard shutdown. there was no subsequent harm to the patient. | [5 empty] | It was further reported that care was withdrawn and the patient expired. the patient's outcome was reviewed by abiomed's medical safety officer (mso) and it was determined that use of the device cannot conclusively be associated with the outcome. the event occurred after the patient's demise. no interruption or compromise to patient care during use based on information at hand. the device malfunctioned at the end of support after the patient had already been assessed as expired. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10592 | 1220648-2025-27172 | 21750819 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21750819&pc=OZD | 03/10/2025 | 04/14/2025 |  | 04/02/2025 |  |  | 03/10/2025 | 04/02/2025 |  |  |  |  | 04/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1881872 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | No Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient on impella cp support. it was reported that during support, the automated impella controller (aic) was unable to detect the installed purge cassette. the console was swapped as a result of the purge disc not detected alarm, and a new cassette was used with the existing pump successfully. it was noted that the purge cassette drawer was wet on the initial aic. there was no patient harm. | [5 empty] | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. the logs from the complaint date revealed that the console issued purge disc not detected alarm several times. the console was examined, and a broken purge disc retainer was identified. the purge disc flag was observed to be broken and was the root cause of the reported purge disc not detected alarm issue. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10593 | 1220648-2025-27173 | 21750936 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21750936&pc=OZD | 03/07/2025 | 04/30/2025 |  | 04/02/2025 |  |  | 03/07/2025 | 04/02/2025 |  |  | 07/31/2025 |  | 04/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024387415 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient undergoing a high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. after approximately one hour after arriving to the unit, a medium size hematoma, that was not immediately present post-operation, was noted at groin site. manual pressure was held. a sandbag was placed and monitored. two units of packed red blood cells were given to the patient, and anticoagulation was withheld. doppler pulses remain present. the decision was made to leave on the impella cp at performance level p-5 overnight with the plan to explant the impella the next day given the hematoma did not grow, and hemoglobin and hematocrit remained stable. the impella cp device was explanted the following day as planned. the pump was removed from the artery with no evidence of biomaterial on the cannula or sheath. access site was closed with two perclose. the post image showed good hemostasis. the patient was transferred back to the unit and was noted to be in stable condition. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of access site bleeding is determined to be anticoagulation management because the activated clotting time (act) values during the time of bleeding are reported to higher than the therapeutic range(act 247). | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10594 | 1220648-2025-27174 | 21751023 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21751023&pc=OZD | 03/15/2025 | 04/22/2025 |  | 04/02/2025 |  |  | 03/15/2025 | 04/02/2025 |  |  |  |  | 04/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025600732 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and approximately nine days after start of the support, the distal part of the sleeve was noted to be torn. the sleeve was cleaned with non-alcohol prep and secured with tegaderm to enclose the opening. the patient continued on support until the impella cp device was successfully weaned three days later and explanted as planned. there was no harm to the patient as a result of the event. | [5 empty] | The investigation of product damage -sleeve has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10595 | 1220648-2025-27175 | 21751102 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21751102&pc=OZD | 03/06/2025 | 05/14/2025 |  | 04/02/2025 |  |  | 03/06/2025 | 04/02/2025 |  |  |  |  | 04/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025579356 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | Y | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | Yes | Malposition of Device | Deformation Due to Compressive Stress | [5 empty] | Intraoperative Cardiac Valve Injury | [12 empty] | Death | Life Threatening | Required Intervention | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during repositioning of the impella cp device, the cannula kinked, and the pump could no longer be adjusted. the patient was taken to the catheterization lab where it was discovered that the inlet was stuck in the mitral valve. attempts to straighten the impella pump with a lasoo catheter were unsuccessful and subsequently pulling on the pump resulted in a papillary muscle injury. following a cardiac computed tomography, the patient was deemed to be brain dead, expired on support. the impella cp device remained in the aorta and was removed post-mortem. | [5 empty] | The investigation of product damage (cannula kink), positioning issues, and ventricle perforation has been completed. the pump was returned for investigation. the returned pump had significant cannula kinks and a pigtail kink. no other physical damage was noted on the returned product. the pump was able to run in a bucket of water. a data log review was not required for this case run. according to the given clinical details, the impella was repositioned for unknown reason, and it was pushed in deep in left ventricle (lv), it caused the cannula to kink. the pump's inlet valve was stuck at the mitral valve, and pulling the pump caused an injury to the papillary muscle. the cause of the cannula kink issue was use-related, since the cannula was pushed during repositioning. the cause of the positioning issue was not determined since sufficient clinical information was not provided. the cause of the ventricle perforation issue was most likely improper positioning of the device, based on given clinical information. d9 date device returned to manufacturer added. | Patient-specific information is unavailable at the time of this mdr writing; therefore, respective fields reflect "unknown" or left blank accordingly. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10596 | 1220648-2025-27176 | 21751151 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21751151&pc=PYX | 03/08/2025 | 09/26/2025 |  | 04/02/2025 |  |  | 03/08/2025 | 04/02/2025 |  |  |  |  | 04/02/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025567894 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced persistent anemia with low haptoglobin which raised concern for possible device related hemolysis. the patient's urine was clear yellow. it was noted that there are many other possible reasons for the issues experienced issues. blood products were administered and the complication was successfully managed. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. since administering volume resolving the complication and data logs showed suction alarms leading up to and on the date of reported hemolysis, the root cause of the hemolysis was determined to most likely be patient condition related. after further clinical review the impella cp pump was erroneously entered on on the initial manufacturer device report number 1220648-2025-27176. the correct pump that correlates with the reported failure mode date is the impella rp flex. see following corrections: d1 brand name and product code d2 common device name d4 model number, catalog number, serial number, lot number, expiration date, and udi number. d6a and d6b g4 pma/510(k)number h4 device manufacture date h10 instructions for use were erroneously entered. | H6: added optical signal issue as part of a retrospective review of optical signal issues in accordance with fda recommendation.
10597 | 1220648-2025-27177 | 21751257 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21751257&pc=OZD | 03/08/2025 | 04/24/2025 |  | 04/02/2025 |  |  | 03/08/2025 | 04/02/2025 |  |  |  |  | 04/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025496344 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Hispanic |  | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient developed hemolysis during support. the pump was found to be in the mitral valve at the time of the noted issue. attempts to reposition were unsuccessful and the pump was inadvertently pulled back across the valve. the pump was replaced in response to the noted issues. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of positioning issues and hemolysis has been completed. the complaint device not returned for analyzing. no data log available to review. based of the provided clinical details: the cause of the hemolysis issue most likely was improper positioning related since the pump was found in the mitral valve at the time of the noted hemolysis issue. the cause of the positioning issue most likely was use related since the attempts to reposition were unsuccessful and the pump was inadvertently pulled back across the valve.
10598 | 1220648-2025-27178 | 21751291 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21751291&pc=OZD | 03/07/2025 | 04/02/2025 |  | 04/02/2025 |  |  | 03/07/2025 | 04/02/2025 |  |  |  |  | 04/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026628363 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | American Indian Or Alaskan Native | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Bradycardia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient undergoing high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support. the support was successful for under three hours, and then the impella cp device was explanted. at the time of explant there was an oozing at the access site, which was treated by contralateral leg ballooning and eventual vascular surgery. the site achieved hemostasis, but in later hours the patient experience bradycardia and expired. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
10599 | 1220648-2025-27179 | 21751376 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21751376&pc=OZD | 03/09/2025 | 04/02/2025 |  | 04/02/2025 |  |  | 03/09/2025 | 04/02/2025 |  |  |  |  | 04/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026613719 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Death | Required Intervention |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was noted that patient was already in resuscitation before impella implantation. after successfully implantation and performed percutaneous coronary intervention and upon transfer of the patient to the intensive care unit, the patient had to be resuscitated again for approximately one hour. due to a hematoma in the groin/puncture site, the pump was removed. it was noted that the patient was on a high dose of heparin and had multiple organ failure and this was the suspected cause of the hematoma. survival outcome at explant noted that the patient expired a few hours after explant and cause of death was suspected as multiple organ failure. however, there is not enough information to exclude the impella device as an associated factor in the patient's demise.. no further information was provided. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10600 | 1220648-2025-27180 | 21751512 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21751512&pc=OZD | 03/09/2025 | 04/28/2025 |  | 04/02/2025 |  |  | 03/09/2025 | 04/02/2025 |  |  |  |  | 04/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026619261 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump placed to support a patient admitted in with coronary artery disease and heart failure. the patient was placed on the impella cp for the cardiogenic shock/coding. the pump was placed but there was limb ischemia due to the retained 14 french sheath. the impella cp remained on despite the ischemia as they plan for veno-arterial extracorporeal membrane oxygenation, left ventricular assist device, or transplant. four days later, the patient's urine was red colored with a cyanokit given a few days ago. the labs were trending up and platelets were trending down. the patient was heparin-induced thrombocytopenia positive. the team gave red blood cells and after two more days the pump was explanted and escalated to an impella 5.5 pump. | [5 empty] | The investigation of limb ischemia and hemolysis has been completed. hemolysis - clinical reported observing red/brown urine (b)(6) 2025, a day after the pump was implanted. red urine observed on (b)(6) 2025 and pump was repositioned under echo. red urine still observed (b)(6) 2025 and imaging confirmed good pump positioning and pump is not contacting the mitral apparatus. the pump and purge cassette were returned and biomaterial was observed on the impeller. after running the pump, there was biomaterial around the impeller. on (b)(6) 2025, data logs show decreasing flow and motor current at p-7. the pump was briefly turned down to p-2 then back up to p-7. flow and motor current continued to decrease and were followed by multiple suction alarms. no motor current spike observed. the root cause of the hemolysis was most likely patient condition related since data logs show motor current trending down when the p-level is turned up. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. h3 should have reflected the device was returned on the initial manufacturer device report number 1220648-2025-27180. | The pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
10601 | 1220648-2025-27181 | 21751593 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21751593&pc=PYX | 03/08/2025 | 04/25/2025 |  | 04/02/2025 |  |  | 03/08/2025 | 04/02/2025 |  |  |  |  | 04/02/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025575363 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | Pumping Stopped | Difficult to Remove |  |  |  |  | Hemolysis | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced hemolysis. the patient was also on dialysis with no urine output. the cause of the hemolysis was unknown however, it may have contributed to the explant of the impella. additionally, the picc line/central line became entrained in the inlet. the impella was removed, and picc line was pulled out with the impella. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of positioning issues and hemolysis has been completed. the impella device was not returned for evaluation. clinical reports that the patient¿s picc line was entrained in the inlet of the impella and pulled out along with the impella. the cause of the removal issue was most likely an interaction with another device as the picc line was entrained in the inlet of the rp flex during explant and was removed along with the flex. fm was changed from positioning issues to removal issues as the pump was planned to be explanted when the picc line issue occurred. product was not returned. there is a field image from explant of the picc line in the flex inlet. it was stated that this was a recreation of how the picc was found in the flex upon explant and is accidentally backwards in the recreation. failure mode temporary pump stop was added per review of the data logs. the cause of the hemolysis was not determined since product was not returned and lack of clinical information. the cause of the temporary pump stop was not determined since product was not returned and lack of clinical information.
10602 | 1220648-2025-27182 | 21751681 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21751681&pc=OZD | 03/09/2025 | 04/02/2025 |  | 04/02/2025 |  |  | 03/09/2025 | 04/02/2025 |  |  |  |  | 04/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026627626 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 26 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 26-year-old male patient was on impella cp support and during support, the patient had bleeding at the access site. heparin was withheld and support continued. additionally, the peel away sheath was left in place. the patient's artery was noted to be damaged due to the peel away. vascular cutdown was performed and blood products were given. the patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10603 | 1220648-2025-27183 | 21751704 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21751704&pc=OZD | 03/10/2025 | 04/21/2025 |  | 04/02/2025 |  |  | 03/10/2025 | 04/02/2025 |  |  |  |  | 04/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025559458 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | Increase in Pressure | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella 5.5 pump supporting a patient admitted in with cardiomyopathy. the pump was supporting when after approximately a month of support, the purge flow decreased prompted the team to troubleshoot with tissue plasminogen activator and to change from sodium bicarbonate to heparin in the purge. the pump remained on for support while awaiting a transplant. approximately two weeks later, there was a purge leak at the pump driveline connection. the pump still remains on for support. | [5 empty] | The investigation of high purge pressure and purge leak ¿ pump has been completed. the cause of the high purge pressure was not determined since product was not returned, data logs or the days in question were not recovered. and lack of clinical information. the cause of the purge leak - pump was a leak from the sidearm but the exact location of the leak was unknown. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
10604 | 1220648-2025-27184 | 21751816 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21751816&pc=OZD | 03/10/2025 | 04/02/2025 |  | 04/02/2025 |  |  | 03/10/2025 | 04/02/2025 |  |  |  |  | 04/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-EU |  |  | ABIOMED, INC. |  |  | US |  | 2026617975 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 70 year old male patient on impella cp support and during support, there was bleeding at the access site. surgical sutures were applied. patient remained on support. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10605 | 1220648-2025-27185 | 21751914 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21751914&pc=OZD | 03/06/2025 | 04/02/2025 |  | 04/02/2025 |  |  | 03/06/2025 | 04/02/2025 |  |  |  |  | 04/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025523071 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was diagnosed with congestive heart failure exacerbation (known ejection fraction of 20%) and was admitted for diuresis and further workup. further cardiac workup revealed that a redo cabg was necessary, and plans were made and executed upon for the redo cabg with impella cp placement pre-operatively. approximately 25 hours later, the patient developed limb ischemia. the following day after the start of impella mechanical circulatory support (mcs), bilateral lower extremities were cool to touch. as the morning progressed the right lower extremity (impella insertion site right femoral) became more mottled; therefore, bilateral lower extremity dopplers was completed and preliminary, the doppler report read thrombus from left common femoral artery to popliteal artery with no flow and no flow from the right superficial femoral artery down. of note, the patient has history of peripheral artery disease with previous occlusion. additionally, the patient had increased bleeding from chest tubes and required numerous blood transfusions: 8 units of red blood cells, 4 units of fresh frozen plasma, 2 units of cryoprecipitate, 1 pack of platelets, and 2mg factor 7. a stat vascular surgery consult was ordered and initiation of systemic heparin infusion. the physician decided to explant the device sooner than initially planned to attempt to restore perfusion to the right lower extremity. the impella was explanted and perclose was placed to obtain hemostasis. additionally, staff held manual pressure for 30 minutes. stat computed tomography angiogram of the abdomen and pelvis was obtained and noted the following: occlusion of bilateral lower extremity arteries starting at the common iliac on the left and the common femoral artery on the right. an emergent bilateral lower extremities/aortic thrombectomy was completed. the patient reportedly survived the explant; however, the patient was noted to be in critical condition following the event. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
10606 | 1220648-2025-27186 | 21751924 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21751924&pc=OZD | 03/07/2025 | 04/02/2025 |  | 04/02/2025 |  |  | 03/07/2025 | 04/02/2025 |  |  |  |  | 04/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026614173 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Material Split, Cut or Torn | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported approximately 18 days after start of the impella mechanical circulatory support to a patient with cardiomyopathy the sterile sleeve of the implanted impella 5.5 pump tore at the distal end. the tear was covered in tegaderm and support continued uninterrupted. there was no patient harm reported as a result of the event. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10607 | 1220648-2025-27187 | 21751939 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21751939&pc=OZD | 03/07/2025 | 05/06/2025 |  | 04/02/2025 |  |  | 03/07/2025 | 04/02/2025 |  |  | 07/31/2025 |  | 04/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025437184 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Cardiac Arrest | Hematoma | [11 empty] | Life Threatening | Death | Required Intervention | [22 empty] | Invalid/Incomplete | Inspection |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock implanted with an impella cp device for mechanical circulatory support experienced an access site hematoma, positioning issue and low flows requiring intervention. however, the patient would subsequently expire. the patient underwent a catheterization which showed 100% occluded left anterior descending artery, 95% occluded circumflex and right coronary arteries with an ejection fraction of 35%. during the percutaneous coronary intervention, ballooning the right coronary artery, the patient had a seizure, and the decision was made to implant an impella cp device. approximately five minutes after the impella cp device implant, a "large" hematoma developed at the access site. manual pressure was applied; however, the patient's blood pressure began to rapidly drop. low pump flows were then encountered, and repositioning proved to be difficult as the left ventricle repeatedly pushed the device out of the heart. cardiopulmonary resuscitation was subsequently initiated; however, the flows kept dropping and a code was called. ultimately, the patient expired. | [5 empty] | The investigation of positioning issues and low pump flow, has been completed. the impella device was not returned for evaluation. the root cause of the positioning issue most likely was patient condition related since patient had left ventricle (lv) aneurysm and the left ventricle repeatedly pushed the device out of the heart. the cause of the low pump flow issue cannot be determined since complaint device not returned for analysis and insufficient clinical data provided. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27187 as it is unknown if the initial reporter also sent the report to the food and drug administration. h6 code 1884 was erroneously entered on the initial manufacturer device report number 1220648-2025-27187. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: general patient care considerations: "assess access site for bleeding and hematoma." section: entering cardiac output: use of the repositioning sheath and peel-away introducer: "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site." section: potential adverse events: "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury" (including ventricular perforation).
10608 | 1220648-2025-27189 | 21752058 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21752058&pc=OZD | 03/10/2025 | 04/28/2025 |  | 04/02/2025 |  |  | 03/10/2025 | 04/02/2025 |  |  |  |  | 04/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026616053 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that the 57-year-old male patient on impella 5.5 support developed a large hematoma at axillary site with numbness and tingling down arm distal to impella insertion. surgical exploration was performed. the patient bridged to transplant. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27189 as it is unknown if the initial reporter also sent the report to the food and drug administration. h6 health effect - clinical code 1884 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27189.
10609 | 1220648-2025-27188 | 21752059 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21752059&pc=OZD | 03/08/2025 | 04/02/2025 |  | 04/02/2025 |  |  | 03/08/2025 | 04/02/2025 |  |  |  |  | 04/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026632243 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs), and the following day limb ischemia developed. the pump was removed, and vascular surgery was called for a thrombectomy. the physician stated regarding the cardiac issue, the patient had improved "a lot" and no longer need impella mcs. therefore, the explanted device was not replaced. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10610 | 1220648-2025-27190 | 21752067 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21752067&pc=OZD | 03/10/2025 | 04/17/2025 |  | 04/02/2025 |  |  | 03/10/2025 | 04/02/2025 |  |  |  |  | 04/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026622564 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Ischemia | Vascular Dissection | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that upon attempting to advance an impella cp device on a 79 year old female patient, the staff met resistance in the distal aorta. the pump was removed and an angiogram revealed severe stenosis in the distal aorta. the implant was aborted and the introducer was removed. after closing the site, the patient began to complain of pain in their right lower extremity (rle). their foot became pale with no pulses present in the rle. another angiogram was performed and a dissection was discovered with no distal flow. the vessel was wired distal to the occlusion and a balloon angioplasty was performed to the common femoral artery. the operation was successful and flow was restored to the extremity. | [5 empty] | The investigation of delivery issues, vascular damage - peripheral vessel, and limb ischemia has been completed. the pump was not returned for investigation. data review was not required for this case run. as per the clinical notes, delivery was aborted after the physician encountered resistance in the distal aorta. after the removal of a 14f sheath and closure of the site using perclose devices, the patient began to complain of pain in the foot. upon examination, no pulses were present in the rle. a femoral angiogram was performed, which revealed a dissection of the cfa. balloon angioplasty was performed. this intervention successfully restored flow to the distal extremity. the cause of the delivery issue was most likely patient condition related to diseased vessels. the cause of the issue was most likely patient condition related to calcified vessel condition. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. instructions for use as follows: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: warnings & cautions: cautions. ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27190 as it is unknown if the initial reporter also sent the report to the food and drug administration. h6 component code 3038 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27190. h10 instruction for use were inadvertently omitted on the initial manufacturer device report number 1220648-2025-27190. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10611 | 1220648-2025-27191 | 21752093 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21752093&pc=OZD | 03/10/2025 | 04/11/2025 |  | 04/02/2025 |  |  | 03/10/2025 | 04/02/2025 |  |  |  |  | 04/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1154856 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported when the console (aic) was started up, the message "controller error" appeared. the team replaced the aic. there was no noted harm or injury. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - controller failure (red alarm) has been completed. the data analysis confirmed the failure mode reproduced on every boot up after the complaint. the console was returned and the failure mode reproduced. the optical connector was found to be contaminated and after cleaning, 5s improved and the controller error did not return. root cause: the cause of the controller error on boot up was contamination of the optical connector. the contamination caused light to not be reflected back as expected which is how light was lost. g1 reporting contact address and country was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27191.
10612 | 1220648-2025-27192 | 21752227 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21752227&pc=OZD | 03/10/2025 | 10/24/2025 |  | 04/02/2025 |  |  | 03/10/2025 | 04/02/2025 |  |  |  |  | 04/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026613446 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient had an attempted impella cp implant and during the implant, the abiomed 0.018" guidewire buckled. the doctor elected to abort the implant and insert a new pump. there was no patient harm. | The complainant reported the patient had an attempted impella cp implant and during the implant, the abiomed 0.018" guidewire buckled. the doctor elected to abort the implant and insert a new pump. there was no patient harm. |  |  |  |  | The investigation of delivery issues and product damage (guidewire kink) has been completed. the guidewire was not returned. the purge cassette and pump were returned and the red lumen was still inside the pump. there was dried dextrose in the purge gap. no damages or abnormalities noted. logs are not applicable. during pump implant, the wire buckled and bent, so the pump was removed and replaced. no vessel conditions were noted. delivery issues: the root cause of the delivery issues was not determined due to insufficient clinical details provided. product damage guidewire: the root cause of the guidewire damage was not determined since the product was not returned and insufficient clinical details were provided. d9 updated. h3 updated. d10 concomitant medical products and therapy dates updated. | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
10613 | 1220648-2025-27193 | 21752958 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21752958&pc=OZD | 03/09/2025 | 05/09/2025 |  | 04/02/2025 |  |  | 03/09/2025 | 04/02/2025 |  |  |  |  | 04/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026624471 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | [6 empty] | 18 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. a suction alarm with flow decline was noted. fluoro showed cannula kinking near the outlet, attributed to a tight aortic arch. repositioning attempts were unsuccessful, prompting replacement of the pump. no patient harm was reported. | [5 empty] | The investigation of low pump flow and positioning issues has been completed. the clinical states suction alarms and low flow while on support. fluoro showed cannula kinking near the outlet which physician attributed to a tight aortic arch. pump stated to be in good position, but unsuccessful attempts to reposition and straight the cannula so pump was replaced. data review: data logs confirmed low pump flow when pump started and remained to the rest of the case. product was not returned for review. based on clinical description and data analysis, the root cause of low flow was most likely kinked cannula since cannula kink was found after low flow occurred. the root cause of positioning issue was most likely patient condition related since the patient had tight aortic arch. the root cause of cannula kink was most likely patient condition related since the patient had tight aortic arch. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10614 | 1220648-2025-27194 | 21756366 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21756366&pc=OZD | 03/03/2025 | 05/02/2025 |  | 04/02/2025 |  |  | 03/03/2025 | 04/02/2025 |  |  |  |  | 04/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025525096 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, the patient began to experience runs of ventricular tachycardia prompting defibrillation. the patient subsequently raised their leg and the impella came out of the left ventricle. the pump was fully removed and rewired into the left ventricle. upon the second placement, the staff noted that the pigtail was unnaturally curled. repositioning was attempted, but the pigtail was believed to be stuck under the chordae tendineae. support proceeded despite the noted positioning issue. it was additionally reported that the groin site was oozing/bleeding. the percloses were tightened and blood products were given in response to the condition. in addition, the dressing was opened and the site was adjusted due to poor angle matching. no further intervention was performed. | [5 empty] | The investigation of positioning issues, vf/vt/af/at, and access site bleeding has been completed. the impella device was not returned for evaluation. positioning issue: clinical states that during initial pump implant, patient had a vt episode and while defibration, the pump was accidently pulled back into aorta when the patient moved the leg. pump was not returned. as per the clinical information, the root cause of the positioning issue was patient movement that happened at the time of defibration process which led to pump being pulled into aorta. access site bleeding: clinical states that after 7 hours of pump implant, the site was bleeding and the preclose was tightened around the sheath and the suture hub was pulled towards the skin in attempt to decrease oozing. pump was not returned. as per the clinical details, the reason for bleeding is unclear, activated clotting time (act) values are unknown. therefore, the root cause of the access site bleeding issue was not determined since the clinical information insufficient and the reason for the bleeding is unclear. as the necessary information was not provided, the root cause of the vt/vf was not determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27194 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10615 | 1220648-2025-27195 | 21757919 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21757919&pc=OZD | 03/09/2025 | 04/22/2025 |  | 04/02/2025 |  |  | 03/09/2025 | 04/02/2025 |  |  |  |  | 04/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026613599 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | Material Twisted/Bent |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was attempted to be implanted with an impella device for mechanical circulatory support. the pump could not be delivered and kinked. the team removed the malfunctioning pump and replaced it with another 5.5 pump. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.\ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controllerrp flex with smartassist system with automated impella controller section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | H.6 medical device problem codes and component codes were added due to evaluation findings of both the catheter and the cannula having kinks. the investigation of delivery issues and product damage (cannula kink) has been completed. the returned complaint device visually inspected. catheter kink was observed near the outlet and cannula kink was observed at the bending area. kinks may have occurred during the same time with insertion attempts. product damage (cannula kink) was added since both catheter and cannula kinks were found during evaluation. the cause of the delivery issue was due to patient tortuosity vessel condition unable to advance impella around axillary into the ascending and led to kinks in the device. the cause of the product damage (cannula kink) was cather kink due to patient tortuosity vessel condition. the cause of the product damage (cannula kink) was cannula kink due to patient tortuosity vessel condition. b.7 information was added as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-27195. e.1 added phone number and fax number as they were inadvertently omitted from the initial manufacturer device report number 1220648-2025-27195. h.6 codes 4118 and 3221 were reported incorrectly on the initial manufacturer device report number 1220648-2025-27195.
10616 | 1220648-2025-27196 | 21757994 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21757994&pc=OZD | 03/10/2025 | 04/22/2025 |  | 04/02/2025 |  |  | 03/10/2025 | 04/02/2025 |  |  |  |  | 04/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025599753 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | Yes | Yes | Positioning Failure | [6 empty] | Tachycardia | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The case was further reviewed and it was discovered that the pump had been repositioned multiple times during the support and the pump positioning may have contributed to the ventricular tachycardia (vt) that occurred and was treated on the 2nd pump. there was no vt on this pump. another pump was placed to support after weaning and explanting the pump that had supported for 24 days. | The user facility reported that the 64 year old female patient with a 5.5 pump placed to support for cardiogenic shock. after a few weeks on support, the pump was removed and replaced due to poor pump positioning. the pump had been repositioned multiple times during the support and the pump positioning may have contributed to the ventricular tachycardia (vt) that occurred and was treated. another pump was placed to support after weaning and explanting the pump that had supported for 24 days. the patient survived the procedure. |  |  |  |  | The investigation of positioning issues and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27196 as it is unknown if the initial reporter also sent the report to the food and drug administration. h6 health effect - impact code 4641 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27196. | B1. revised on the initial manufacturer device report number 1220648-2025-27196 based on corrected information. b5. revised on the initial manufacturer device report number 1220648-2025-27196 based on corrected information. h1. type of reportable event revised on the initial manufacturer device report number 1220648-2025-27196 based on corrected information. h6. health effect - clinical code revised on the initial manufacturer device report number 1220648-2025-27196 based on corrected information. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed.
10617 | 1220648-2025-27197 | 21758019 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21758019&pc=OZD | 03/08/2025 | 04/29/2025 |  | 04/02/2025 |  |  | 03/10/2025 | 04/02/2025 |  |  |  |  | 04/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026612942 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Sepsis | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that a 5.5 pump placed in a 75 year old male patient to support during high risk cardiac procedure. information was received via impact study that noted arrhythmia with a "probable" relationship to the pump had occurred. medical therapy was done for the harm. p level adjustment, and lidocaine and levophed given. weeks later the impact study noted there was fluid overload with possible relationship to the impella. | [5 empty] | The investigation of infection/sepsis and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-27197. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27197 as it is unknown if the initial reporter also sent the report to the food and drug administration. g2 report source; study was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27197. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smartassist for use during cardiogenic shock states the following: section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
10618 | 1220648-2025-27198 | 21758072 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21758072&pc=OZD | 03/11/2025 | 04/14/2025 |  | 04/02/2025 |  |  | 03/11/2025 | 04/02/2025 |  |  |  |  | 04/02/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  |  |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that impella implantation was planned but after starting the console (aic), a controller failure occurred. the aic was successfully exchanged before the procedure. there was no patient harm. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. | The investigation of automated impella controller (aic) - system self-check error has been completed. the case associated with the reported complaint was started on 10th march 2025. log entries indicating "imc-kbd error: 0xa4 0x01 0x00" (imc_ic2659.pdf) appeared, suggesting that the issue occurred during the reading of the chip id (read failure of chip id). · after this, the controller error (loss of comm with keypad or between the kbd and i2c for kbd hw revision 1) - alarm 24 issued. device analysis: ·the console was tested and after performing 100 power cycles, the failure could not be reproduced. no damage was found to the keyboard or the keyboard- 4in1 cable. root cause: the root cause for the intermittent keyboard communication issue could not be determined due to the inability to reproduce. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27198 as it is unknown if the initial reporter also sent the report to the food and drug administration.
10619 | 1220648-2025-27126 | 21737729 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21737729&pc=OZD | 03/05/2025 | 04/01/2025 |  | 04/01/2025 |  |  | 03/05/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025590569 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that upon evaluation post impella cp placement, there was distal limb ischemia to the left arm. the decision was made to place a radial ipsilateral reperfusion circuit with a six french sheath in the left radial. good flow was obtained to the arm. additionally, there was bleeding from the insertion site. gauze was packed around the site and two units of packed red blood cells were administered. bleeding was resolved. it was further reported that the family decided to withdrew care and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | B.2 required intervention received prior to death was inadvertently omitted from the initial manufacturer device report number 1220648-2025-27126.
10620 | 1220648-2025-27127 | 21737823 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21737823&pc=OZD | 02/20/2025 | 04/01/2025 |  | 04/01/2025 |  |  | 03/06/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025587421 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Life Threatening | Required Intervention |  |  |  |  | INTRA-AORTIC BALLOON PUMP | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella cp device as escalation in care from an intra-aortic balloon pump (iabp), pre-operative placement for coronary artery bypass graft (cabg). the cabg was started, and bleeding occurred due to vascular injury at the forward insertion site. volume loss prevented the impella cp device from effectively operating. therefore, conversion to cardiopulmonary bypass was required. it was noted that the injury may have been due to catheter manipulation during iabp or impella cp device insertion. a blood transfusion was performed, but no information was obtained regarding the amount of blood transfused. the patient was noted to be "stable and nimble" following the event. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
10621 | 1220648-2025-27128 | 21738016 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21738016&pc=OZD | 03/06/2025 | 04/14/2025 |  | 04/01/2025 |  |  | 03/06/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1390968 |  | 0042-0010-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Crack | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported an automated impella controller loaner was shipped out for preventative maintenance console to come in for service. the loaner was received damaged with a cracked and glued monitor bezel (housing). the issue was found outside of patient use. | [5 empty] | The investigation of automated impella controller (aic) - damage before therapy has been completed. log analysis was unnecessary for this complaint. the console's cracked and glue repair was confirmed. root cause: the repair that complained about was most likely perform by field service. | The automated impella controller device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
10622 | 1220648-2025-27129 | 21738173 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21738173&pc=OZD | 03/06/2025 | 04/01/2025 |  | 04/01/2025 |  |  | 03/06/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025612223 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Failure to Advance | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | A |  |  |  | Initial submission | [9 empty] | The complainant reported the insertion of an impella 5.5 device in a patient with postcardiotomy cardiogenic shock was unsuccessful with anastomosis bleeding. the patient would ultimately expire but not before implanting an impella cp device that experienced low pump flow. the patient was admitted for a pericardiectomy and during this procedure, the patient arrested. it was suspected that the left anterior descending artery had occluded and the patient was taken for an emergent coronary artery bypass graft surgery. however, the patient would not separate from the bypass. the decision was made to implant an impella 5.5 device. during the implant, the impella 5.5 made it into the graft, past anastomosis but was unable to make the turn into the heart after multiple attempts. multiple blood products were used due to axillary anastomosis bleeding a "great deal." what was infused and the number of units is unknown. the physician felt there was a stricture preventing the impella 5.5 from entering the heart. the impella 5.5 implant was aborted, and an impella cp was called for to attempt to help separate from the patient from the bypass. the impella cp device was implanted without incident. however, it was noted that there were "great flows at the beginning." the blood pressure would increase and decrease with blood products being given. however, despite "aggressive" treatment, the patients blood pressure fell despite volume, and the patient expired on the operating room table. medical safety review of the event noted the impella cp to have "great flows at beginning¿ reasonably suggesting optimal flows were not sustained resulting in an outcome of serious injury. however, the impella 5.5 bleeding from the anastomosis requiring multiple blood products as the onset of serious adverse complications and likely to have contributed to the patient¿s demise. thus, cannot be excluded as a possible contributing factor. | [5 empty] | Medical safety review of the event noted the impella cp to have "great flows at beginning¿ reasonably suggesting optimal flows were not sustained resulting in an outcome of serious injury. however, the impella 5.5 bleeding from the anastomosis requiring multiple blood products as the onset of serious adverse complications and likely to have contributed to the patient¿s demise. thus, cannot be excluded as a possible contributing factor. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ this is one of two devices associated with the event. refer to medical device report for the impella cp serial number (b)(6) with date of event 06-mar-2025 and patient identifier ending in (b)(6).
10623 | 1220648-2025-27130 | 21738217 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21738217&pc=OZD | 04/17/2023 | 04/01/2025 |  | 04/01/2025 |  |  | 03/05/2025 | 04/01/2025 |  |  | 02/28/2025 |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024313260 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 47 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A notification was received via abiomed's long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured ventricular tachycardia with a "possible" relationship to the impella device used: " on (b)(6) 2023: va ecmo decannulation w l cfa and cfv repairs, protek duo w oxygenator inserted c/b vtach in or which was treated w amio". the patient recovered with sequelae. if was further reported that the patient expired while on impella 5.5 support. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)."
10624 | 1220648-2025-27131 | 21738246 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21738246&pc=OZD | 03/07/2025 | 04/24/2025 |  | 04/01/2025 |  |  | 03/07/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1603745 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Hispanic |  | Death | Yes | Yes | Application Program Problem | [6 empty] | Cardiovascular Dysfunction/Insufficiency | [12 empty] | Death | [5 empty] | IMPELLA CP | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainants reported a patient came in with chest pain and cough and was sent to the medical intensive care unit for pneumonia treatment. the patient was hypotensive and had shortness of breath and so the patient was intubated and troponin was increased. the patient was found to have microvascular disease and three stents to the right coronary artery. the patient, decompensated during percutaneous coronary intervention (pci). it was reported that during support, the automated impella controller (aic) had a controller failure error when the pump stopped. the pump stop alarm happened simultaneously as the controller error. the pump was explanted and the aic was removed. outcome at the end of unit support was noted that the patient expired on support. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - controller failure (red alarm) has been completed. the root cause was unable to be determine due to the inability to be reproduced in the device analysis.
10625 | 1220648-2025-27132 | 21738336 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21738336&pc=OZD | 03/07/2025 | 04/22/2025 |  | 04/01/2025 |  |  | 03/07/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026617651 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Female |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. the pump had high purge pressures. no action was required. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation of high purge pressure has been completed. after 24 hours on support, high purge pressure was noted. troubleshooting to resolve the problem was unknown. however, the pump remained implanted. the product was not returned. data logs were not recovered. the cause of the high purge pressure was not determined since the product was not returned, data logs were not recovered, and lack of clinical information. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure, which triggers the ¿purge pressure high¿ alarm message, could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
10626 | 1220648-2025-27133 | 21738414 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21738414&pc=OZD | 03/06/2025 | 05/05/2025 |  | 04/01/2025 |  |  | 03/06/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1483145 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | It was reported during a "quick" inspection the optical bench check score was below 2000 (signal-to-noise ratio out of specification). the issue was found during servicing, outside of patient use. | [5 empty] | The investigation of optical signal issue has been completed. data analysis: upon connecting a test pump, the console displayed 'impella defective ¿ alarm,' event #3, and the signal not reliable (snr) was 1729. on the complaint date, march 6, 2025, event #3 occurred again, and the snr was 1682.1, confirming the reported failure mode . device analysis: the reported failure mode was identified during the service. the low snr is persistent. the bulb power measured 2.65 w, which is supposed to be greater than 2.7 w according to the preventative maintanance procedure. root cause: the root cause of the low snr was a defective lamp assembly. | The automated impella controller was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
10627 | 1220648-2025-27134 | 21738632 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21738632&pc=OZD | 02/28/2025 | 04/01/2025 |  | 04/01/2025 |  |  | 03/06/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025537465 |  | 1000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | [12 empty] | Life Threatening | Required Intervention | Disability | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was implanted with an impella 5.5 device as pre-operative placement for surgery for mechanical circulatory support and experienced a perforation. it was noted when checking for bleeding before de-clamping the aorta, the tip of impella 5.5 device was found to be protruding from the heart. the surgical repair was performed on the perforated site. the patient was noted to be in serious condition following the event. it was further stated the surgical repair was completed on the next day after explant of the impella 5.5 device. however, it is unclear when the perforation occurred; timing is unknown. | [5 empty] | The perforation surgical repair date is capture as the provisional date of event as the actual date is unknown. the impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
10628 | 1220648-2025-27135 | 21738868 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21738868&pc=OZD | 03/06/2025 | 04/17/2025 |  | 04/01/2025 |  |  | 03/06/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1302102 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller experienced a system self-check failure during preparation for patient support. the aic was exchanged. there was no report or indication of harm to the patient. | [5 empty] | The investigation of automated impella controller (aic) - system self-check error has been completed. the console logs showed a communication issue with the pbm (power battery manager) during the initial bootup. due to the communication issue, the aic console rebooted automatically (as designed) to attempt another power on. after rebooting, the console displayed the "system self check failed" screen, and the console was then forcefully shut down. device analysis: the reported failure mode was reproduced during testing. upon bootup, the console rebooted automatically and showed the "system self check failure" screen. visual inspection confirmed the pbm contamination which has impacted a neighboring rs232 communication channel. root cause: the root cause of the ¿system self check failure¿ was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps. | Patient-specific information is unavailable at the time of this mdr writing; therefore, respective fields reflect "unknown" or left blank accordingly. the automated impella controller device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
10629 | 1220648-2025-27136 | 21739218 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21739218&pc=OZD | 03/06/2025 | 04/01/2025 |  | 04/01/2025 |  |  | 03/06/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1606055 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | It was reported that during training, a battery fault alarm occurred on the automated impella controller (aic). troubleshooting was done but unsuccessful. the aic was exchanged to resolve the issue. there was no patient involvement. | [5 empty] | The automated impella controller data logs and console were returned for evaluation and analysis. data analysis demonstrated that logs were consistent with what was reported in the complaint. instances of battery failure (transport connection blown/missing) alarm and battery failure (low voltage) alarm were observed in the imc logs. the console was tested after arriving in field service, and the reported battery failure was able to be reproduced. results of running batttest utility revealed that cell 3 of battery 2 had a voltage of 0mv. isolative testing was done by swapping battery 2 with a known good one which resolved the alarm. a product history review was completed, and no action was required as a result of this review. in conclusion, the root cause of the battery failure was determined to be cell imbalance in battery 2.
10630 | 1220648-2025-27137 | 21739351 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21739351&pc=OZD | 03/06/2025 | 04/24/2025 |  | 04/01/2025 |  |  | 03/06/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025599045 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | Infusion or Flow Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support. ten days into support, a motor current high alarm appeared and the pump stopped. evidence of saturated flow was noted by the staff. the pump was restarted successfully, but was ultimately explanted the following day. upon removal, the initial percloses broke. new percloses were deployed to achieve hemostasis and manual pressure was held. the procedural outcome was the patient survived the surgery. | [5 empty] | The pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of temporary pump stop, saturated flow, and access site bleeding has been completed. the pump was returned for investigation. the data log indicates that the pump stopped on the 9th day of use due to a high motor current alarm. after this, the pump was able to restart with increased purge pressure, and a saturated pump flow was observed. several motor current spikes were noted over the next 28 hours before the pump was explanted. upon returned pump examination, no physical abnormalities were found on the returned product. the pump was then tested in a bucket of water, where a saturated pump flow was again observed. further investigation revealed a large purge gap, which is contributing to bearing wear due to the duration of use exceeding the intended design life of the cp pump. the cause of the temporary pump stop was bearing wear, based on data log review and returned product investigation. the cause of the saturated pump flow was bearing wear since the bearing wear create extra resistance on motor cause the inaccurate/saturated flow based on motor current. the clinical report indicates that the access site bleeding was controlled after applying a new parclose with adjustment and 20 minutes of manual pressure. the cause of the access site bleeding issue was most likely related to the procedure based on the provided clinical details. instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: entering cardiac output. ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention. section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings & cautions: warnings. ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ h10 instructions for use was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27137.
10631 | 1220648-2025-27138 | 21739454 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21739454&pc=OZD | 03/06/2025 | 04/01/2025 |  | 04/01/2025 |  |  | 03/06/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025599136 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | Thrombosis/Thrombus | [11 empty] | Required Intervention | Death |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient had no flow down the right leg past the sheath. peel away was removed and repo sheath positioned. penumbra thrombectomy was performed and flow restored. the physician also placed a repercussion catheter and two units of packed red blood cells were administered. the "patient's" general outcome was noted to be critically ill and grim. the pump supported the patient for two more days and then care was withdrawn and patient subsequently expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smartassist catheter insertion section: axillary insertion of the impella cp with smartassist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
10632 | 1220648-2025-27139 | 21739614 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21739614&pc=OZD | 10/16/2024 | 04/01/2025 |  | 04/01/2025 |  |  | 03/06/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025555802 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient who experienced "acute worsening of chronic anemia" during impella cp support. it was noted that hemolysis may have been due to impella cp. the patient was administered packed red blood cells (prbc) with dual therapy of clopidogrel and rivaroxaban. lasix after prbc transfusion showed good response. the patient outcome was noted as stable and the procedure was successful. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
10633 | 1220648-2025-27140 | 21739678 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21739678&pc=OZD | 03/07/2025 | 10/01/2025 |  | 04/01/2025 |  |  | 03/07/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025589205 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Positioning Failure | Restricted Flow rate | Increase in Pressure | Device Sensing Problem |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that placement signal was not correlating with the arterial line pressures. frequent position unknown alarms appeared as a result of the issue. support continued uninterrupted. | The complainant reported that multiple complications were encountered during impella 5.5 support. as support was raised above p4, suction was encountered. an echocardiogram showed the inflow in close proximity to the papillary muscle. multiple attempts to reposition the device were unsuccessful and the staff had to settle for lowering support to p4. twenty-seven days into support, high purge pressure developed in the system coinciding with a purge blocked alarm. tissue plasminogen activator was run through the purge to resolve the issue. there was no patient harm. |  |  |  |  | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. optical signal issues: device implanted (b)(6) without issue. data log review showed placement signal failure on (b)(6) 2025, 54 days on support. pump was not returned. data log review showed placement signal failure that aligns with sensor wear patterns of drifting down and failing within 24 hours. drift begins (b)(6) 2025 at 1415 when ps is around 25mmhg. then at 2353 on (b)(6) 2025, ps is between -40 and -55 mmhg when it rails to 3000. the cause of the optical signal issue was most likely sensor silicone wear due to data log review showing known sensor wear data fingerprint where placement signal drifts down and fails within 24 hours. a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical signal issue "2917" as part of a retrospective review of optical signal issues in accordance with fda recommendation. | The investigation of positioning issues and high purge pressure has been completed. the pump was not returned. the cause of the positioning issue was not determined since product was not returned and lack of clinical details. high purge pressure: device implanted 13feb without issue. after 38 days on support, purge pressure high alarms occurred. tpa was run in the purge after recommended by csc. the pump was not returned. data log review showed purge pressure build up just before high purge pressure alarms were triggered. after several bag change attempts showed unsuccessful, the purge pressure eventually returned to nominal parameters, the same day tpa was instructed to be used. the cause of the issue was most likely biomaterial due to the purge pressure build up over the course of 2 hours followed by the issue resolving the same day tissue plasminogen activator (tpa) was instructed to be placed in the the purge fluid. low pump flow was added as a failure mode: data log review showed low pump flows of3.6l/min at p-7 (expected flows are 3.9-4.5l/min) with no mc spike occurred but suction alarm 14 was triggered and no observable trend in ps. the cause of the low pump flow was not determined since product was not returned and lack of clinical details. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist. section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
10634 | 1220648-2025-27141 | 21739828 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21739828&pc=OZD | 03/08/2025 | 04/10/2025 |  | 04/01/2025 |  |  | 03/08/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026617718 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support. the patient had pre-existing hemolysis with an extracorporeal membrane oxygenation (ecmo) and after the impella 5.5 was delivered the hemolysis worsened. the pump positioning was assessed and hemodialysis began. the patient developed disseminated vascular coagulation and the impella 5.5 pump was explanted after 46 hours. the patient continued on support with ecmo. | [5 empty] | The investigation of hemolysis has been completed. the patient was on extracorporeal membrane oxygenation for days prior to impella insertion with pre-existing hemolysis. insertion was tough. server hemolysis was noted for 24-36 hours after impella placement. the data logs showed multiple motor current spikes after the pump was started with multiple auto suction response. logs also showed some positioning issues but resolved. no other issue was found on the logs. visual inspection found biomaterial inside the pump housing and wrapped around the impeller. a kink on cannula was found. no other issues were found on the rest of the pump. the root cause of hemolysis was due to biomaterial ingestion based on return product and log analysis. | The pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
10635 | 1220648-2025-27142 | 21740009 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21740009&pc=OZD | 11/27/2024 | 04/28/2025 |  | 04/01/2025 |  |  | 03/06/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571996 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | [12 empty] | Death | Required Intervention |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient who experienced ventricle perforation during impella cp support. it was noted that the patient's course was complicated by pericardial effusion and left ventricle anterior wall perforation as well as right ventricle perforation. the field representative noted that this occurred upon aspiration of a prior thrombus by the physician. the patient was taken to operating room emergently for repair with suture and packing as well as transcatheter aortic valve replacement. blood loss was significant and patient received six units of packed red blood cells, two units of placements, twenty units of prothrombin complex concentrate and desmopressin. the patient recovered with sequelae. | The family did make choice to withdraw care and the patient expired. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The investigation into the reported issue has been completed. the pump was not returned for investigation. data log review was not required for this case run. as per the clinical details, the left ventricle was perforated during impella use. the cause of the ventricle perforation issue was not determined since product was not returned and sufficient clinical information was not provided. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. | B2 outcome revised to reflect death and date of death which was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27142. b5 revised to reflect the patient outcome death as it was erroneously documented on the initial manufacturer device report number 1220648-2025-27142. h1 type of reportable event was erroneously on the initial manufacturer device report number 1220648-2025-27142. h6 revised to reflect death which was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27142.
10636 | 1220648-2025-27143 | 21740056 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21740056&pc=OZD | 08/19/2024 | 11/18/2025 |  | 04/01/2025 |  |  | 08/23/2024 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1684102 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported the automated impella controller had a ¿defective purge cassette¿ alarm. staff then primed and tried a second new purge cassette, which did not clear the alarm. the decision was made by nursing to change console. there was no patient harm. | [5 empty] | Corrections to the initial mfr report: g1 mdr reporting contact email. h6 type of investigation code added. | The investigation has been completed. data analysis: on 8/23, 8 days into running pump 498216 the user attempted a cassette change. after removing purge cassette 1786350, a new purge cassette was inserted. unfortunately, the purge cassette was not able to be read causing the aic to log "purgeadmin: rfid crc failure 0xc80023 44735" which resulted in a defective purge cassette alarm. unidentified cassettes are inserted and removed repeatedly over the following 30 minutes which were not read properly. just before the pump is removed, cassette 107320 is finally read correctly, removed, then read correctly again. device analysis: the console was returned for testing but multiple purge cassettes were inserted with no ability to reproduce the failure mode. decoding the crc failure, the rfid was looking to read 0xc80023 but read a different value. this could be force reproduced by changing the rfid on a purge cassette which will output what the console thinks the rfid should read. root cause: the defective purge cassette alarms could have been caused by an intermittently failing rfid pca (0042-3070). d10 concomitant medical products and therapy dates updated. e4 should have been left blank on the initial report. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
10637 | 1220648-2025-27144 | 21740283 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21740283&pc=OZD | 03/07/2025 | 04/14/2025 |  | 04/01/2025 |  |  | 03/07/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1715019 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The automated impella controller was noted to have optical bench issues. there was no patient use at the time or any chance of harm. the aic will be removed from service and repaired. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - loss of opm data has been completed. the console logs from the reported event day correspond with the complaint, indicating that the contrast value was out of specification. upon testing the console, the, voltage measurements were conducted from the power battery manager (pbm) to the optical bench, as well as across the lamp. it was noted that the voltage across the lamp measured approximately 0.227v. further voltage across the impellatronic were also measured and no issue were found. additionally, the external optical connector was found contaminated and remained unclean despite efforts to clean it. the root cause for reported complaint of the contrast value outside the specification range and later found controller error (defective optical sensor lamp) alarm was due to the defective optical bench. h3 device evaluation was erroneously selected yes on the initial manufacturer device report number 1220648-2025-27144.
10638 | 1220648-2025-27145 | 21740438 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21740438&pc=OZD | 10/11/2024 | 04/01/2025 |  | 04/01/2025 |  |  | 10/11/2024 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025543814 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 25 YR |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was attempted to be implanted with an impella 5.5 device for mechanical circulatory support. the native anatomy did not allow for the pump to be delivered. the team instead placed a cp for the support. no patient harm was reported. | [5 empty] | The investigation into the reported issue has been completed. no device was received for evaluation. the root cause of the delivery issue was most likely patient condition related, the doctor attempted to insert impella 5.5 but the vessel size was too small as per the clinical description provided. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
10639 | 1220648-2025-27147 | 21740666 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21740666&pc=OZD | 02/17/2025 | 10/01/2025 |  | 04/01/2025 |  |  | 02/17/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025605483 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | Stroke/CVA | Ischemia Stroke | Thrombosis/Thrombus | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | Followup | [6 empty] | Additionally, it was reported that the patient had a thrombus around the cannula. recommendations were made to exchange the impella 5.5 with a new one, however the heart failure team opted to leave existing impella 5.5 in and increase the ptt goal. it was noted that the patient had signs of a cva/stroke and an ischemia stroke was noted. a thrombectomy was performed. a week later the pump was explanted. | The complainant reported the patient was on impella 5.5 support and during support, the patient had a computed tomography scan that revealed a subacute left posterior cerebral artery infarct, no hemorrhage nor neuro deficits noted. heparin was decreased and support continued. additionally, there were low purge flows noted. purge flow decreased significantly from 10.5 cc/hr to 5.8 cc/hr. the purge was changed from heparin to bicarb fluid. furthermore, there as a clot noted in the right coronary artery. the relationship to the clot and the impella is unknown. support continued. | The customer reported that the device exhibited an optical sensor issue. |  |  |  | The high purge pressure failure mode was removed because there was no complaint of a high purge pressure occurrence neither was there a high purge pressure alarm/event in the data logs. the issue reported, low purge flow, was resolved during the case. the investigation of stroke/cva and thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6b explant date added e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27147 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The high purge pressure failure mode was removed because there was no complaint of a high purge pressure occurrence neither was there a high purge pressure alarm/event in the data logs. the issue reported, low purge flow, was resolved during the case. as this clinical information was not provided, the true root cause of the stroke/cva and thrombus could not be determined. b5 updated with additional information. d6b explant date added. h6 health effect - clinical code was revised to reflect ischemic stroke and health effect - impact code 4624 was added e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27147 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the device was not returned for evaluation, the root cause of the optical signal issue was most likely due to a sensor wear of the optical sensor as evidenced by the log pattern and duration of support. however, the root cause of this issue is unable to be determined. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical sensor code 2917 and the codes for the results of the optical sensor investigation.
10640 | 1220648-2025-27146 | 21740679 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21740679&pc=OZD | 02/11/2025 | 04/02/2025 |  | 04/01/2025 |  |  | 02/11/2025 | 04/01/2025 |  |  | 04/30/2025 |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024336955 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient was taken to cvor for escalation of support with impella 5.5 placement. it was noted that the 5.5 was inserted through axillary however was unable to pass through the innominate artery due to vascular tortuosity and calcification. decision was made to abort case and remove 5.5. the patient survived the procedure. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the impella 5.5 was inserted through the axillary but was unable to pass through the innominate artery due to vascular tortuosity and calcification. the cause of the delivery issues was most likely patient condition related due to patient's vascular tortuosity and calcification. d.4 revised expiration date and unique identifier (udi) # as they were previously entered incorrectly on the initial manufacturer device report number. e.1 revised us phone number as it was previously entered incorrectly on the initial manufacturer device report number. e.4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration. h.3 revised as the investigation was completed at the time the initial manufacturer device report number was submitted. h.6 codes 3221 and 4315 were previously entered incorrectly on the initial manufacturer device report number as the investigation was completed at the time the initial report was submitted. new codes were added that should of been reflected on the initial manufacturer device report number. h.10 investigation results were added as they were inadvertently omitted from the initial manufacturer device report number. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.
10641 | 1220648-2025-27148 | 21740699 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21740699&pc=OZD | 03/28/2025 | 07/03/2025 |  | 04/01/2025 |  |  | 03/03/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1336513 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that when attempting to prime an impella cp pump, a "please connect the red plug" message was displayed on the automated impella controller screen. as it was not possible to proceed with the priming process, the console was swapped. there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | Additional information provided in h3, h6 and h10. console logs from the reported day of event are consistent with complaint and show that during case start, after step 3 prompt to connect pump is displayed on console screen, pump is not detected. after 20 seconds from the prompt, the epm circuitry is power-cycled, but because pump connection is still not detected, a message ¿connect grey connectors¿ is displayed on aic screen. after case start wizard was launched again, pump was still not detected after step 3 prompt, epm was power-cycled, and a prompt ¿connect grey connectors¿ was displayed again, indicating no pump connection was detected. third attempt after re-launching case start wizard was unsuccessful. console was tested on-site 5 hours later with a demo pump (sn: 6) which was detected, without issues, twice. device logs from the backup console, show pump correctly detected and utilized for support for several days. console was tested, but the issue was not reproduced, and service pump was able to be detected each time. inability to detect pump, as evidenced in device logs, was consistent with an ¿epm crystal issue¿ failure mode, where i2c communication line between msc1210 microcontroller (epm) and eeprom in the pump red handle is not correctly established, due to impellatronic start-up issue. this intermittent pump detection failure occurs infrequently and has a mitigation in place. however, it appears that the mitigation mechanism was unable to resolve the issue, most likely due to some additional intermittent malfunction of impellatronic board. replacement of impellatronic assembly was completed and functional check of the console was performed.
10642 | 1220648-2025-27149 | 21740742 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21740742&pc=OZD | 03/07/2025 | 05/12/2025 |  | 04/01/2025 |  |  | 03/07/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026622801 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Device-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support using the single access technique. the physician noted that there was a lot of interaction when advancing the guide and the sheath got bunched up. the device was successfully placed. no patient harm was reported. | [5 empty] | The investigation has been completed. no product was returned and data logs are not applicable. it was reported that resistance was met when advancing guide catheter through companion sheath. the physician also felt a lot of interaction with impella pump when advancing the guide catheter. the root cause of the reported issue was most likely interaction with other device since the clinical details state the guide catheter and companion sheath getting bunched up and interacting with the impella. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27149 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10643 | 1220648-2025-27150 | 21740797 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21740797&pc=OZD | 03/07/2025 | 04/01/2025 |  | 04/01/2025 |  |  | 03/07/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025581142 |  | 1000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 81 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was attempted to be implanted with an impella device for mechanical circulatory support. the device was unable to fully advance through the patient's vasculature. resistance was met where the subclavian artery joined the brachiocephalic artery and an angiography revealed a tortuous/small vessel size. due to the noted resistance, the decision was made to abort the implant. an impella cp pump was subsequently placed for planned support. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controllerrp flex with smartassist system with automated impella controller section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
10644 | 1220648-2025-27151 | 21740810 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21740810&pc=OZD | 03/07/2025 | 05/01/2025 |  | 04/01/2025 |  |  | 03/07/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025601916 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced bleeding from the access site. heparin was withheld and blood products were administered. support continued until the patient was bridged to left ventricular assist device (lvad). | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the access site bleeding was determined to be anticoagulation management because since hemostasis was achieved by withholding anticoagulants. d9 was erroneously selected yes on the initial manufacturer device report number 1220648-2025-27151. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27151 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10645 | 1220648-2025-27152 | 21740935 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21740935&pc=OZD | 03/07/2025 | 04/30/2025 |  | 04/01/2025 |  |  | 03/07/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026617698 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 84 YR |  | Unknown |  |  |  | Malfunction | No | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that following impella cp implant, flows were found to be insufficient. the cannula was found to be kinked and the pump was replaced. there was no patient harm. | [5 empty] | The investigation of product damage (cannula kink) and low pump flow has been completed. the pump was returned for investigation. cannula kink was reported on returned product. no other physical damage was noted. pump was ran per the specification during the test in bucket of water. the data log shows low pump flow throughout the case run of 47 minutes with an active suction alarm. as per the clinical team, the aortic arch was small and difficult to pass but it was managed. during the case, a kink was reported with low pump flow, and the pump was exchanged for a new one. the cause of the low pump flow issue was a kinked cannula, based on returned product investigation. the cause of the product damage issue was a kinked cannula, which most likely related to the patient's anatomy. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
10646 | 1220648-2025-27153 | 21741234 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21741234&pc=OZD | 03/07/2025 | 05/06/2025 |  | 04/01/2025 |  |  | 03/07/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1177183 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | Yes | Pumping Stopped | [6 empty] | Cardiovascular Dysfunction/Insufficiency | No Clinical Signs, Symptoms or Conditions | [11 empty] | Death | [5 empty] | IMPELLA CP. | [18 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | At the time of explant, the patient was deemed hemodynamically stable and hemostasis was successfully achieved. unfortunately the patient expired in the procedure area. there is not enough evidence to disassociate the patient's passing with impella support. | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. pump stopped unexpectedly during support. it was noted that the aic screen returned to the home screen and showed a prompt stating "press ok to resume displayed performance level" at the time of the failure. the pump restarted on its own but was explanted following pci completion. there was no patient harm reported. |  |  |  |  | The investigation of automated impella controller (aic) - pump stop has been completed. data analysis: cp pump was connected to console at 13:38 on (b)(6) 2025. at 15:04 on (b)(6) the imcgui process stopped triggering the watchdog and a partial reset. the epc rebooted and the pump returned hemodynamic support through the conclusion of the procedure. device analysis: the suspect console was tested but the failure mode did not reproduce. the imc gui process could be forced to crash reproducing the symptoms in the field but the process did not crash on its own. root cause: the cause of the unexpected restart during support was the software imcgui process crashing. it is unknown what triggered the process to crash as the failure mode was not reproduced. b1 adverse event was inadvertently omitted on the initial manufacturer device report number. b2 the selection death and date of death was inadvertently omitted on the initial manufacturer device report number. b5 information regarding patient outcome; death was inadvertently omitted on the initial manufacturer device report number. h1 type of report was erroneously selected on the initial manufacturer device report number. h6 codes 4582 and 2199 were erroneously entered on the initial manufacturer device report number. new codes were added to health effect - clinical code and health effect - impact code. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
10647 | 1220648-2025-27154 | 21741301 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21741301&pc=OZD | 03/06/2025 | 04/01/2025 |  | 04/01/2025 |  |  | 03/06/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026617503 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Hemodynamic instability | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the insertion of impella 5.5 device to a patient in postcardiotomy cardiogenic shock was unsuccessful with anastomosis bleeding. the patient would ultimately expire but not before implanting an impella cp device that experienced low pump flow. the patient was admitted for pericardiectomy and during this procedure, the patient arrested. it was suspected that the left anterior descending artery had occluded and the patient was taken for an emergent coronary artery bypass graft. however, the patient would not separate from the bypass. the decision was made to implant an impella 5.5 device. during the implant, the impella 5.5 made it into the graft, past anastomosis but was unable to make the turn into the heart after multiple attempts. multiple blood products were used due to axillary anastomosis bleeding a "great deal." what was infused and the number of units were unknown. the physician felt there was a stricture preventing the impella 5.5 from entering the heart. the impella 5.5 implant was aborted, and an impella cp was called for to attempt to help separate from the patient from bypass. the impella cp device was implanted without incident;. however, it was noted that there were "great flows at the beginning." the blood pressure would increase and decrease with blood products being given. however, despite "aggressive" treatment, the blood pressure fell despite volume, and the patient expired on the operating room table. medical safety review of the event noted the impella cp have "great flows at beginning¿ reasonably suggesting optimal flows were not sustained resulting in an outcome of serious injury. however, the impella 5.5 bleeding from the anastomosis requiring multiple blood products at the onset of serious adverse complications and was likely to have contributed to the patient¿s demise. thus, cannot be excluded as a possible contributing factor. | [5 empty] | Medical safety review of the event noted the impella cp noted to have "great flows at beginning" reasonably suggesting optimal flows were not sustained resulting in an outcome of serious injury. however, the impella 5.5 bleeding from the anastomosis requiring multiple blood products is the onset of serious adverse complications and likely to have contributed to the patient¿s demise. thus, cannot be excluded as a possible contributing factor. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. this is one of two devices associated with the event. refer to manufacturing report number 1220648-2025-27129 for the impella 5.5 device report.
10648 | 1220648-2025-27155 | 21741329 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21741329&pc=OZD | 02/28/2025 | 04/01/2025 |  | 04/01/2025 |  |  | 03/07/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025577929 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 86 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Cardiac Arrest | [12 empty] | Required Intervention | Death |  |  |  |  | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient who experienced pulseless electrical activity (pea) cardiac arrest during impella cp support. the patient required around 25 minutes of cardiopulmonary resuscitation (cpr) with multiple doses of epinepherine, 2 amps of bicarb and also calcium. the patient eventually regained pulse. two additional units of bicarb were administered with substantial improvement in heart rate and stabilized hemodynamics noted. impella was turned down to p2 during cpr and back up to p8.¿ the patient recovered with sequelae. as support continued, pea arrest occurred again. six rounds advanced cardiovascular life support (acls) were provided while the impella was at p2 and code was called. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10649 | 1220648-2025-27156 | 21743003 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21743003&pc=OZD | 03/07/2025 | 05/01/2025 |  | 04/01/2025 |  |  | 03/07/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025599659 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the impella cp was opened and placed on the sterile field. vessel size and complication occurred with closure device which made the impella cp not able to be used. the physician was obtaining groin access and noticed the left groin was borderline on its size for impella cp use. the physician opted to move forward and loaded two perclose and while the second one was being loaded, it broke and was stuck in the vessel. vascular surgeon assisted and impella cp was removed and the intra-aortic balloon pump was placed. there was no harm to the patient. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since the implant of pump was aborted due to patient's anatomy. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
10650 | 1220648-2025-27157 | 21745441 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21745441&pc=OZD | 03/07/2025 | 04/08/2025 |  | 04/01/2025 |  |  | 03/07/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1786119 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Female |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) had a purge disk not recognized alarm while on patient support. the aic was exchanged, and the alarm was resolved. there was no patient harm. | [5 empty] | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. the logs from the complaint date revealed that the console issued purge disc not detected alarm several times. the console was examined and a broken purge disc flag was identified. the purge disc flag was observed to be broken and was the root cause of the reported purge disc not detected alarm issue. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
10651 | 1220648-2025-27158 | 21746253 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21746253&pc=OZD | 03/08/2025 | 05/09/2025 |  | 04/01/2025 |  |  | 03/08/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025594751 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Pumping Stopped | [5 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the pump developed frequent suction issues during patient support. volume and positioning were both verified, but the issue persisted. the pump was exchanged as a result of the noted issue. although there was no thrombus in the left ventricle, biomaterial was observed on the device. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." . in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation of low pump flow and thrombus has been completed. the impella device was not returned for evaluation. pump ran with low flow and suction alarms for entire case. biomaterial observed per clinical description. temporary pump stop was found on data logs. data review: data logs confirmed the failure mode & clinical narrative. logs showed pump started & ran with slightly low flow on p-8 followed mc spikes. about 269 hours into support, pump stopped temporarily & could restarted by itself that triggered alarm impella stopped-mc high. pump continued running on p-8 with low flow & multiple suction alarms. about 1347 hours into the case, flow dropped to 1.8 l/min & remained to the rest of the case. product was not returned for review. based on clinical description and data analysis, the root cause of low flow was most likely biomaterial ingestion since biomaterial observed per clinical team and on data logs. the root cause of temporary pump stop was most likely biomaterial ingestion since biomaterial observed per clinical team and on data logs. as the necessary information was not provided, the root cause of the thrombus was not determined.
10652 | 1220648-2025-27159 | 21746683 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21746683&pc=OZD | 01/22/2025 | 04/01/2025 |  | 04/01/2025 |  |  | 03/05/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025587861 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant had a impella cp support for a high-risk percutaneous coronary intervention patient. the pump was explanted after the percutaneous coronary intervention in roughly 3 hours. the impact study crf form later in 2/2025 reported upon a "closure of the pelvic axis" with possible relationship to the pump. the team used the pump without abiomed support present, so the need to perform percutaneous transluminal coronary angioplasty to close the access site, and the presence of an access site bleed or vascular damage is not known. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10653 | 1220648-2025-27160 | 21746690 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21746690&pc=OZD | 03/05/2025 | 04/01/2025 |  | 04/01/2025 |  |  | 03/05/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025594755 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Male |  |  |  | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported the patient had an attempted impella cp implant and during the implant, the pump was unable to pass through the ostial/prox common iliac artery due to stenosis. the implant was aborted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk pci section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions, ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ a notification was received from abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured respiratory failure and sepsis both with a "possible" relationship to the impella 5.5 device used.
10654 | 1220648-2025-27161 | 21746698 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21746698&pc=OZD | 03/06/2025 | 04/01/2025 |  | 04/01/2025 |  |  | 03/06/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1578302 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | It was reported during routine preventive maintenance, field service discovered corrosion in the power battery manager board of the automated impella controller (aic). consequently, the aic was sent in for repair. | [5 empty] | The automated impella controller was returned for evaluation and analysis. while inspecting the aic during preventative maintenance, field service found a corroded power battery manager (pbm) board. visual inspection confirmed the pbm contamination which impacted a neighboring communication channel. the pbm was replaced, and the console was able to perform as expected. a product history review was competed, and no action was required as a result of this review. in conclusion, the root cause of the "system self check failure" was determined to be pbm corrosion due to leakage of the electrolyte from the battery failure alarm super caps.
10655 | 1220648-2025-27162 | 21746706 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21746706&pc=OZD | 03/06/2025 | 04/01/2025 |  | 04/01/2025 |  |  | 03/06/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026622041 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Injury | No | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia Stroke | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy implanted with an impella cp device for mechanical circulatory support (mcs) experienced a stroke at explant. patient successfully placed on impella cp in the setting of cardiogenic shock. after three days of mcs, the patient was successfully explanted at bedside and manual pressure applied followed by femostop for hemostasis. however approximately 15 minutes later, the patient developed dysphasia and word finding difficulties. a stroke was suspected, and code stroke was called. the patient was taken for a computed tomography scan which showed left middle cerebral artery occlusion. interventional radiologist performed an urgent mechanical thrombectomy. occlusion was aspirated. the patient reportedly survived and was noted to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
10656 | 1220648-2025-27163 | 21746718 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21746718&pc=OZD | 03/06/2025 | 04/16/2025 |  | 04/01/2025 |  |  | 03/06/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025525061 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced bleeding from the groin site and mouth. the groin was treated by manual pressure, site adjustment, femstop, and stopping all systemic anticoagulation. of note the patient had been on heparin and eliquis which may have been contributing factors to the bleed. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27163 as. it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10657 | 1220648-2025-27164 | 21746800 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21746800&pc=OZD | 03/07/2025 | 05/15/2025 |  | 04/01/2025 |  |  | 03/07/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1376689 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The automated impella controller (aic) had a controller failure alarm. this issue occurred during servicing. no patient was involved. | [5 empty] | The data analysis confirmed the reported failure mode. during the device analysis the controller failure was intermittent and could not be reproduced after the initial occurrence. the root cause of the reported controller failure was not determined due to the inability to reproduce it. it could be due to either a pud board malfunction or a purge pressure sensor malfunction. d2 common device name was erroneously entered on the initial manufacturer device report number 1220648-2025-27164. d9 date device returned to manufacturer was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27164. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27164 as it is unknown if the initial reporter also sent the report to the food and drug administration. h3 was erroneously selected on the initial manufacturer device report number 1220648-2025-27164, as the device had been received at time of the report h6 type of investigation 4114 was erroneously entered on the initial manufacturer device report number 1220648-2025-27164. h10 narrative stated device was not received on the initial manufacturer device report number 1220648-2025-27164, as the device had been received at time of the report. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10658 | 1220648-2025-27165 | 21746857 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21746857&pc=OZD | 03/06/2025 | 04/17/2025 |  | 04/01/2025 |  |  | 03/07/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1395260 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the automated impella controller displayed the ¿battery fault¿ alarm after maintenance, despite the battery switch-in being performed using the button at the base of the console. there was no patient harm. | [5 empty] | The investigation of automated impella controller (aic) - battery charging/other issues has been completed. on the complaint date, the console was rebooted multiple times, and each time it displayed ¿battery failure (transport connection blown/missing). this confirms the reported failure mode. likely after the third reboot, when the battery switch was in a disengaged state, the ac power supply was disconnected, leading to an improper shutdown. subsequently, after reconnecting the ac power, the console was manually rebooted. after this reboot, the console displayed ¿unexpected controller shutdown ¿ alarm,¿ due to an improper shutdown, in addition to the battery failure alarm. device analysis: this console was tested and the battery switch was in a disengaged state when received. upon turning on the console, a battery failure alarm was triggered. after engaging the battery switch, the battery failure alarm was resolved. no issues were found with the batttest. root cause: the root cause of the battery failure alarm was most likely due to the battery switch being disengaged, which was caused by use issue. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-27165. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27165 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
10659 | 1220648-2025-27166 | 21746908 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21746908&pc=OZD | 03/08/2025 | 04/01/2025 |  | 04/01/2025 |  |  | 03/08/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026622610 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was attempted to be implanted with an impella device for mechanical circulatory support. implant was unsuccessful due to size of artery. the case was aborted. no patient harm was reported. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ impella cp with smartassist for use during cardiogenic shock and high risk pci & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smartassist system with automated impella controller. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. impella cp with smartassist system & impella 5.5 with smartassist for use during impella rp smartassist system with automated impella controller & rp flex with smartassist system with automated impella controllerrp flex with smartassist system with automated impella controller. section: warnings & cautions: warnings: ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
10660 | 1220648-2025-27167 | 21747030 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21747030&pc=PYX | 03/09/2025 | 04/01/2025 |  | 04/01/2025 |  |  | 03/09/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025569700 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Increase in Pressure | Failure to Pump | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported that several purge pressure high alarms were encountered during support followed by multiple pump stops. the patient was being repositioned at the time of the failures. the issues self resolved and support continued uninterrupted. there was no patient harm. the patient was eventually weaned and explanted. | [5 empty] | The investigation for the temp pump stop and the high purge pressure has been completed. no product was returned for analysis. the data logs from the case show a motor current spike coincident with a drop in placement signals, motor speed and a pump stop. the motor current pump stop was confirmed with alarm #13. the log pattern was characteristic of biomaterial ingestion. there were multiple failed restart attempts during an approximate thirty minute window before the pump was able to restart during which purge pressure started trending up. high purge pressure and purge system blocked alarms were triggered. after restarting, the pump was able to run without any issues and the purge pressure returned to normal values. the root cause of the temporary pump stop was most likely biomaterial as evidenced by the ingestion log patterns. the root cause of the high purge pressure was most likely biomaterial as evidenced by the ingestion log pattern. the instructions for use for the impella rp with smartassist system states the following: section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
10661 | 1220648-2025-27168 | 21747103 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21747103&pc=OZD | 03/09/2025 | 04/29/2025 |  | 04/01/2025 |  |  | 03/10/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2025611783 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported the patient was on impella cp support and during support, there was a leak observed from the purge side arm. there was a purge system open alarm that could not be reset. the pump was explanted. the patient survived the procedure. | [5 empty] | The investigation of purge leak ¿ pump has been completed. the pump was returned for investigation. the data log shows a purge leak in the purge pressure trend after 16 hours of case start. the returned product was examined, and the purge leak was noted from a punctured hole in the blue membrane on the sidearm. the cause of the purge leak was damage to the sidearm reservoir, which was most likely due to the use of the device. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-27168. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27168 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high-risk cpi states the following: section: cleaning, ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir."
10662 | 1220648-2025-27169 | 21747143 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21747143&pc=OZD | 03/10/2025 | 04/01/2025 |  | 04/01/2025 |  |  | 03/10/2025 | 04/01/2025 |  |  |  |  | 04/01/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025521496 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that a patient developed ischemia in their impella cp limb following implant. the patient's foot was marbled and there were no distal pedal pulse. vascular surgery was made aware of the condition. the pump was explanted less than thirty six hours later. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
10663 | 1220648-2025-27074 | 21729509 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21729509&pc=OZD | 02/25/2025 | 03/31/2025 |  | 03/31/2025 |  |  | 03/04/2025 | 03/31/2025 |  |  | 12/31/2025 |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025449543 |  | 1000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was initially implanted with an impella cp and supported for approximately five days before being escalated to an impella 5.5 for continued mechanical circulatory support. soon thereafter, there was high purge pressure. lines were confirmed to not be kinked. troubleshooting, there was temporarily release of the purge pressure which confirmed that the problem was with the patient rather than the purge cassette. the pump was explanted. there was no harm to the patient as a result of the event. | [5 empty] | The impella device was discarded by the customer due to infection and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
10664 | 1220648-2025-27075 | 21729525 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21729525&pc=OZD | 03/04/2025 | 05/08/2025 |  | 03/31/2025 |  |  | 03/04/2025 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025594183 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and experienced access site bleeding requiring intervention. the peel away 14 french peel-away sheath was removed and repositioning sheath was advanced. however, bleeding occurred at the groin site with difficulty controlling hemostasis. pulsatility was flattening on arterial line, and the mean arterial pressure dropped. epinephrine drip was started with pushes of neo-synephrine. there were suction alarms on the impella cp. blood products were ordered and integrillin drip was stopped. the patient was intubated. the left femoral artery was accessed and guide catheter up and over. the impella cp was turned off and explanted. retrograde balloon was deployed, two perclose cinched and manual pressure with good hemostasis and immediately place of a new impella cp through left groin. the patient was noted to be in critical condition following the event and last noted on support. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the access site bleeding issue was most likely use related as groin bleeding from under deployed per close. unable to cinch perclose after repositioning sheath advanced.
10665 | 1220648-2025-27076 | 21729548 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21729548&pc=OZD | 03/05/2025 | 03/31/2025 |  | 03/31/2025 |  |  | 03/05/2025 | 03/31/2025 |  |  | 11/30/2025 |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | OZD |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024435319 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on impella rp flex support and during support, the patient's distal right extremity was cold and no pulses were present. the patient was taken to the hybrid room to do an abdominal angiogram with runoff of the right lower extremity. when runoff was completed, there was no distal flow to the right femoral artery or right femoral vein. the impella rp flex was weaned. a fasciotomy was completed to the right lower extremity through the calf muscle due to swelling and blisters noted to the skin. the patient survived the procedure. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smart assist system with automated impella controller section: contraindications (united states) ¿the impella rp flex with smartassist system is contraindicated for use with patients experiencing any of the following conditions: disorders of the pulmonary artery wall that would preclude placement or correct positioning of the impella rp flex with smartassist device; mechanical valves, severe valvular stenosis or valvular regurgitation of the tricuspid or pulmonary valve; mural thrombus of the right atrium or vena cava; anatomic conditions precluding insertion of the pump; presence of a vena cava filter or caval interruption device, unless there is clear access from the internal jugular vein to the right atrium that is large enough to accommodate a 22 fr catheter.¿
10666 | 1220648-2025-27077 | 21729593 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21729593&pc=OZD | 03/05/2025 | 03/31/2025 |  | 03/31/2025 |  |  | 03/05/2025 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1765949 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the console would not turn on and until the power cord was plugged in. the bottom power button was engaged on. upon power up, there was a battery failure and battery communication failure alert. there was no patient involvement. | [5 empty] | The investigation has been completed. the automated impella controller (aic) logs confirmed the reported failure. there was a battery failure alarm, battery 1 communication failure alarm and a battery 2 communication failure alarm due to loss of communication to the battery. further investigation into the aic logs revealed that the unit had been unused for four months prior to the reported issue. the console was tested on a lab bench, where the failure mode was successfully reproduced. the aic would not turn on using only battery power. upon arrival, the batteries were found to be dead, and the logs confirmed lack of use during the four-month period. upon replacing the original batteries with known good ones, there was no battery failure alarm, and the batteries were able to charge and the power battery manager was working as per the specification. the battery failure alarm and the battery communication failure alarms were due to the depleted batteries. the underlying cause of the battery depletion was likely the extended period of inactivity, as the aic was kept unplugged for four months.
10667 | 1220648-2025-27078 | 21729611 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21729611&pc=OZD | 03/05/2025 | 03/31/2025 |  | 03/31/2025 |  |  | 03/05/2025 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025603998 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient lost pulses in right lower extremity during support. it was decided to explant the pump. the extremity remained pulseless following the explant and vascular surgery was consulted. the patient was reported to be in critical condition. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10668 | 1220648-2025-27079 | 21729660 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21729660&pc=OZD | 03/04/2025 | 04/29/2025 |  | 03/31/2025 |  |  | 03/04/2025 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1155606 |  | 0042-0000-US |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) had a controller failure alarm which resolved on its own. it was not reported if there was any patient involvement. | [5 empty] | After further review manufacturer device report 1220648-2025-27079 is a duplicate of manufacturer device report 1220648-2025-27041. no further information will be provided in manufacturer device report 1220648-2025-27079. should any additional information become available it will be reported in manufacturer device report 1220648-2025-27041. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10669 | 1220648-2025-27080 | 21729707 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21729707&pc=OZD | 03/03/2025 | 03/31/2025 |  | 03/31/2025 |  |  | 03/03/2025 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025610276 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 67 year old male patient with a cp pump placed for support due to acute myocardial infarction. it was reported that during support, the patient developed hemolysis. the patient had urine color change to pink and so the p level was lowered. plan was to remain on support despite the color change. the patient was eventually explanted. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction. ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
10670 | 1220648-2025-27081 | 21729736 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21729736&pc=PYX | 02/13/2025 | 03/31/2025 |  | 03/31/2025 |  |  | 03/04/2025 | 03/31/2025 |  |  | 11/30/2025 |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2024425653 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | Information was received via long-term outcome and quality indicator (loqi) impella registry database reporting two adverse events. "first ischemic stroke middle cerebral artery and second mediastinal bleeding. cerebral bleed of unknown relatedness and bleeding of possible relationship. the loq states the following: after the patient was extubated, he developed right-sided weakness, expressive and receptive aphasia. initial nihss was 23 points for right facial droop, aphasia, right-sided weakness. nihss stabilized at 8 points on (b)(6) 2025 at 0730. cta was negative for large vessel occlusion, and he is not a candidate for mri due to durable lvad. the patient had three chest tubes placed on (b)(6) 2025 during procedures (aortic valve replacement, redo sternotomy, tavr explant; impella rp flex implant). chest tube output between (b)(6) equated to greater than 2l in a 24 hour period. the site was also noted to be "oozy". blood loss also attributable to repeated wiring catheters. admission (baseline) hemoglobin 13.7 g/dl, nadir hemoglobin 6.6 g/dl (b)(6) 2025." | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
10671 | 1220648-2025-27082 | 21729785 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21729785&pc=OZD | 03/03/2025 | 04/28/2025 |  | 03/31/2025 |  |  | 03/03/2025 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | UNKNOWN |  | UNKNOWN |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | Yes | Patient-Device Incompatibility | [6 empty] | Great Vessel Perforation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported the patient had an attempted impella 5.5 implant and during the implant, they had difficulty advancing the wire in the ascending aorta. staff attempted through the left femoral artery and this was unsuccessful. the implant was aborted and an axillary artery dissection was suspected. the only impella product used was the introducer, the guidewire used was not an abiomed product. the patient outcome was unknown. | [5 empty] | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. the instructions for use for the impella cp with smartassist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. missing details: a.1, a2, a.4 and a.5 are unknown. d.4 catalog number, serial number, lot number, expiration date and unique identifier (udi) # are unknown. d.6a and d.6b implant and explant dates are unknown. e.1 name prefix/title, first name and last name are unknown. h.4 device manufacture date is unknown. | The investigation of delivery issues and vascular damage - peripheral vessel has been completed. only short 23fr x 6cm introducer and guidewire was returned for investigation. data log was not required for this case run. the returned guidewire was not being used. and no damage was noted on short 23fr x 6cm introducer. the cause of the delivery issue was most likely patient condition related to challenging anatomy. the cause of the vascular damage issue was not determined since sufficient clinical information was not provided. g1 reporting contact facility name and address was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27082.
10672 | 1220648-2025-27083 | 21729824 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21729824&pc=OZD | 03/04/2025 | 03/31/2025 |  | 03/31/2025 |  |  | 03/04/2025 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025610441 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 86 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and experienced hemolysis and thrombus requiring intervention. the patient in st-segment elevation myocardial infarction (stemi) was taken to the catheterization lab for an emergent percutaneous coronary intervention (pci). the left main was engaged for angiogram and the patient became acutely hypotensive requiring vasopressors. the decision was, therefore, made to place the patient on impella mcs before proceeding with the pci. an impella cp device was implanted, and the patient underwent "successful" pci for anterior stemi. the patient was sent to the unit on support. the following day, the patient was noted to have hemolysis. the patient had continuous red urine. the impella was repositioned, and the adverse event was being monitored for improvement. additionally, it was reported, three days later, the impella cp device was explanted, and an angiogram of the right femoral artery was performed from the contralateral side revealing thrombus in the right common femoral artery at the sheath insertion site and slightly above. several attempts to aspirate the thrombus were done using the penumbra device. the physician stated results were not "fully satisfactory," and further commented that the patient may lose the foot if survives. last noted, the patient survived the explant and was in stable condition. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
10673 | 1220648-2025-27085 | 21730144 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21730144&pc=OZD | 03/03/2025 | 06/26/2025 |  | 03/31/2025 |  |  | 03/03/2025 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026622870 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | Atrial Flutter | Hematoma | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | Gastrointestinal Hemorrhage | [7 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | It was further reported the patient experienced runs of atrial fibrillation and atrial flutter. | The complainant reported that there was some oozing and a medium sized hematoma at the impella cp access site. the medical doctor adjusted repositioning sheath and manipulated the perclose and oozing/hematoma resolved. additionally, the patient developed thrombocytopenia and blood products were administered. the patient also developed a gastrointestinal bleed. the family decided to withdrew care and the patient expired. |  |  |  |  | The investigation of access site bleeding, thrombocytopenia, vf/vt, and vascular damage - peripheral vessel has been completed. access site bleeding: product was not returned for investigation. as per the clinical information provided, the reposition sheath and the perclose had to be adjusted to stop the oozing. therefore, the root cause of the access site bleeding issue was most likely use related to improper technique. thrombocytopenia: the root cause of the thrombocytopenia issue was not determined due to insufficient clinical details. vascular damage: pump was not returned. clinical information provided was very limited and no information provided about the reason of the gi bleed. therefore, the root cause of the vascular damage issue was not determined due to insufficient clinical information and unknown relation to impella with gi bleed. vf/vt: as the necessary information was not provided, the root cause of the vt/vf was not determined. b5 revised to include additional information inadvertently omitted on the initial manufacturer device report number: 1220648-2025-27085. h6 health effect - clinical code 1729 and 1730 were inadvertently omitted on the initial manufacturer device report number: 1220648-2025-27085. | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10674 | 1220648-2025-27084 | 21730145 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21730145&pc=OZD | 03/03/2025 | 04/18/2025 |  | 03/31/2025 |  |  | 03/03/2025 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025611389 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Perforation of Vessels | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 58 year old male patient on impella 5.5 support experienced ventricular fibrillation/ventricular tachycardia (vf/vt). attempted to give lidocaine and amiodarone was unsuccessful. the team ended up cardioverting the patient. impella with suction alarms during this whole episode. echo performed, impella appeared to be interacting with pap muscles of mitral valve. impella was repositioned. vf subsided after cardioversion and impella retraction. it was further reported that the patient had a drop in blood count due to gi bleed. heparin is on hold. patient eventually bridged to transplant. | [5 empty] | The investigation of vf/vt/af/at and vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. g1 reporting contact facility name was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27084. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. section: warnings & cautions: cautions: ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
10675 | 1220648-2025-27086 | 21730146 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21730146&pc=OZD | 12/19/2024 | 03/31/2025 |  | 03/31/2025 |  |  | 03/03/2025 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025534515 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 40 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured right lower extremity (rle) ischemia with "definitely related" to the impella. loqi states: patient showed mottling in the rle, and no signal found in rle below groin. required emergent distal perfusion cannula placement at bedside. additionally, care was withdrawn and patient expired. there is not enough information to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10676 | 1220648-2025-27087 | 21730223 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21730223&pc=OZD | 03/04/2025 | 05/21/2025 |  | 03/31/2025 |  |  | 03/04/2025 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025501599 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombocytopenia | [11 empty] | Required Intervention | Death |  |  |  |  | HEMASHIELD PLATINUM GRAFT. | VENO-ARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION. | [17 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was initially implanted with an impella cp device with veno-arterial extracorporeal membrane oxygenation (va ecmo) support for approximately 10 days until escalation in therapy to an impella 5.5 device for continued mechanical circulatory support (mcs). during impella 5.5 support, the patient experienced bleeding at the graft and thrombocytopenia requiring blood products transfusion. the impella cp device was successfully weaned and explanted by utilizing the double barrel method inserting the impella 5.5. the impella 5.5 pump was started at performance level p-2. there was an attempt to wean off of ecmo, but the patient's hemodynamics were not tolerating this wean. simultaneously, the walls of the graft were sweating noticeable amounts of blood. consequently, performance level was left at p-2 and maintained support on ecmo, with plan to return at a later date for removal of ecmo support. for the graft, this was known to be coagulaopathic. pressure was held, hemostatic agent was applied, and blood product were given with eventual hemostasis on the graft. the site was then closed and dressed appropriately. three-point fixation was applied, and the patient in stable manner was transferred to unit on mcs. the following day, one unit of platelets was given for thrombocytopenia (unclear relationship to the impella 5.5). the patient continued on support and was last noted to be in stable condition. | It was further reported the family decided to withdraw care and the patient expired. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | The device remains implanted and therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | B2 updated to include death and date of death b5 updated with additional information received from the clinical site d6b explant date added h1 type of report revised based on the additional information received from the clinical site h6 revised to reflect death. | The investigation of access site bleeding and thrombocytopenia has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding was most likely use issue since non-abiomed product was used which is not recommended per instructions for use. the root cause of thrombocytopenia cannot be determined since the product was not returned and insufficient clinical details were provided.
10677 | 1220648-2025-27088 | 21730294 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21730294&pc=OZD | 03/04/2025 | 12/09/2025 |  | 03/31/2025 |  |  | 03/04/2025 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025611529 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp stopped when performing an automated impella controller (aic) to aic transfer. the impella did not restart again even when it was reconnected to the original aic. the pump was explanted and the patient was reported to be stable on extracorporeal membrane oxygenation (ecmo). | [5 empty] | Clinical details note that after 1 day on support, a pump stop occurred when performing an aic to aic transfer, with pump not working on either console after transfer was initiated. patient was stable on ecmo. the console numbers are not known so data logs were not recovered. the pump was returned and evaluated. during testing, the pump function was intermittent. the pump could connect to the aic and run without issue; however, when plugged into the aic at another point, an impella defective alarm triggered. the pump was also plugged into the console without being detected at all. the patient cable plug had a raised area that made it difficult to push the plug fully into the receptacle. there was also an abnormal gap observed on the patient cable plug. the motor cables passed all electrical testing. the 4 communication cables from the patient cable plug to the red handle pcb were electrically tested. there were no abnormalities with cable connections to pin 1, 5, and 6 of the patient cable plug. the connection to pin 11 was an electrical open; however, when applying force to "close the gap" on the patient cable plug, the electrical connection could be re-established temporarily. this indicated an intermittent electrical open line inside the patient cable plug. an x-ray of the patient cable plug showed clear visual narrowing at the pin 11 connection in the plug. the cause of the pump not being detected was an electrical open based on returned product analysis but the cause of the electrical open is unknown. | The impella pump was discarded. should new information be received, a supplemental manufacturer device report will be filed.
10678 | 1220648-2025-27090 | 21730477 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21730477&pc=OZD | 03/04/2025 | 05/19/2025 |  | 03/31/2025 |  |  | 03/04/2025 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026612842 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 82 YR |  | Male |  |  |  | Malfunction | No | Yes | Difficult to Remove | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was to undergo high risk percutaneous coronary intervention was to be implanted with impella cp for mechanical circulatory support. the impella cp was delivered to the left ventricle, however the wire could not be retracted out. the impella was therefore explanted, and no harm or injury was noted. | [5 empty] | The investigation of delivery issues, product damage (cannula kink), and introducer damage ¿ mechanical has been completed. introducer damage added since damage was found on the returned product. delivery issue - reported implanting the pump with no issues but had issues retracting the guidewire. the guidewire was unable to be removed and the pump was explanted through the introducer. impella cp, introducer and guidewire were returned. the guidewire was still in the cannula. the end of the guidewire coming out of the outflow cage was wrapped around the cannula, which was kinked. the cause of the guidewire delivery issue was most likely due to a kinked cannula as the cannula kink, observed on the returned product, likely made it difficult to remove the guidewire. cannula kink - clinical reported a cannula kink. no issues delivering the pump and no resistance crossing the valve. cannula kink was found on the returned product. the root cause of the cannula kink was not determined due to insufficient clinical details. introducer damage - reported being unable to removed the guidewire resulting in explanted the pump via the sheath. the introducer was returned with the guidewire and pump catheter within the sheath. introducer damage was found on the returned introducer. there was buckling of the sheath near the introducer hub and damage to the tip of the sheath. the root cause of the introducer damage - mechanical was most likely an introducer sheath kink since clinical reporting having issues removing the pump and guidewire via the sheath and there was buckling found near the introducer hub on the returned product. h6 component code 4742 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27090. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance."
10679 | 1220648-2025-27091 | 21730687 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21730687&pc=OZD | 03/04/2025 | 05/06/2025 |  | 03/31/2025 |  |  | 03/04/2025 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025591637 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | Deformation Due to Compressive Stress |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the insertion of impella cp device to a patient undergoing a high-risk percutaneous coronary intervention was unsuccessful. the device was unable to pass the aortic valve. impella cp crossed but would deflect lateral. multiple attempts made with different buddy wires and pigtail. the physician confident there was no involvement with mitral structure. the physician stated there was ¿tons of resistance¿ and while advancing the guidewire it would buckle and/or move toward free wall causing concern for patient harm if continued attempts. therefore, unable to pass the aortic valve, the impella cp implant was aborted. the patient's outcome following the event was unknown; however, the patient was indicated to have survived the event. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since resistance was met due to patient's anatomy. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
10680 | 1220648-2025-27092 | 21730807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21730807&pc=OZD | 03/04/2025 | 04/15/2025 |  | 03/31/2025 |  |  | 03/04/2025 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026618413 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Hispanic |  | Malfunction | No | Yes | Infusion or Flow Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella cp device preoperatively for mechanical circulatory support and experienced saturated flows that led to device replacement. the following day after start of impella mcs, coronary artery bypass graft surgery was completed with the impella cp in the aorta. post bypass graft surgery, the impella cp was repositioned in the left ventricle. following, placement signal showed saturated flows, and consequently, the impella cp was explanted and replaced without incident. the patient sent to the unit on mcs in stable condition. | [5 empty] | The pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of saturated flow has been completed. based on data log review and the biomaterial on returned product, the root cause of the saturated pump flows was determined to be biomaterial.
10681 | 1220648-2025-27093 | 21730860 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21730860&pc=OZD | 03/03/2025 | 05/05/2025 |  | 03/31/2025 |  |  | 03/03/2025 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP-JP |  |  | OZD |  |  | IMPELLA CP-JP |  |  | ABIOMED, INC. |  |  | US |  | 2025597525 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Death | Yes | Yes | Device Sensing Problem | [6 empty] | Great Vessel Perforation | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient on the impella cp for acute myocardial infarction. it was reported that the pump had suction, position unknown, and position in aorta alarms while on patient support. the positioning of the impella was checked and no problems were noted. it was further reported that transthoracic echocardiogram revealed pericardial effusion from ventricular septal perforation (vsp) and oozing rupture. the relationship between oozing rupture/pericardial effusion and impella is unknown. the patient eventually expired on support. | [5 empty] | The investigation of optical signal issue and ventricle perforation has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27093 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27093. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10682 | 1220648-2025-27094 | 21730871 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21730871&pc=OZD | 03/04/2025 | 05/08/2025 |  | 03/31/2025 |  |  | 03/04/2025 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025597758 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 78 YR |  | Female |  |  |  | Death | Yes | Yes | Failure to Advance | Malposition of Device | [5 empty] | Cardiac Arrest | [12 empty] | Death | Life Threatening | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported the insertion of impella cp device to a patient (unknown race and ethnicity) in acute myocardial infarction/cardiogenic shock was unsuccessful and the patient would expire on the procedure table in the cardiovascular lab. the patient presented to an outside facility and was found to be having a st-segement evaluation myocardial infarction (stemi). the patient was given tissue plasminogen activator (tpa) and emergently transferred in. on arrival, a repeat electrocardiogram confirmed stemi; the patient was emergently brought to the cardiovascular lab. a left heart catheterization found 100% occluded left anterior descending artery (lad). the physician utilized intracoronary thrombolysis to dissolve a thrombus in the lad which restored thrombolysis in myocardial infarction 3 flow. upon reperfusion, the patient went pulseless and cardiopulmonary resuscitation (cpr) was initiated. mid-cpr, the impella was emergently prepped and primed. the peel-away sheath was inserted. due to cpr, the placement guidewire slipped out of the left ventricle while advancing the impella and the physician was unable to successfully pass the pigtail across the aortic valve. cpr continued while the physician attempted placement. a pulse check was done, and the physician made the decision to abort both cpr and impella placement after inability to achieve return of spontaneous circulation. there is not enough information to exclude the impella cp as an associated factor to the event. however, it should be noted the patient was in a state of decline without pulse at the onset of impella cp placement attempt. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be use issue because the issue happened during performing the cpr. instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings section: pre-support evaluation section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ section: warnings & cautions: warnings, section: pre-support evaluation, section: direct aortic insertion, section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ d6a and d6b were inadvertently omitted on the initial manufacturer device report number 1220648-2025-2709. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27094 as it is unknown if the initial reporter also sent the report to the food and drug administration. h10 instructions for use were inadvertently omitted on the initial manufacturer device report number 1220648-2025-2709.
10683 | 1220648-2025-27095 | 21730877 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21730877&pc=OZD | 03/04/2025 | 03/31/2025 |  | 03/31/2025 |  |  | 03/04/2025 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025605355 |  | 1000080 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | A |  |  | No | Initial submission | [9 empty] | The user facility reported a patient was implanted with an impella cp device for mechanical circulatory support and experienced limb ischemia. the patient was status post coronary artery bypass graft surgery x3 vessels with maze on intra-aortic balloon pump mcs three days prior to the start of impella mcs. the patient coded 30 minutes after surgery with return of spontaneous circulation achieved. however, the patient continued to decompensate over the weekend and was placed on the impella cp device. the following day after start of the impella mcs. the patient lost radial pulse in the arm. consequently, the impella cp device was explanted and replaced. the patient's status following the event and post explant and replacement of the impella cp device is unknown. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
10684 | 1220648-2025-27096 | 21730887 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21730887&pc=OZD | 03/04/2025 | 05/13/2025 |  | 03/31/2025 |  |  | 03/04/2025 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025592152 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematuria | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support as the patient became unstable during percutaneous coronary intervention to the left anterior descending and circumflex arteries. following the impella cp device implant, the patient developed hematuria, and the decision was made to explant the pump as a result of the condition. the patient was reported to be in stable condition following the event and explant of the device. | [5 empty] | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. the root cause of the hemolysis issue most likely was improper positioning related since the clinical attachment confirmed hematuria occurred due to the fact that impella needed to be repositioned and the physician refused to get echocardiogram and pull back. instructions for use: impella cp with smartassist system: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ e4 should have been left blank on the initial manufacturer device report number: 1220648-2025-27096 as it is unknown if the initial reporter also sent the report to the food and drug administration. h10 instructions for use was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27096. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10685 | 1220648-2025-27097 | 21730893 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21730893&pc=OZD | 03/05/2025 | 03/31/2025 |  | 03/31/2025 |  |  | 03/05/2025 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025550415 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Female |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, the repositioning sheath was occlusive in the right femoral artery (rfa). impella was removed in cath lab. rfa angiogram was performed after 1 perclose was deployed and second perclose tightened down around micro puncture catheter. angiogram revealed blood flow to the rfa, proximal superficial femoral artery and proximal profunda. blood flow to the distal extremity was not visualized while in cath lab. venoarterial extracorperal membrane oxygenation (ecmo) access in the left femoral artery with distal perfusion catheter in place to lower left extremity. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
10686 | 1220648-2025-27098 | 21730896 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21730896&pc=OZD | 03/05/2025 | 05/01/2025 |  | 03/31/2025 |  |  | 03/05/2025 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025591315 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump placed to support a acute myocardial infarction and cardiogenic shock patient. the patient had an access site hematoma and bleeding which prompted the team to explant the pump after just over 2.5 hours. the patient survived the event. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details, the root cause of access site bleeding was most likely patient movement since the bleeding and hematoma was observed due to patient movement. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27098 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10687 | 1220648-2025-27099 | 21730961 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21730961&pc=OZD | 03/05/2025 | 04/16/2025 |  | 03/31/2025 |  |  | 03/05/2025 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025603509 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that during placement of a impella 5.5 pump on a 59 year old male patient, the pump was unable to advance into the patient's vessel during attempted delivery. after multiple failed attempts, the decision was made to abort the implant. an intra aortic balloon pump was placed for planned support. there was no patient harm. | [5 empty] | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details, the root cause of the delivery issue was most likely patient condition since device was not able to advance into the vessel after multiple attempts and troubleshooting due to axillary artery stenosis. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27099 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27099. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.
10688 | 1220648-2025-27100 | 21730994 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21730994&pc=PYX | 03/04/2025 | 05/20/2025 |  | 03/31/2025 |  |  | 03/05/2025 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | PYX |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1752054 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Decrease in Pressure | Application Program Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported approximately 12 days after start of impella mechanical circulatory support, the automated impella controller (aic) sounded for a controller failure alarm. consequently, the team switched the aic in response to the failure. there was no patient harm as a result of this event. | [5 empty] | The investigation of automated impella controller (aic) - purge issue ¿ other, aic - controller error (yellow alarm), aic - controller failure (red alarm) has been completed. the console was received and data and device analysis was performed. there was no controller failure (red alarm) recorded in the log on the complaint date (mar 04, 2025). most likely, the controller error was misreported as a controller failure. the root cause for the controller error was not determined due to the inability to reproduce it. the root cause of the purge system blocked alarm was most likely due to fluid ingress towards the purge motor. (this is an additional analysis finding related to the clinical procedure). | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | D2 common device name was erroneously entered on the initial manufacturer device report number 1220648-2025-27100. g4 pma/510(k)number was erroneously entered on the initial manufacturer device report number 1220648-2025-27100.
10689 | 1220648-2025-27101 | 21731075 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21731075&pc=OZD | 03/04/2025 | 03/31/2025 |  | 03/31/2025 |  |  | 03/04/2025 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1368535 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported during preparation for a case, the automated impella controller (aic) was unable to recognize the purge cassette. consequently, the team changed over to new aic without issue for successful case using the same purge cassette and impella pump. it was further reported there was no harm to the patient. | [5 empty] | Patient-specific information is unavailable at the time of this mdr writing; therefore, respective fields reflect "unknown" or left blank accordingly. the automated impella controller data logs and console were returned for evaluation and analysis. data analysis demonstrated logs were inconsistent with what was reported in the complaint. although it was reported that the console could not recognize the purge cassette, no evident purge cassette recognition issues were observed. however, there were multiple purge fluid not detected errors observed in the logs. rtlog data that was available showed that at the time the purge fluid not detected errors occurred, the purge pressure was at 0mmhg. the console was tested after arriving in field service, and the purge cassette recognition issue was unable to be reproduced. the observed purge fluid detection issues also could not be reproduced. the console was tested with a purge cassette kit and pump and operated as intended. a product history review was completed, and no action was required as a result of this review. in conclusion, the reported purge cassette recognition issue and the observed purge fluid not detected issue could not be determined due to the inability to reproduce. the purge pressure was at 0mmhg at the time of the purge fluid not detected event. this absence of purge pressure could be indicative of a kink in the purge line or a possible leak but no there was no conclusive evidence. the aic operated as intended when testing in the lab.
10690 | 1220648-2025-27102 | 21731149 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21731149&pc=OZD | 03/05/2025 | 04/19/2025 |  | 03/31/2025 |  |  | 03/05/2025 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026619790 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Positioning Failure | Restricted Flow rate | Material Twisted/Bent |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, the proper position could not be achieved. the cannula had a kink in it that was caused by patient anatomy. multiple attempts to reposition were made, including removing the cannula from the ventricle. the pump was explanted and there was no reported patient harm. | [5 empty] | The investigation of positioning issues has been completed. the returned complaint cp device visually inspected. cannula kink observed near outlet. no biomaterial, no broken blade observed. no other abnormality found. failure mode positioning issues changed to failure mode low pump flow issue since low flow observed after insertion and kink occurred and pump reported to be in good position. the cause of the low pump flow issue was cannula kink due to patient anatomy . failure mode product damage (cannula kink) added since kink occurred during insertion and positioning due to patient anatomy. the cause of the product damage (cannula kink) was cannula kink due to patient anatomy condition. | The pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
10691 | 1220648-2025-27103 | 21731335 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21731335&pc=OZD | 03/05/2025 | 04/01/2025 |  | 03/31/2025 |  |  | 03/05/2025 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026617585 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 89 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella cp support developed a hematoma at the insertion site. the patients¿ activated clotting time was over 300. integrilin was stopped, pressure was held for 10 minutes with sandbag placed after, and the hematoma was resolved. additionally, the patient developed an gastrointestinal bleed and blood products were administered. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | H.6 code 1884 was inadvertently omitted from the initial manufacturer device report number 1220648-2025-27103.
10692 | 1220648-2025-27104 | 21731784 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21731784&pc=OZD | 03/06/2025 | 04/25/2025 |  | 03/31/2025 |  |  | 03/06/2025 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026620233 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, the patient experienced limb ischemia so there was placement of a distal perfusion catheter. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
10693 | 1220648-2025-27105 | 21732188 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21732188&pc=OZD | 03/04/2025 | 03/31/2025 |  | 03/31/2025 |  |  | 03/04/2025 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025533150 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Cardiac Arrest | Hemodynamic instability | [11 empty] | Required Intervention | Death | Life Threatening | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient underwent percutaneous coronary intervention: drug-eluting stent to the left anterior descending artery (lad) and right coronary artery (rca). the patient began feeling "bad" throughout the night and an electrocardiogram showed anterior st-segment elevation the following morning. the patient was taken to the catheterization lab and in-stent thrombosis of the lad was noted. upon wiring the lad the patient went into ventricular fibrillation arrest and was shocked five times. the decision was made to implant an impella cp device, which was placed via the right femoral artery without incident. the patient was intubated. penumbra and percutaneous transluminal angioplasty restored flow to the lad. it was reported, however, after start of the impella mcs, the pump began to have suction. a fluid bolus was given. the patient became very acidotic and multiple pressors were added. multiple bicarb pushes and drip were started. an echocardiogram was completed and showed the impella cp pump pointed toward mitral instead of toward apex. the physician was unable to get better positioning with the device. the patient continued to decline and ultimately expired. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10694 | 1220648-2025-27106 | 21732344 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21732344&pc=OZD | 12/10/2022 | 03/31/2025 |  | 03/31/2025 |  |  | 12/10/2022 | 03/31/2025 |  |  | 07/31/2024 |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023255173 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 71 YR |  | Male |  | Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. the patient was then escalated from the impella cp to an impella 5.5 for increased circulatory support after 4 days of support on the impella cp. when the patient was brought to the operating room (or) to replace the impella cp with impella 5.5 the decision was made for the patient to have coronary artery bypass (cabg). the patient was placed on bypass, and the impella was place directly through the aorta after an aortic valve replacement was performed. cabg was performed successfully and without complications. the impella 5.5 placement was confirmed with transesophageal echocardiography (tee), support was initiated and the patient was weaned off bypass. tee of the right ventricle revealed worsening right heart function. the patient was given an increased inotrope support and transported back to the icu. the following day, on (b)(6) 2022, the patient was taken back to the or for venoarterial extracorporeal membrane oxygenation cannulation (va ecmo) due to the failing right ventricle and associated decompensation. during this time, it was reported that the impella 5.5 had alarms for high purge pressure. the medical team decided to place tissue plasminogen activator in the purge solution. the high purge pressure alarm was resolved; however, it was noted that the pressure remained high, with the mean motor current within expected ranges. there were no further impella complications and the patient was transported back to the icu. on (b)(6) 2022, the patient's right lung was discovered to be necrotic. the surgical team made plans for resection at a later time, however the patient continued to decline and required increased support despite the patient being supported by impella 5.5, va ecmo, and hemodialysis. the family made the decision to withdraw care, and the patient expired. although the patient's presenting condition may be more likely have impacted the event the impella cannot be excluded as a possible contributing factor due to the high purge pressure which is a product malfunction and cannot exclude this from the outcome of demised. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ this report is being filed as part of a retrospective review of historical records.
10695 | 1220648-2025-27107 | 21732384 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21732384&pc=OZD | 12/05/2021 | 03/31/2025 |  | 03/31/2025 |  |  | 12/05/2021 | 03/31/2025 |  |  | 07/31/2023 |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2022113536 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Bradycardia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient in non-st elevation myocardial infarction with hypotension, acute severe mitral regurgitation, and low ejection fraction was implanted with an impella 5.5 device for mechanical circulatory support to stabilize and unload the patient's left heart. during delivery of the impella into the left ventricle, the impella 5.5 was placed directly info the papillary wall and was unable to be moved. the medical team was advised to pull the guidewire, and then turn the impella as there was suspicion the guidewire was in the chordae. the patient then went into bradycardia and lost blood pressures. the impella 5.5 was successfully turned towards the apex after removal of the guidewire and support was then initiated. it was noted by the physician they believe that the patient's pressures dropped due to the right coronary artery being completely blocked and having no perfusion. during support, there were suction alarms. the power level of the impella 5.5 was dropped from p9 to p8 and the patient was given fluid volume to resolve the alarms. the patient remained on support for 5 days, and it was determined the patient was unlikely to survive any further revascularization and the decision was made to withdraw care. care was withdrawn, and the patient expired. although the patient's presenting condition may be more likely have impacted the event the impella cannot be excluded as a possible contributing factor in the patient's demise. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)." this report is being filed as part of a retrospective review of historical records.
10696 | 1220648-2025-27109 | 21732710 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21732710&pc=OZD | 03/04/2025 | 09/30/2025 |  | 03/31/2025 |  |  | 03/04/2025 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025612245 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant also reported that the pump had placement signal issues. the audio was disabled, and position was verified. support continued. | The complainant reported that a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support. it was reported that the patient developed access site bleeding and anticoagulation was held while the support continued. the decision was made to withdraw care and the patient expired. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  |  | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue and access site bleeding. the pump was not returned. data logs were not recovered. the cause of the optical signal issue was not determined since product was not returned, data logs were not recovered, and lack of clinical information. the root cause of the access site bleeding - major was not determined since product was not returned and lack of clinical details. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include placement signal issue description. h6 updated to include placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-27109. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated. | The impella pump was not returned for investigation. should new information be received, a supplemental manufacturer device report will be filed. section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10697 | 1220648-2025-27108 | 21732728 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21732728&pc=OZD | 03/05/2025 | 04/04/2025 |  | 03/31/2025 |  |  | 03/05/2025 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026622081 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump placed to support a patient admitted in acute myocardial infarction/cardiogenic shock. as the device was explanted, it was noted that there was a clot in the peel away sheath. manta closure was utilized. after closure, the leg was assessed to confirm no clot migrated distally. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation of thrombus has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
10698 | 1220648-2025-27110 | 21732762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21732762&pc=OZD | 12/08/2022 | 03/31/2025 |  | 03/31/2025 |  |  | 12/08/2022 | 03/31/2025 |  |  | 07/31/2024 |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023248881 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 45 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Failure to Advance | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A patient in shock from doxorubicin induced cardiomyopathy was admitted for medical management. the patient's presenting oxygen saturation (o2 sat) results were 30-40% with a cardiac index of 1.7. an intra-aortic balloon pump was placed where the patient had no major improvements to the o2 sat or the cardiac index results. the decision was then made to place an impella 5.5 for increased mechanical circulatory support. during the delivery of the impella 5.5 to the left ventricle, it was reported multiple unsuccessful attempts were made to advance the impella 5.5 through the graft anastomosis site. during these attempts, the 0.018 wire came out of the left ventricle. the decision was made to remove the impella from the graft and tunnel the graft more laterally and less inferior. the graft was also trimmed before the sheath and graft locks were placed in the graft. an angled pigtail and j wire utilized to re-cross the aortic valve and placed in the left ventricle apex. the j wire was exchanged for an 0.018 steel core wire. viper slide was then coated on the impella cannula to allow for smoother advancement of the impella through the graft and anastomosis. the impella 5.5 was placed on the steel core wire and advanced into the left ventricle with minimal difficulty and support was initiated. the patient remained on support until successfully weaned and explanted. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issues was most likely due to patient anatomy. impella 5.5 with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: warnings & cautions: warnings: ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/ descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels¿. ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with echocardiography guidance.¿ section: warnings & cautions: cautions: never advance the guidewire or sheath when resistance is met. determine the cause of resistance using fluoroscopy and take remedial action. this report is being filed as part of a retrospective review of historical records.
10699 | 1220648-2025-27111 | 21732790 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21732790&pc=OZD | 03/04/2025 | 09/26/2025 |  | 03/31/2025 |  |  | 03/04/2025 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026613662 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Positioning Problem | [5 empty] | Hemolysis | [12 empty] | Other | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that a patient undergoing aortic and mitral valve replacement surgery was implanted with an impella 5.5 device for mechanical circulatory support. it was noted that the patient was severely vasoplegic after coming off cardiopulmonary bypass, multiple drips initiated, and blood products given. additionally, it was noted that hemolysis had occurred post implant. waveforms were indicated position with separated placement and left ventricle (lv) signals. echocardiogram revealed mid lv not touching any walls. consideration of thrombus ingestion. creatine elevated, lactate dehydrogenase (ldh) elevated, creatine kinase - mb elevated, dialysis initiated, patient developed shock liver which resolved. the impella was successfully weaned and explanted. | [5 empty] | Based on the investigation, the failure mode (fm) of positioning issue was changed to optical signal issue as they report waveform issues with no true positioning failure event. return product was visually inspected and biomaterial was observed at outflow under the impeller. light continuity test was done and light passed to sensor face. no other abnormalities were seen. pump was ran at p-9 in lab in bucket of water for around 20 minutes. all optical 5s parameters were in normal range. no psnr triggered and no issue with optical signals relative to the lab run. the root cause of the optical signal/waveform issue was likely biomaterial ingestion triggering the false position aorta/unknown alarms based on log, clinical and return product review. hemolysis: data logs were reviewed and first 8 minutes on rt log since pump start showed some small motor current mimics/spike without any speed deviation or change in p-level. impella in aorta alarms trigger relative to those small mc spike events on 03/04/25. rt log showing pulsatility and mc dampening due to likely ingestion triggering false aorta/position alarms. no true positioning alarms were seen as pump position was confirmed it was not touching any walls per clinical information reported. return product was visually inspected and biomaterial was observed at outflow stuck under the impeller. no damaged impeller blades or other abnormalities were seen. the root cause of the hemolysis issue was likely due to biomaterial wrapped under impeller based on return product analysis and ingestion trend seen relative to reported hemolysis event per log review. d9 date device returned to manufacturer added. | The impella pump was not returned for investigation. should new information be received, a supplemental manufacturer device report will be filed. instructions for use as follows: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6. updated to capture optical signal problem code. d10 concomitant medical products and therapy dates.
10700 | 1220648-2025-27112 | 21732935 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21732935&pc=OZD | 03/06/2025 | 04/16/2025 |  | 03/31/2025 |  |  | 03/06/2025 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026619225 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced bleeding from the impella access site. the perclose were tightened and purse string suture was added. site redressed. the bleeding was resolved. the patient continued on support until the device was successfully weaned. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details, the root cause of the access site bleeding issue was most likely use related, as hemostasis was achieved tightened perclose and added purse string suture. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27112 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10701 | 1220648-2025-27113 | 21732938 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21732938&pc=OZD | 08/14/2024 | 03/31/2025 |  | 03/31/2025 |  |  | 03/06/2025 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD |  |  | IMPELLA |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Atrial Fibrillation | [12 empty] | Other | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient who experienced atrial fibrillation during impella support. additional information pertaining to the device and/or the event was not provided. | [5 empty] | . d.6a and d.6b implant and explant dates are unknown. h.4 device manufacture date is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10702 | 1220648-2025-27115 | 21732985 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21732985&pc=OZD | 03/06/2025 | 09/22/2025 |  | 03/31/2025 |  |  | 03/06/2025 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026620821 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Cardiovascular Dysfunction/Insufficiency | [9 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, the patient had access site bleeding. pressure dressing was applied and heparin was withheld. additionally, the pump was noted to be mispositioned in a folded position causing significant aortic insufficiency and pink/red urine was observed. the pump was repositioned successfully. physician stated that significant hemolysis resulted in damaged kidneys. impella support continues. | The complainant also reported that there were placement signal issues. staff monitored impella flows and motor current. |  |  |  |  | The investigation of access site bleeding, hemolysis, renal failure, and heart valve damage has been completed. the device and data logs were received for investigation. access site bleed: clinical details report that on the 3rd day of support, the access site was oozing. the complication was managed by pausing systemic heparin and applying a pressure dressing. returned product was investigated and there were no visual abnormalities with the repositioning unit or catheter anchor. based on limited clinical details and no abnormalities with returned product, the root cause of the access site bleed could not be determined. hemolysis: clinical details reported that the pump was noted to be deep in the lv and the cannula was folded on echocardiogram. the patient also began to test positive for hemolysis through pfhgb. the pump was repositioned, and the complication resolved. data logs leaded up to the reported date of the issue did not reveal any abnormalities. pump flows were within specification, there were no suction alarms, and there were no positioning alarms, but psnr was active for the majority of (b)(6) 2025. returned product had a distinct kink in the cannula, confirming the clinical report. there were no other visual abnormalities or damages noted. based on clinical details provided and the kink in returned product, the root cause of the hemolysis was determined to most likely be improper positioning. renal failure: clinical details reported that as a result of the experienced hemolysis, the patient had kidney damage, ultimately requiring that they get a transplant. data logs nor product are relevant to the issue. the cause of the renal failure was not determined due to insufficient clinical details. heart valve damage: clinical details report that at the same time that the patient was experiencing hemolysis symptoms, aortic insufficiency (ai) was also occurring. during this time, the pump was seen to be positioned deep in the lv and the cannula was folded over on echo. data logs were reviewed and there were no positioning alarms leading up to the complication. this could be due in part to the fact that the psnr alarm was active for a large portion of the case, disabling position monitoring. returned product was investigated and there was a distinct kink in the cannula, confirming the clinical report. based on the clinical details provided and the kink in returned product, the root cause of the heart valve damage was determined to most likely be improper positioning. d6b explant date was added. d9 date device returned to manufacturer added. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27115 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. data log review confirmed the optical signal issues. the signal-to-noise ratio (snr) was below 2000 for the majority of the case, however, there were also several instances where the optical signal metrics suddenly worsened/improved. these are all indications of an airgap between the console and patient cable plug. returned product was inspected and the only visual abnormality was damage to the patient cable plug. based on data log review, clinical details reporting that the issue would resolve when the pump plug was pushed in properly, and optical signal metrics being normal on returned product, the root cause of the optical signal issue was determined to be an airgap between the console and patient cable plug. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b1 product problem updated. b5 updated to include placement signal issue description. h6 updated to include placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-27115. b7 other relevant history, including preexisting medical conditions updated. d10 therapy dates updated.
10703 | 1220648-2025-27116 | 21733100 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21733100&pc=OZD | 03/05/2025 | 03/31/2025 |  | 03/31/2025 |  |  | 03/05/2025 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1713948 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No |  | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | NA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported that the automated impella controller (aic) had a white line in the screen. there was no patient involvement. | [5 empty] | The investigation into the reported issue has been completed. console logs from the reported day of event do not contain any information relevant to this failure mode. the console was sent back to field service for further investigation and was tested. the reported issue was reproduced, and intermittent lines with varying intensity were observed on lcd screen. further testing aimed at determining whether the issue was caused by the 4-in-1 assembly or the lcd were inconclusive. however the backlight cable had a visible kink. the root cause of the intermittent lines was mostly due to the 4-in-1 assembly or the lcd.
10704 | 1220648-2025-27117 | 21733278 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21733278&pc=OZD | 03/04/2025 | 05/02/2025 |  | 03/31/2025 |  |  | 03/05/2025 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025607872 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient developed a hematoma at the axillary site. the patient had been laying with arms above the head. the hematoma prompted the team pause the heparin for 24 hours. the next day the site was flat and the patient had no pain. the patient continued on support until the patient was bridged to a transplant. | [5 empty] | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of access site bleeding is determined to be patient movement because the hematoma was noted after patient napping his hands. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-27117. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10705 | 1220648-2025-27118 | 21733373 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21733373&pc=OZD | 03/05/2025 | 03/31/2025 |  | 03/31/2025 |  |  | 03/05/2025 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025607712 |  | 1000100 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [5 empty] | AUTOMTED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. there was increased oozing from right axillary site and in drain. dressing was changed and one unit of blood was administered. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10706 | 1220648-2025-27119 | 21733402 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21733402&pc=OZD | 02/13/2023 | 03/31/2025 |  | 03/31/2025 |  |  | 02/14/2023 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1325322 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported an automated impella controller would not save its date and time after power cycling. it was reported that after powering the unit off, the date would reset to 05-dec-2011. there was no patient harm reported for this issue. | [5 empty] | The investigation for the reported issue has been completed. the device event logs were reviewed and confirmed the reported issue. the device was inspected and it was determined the root cause of the issue was a defective coin cell battery. this report is being filed as part of a retrospective review of historical records.
10707 | 1220648-2025-27120 | 21733412 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21733412&pc=OZD | 02/28/2025 | 05/12/2025 |  | 03/31/2025 |  |  | 03/05/2025 | 03/31/2025 |  |  | 12/31/2025 |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025470428 |  | 0048-0024-JP |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that the patient developed cardiac tamponade due to acute myocardial infarction, and impella cp was inserted to stabilize hemodynamics. after percutaneous coronary intervention, it was difficult to control the bleeding due to cardiac tamponade. the patient was moved to the operating room and thoracotomy was performed to stop the bleeding. when the thoracotomy was performed, it was visually confirmed that the impeller had perforated the ventricle. bleeding was not able to be controlled, and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system and impella 5.5 with smartassist for use during section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions. ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ | B3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-27120. | The investigation of ventricle perforation has been completed. the pump was not returned for investigation. a data log review was not required. according to the clinical report, the patient was reported to have a fragile area of the myocardium. the doctor confirmed, based on visual inspection, that the impella had perforated the ventricle. the cause of the ventricle perforation issue was most likely patient condition related to poor heart condition, based on provided clinical details.
10708 | 1220648-2025-27121 | 21733562 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21733562&pc=PYX | 03/05/2025 | 04/22/2025 |  | 03/31/2025 |  |  | 03/05/2025 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025600672 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that high purge pressure was encountered during impella rp flex support. roughly thirteen days into support, purge flows began to decrease while the purge pressure rose. multiple infusions of tissue plasminogen activator were performed, but the issue persisted. the pump was replaced as a result of the high purge pressure and climbing motor current. there was no patient harm. | [5 empty] | The investigation of high purge pressure has been completed. the impella rp flex returned for evaluation. upon visual inspection, biomaterial was present in the inflow cage around the purge gap and impeller blade. high purge pressure reproduction testing was performed and pump was connected to lab console and purged for approximately 45 minutes. purge pressure was high but no alarms were triggered. pump was turn on to p-9 and was able to run without issue. again, no "high purge pressure" alarms were reproduced. biomaterial was present on the impeller blade after pump run. returned data logs were reviewed and long-term log at time of reported issue showed no "high purge pressure" alarms being triggered with a stable purge pressure and purge flow. clinical details noted that patient had had previous indwelling lines and extracorporeal membrane oxygenation prior to impella rp flex insertion. additionally, tissue plasminogen activator infusion was attempted twice for decreased purge flows. the root cause of the high purge pressure was most likely due to biomaterial as evidenced with the returned product. b.3 revised date of event as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-27121. b.7 added information that was inadvertently omitted from the initial manufacturer device report number 1220648-2025-27121. d.1 brand name was revised as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-27121. d.4 model number was revised as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-27121. g.4 added pma/510(k) number as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-27121. g.6 30-day should of been selected on the initial manufacturer device report number 1220648-2025-27121. h.6 code 3221 was repeated twice on the initial manufacturer device report number 1220648-2025-27121 by mistake. code 3233 should of been reported along with code 3221. h.10 instructions for use were mistakenly reported on the initial manufacturer device report number 1220648-2025-27121 for an impella rp device and this report is for an impella rp flex issue. | The impella device was received from the customer. upon completion of the investigation, a supplemental mdr will be filed. as noted in the impella instructions for use: impella rp with smartassist system section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿
10709 | 1220648-2025-27122 | 21733762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21733762&pc=OZD | 03/05/2025 | 04/19/2025 |  | 03/31/2025 |  |  | 03/05/2025 | 03/31/2025 |  |  | 11/30/2025 |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024431183 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Material Twisted/Bent | Adverse Event Without Identified Device or Use Problem | [5 empty] | Perforation of Vessels | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 76 year old male patient had a 5.5 pump placed for support during an emergent mitral valve surgery. the patient was supported by ecmo and cp at the time. later after the explant of the cp, the patient condition deteriorated. the team made choice to place a 5.5 pump via open chest and direct insertion of the 5.5 pump. the team noted a lv perforation from the 5.5 pump in the or. the team felt this occurred during cannula decannulation or during chest closure. the lv was repaired and the patient survived. | [5 empty] | The investigation of ventricle perforation has been completed. the impella 5.5 was returned for evaluation. ventricle perforation: upon visual inspection, a kink in the pump cannula was observer near the distal tip of the pump. data logs were reviewed and lt log of full case shows some repeated ¿impella position unknown¿ and ¿suction¿ alarms occurring during case. at time of reported issue, some ¿impella position unknown alarms were triggered. no ¿impella in ventricle¿ alarms observed at all during case. clinical details noted that ventricle perforation occurred during cannulation relocation of extracorporeal membrane oxygenation (ecmo). additionally, patient had a diseased/damaged ventricle prior to impella placement. the root cause of the ventricle perforation was most likely due to a use related issue as the clinical details noted that the ventricle perforation occurred during relocation from central to peripheral ecmo. product damage (cannula kink): an additional failure mode of "product damage (cannula kink)" was added to the investigation as a cannula kink was found on the returned product. clinical details noted that the during ecmo relocation, a ventricle perforation occurred and was likely caused a canula kink. the root cause of the product damage was most likely due to a cannula kink however the exact cause of the cannula kink was not clear. | The impella device was received from the customer. upon completion of the investigation, a supplemental mdr will be filed. impella 5.5 with smartassist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10710 | 1220648-2025-27123 | 21733788 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21733788&pc=OZD | 07/18/2023 | 11/14/2025 |  | 03/31/2025 |  |  | 07/18/2023 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1541271 |  | 0042-0040-US |  | No |  |  | Returned to Manufacturer |  | No | No |  |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 82 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | Device Sensing Problem | No Pressure |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that during impella support on automated impella controller (aic) with serial number (b)(6), the placement signal was noted not to be present, followed by a controller error alarm. the issue self-resolved, but later repeated. the console was exchanged for a backup aic with serial number (b)(6), however during the switch purge disk clip was noted to be loose and a low purge pressure alarm occurred; to resolve this issue, pump was then transferred to the third aic with serial number (b)(6), where issue was resolved. there was no patient harm reported, the patient remained on the third aic for support until successfully weaned and removed from impella support. | [5 empty] | This is one of two medical device reports associated with the event. refer to manufacturing report number 1220648-2023-02683 for automated impella controller with serial number (b)(6). the automated impella controller event logs have been received and an evaluation is underway. upon completion of the investigation, a supplemental report will be filed. this report is being filed as part of a retrospective review of historical records. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. console logs from the reported day of event are consistent with complaint and show that after transferring pump (430956) and purge cassette (24481) from (b)(4) to the complaint console (b)(4) , console presented with alarms ¿purge system open¿ and ¿purge pressure low.¿ during console inspection in danvers, damage of the right post of purge disk retainer was confirmed, and it explains purge-related alarms during case (purge disk membrane was not held firmly against purge pressure sensor). a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. d2 common device name updated. d10 concomitant medical products and therapy dates updated.
10711 | 1220648-2025-27124 | 21734034 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21734034&pc=OZD | 03/05/2025 | 03/31/2025 |  | 03/31/2025 |  |  | 03/05/2025 | 03/31/2025 |  |  |  |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025589927 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a impella cp pump placed to support a coronary artery bypass graft patient. the pump was supporting when after just over 4 hours the pump was explanted due to limb ischemia. the patient survived the event. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
10712 | 1220648-2025-27125 | 21734186 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21734186&pc=OZD | 03/15/2022 | 03/31/2025 |  | 03/31/2025 |  |  | 03/15/2022 | 03/31/2025 |  |  | 11/30/2023 |  | 03/31/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2023168293 |  | 0550-0008 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Advance | [6 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient, who was previously implanted with an impella cp was being worked up for a durable left ventricular assist device when they started decompensating. the patient went into atrial fibrillation. and quickly increased decompensating and was administered amiodarone with nitroprusside and milrinone. the patient did not respond to the treatment and the decision was made to place an impella 5.5 for mechanical circulatory support. during impella 5.5 delivery, the physician was unable to advance the impella from the axillary artery into the aorta. the impella was removed and washed in a sterile canister. an amplase wire placed and the impella implanted over the impella wire. with noted difficulty getting the impella around the bend in the artery, the impella was eventually successfully delivered to the left ventricle. support was initiated and the impella was set to power level p9. the patient remained on impella support for 6 days. after further declining in condition, the family decided to withdraw care. care was withdrawn and the patient expired. although the patient's presenting condition may be more likely have impacted the event the impella cannot be excluded as a possible contributing factor in the patient's demise. | [5 empty] | The investigation for the reported delivery issue has been completed. the impella device was not received from the customer and therefore, an evaluation of the device was not possible. however, clinical details provided were sufficient to determine a root cause. the root cause of the delivery issue was most likely due to patient condition. impella 5.5 with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
10713 | 1220648-2025-27017 | 21717596 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21717596&pc=OZD | 03/02/2025 | 04/10/2025 |  | 03/28/2025 |  |  | 03/02/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025507091 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella 5.5 was unable to be implanted due to the patient¿s anatomy. the impella 5.5 insertion was aborted. the patients outcome is unknown. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. the root cause of delivery issue was determined to be patient condition because the pump was unable to pass through the vasculature.
10714 | 1220648-2025-27018 | 21717692 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21717692&pc=OZD | 03/02/2025 | 04/03/2025 |  | 03/28/2025 |  |  | 03/02/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026616755 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 44 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on implanted with an impella cp due to ventricular tachycardia and acute myocardial infarction. the impella cp was supporting for just over 28 hours when explanted with care escalation to an impella 5.5 pump. the impella cp limb was ischemic and the patient had low urine output. the impella 5.5 was placed. there was no patient harm reported. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of limb ischemia and hemolysis has been completed. based on the patient update, good flow was maintained during impella cp support at p-9, but the patient started having hemolysis, blood in urine and pfhgb was 340. echocardiogram was done and position was confirmed. no distal pulses were found in the right foot and the vascular team assessed the patient. unclear details were provided to confirm the cause of the hemolysis issue. complaint device was not available to investigate and data logs were not returned to review. the cause of the hemolysis issue could not be determined since the complaint device was not returned for analysis and insufficient clinical data was provided. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
10715 | 1220648-2025-27019 | 21717751 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21717751&pc=OZD | 03/04/2025 | 06/02/2025 |  | 03/28/2025 |  |  | 03/04/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp stopped when performing an automated impella controller (aic) to aic transfer. the impella cp did not restart again even when it was reconnected to the original aic. the patient was reported to be stable on extracorporeal membrane oxygenation. | [5 empty] | D.4 catalog number, serial number, lot number, and unique identifier (udi) # are unknown. h.4 device manufacture date is unknown. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | Console logs were not provided for analysis. consoles involved in the complaint were not returned. pump 568751 has been returned and was investigated under (b)(4).
10716 | 1220648-2025-27020 | 21717811 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21717811&pc=PYX | 12/13/2024 | 04/16/2025 |  | 03/28/2025 |  |  | 03/03/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025502679 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 75 YR |  | Female |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with an impella rp flex for support of st elevated myocardial infarction and right ventricle (rv) failure. the patient was taken to the operating room and had the impella rp flex placed but, the patient expired due to prolonged down time and stopping the resuscitation efforts. the long-term outcome and quality indicator (loqi) database later reported upon a rv perforation with a "possibly" relatedness to the pump. the perforation was found after emergency sternotomy and the native tissue was poor and unable to be repaired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella rp flex with smartassist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: potential adverse events (united states) ¿arrhythmia, atrial fibrillation, bleeding, cardiac tamponade, cardiogenic shock, death, device malfunction, hemolysis, hepatic failure, insertion site infection, perforation, phlegmasia cerulea dolens (a severe form of deep venous thrombosis), pulmonary valve insufficiency, respiratory dysfunction, sepsis, thrombocytopenia, thrombotic vascular (non-central nervous system complication, tricuspid valve injury, cardiac or vascular injury (including ventricular perforation), venous thrombosis, ventricular fibrillation and/or tachycardia.¿ section: warnings & cautions: warnings ¿do not advance or withdraw the impella rp flex with smartassist system catheter against resistance without using fluoroscopy to determine the cause of the resistance. doing so could result in separation of the catheter or guidewire tip, damage to the catheter or vessel, or perforation.¿ | The investigation of ventricle perforation has been completed. the cause of the ventricle perforation was not determined since product was not returned and lack of clinical details. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-27020.
10717 | 1220648-2025-27021 | 21717873 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21717873&pc=OZD | 03/03/2025 | 04/02/2025 |  | 03/28/2025 |  |  | 03/03/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025597616 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 42 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Other | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient on impella 5.5 support developed bleeding at the insertion site. the site has improved. it is unknown how the bleeding was resolved and if intervention was done. no additional information has been received. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
10718 | 1220648-2025-27022 | 21717932 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21717932&pc=OZD | 03/03/2025 | 03/28/2025 |  | 03/28/2025 |  |  | 03/03/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025522256 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that after an impella cp pump was inserted in a patient, the pump stopped. a high motor current high alarm was triggered. the pump was replaced. there was no patient harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
10719 | 1220648-2025-27023 | 21718160 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21718160&pc=OZD | 03/02/2025 | 03/28/2025 |  | 03/28/2025 |  |  | 03/02/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025584552 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation of Vessels | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient on impella cp support expired after the patient¿s condition worsened due to vascular damage that occurred when extracorporeal membrane oxygenation was removed. vascular repair procedure was performed but the patient continued with bleeding at the site of the vascular injury. heparin was used in the impella purge but it could not be determined if the heparin in the purge exacerbated the situation for the patient. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
10720 | 1220648-2025-27024 | 21718252 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21718252&pc=OZD | 02/26/2025 | 03/28/2025 |  | 03/28/2025 |  |  | 03/03/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025584621 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 86 YR |  | Unknown |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Vascular Dissection | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient had an impella cp placed for support due to acute myocardial infarction. the patient was supported by extracorporeal membrane oxygenation additionally. the pump was placed but due to aortic dissection, the patient expired. the contribution of the impella cp to dissecting the ascending aorta is not known. | [5 empty] | The impella device was discarded by the customer and therefore,.an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings & cautions: cautions ¿physicians should exercise special care when inserting the impella catheter in patients with known or suspected unrepaired abdominal aortic aneurysm or significant descending thoracic aortic aneurysm or dissection of the ascending, transverse, or descending aorta.¿ caution: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
10721 | 1220648-2025-27025 | 21718321 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21718321&pc=OZD | 03/01/2025 | 05/16/2025 |  | 03/28/2025 |  |  | 03/01/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025588389 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 31 YR |  | Male |  |  | White | Injury | Yes | Yes | Pumping Stopped | Infusion or Flow Problem | [5 empty] | Embolism/Embolus | Thrombosis/Thrombus | No Clinical Signs, Symptoms or Conditions | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was in impella cp support for over 11 days when there was saturated flows and a pump stop. the pump had concerns for thrombus burden as the dialysis line had clotted. the team replaced the pump without harm or injury. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of saturated flow and pump stop has been completed. the impella device was not returned for evaluation. the failure mode (fm) thrombus was added to address the voice of the customer (voc). the fm pump stop was changed to temporary pump stop to address the voc and based on data log analysis. pump stop and saturated flow on the 28th, the purge pressure slightly dropped and there was an increase in the purge flow. purge flow was seen to be at increased values until the 1st when the pump was explanted. on the day of the reported event (march 1st), the pump had been running for 274 hours when the motor current spikes and the pump stops with saturated flows with no increase in purge pressure or spike in placement signal. the motor current remained high for the rest of the case. the pump is able to be restarted until it stopped again 3 hours later. it was able to be restarted and support was used for 2 hours longer until the decision was made to explant the pump. without the returned product, the root cause of the saturated flow and pump stop could not be determined. as the necessary information was not provided, the root cause of the thrombus was not determined. based on the investigation findings the following fields were updated. b1 updated to both adverse event and product problem. b2 updated. h1 type of report. h6 health effect - clinical code updated to reflect thrombus. instructions for use (thrombus). impella cp with smart assist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
10722 | 1220648-2025-27026 | 21718400 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21718400&pc=OZD | 01/09/2025 | 03/28/2025 |  | 03/28/2025 |  |  | 01/09/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025585560 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that a patient on impella cp support was taken to the operating room for a coronary artery bypass graft and mitral valve regurgitation repair. aorta cross was clamped with the impella cp in the left ventricle (lv) after shallowing position out to three centimeters using transesophageal echocardiogram. the impella cp was placed in surgical mode for case duration periodically turning up p level for ante grade cardioplegia and venting the lv. at the conclusion of the case, the impella cp was repositioned across the aortic valve (av) with pigtail toward the mitral apparatus. the surgeon attempted to torque the impella cp away from the mitral and pulled the impella cp back across the av. they were unable to advance the impella cp so a decision was made to pull back into the descending aorta on p1 and remove the impella at the conclusion of the case. the patient outcome is unknown. | [5 empty] | The investigation has been completed. the product was not returned. the root cause of the positioning issue was determined to be most likely use issue because the surgeon while attempting to torque the impella cp away from the mitral pulled the impella cp back across the aortic valve.
10723 | 1220648-2025-27027 | 21718413 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21718413&pc=OZD | 03/02/2025 | 04/14/2025 |  | 03/28/2025 |  |  | 03/02/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025585482 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that an impella cp had less than adequate flows at implant. echocardiogram was reviewed, fluids were given and placement was checked. a decision was made to remove the impella cp and replace with a new one. the patient remained stable. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of low pump flow has been completed. the root cause of low flow was unable to be determined as limited clinical details were provided and no product was returned for review.
10724 | 1220648-2025-27028 | 21718755 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21718755&pc=OZD | 03/01/2025 | 04/11/2025 |  | 03/28/2025 |  |  | 03/01/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026617610 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella 5.5 support and during support, there was oozing at the access site. the dressing was changed and heparin was held. support was continued. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of access site bleeding is determined to be patient movement because the patient was up and sitting in the chair and doing rounds during support which caused site bleeding.
10725 | 1220648-2025-27029 | 21718760 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21718760&pc=OZD | 08/10/2022 | 03/28/2025 |  | 03/28/2025 |  |  | 08/10/2022 | 03/28/2025 |  |  | 04/30/2024 |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2023226056 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with cardiomyopathy was implanted with an impella cp for mechanical circulatory support. the patient had previously been implanted with an impella cp via the right femoral artery and was converted to an impella cp implanted into the right axillary artery for extracorporeal membrane oxygenation decannulation. it was reported while on impella cp support, the patient experienced bleeding from the mouth and as a result heparin was discontinued until resolved. the following day, on (b)(6) 2022 a hematoma was noted at the impella access site. it was noted as a result of the hematoma, systemic heparin was withheld. the patient was noted to have had the impella successfully explanted, however the patient's outcome was unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. impella cp with smartassist for use during cardiogenic shock and high-risk pci instructions for use for the related event are as follows: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ this report is being filed as part of a retrospective review of historical records.
10726 | 1220648-2025-27030 | 21718956 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21718956&pc=OZD | 02/02/2025 | 03/28/2025 |  | 03/28/2025 |  |  | 02/02/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025580291 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Fluid/Blood Leak | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was on impella 5.5 support and during support, the patient had a gastrointestinal bleed (gib). the patient was switched from systemic heparin to heparin in the purge. blood products were given. the gib continued for a couple days with heparin in the purge. additionally, the pump was attempted to be weaned and the patient went hypotensive with runs of ventricular tachycardia (vt). the pump was turned back to p5. furthermore, there was purge fluid leaking noted at the transition from the clear plastic to the white cable. no change in purge flow/purge pressure was noted. support was continued. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿
10727 | 1220648-2025-27031 | 21719095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21719095&pc=OZD | 02/13/2025 | 04/08/2025 |  | 03/28/2025 |  |  | 02/19/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025541969 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Hispanic |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Disseminated Intravascular Coagulation (DIC) | Renal Failure | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a cp pump placed for 65-year-old male patient admitted in with st-elevated myocardial infarction and support during cardiopulmonary resuscitation (cpr). cpr was performed was for 90 minutes and the patient was taken for external jugular vein cannulation and placed on ecmo. the patient had severe hypovolemia and lactic acidosis when mass transfusion protocol was initiated. however, the patient developed disseminated intravascular coagulation (dic) and renal failure and a decision was made by family to withdraw care. the patient subsequently expired. | [5 empty] | The investigation of renal failure has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-27031. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27031 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10728 | 1220648-2025-27033 | 21719146 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21719146&pc=OZD | 02/04/2025 | 03/28/2025 |  | 03/28/2025 |  |  | 02/04/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025599866 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 28 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Infection | Ectopic Heartbeat | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that an impella 5.5 needed to be adjusted because the patient was experiencing constant ectopy. the nurse reported that premature ventricular contractions have been minimal since readjusting the impella. it was further reported that the patient was taken for a washout of the impella site and wound vac placement. impella site was cultured showing klebsiella. the infection is localized and being treated with vancomycin. the patient has been moved to status one on the transplant list. | [5 empty] | The investigation has been completed. no product was returned. as the necessary clinical information was not provided and product was not returned, the root cause of the infection/sepsis, bleeding, and ventricular contractions/ectopic heartbeat was not determined. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
10729 | 1220648-2025-27032 | 21719300 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21719300&pc=OZD | 01/15/2025 | 03/28/2025 |  | 03/28/2025 |  |  | 01/15/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025571949 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 63 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that a patient was implanted with an impella 5.5 pump and began to experience suction events beyond support level p4. as the team wanted to increase support, they attempted to reposition the pump, but it remained obstructed by the papillary muscle. when attempts to free from the papillary muscle failed, the decision was ultimately made to remove/replace the pump. there was no resulting patient harm. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
10730 | 1220648-2025-27034 | 21719333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21719333&pc=OZD | 02/11/2025 | 03/28/2025 |  | 03/28/2025 |  |  | 02/11/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025595571 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that a patient encountered access site bleeding, noted by the oozing, as a result, two sutures were placed and the site was reclosed. the patient is stable and the procedure outcome was successful. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10731 | 1220648-2025-27035 | 21719526 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21719526&pc=OZD | 03/01/2025 | 04/09/2025 |  | 03/28/2025 |  |  | 03/01/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025611753 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella cp support and developed a hematoma and bleeding from the impella access site. fem stop was placed but given the severity of the hematoma, the patient was taken to the operating room and impella cp was removed with vascular repair. the impella cp was exchanged for impella 5.5, but patient expired soon after. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
10732 | 1220648-2025-27036 | 21719600 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21719600&pc=OZD | 03/01/2025 | 04/02/2025 |  | 03/28/2025 |  |  | 03/01/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025595738 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella cp support and developed bleeding at the impella access site. heparin systemic was stopped and heparin through the purge was initiated. the physician placed another stitch to the site and fem stop was applied. bleeding was resolved. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of access site bleeding was most likely patient movement since the hematoma occurred after the patient turning in the bed.
10733 | 1220648-2025-27037 | 21719753 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21719753&pc=OZD | 02/26/2025 | 03/28/2025 |  | 03/28/2025 |  |  | 02/26/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025601371 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs). the patient was in critical condition when the decision was made to implant the impella cp device (pump 1). low pump flow and positioning issue were encountered. impella cp device (pump 1) would be explanted and replaced with another impella cp device (pump 2), which experienced low pump flow as well. the patient would eventually pass away. the patient was admitted to the hospital with complaints of chest pain and shortness of breath. initial test did not show any acute event. therefore, the patient was sent to the unit and returned to the catheterization lab the next day. left and right heart catheterizations were completed. the right heart showed normal pressures. the coronary angiogram revealed a large patent right coronary artery. the left side was severely calcified and revealed left main. the ejection fraction was estimated at 25%. the physician made the decision to try and fix the left main. a guiding catheter was delivered radially, and as the physician attempted to seat the guide, a dissection to the left main artery occurred due to the severe calcification. the patient rapidly declined, requiring intubation, maximum pressors and cardiopulmonary resuscitation (cpr). the decision was then made to implant an impella cp device (pump 1). cpr was being done on the patient for the entirety of the procedure. the impella cp device (pump 1) was implanted, and suction was immediately present on the automated impella controller (aic) as the impella was started in auto. a bolus of saline was given, and the placement of the device was confirmed under fluoroscopy. flows of 1.2 liters per minute were the best that could be obtained due to constant suction events. the physician utilized the 7 french companion sheath/single access to deliver the guide and wire across the left main. angioplasty and stenting (with two drug-eluting stents) of the left main and the left main ostium were done. the dissection was no longer present per the report. however, the patient continued to decline requiring cpr. in an attempt to reposition the impella cp (pump 1), the catheter was unintentionally pulled back across the valve. the device was explanted, removed from the body, and the physician decided to use a new impella cp device (pump 2) to avoid any possible thrombus formation within the previous impella cp device (pump 1) and hopes of better flows. the new impella cp (pump 2) catheter was prepped and placed into the left ventricle in standard fashion. as cpr continued, the flows on the new impella cp device (pump 2) could not rise above 1.2 liters per minute and continued suction was seen. eventually, after all heroic efforts were made, the patient did not recover. the decision was made to discontinue further efforts and the patient expired. medical safety review of the event noted the patient likely expired from the left main artery dissection complication and cardiac condition as the patient was in a state of decline from the very beginning of the impella support and unable to recover. however, given the impella cp device (pump 1) inability to deliver (with delay in) optimal flow for support at onset and malpositioning as well, the impella cp device (pump 1) cannot be dissociated for the event outcome. the impact to the overall outcome is unknown. note: impella cp device pump 2 was implanted as cpr continued. flows could not rise above 1.2 liters per minute. however, impella cp device pump 1 has been associated with the patient's demise (for reason noted above). | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. this is one of two device manufacturing reports associated with this event. refer to initial medical device report for impella cp serial number (b)(6) with date of event 26-feb-2025 and patient identifier ending in (b)(6).
10734 | 1220648-2025-27038 | 21719808 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21719808&pc=OZD | 03/01/2025 | 04/17/2025 |  | 03/28/2025 |  |  | 03/01/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026614325 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a patient was on impella cp support and had dark urine. the impella cp was repositioned. the hemolysis started to clear but the impella cp came into the aorta shortly before the impella explant took place. the patient survived the procedure. | [5 empty] | The investigation of positioning issues and hemolysis has been completed. the impella device was not returned for evaluation. the clinical confirmed impella pulled out to aorta before explanting and hemolysis noticed with 18hrs prior to aorta alarms. from the clinical details confirmed impella position was deep and hemolysis issue resolved after repositioning. data log reviewed: impella position in aorta alarms observed at p-2 to end of the case. a mc spike observed during imeplla in aorta alarms followed by gradually rising in flow to 3l/min at p-2. no purge pressure issue observed. no suctions of placement signal issue seen. the cause of the positioning issue most likely was use related since the impella position was too deep. the cause of the hemolysis issue most likely was improper positioning related since urine cleared after repositioning. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
10735 | 1220648-2025-27039 | 21720238 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21720238&pc=OZD | 02/14/2025 | 03/28/2025 |  | 03/28/2025 |  |  | 02/14/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025603567 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was on impella cp support and during support, the pump had positioning issues at the start of the case. the patient was advised to lay down. the positioning issues were resolved by explanting the pump. the patient survived the event. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10736 | 1220648-2025-27040 | 21720789 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21720789&pc=OZD | 03/01/2025 | 03/28/2025 |  | 03/28/2025 |  |  | 03/01/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025559433 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient experienced complications following impella cp implant. it was initially noted that the staff was unable to optimize positioning of the pump following implant. this resulted in a low placement signal and impella in aorta positioning alarms. the patient also developed hemolysis noted by bloody urine and plasma free hemoglobin levels. the pump was subsequently explanted as a result of the noted issue. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
10737 | 1220648-2025-27041 | 21721007 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21721007&pc=OZD | 02/26/2025 | 03/28/2025 |  | 03/28/2025 |  |  | 03/01/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1155606 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The user facility reported that an automated impella controller (aic) experienced a controller failure alert for three seconds during patient support. there was no patient harm and support continued uninterrupted. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10738 | 1220648-2025-27042 | 21721104 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21721104&pc=OZD | 03/01/2025 | 04/08/2025 |  | 03/28/2025 |  |  | 03/01/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025580259 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  |  | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient had multiple complications during impella 5.5 support. following implant, the patient experienced episodes of ventricular tachycardia. an echocardiogram showed that the pump was positioned against the mitral leaflet prompting repositioning. it was additional noted that a hematoma had developed at the site, resulting in some swelling and pain in the shoulder area. the patient was given a blood transfusion and a washout was performed at the site to manage the condition. support was continued uninterrupted and a decision to bridge was made. | [5 empty] | The investigation of access site bleeding and vf/vt/af/at has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. d6b added. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10739 | 1220648-2025-27043 | 21721138 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21721138&pc=OZD | 01/14/2025 | 04/01/2025 |  | 03/28/2025 |  |  | 01/14/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1291633 |  | 0042-0010-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 39 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported the automated impella controller issued a controller error alarm around 08:56am during patient impella mechanical circulatory support. it was noted there were no issues with motor current. placement and left ventricle signals could be seen on impella connect, and the pump was running without any changes. approximately five minutes later the alarm resolve on its own and did not reoccur until later in the evening at approximately 17:33pm. there is no information that indicates the aic was exchanged. reportedly there was no patient harm and no report or indication of interruption in therapy. | [5 empty] | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. the logs confirmed that on the day of the event, there was a controller error indicated by an out-of-spec power battery manager (pbm) voltage. this was associated with an alarm, which lasted for approximately five minutes after four and a half days. following this incident, the pump operated without any issues for the next two days. the console was returned for investigation. a thorough visual inspection of the internal components was conducted, but no physical problems were found. the power supply from the pbm to the impellatronic pca and the 4-in-1 unit was stable. the console was then tested with a known working pump and purge system, and everything functioned normally with no issues. the root cause of the controller failure alarm (pbm voltage out-of-spec) was not determined as the error was not reproduced. e.4 should have been left blank on the initial manufacturer device report number as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 codes 3233 and 11 were reported incorrectly on the initial manufacturer device report number 1220648-2025-27043 as the investigation was completed. h.10 investigation results was inadvertently omitted on the initial manufacturer device report number. | The automated impella controller has been received from the customer; however, the evaluation of the device is not complete. upon investigation closure, a supplemental mdr will be filed.
10740 | 1220648-2025-27044 | 21721287 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21721287&pc=OZD | 03/01/2025 | 03/28/2025 |  | 03/28/2025 |  |  | 03/01/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025564542 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | Life Threatening | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that a potential clot was noted in the left ventricle during impella cp support for a patient due to high risk cardiac procedure. heparin was added to the intravenous drip and plans were made to explant the pump as soon as the patient's cardiac function allowed. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | B.2 revised to add required intervention that was inadvertently omitted on the initial manufacturer device report 1220648-2025-27044. h.6 added health effect - impact code that was inadvertently omitted on the initial manufacturer device report 1220648-2025-27044.
10741 | 1220648-2025-27045 | 21721445 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21721445&pc=OZD | 03/03/2025 | 05/14/2025 |  | 03/28/2025 |  |  | 03/03/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1889477 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Other |  |  | Manufacturer report |  | Non-Healthcare Professional | Foreign | Company representation |  |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that during training with a demo pump, the automated impella controller experienced a controller failure alarm and subsequently failed to start the attached demo pump. there was no patient involvement during this training. | [5 empty] | The investigation of automated impella controller (aic) - controller failure (red alarm) has been completed. data analysis: the console logs from the reported day of the event align with the complaint received. it was noted that the issue associated with the complaint began on february 4, 2025. during this period, the test pump was connected and reconnected a total of five times. on the fifth reconnection, an error message was recorded, problems while reading pump limits data." this triggered the "impella defective (error reading eeprom)" alarm. after this incident, the pump was subsequently disconnected and reconnected for a sixth time, but the same alarm was triggered again. additionally, a controller error related to a "defective optical sensor lamp" was issued after the pump was disconnected, and this alarm persisted even after the console was rebooted twice. device analysis: the console was returned for further investigation. testing was conducted using the test pump and purge cassette. during this process, the impella defective alarm was triggered after the pump was removed and then reinserted. additionally, a "controller error (optical lamp defective)" alarm was also observed. voltage measurements taken from the power battery manager (pbm) to the impellatronic pca were within the required specifications, and no loose cable connections were found. further the aic was examined for the 5s parameter values; however, the results indicated that these values did not meet the specified criteria as mentioned in the aic preventative maintenance procedure. a visual inspection of the console¿s internal circuitry revealed no apparent issues. root cause: the root cause for the impella defective (error reading eeprom) - alarm issued was due to an impellatronic board malfunction and the controller error (defective optical sensor lamp) - alarm issued was due to optical bench malfunction. | The investigation of automated impella controller (aic) - controller failure (red alarm) has been completed. the console logs from the reported day of the event align with the complaint received. testing was conducted using the test pump and purge cassette. during this process, the impella defective alarm was triggered after the pump was removed and then reinserted. additionally, a "controller error (optical lamp defective)" alarm was also observed. voltage measurements taken from the power battery manager (pbm) to the impellatronic pca were within the required specifications, and no loose cable connections were found. further the aic was examined for the 5s parameter values; however, the results indicated that these values did not meet the specified criteria as mentioned in the aic preventative maintenance procedure. a visual inspection of the console¿s internal circuitry revealed no apparent issues. the root cause for the impella defective (error reading eeprom) - alarm issued was due to an impellatronic board malfunction and the controller error (defective optical sensor lamp) - alarm issued was due to optical bench malfunction. | The automated impella controller was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
10742 | 1220648-2025-27046 | 21721571 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21721571&pc=OZD | 02/26/2025 | 03/28/2025 |  | 03/28/2025 |  |  | 02/26/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025576303 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Cardiac Arrest | Hemodynamic instability | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs), the patient was in critical condition when the decision was made to implant the impella cp device (pump 1). low pump flow and positioning issue were encountered. impella cp device pump 1 would be explanted and replaced with another impella cp device (pump 2), which experienced low pump flow as well. the patient would eventually pass away. the patient was admitted to the hospital with complaints of chest pain and shortness of breath. the initial test did not show any acute event. therefore, the patient was sent to the unit and returned to the catheterization lab the next day. left and right heart catheterizations were completed. the right heart showed normal pressures. the coronary angiogram revealed a large patent right coronary artery. the left side was severely calcified and revealed left main. the ejection fraction was estimated at 25%. the physician made the decision to try and fix the left main. a guiding catheter was delivered radially, and as the physician attempted to seat the guide, a dissection to the left main artery occurred due to the severe calcification. the patient rapidly declined, requiring intubation, maximum pressors and cardiopulmonary resuscitation (cpr). the decision was then made to implant an impella cp device (pump 1). cpr was being done on the patient for the entirety of the procedure. the impella cp device (pump 1) was implanted, and suction was immediately present on the automated impella controller (aic) as the impella was started in auto. a bolus of saline was given, and the placement of the device was confirmed under fluoroscopy. flows of 1.2 liters per minutes (l/min) were the best that could be obtained due to constant suction events. the physician utilized the 7 french companion sheath/single access to deliver the guide and wire across the left main. angioplasty and stenting (with two drug-eluting stents) of the left main and the left main ostium were done. the dissection was no longer present per the report. however, the patient continued to decline requiring cpr. in an attempt to reposition the impella cp (pump 1), the catheter was unintentionally pulled back across the valve. the device was explanted, removed from the body, and the physician decided to use a new impella cp device (pump 2) to avoid any possible thrombus formation within the previous impella cp device (pump 1) and hopes of better flows. the new impella cp (pump 2) catheter was prepped and placed into the left ventricle in standard fashion. as cpr continued, the flows on the new impella cp device (pump 2) could not rise above 1.2l/min and continued suction was seen. eventually, after all heroic efforts were made, the patient did not recover. the decision was made to discontinue further efforts and the patient expired. medical safety review of the event noted the patient likely expired from the left main artery dissection complication and cardiac condition as the patient was in a state of decline from the very beginning of the impella support and unable to recover. however, given the impella cp device (pump 1) inability to deliver (with delay in) optimal flow for support at onset and malpositioning as well, the impella cp device (pump 1) cannot be dissociated for the event outcome. the impact to the overall outcome is unknown. note: impella cp device pump 2 was implanted as cpr continued. flows could not rise above 1.2l/min. however, impella cp device pump 1 has been associated with the patient's demise (for reason noted above). | [5 empty] | The impella cp device was discarded by the customer. the clinical details mentioned the pump was pulled back inadvertently past the valve. the root cause of the positioning issue was most likely use related as evidenced by clinical details. the data logs confirmed low flows for matching p-level. there were suction alarms and an impella position in ventricle alarm. there were no motor current spikes outside p-level changes. the clinical details do not mention any patient anatomy or volume complications. the root cause of the low pump flow was not determined as no product was returned and insufficient clinical information was provided. this is one of two devices associated with the event. refer to manufacturing report number 1220648-2025-27037 for impella cp (pump 2).
10743 | 1220648-2025-27047 | 21721790 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21721790&pc=OZD | 02/28/2025 | 04/14/2025 |  | 03/28/2025 |  |  | 02/28/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025599884 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Unspecified Infection | Tachycardia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella 5.5 support and during support, the patient's access site was warm to the touch, tender and the patient's white blood cells were elevated. the plan was to possibly start on antibiotics. support was continued. additionally, there was difficultly repositioning the pump. it appeared deep and was unable to retract or advance. the doctor stated they felt like something was wrong with the locking mechanism. furthermore, the patient had tachyarrhythmia and was shocked a few times. the pump was still deep and they were unable to reposition. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock & cp with smartassist section: table 3.2 impella catheter components ¿the infusion filter prevents bacterial contamination and air from entering the purge lumen.¿ section: warnings & cautions: warnings ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of positioning issues, infection/sepsis, and vf/vt/af/at has been completed. the cause of the positioning issue, infection, and vt cannot be determined since complaint device not returned for analysis and insufficient clinical data provided.
10744 | 1220648-2025-27050 | 21722053 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21722053&pc=OZD | 07/10/2024 | 03/28/2025 |  | 03/28/2025 |  |  | 07/10/2024 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025501636 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | [17 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the site had bleeding and a washout was done. additionally the patient had 12 units of blood administered and the patient found to have a coagulopathy. during support, the patient also experienced a run of ventricular fibrillation treated with multiple shock/defibrillation. | [5 empty] | The investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post-sterile inspection checks. the root cause of the access site bleeding was most likely patient condition since the patient was coagulopathic. to determine the true root cause of the ventricular tachycardia/ventricular fibrillation (vt/vf), sufficient clinical information would need to be assessed in conjunction with evidence from the returned device. as the necessary clinical information was not provided and the device was not returned, the root cause of the vt/vf was not determined.
10745 | 1220648-2025-27048 | 21722077 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21722077&pc=OZD | 10/26/2023 | 03/28/2025 |  | 03/28/2025 |  |  | 10/26/2023 | 03/28/2025 |  |  | 09/30/2025 |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024397621 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Renal Failure | Ischemia Stroke | [10 empty] | Disability | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported the patient had an impella 5.5 implant and after the implant, same day, there was oozing at the access site. one unit of blood products was given to the patient and mechanical circulatory support (mcs) continued. additionally, a notification was received from abiomed¿s long-term outcome and quality indicator impella registry regarding a patient that endured stroke and renal insufficiency both with a "possible" relationship to the impella device used. approximately, a little over, nine months after end of the impella mcs and the left ventricular assist device (lvad) was implanted, the lvad was explanted and patient received a heart transplant. prior, the hospital course was complicated by vasodilatory/cardiogenic shock requiring extracorporeal membrane oxygenation (ecmo) and vasopressors, stroke, and hematania. the patient transferred to unit around 11:45pm post procedure. pressors weaned off and sedation fully off at 6:30am. team noted left hemiplegia, prompting stroke code. reportedly the stroke was due to an embolism of right middle cerebral artery (mca); stroke with cerebral ischemia. a right mca thrombectomy procedure was completed, and the patient recovered with no sequelae. additionally, during the same timeframe, the patient experienced acute renal insufficiency. the patient put on potassium restrictive diet and monitored and recovered with no sequelae from this event as well. | [5 empty] | The investigation into the reported issue has been completed. no product was received for evaluation. limited clinical information was provided on patient condition and reported failures. the root cause of access site bleeding - major was not determined as no product was returned and insufficient clinical information was provided. the data logs do not show any motor current spikes or trending placement signal. the root cause of the renal failure was not determined as no product was returned and insufficient clinical information was provided. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined.
10746 | 1220648-2025-27051 | 21722087 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21722087&pc=OZD | 10/28/2024 | 03/28/2025 |  | 03/28/2025 |  |  | 10/28/2024 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025546799 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemolysis | Thrombocytopenia | [10 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | A |  |  | No | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during the support there were multiple blood products given and labs drawn for hemolysis. furthermore, the patient had ventricular tachycardia while on support. the pump was repositioned and no further arrhythmia was reported post repositioning of pump. later, a notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient who experienced thrombocytopenia during impella cp support. the loqi noted multiple platelet infusions were administered. aspirin was held and heparin purge solution was changed to bicarbonate. the patient outcome was noted as not recovered/not resolved. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis.¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
10747 | 1220648-2025-27052 | 21722359 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21722359&pc=OZD | 02/15/2025 | 04/14/2025 |  | 03/28/2025 |  |  | 02/15/2025 | 03/28/2025 |  |  | 10/31/2025 |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024419488 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | Pumping Stopped | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. it was reported that the patient inadvertently ripped the impella catheter apart from the red plug. the patient was stabilized and taken to the operating room for impella replacement. there was no patient harm. the patient condition was noted to be stable. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. positioning issues: data logs were reviewed and lt log of case shows pump was running without issue until sudden pump stop alarm occurred. no impella positioning alarms were observed. clinical details noted patient was confused and not restrained when they pulled the impella out of position and tore the pump at the red handle. the root cause of the positioning issues was most likely patient movement as the patient was agitated and pulled on the impella catheter causing mal-positioning and separation of the catheter and the red pump plug. pump stop: an additional failure mode of "pump stop" was added to the investigation as a pump stop alarm was observed in returned logs at time of positioning issues. clinical details noted that when patient was not restrained, they pulled the pump out and ripped the impella catheter apart at the red handle. data logs were reviewed and lt log shows an ¿impella stopped (rpm deviations)¿ alarm was observed at time of reported positioning issues. pump was running without issue prior to alarm. at time of alarm, motor current spiked to 3000 and motor speed dropped to 0, pump flow immediately dropped to 0, placement signal spiked to 3000, and purge pressure dropped with a sudden increase in purge flow. the root cause of the pump stop was most likely due to an electrical open when patient pulled on impella catheter and separated pump at red handle. instructions for use as follows: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist & impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
10748 | 1220648-2025-27053 | 21722819 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21722819&pc=OZD | 11/14/2024 | 03/28/2025 |  | 03/28/2025 |  |  | 11/14/2024 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED. INC. |  |  | US |  | 2025517870 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. there was a noted increase in purge pressure/decrease in purge flow rate. a check was made to see if there were any kinks or other abnormalities, but no abnormalities were found. anew purge set was opened and replaced. after replacement, the purge set operated normally without any issues. the purging set that was deemed defective was discarded, and it was also difficult to retrieve logs. there was no direct impact on impella operation, so the device was weaned and removed. | [5 empty] | The investigation into the reported issue has been completed. data logs were reviewed, and there were no purge related issues or alarms on the reported date (b)(6) . there was one instance of a high purge pressure alarm being triggered on (b)(6) . in this case, there was a rise in purge pressure to 1140 mmhg in less than a minute with a severe drop in purge flows. the purge pressure resolved within 4 minutes without any apparent troubleshooting as there are no indications that the purge cassette or bag were changed during this time in the logs. additionally, there were no motor current spikes leading up to the event. this spike could have been due to a kink/block of flow in the purge line, but clinical details report that no kinks in the purge line were noted and this event isn't on the date of the reported issue. product was not returned for investigation. due to limited clinical details, no product return, and data log review not providing conclusive evidence, the root cause of the high purge pressure could not be determined. as noted in the impella ifu: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
10749 | 1220648-2025-27054 | 21723101 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21723101&pc=OZD | 12/24/2024 | 09/22/2025 |  | 03/28/2025 |  |  | 12/24/2024 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1207017 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | IMPELLA 5.5 | [18 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported that there was an issue where the purge disc was not detected while changing the purge cassette prior to the case. the complainant planned on getting another automated impella controller to resolve the issue. | Additional information indicated that, during patient support, the pump experienced placement signal unreliable alarm. |  |  |  |  | The investigation of automated impella controller (aic) - purge disc/flag issues has been completed. there was no purge disk/flag issue found during the case. the device functioned/operated to specification. d6a and d6b should have been left blank on the initial manufacturer device report number 1220648-2025-27054. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation of the optical signal issue has been completed. console logs from the reported day of event are consistent with complaint. console was tested in danvers but the issue was not reproduced. the pattern of optical bench data - intermittently oscillating contrast - is consistent with temporary loss of light, caused by failure of either optical bench or its lamp assembly. the root cause of the console issue was a defective optical bench. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated with additional information from the initial manufacturer device report number 1220648-2025-27054.
10750 | 1220648-2025-27055 | 21723433 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21723433&pc=OZD | 02/12/2025 | 09/22/2025 |  | 03/28/2025 |  |  | 02/12/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025603479 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Stroke/CVA | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed a impella 5.5 pump to support the patient along with extracorporeal membrane oxygenation. it was reported that the patient had an embolic stroke on support. no further details were provided. the procedural outcome was the patient survived surgery. | The complainant also reported that the pump alarmed for a false in aorta position on the automated impella controller. the issue self-resolved. neither the pump nor the console were replaced. |  |  |  |  | The investigation has been completed. no product was returned. as this clinical information was not provided, the true root cause of the stroke/cva could not be determined. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. no product was returned. the data logs were evaluated and the optical 5s parameters were within specification and no hard failures of the optical sensor were seen. there were no placement signal not reliable (psnr) alarms seen. there were no abnormalities with the pumps function. the root cause of the optical signal issue is most likely due to the patient¿s condition based on data log analysis and clinical information provided. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b.1. product problem updated. b.5. updated to include optical signal issue description. corrections that were made from the initial manufacturer device report number 1220648-2025-27055. d.10. concomitant medical products and therapy dates updated.
10751 | 1220648-2025-27056 | 21723807 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21723807&pc=OZD | 10/28/2024 | 03/28/2025 |  | 03/28/2025 |  |  | 10/28/2024 | 03/28/2025 |  |  | 11/30/2025 |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024426115 |  | 1000100 |  | No |  |  | No |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Stroke/CVA | Hemorrhage/Blood Loss/Bleeding | [10 empty] | Death | Required Intervention |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, the patient experienced a drop in hemoglobin. the patient reportedly had a gastrointestinal bleed (gib) and two units of packed red blood cells were administered to the patient. heparin was in the purge prior to the gib. a notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient who experienced ventricular tachycardia (vt) during impella cp support. the device tip was found to be shallow, so it was advanced in position it was noted that the patient recovered with no sequelae. additionally, the patient experienced an acute left anterior cerebral artery cerebral vascular accident with a "unknown (undetermined)" relationship to the impella device used. it was noted that the patient was not moving bilateral lower extremities and right upper extremity. computed tomography of the patent's head and neck was done which revealed left common carotid artifact versus occlusion. it was also noted that the patient did not qualify for acute stroke therapy. the patient's stroke deficits showed no improvements. additional information noted care was withdrawn and the patient expired. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
10752 | 1220648-2025-27058 | 21723830 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21723830&pc=OZD | 11/18/2024 | 03/28/2025 |  | 03/28/2025 |  |  | 11/18/2024 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025534557 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 41 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Death |  |  |  |  | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | A notification was received via abiomed's long-term outcome and quality indicator (loqi) impella registry regarding a patient who experienced limb ischemia during impella cp support. the patietn experienced a loss of pulses in the right lower extremity and it required emergent placement of a distal perfusion cannula. it was reported that care was eventually withdrawn due to the patient's neuro status. there is not enough information to exclude the impella device as an associated factor in the patient's demise. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system: section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation."
10753 | 1220648-2025-27057 | 21723922 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21723922&pc=OZD | 01/23/2025 | 07/02/2025 |  | 03/28/2025 |  |  | 01/23/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025589371 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 47 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | No Clinical Signs, Symptoms or Conditions | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was on impella 5.5 as a bridge to therapy. during support there multiple suction alarms, which were resolved after giving bolus. it was also reported that the optical sensor failed, trending motor current which was stable with associated p level. the team opted to exchange pump and the patient survived the event. | During support it was noted that heparin was off due to bleeding from internal jugular line placement and teeth to be pulled later in the day. low platelets so plan to leave heparin off for the day to minimize bleeding from teeth pull. there was also multiple suction alarms, which were resolved after giving bolus. additionally, the nurse called about "air in purge alarm" they de-aired system twice with no resolution, then urgently changed purge cassette and alarms resolved. it was also reported that the optical sensor failed, trending motor current which was stable with associated p level. the team opting to exchange pump. |  |  |  |  | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of optical signal issue, low pump flow, and vascular damage - peripheral vessel has been completed. no data logs were returned. optical signal issue: there were multiple ao adjustments through support and towards the last day of the case, the placement signal (ps) was seen to be drifting down and then it went out and triggered placement signal not reliable (psnr) alarms. the product was returned and the light exited the sensor face and produced psnr alarms when connected to the lab console. the 5s parameters were stable with a slightly lower snr but the ps was completely out. the factory and calibrated g0 were 16078 and the cavity length on the returned product was 17642. based on the clinical details provided, and the returned product analysis, the root cause of the optical signal issue is most likely due to sensor wear. ¿low pump flow: the clinical states that there were multiple suction alarms throughout the case that resolved by adding volume or turning the p-level down. the returned product had no indications of thrombus or a cannula kink. the pump ran and low flow was not able to be reproduced. based on the returned product and the clinical details, the root cause of the low pump flow is most likely due to the patients condition. vascular damage - peripheral vessel: due to a lack of clinical information, the root cause of the vascular damage - peripheral vessel cannot be determined. d9 device returned to manufacturer date added b1 adverse event was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27057. b2 selection was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27057. b5 additional information was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27057. h1 type of report was erroneously selected on the initial manufacturer device report number 1220648-2025-27057. h6 health effect - clinical code 1888 and health effect - impact code 4641 were inadvertently omitted on the initial manufacturer device report number 1220648-2025-27057.
10754 | 1220648-2025-27059 | 21723937 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21723937&pc=OZD | 12/13/2024 | 03/28/2025 |  | 03/28/2025 |  |  | 12/13/2024 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025570007 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | Infusion or Flow Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. during support, there was an increase in motor current and the pump stopped. the pump was restarted with purge pressure in range. there was intermittent saturated flow with increased motor current. it was recommended to exchange the pump however support continued. the physician wanted to do ultrasound to diagnose a possible clot. there was no information on ultrasound to diagnose a possible clot. | [5 empty] | The investigation into the reported issue has been completed. no product was received for analysis. device history record review confirmed that the pump passed all post-sterile inspection checks. the data logs confirm impella stopped motor current high alarms. the logs show there was a motor current spike with speed deviation and shift in purge pressure at the time of temporary pump stop. the logs show the pump restarted and stopped again before restarting with saturated flow, flowing at higher-than-expected range for p7. there were also suction alarms. it is unknown if biomaterial was observed on explant. the root cause of the temporary pump stop was most likely due to biomaterial as evidenced by ingestion pattern in the logs. the root cause of the saturated flows was most likely due to biomaterial as evidenced by ingestion pattern in the logs as noted in the impella ifu: impella cp with smartassist for use during cardiogenic shock. section: using the automated impella controller with the impella catheter, ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressur, which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing the impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure which triggers the ¿purge pressure high¿ alarm message could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped, ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
10755 | 1220648-2025-27060 | 21723996 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21723996&pc=OZD | 12/18/2024 | 09/24/2025 |  | 03/28/2025 |  |  | 12/18/2024 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025574871 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Difficult to Remove | Device Sensing Problem | [5 empty] | Thrombosis/Thrombus | [12 empty] | Other | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. upon explant, there was difficulty noted with pump removal. the physician noted that it felt stuck. the patient's chest was re-opened and pump started to pull back. the physician retracted the graft back into open chest and cut the white cable out and the pump was removed via graft. the physician stated that the graft was little hardened and dry clots were sitting at anastomosis site. the patient outcome was noted as critical. the patient continued on extracorporeal membrane oxygenation (ecmo) support. | The complainant also reported that there was loss of optical signal. |  |  |  |  | The investigation into the reported issues has been completed. an inadequate pump was received for evaluation. it was found to be missing the pump plug. based on the clinical details, the root cause of the removal issue is most likely due to the patients condition as there was clots along the graft. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed on optical signal issue: the pump was returned, and it was cut at the catheter. no optical testing could be performed. looking at the 5s parameters on the data logs, the gain and contrast are slightly spiking with the placement signal throughout the case. on the 17th, the placement signal spikes to 3000, gain decreases, light decreases, contrast oscillates, and lamp was at 100. it remains out for the rest of support. this trend does not align with a specific root cause. since insufficient product was returned and no clear pattern was identified, the root cause of the optical signal issue cannot be determined. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number. b7 other relevant history, including preexisting medical conditions updated.
10756 | 1220648-2025-27061 | 21724310 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21724310&pc=OZD | 03/02/2025 | 04/14/2025 |  | 03/28/2025 |  |  | 03/02/2025 | 03/28/2025 |  |  | 01/31/2026 |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025464722 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 14 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | Hematuria | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a cp pump placed to support a 14-year-old male pediatric patient admitted in with congenital history and postop from the fontan surgery. the cp was placed due to cardiogenic shock. the pump positioning was poor at the ventricle wall due to the native anatomy and congenital history. all measures could not rectify the positioning. the pump was placed however there was signs of hemolysis. there was hematuria and reduced urine output. the pump was explanted after 42 hours. the patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smartassist system stated the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ | The investigation of hemolysis and positioning issues has been completed. the pump was in improper position but was unable to reposition of the pump away from ventricular wall due to patient's anatomy (congenital heart defect with dorv- double outlet right ventricle). it was unclear if hemolysis was caused by improper position of the pump or other causes since insufficient information were provided. the product and data were not returned for review. the root cause of hemolysis was unable to be determined as limited clinical details were provided and no product was returned for review. the root cause of positioning issue was most likely patient condition related since the patient's anatomy (congenital heart defect with dorv- double outlet right ventricle) causes the pump was unable to be repositioned away from ventricular wall. as noted in the impella instructions for use: impella cp with smartassist system section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-27061 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 reporting contact facility name and fax number were inadvertently omitted from the initial manufacturer device report number 1220648-2025-27061 and were added. h.6 code 2558 was inadvertently omitted from the initial manufacturer device report number 1220648-2025-27061 and was added. h.10 incomplete instructions for use were reported on the initial manufacturer device report number 1220648-2025-27061. the remaining portion of the instructions for use was added.
10757 | 1220648-2025-27062 | 21724317 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21724317&pc=OZD | 03/03/2025 | 05/02/2025 |  | 03/28/2025 |  |  | 03/03/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025611049 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Male |  |  |  | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 55-year-old male patient on an impella 5.5 for cardiomyopathy. it was reported that high purge pressure was encountered roughly 18 days into impella 5.5 support. tissue plasminogen activator was added to the purge, but the issue persisted. a lysis protocol was discussed, but the decision was ultimately made to replace the pump. there was no patient harm. | [5 empty] | The investigation of high purge pressure has been completed. the pump was returned and evaluated. there was some biomaterial in the purge gap and on the impeller. no other abnormalities were found. high purge pressure did not reproduce during 30 minutes of purging and running the pump for a portion of the time. purge pressure remained steady around 500 mmhg. there was also no blood ingress found in the purge line. on 02-mar-2025, field data logs show purge pressure began to gradually build up over the next 18 hours to 800 mmhg. then, there was quicker rise within 20 minutes from 900 mmhg to greater than 1000 mmhg. there were no related motor current spikes and therefore, no evidence of ingestion. the purge pressure did not resolve by the end of the case. the root cause of the high purge pressure was most likely biomaterial buildup since the data logs show a gradual rise in purge pressure. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27062 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smartassist states the following: section: using the automated impella controller with the impella catheter. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
10758 | 1220648-2025-27063 | 21724320 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21724320&pc=OZD | 03/03/2025 | 05/02/2025 |  | 03/28/2025 |  |  | 03/03/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026617663 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  |  | White | Injury | Yes | Yes | Restricted Flow rate | Material Twisted/Bent | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 57-year-old female patient with impella cp implant for high-risk percutaneous cardiac intervention. it was reported that the flows steadily decreased and would not go above 1.4 l/m. the scheduled pci was completed successfully and follow up angiograms discovered that the catheter was kinked near the outlet on the cannula. the pump was explanted as a result of the noted issue. upon removal, significant oozing persisted at the site and an additional perclose was required to achieve hemostasis. the patient survived the procedure. | [5 empty] | The device was received for investigation on 4/29/2025 d9 revised to reflect date. the results of the investigation did not change from previous manufacturer device report number 1220648-2025-27063-1. the investigation results revealed low flow occurred and kinked cannula was noted under angiography. oozing occurred while removal of the pump. failure mode product damage (cannula kink) was added since a cannula kink was reported. data was not returned for review. product analysis: visual inspection found a kink on can about 15 mm from of cage and can bonding site. no other issues were found on the rest of the pump. conclusion: the root cause of low flow was most likely kinked cannula since the clinical team confirmed kink was found near the outlet on the cannula under fluoro after percutaneous coronary intervention (pci) procedure and under product analysis. the root cause of access site bleeding was unable to be determined as limited clinical details were provided and no problem was found on the pump. the root cause of cannula kink was unable to be determined as limited clinical details of how the kink occurred were provided. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of low pump flow and access site bleeding has been completed. product and data were not returned for review. based on clinical description, the root cause of low flow was most likely kinked cannula since the clinical team confirmed kink was found near the outlet on the cannula under fluoroscopy after percutaneous coronary intervention (pci) procedure. the root cause of access site bleeding was unable to be determined as limited clinical details were provided and no product was returned for review. the root cause of cannula kink was unable to be determined as limited clinical details were provided and no product was returned for review. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27063 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27063.
10759 | 1220648-2025-27064 | 21724328 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21724328&pc=PYX | 03/03/2025 | 04/15/2025 |  | 03/28/2025 |  |  | 03/03/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025540893 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Sense | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that a 62 year old female patient was on impella rp flex support and during support, there was no waveform for the pa present and the impella flows were missing as well. the only thing that was present on the screen was the motor current which was in accordance to the p level we were at. the decision was made to remove this pump and exchange for a new one. the new one was placed and flowing without issue. the patient survived the procedure. | [5 empty] | The investigation of optical signal issue has been completed. the data log review confirmed the placement signal issue. both the dp and optical sensor were working, measuring reasonable pressure values until one minute before the pump was started up. at that time, the dp signal suddenly went out of range, triggering alarms. since this occurred before the pump started up, flow calculations were disabled during the entire time the pump was running. at the time the dp signal went out of range, all the pressure measurements based on the dp also went out of range. the optical signal remained stable for the entire case. there were very short intermittent periods where pap and cvp calculations would return during the case, but always railed again. the returned product was investigated and there were several abnormalities. light exited the optical sensor face in the laser continuity test, however, when connected to the console all 5s parameters were out of specification. additionally, the peak signal-to-noise ratio alarm reproduced. the dp sensor failed electrical tested, as both lines were open. when visually inspecting the dp and optical sensor faces, both showed abnormalities. there was cracking on the optical sensor face and a skive mark on the silicone of the dp sensor. since data logs only showed alarms and invalid measurements related to the dp sensor during the case, the failure mode was changed to placement signal issues. based on data log review and returned product, the root cause of the placement signal issue was determined to most likely be an electrical open. the cause of the damage to the sensor is unknown as limited clinical details were provided. d.1 revised brand name as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-27064. d.4 revised model number as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-27064. e.1 added fax number and us phone number as they were inadvertently omitted from the initial manufacturer device report number 1220648-2025-27064. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-27064 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 added reporting contact facility name and fax number as they were inadvertently omitted from the initial manufacturer device report number 1220648-2025-27064. h.6 investigation codes 3221 and 4315 previously entered incorrectly on the initial manufacturer device report number 1220648-2025-27064 as the device was returned for evaluation. codes 3233 and 11 should of been reflected on the initial manufacturer device report number 1220648-2025-27064. investigation findings found that code 2917 was incorrect on the initial manufacturer device report number 1220648-2025-27064. code 1559 was added to reflect what was determined during investigation results. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.
10760 | 1220648-2025-27065 | 21724332 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21724332&pc=OZD | 03/03/2025 | 03/28/2025 |  | 03/28/2025 |  |  | 03/03/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026620908 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that the patient on continuous renal replacement therapy and pacer dependent with asystole was implanted with impella cp for left ventricle support. it was reported that multiple times the patient experienced ventricular tachycardia, requiring defibrillation. the patient remained on support until successfully weaned and explanted. | [5 empty] | The impella pump was not returned for investigation. should new information be received, a supplemental manufacturer device report will be filed.
10761 | 1220648-2025-27066 | 21724337 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21724337&pc=OZD | 03/03/2025 | 04/09/2025 |  | 03/28/2025 |  |  | 03/03/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026619102 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 87 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported 87-year-old male patient had no deplorable pulse in their right leg following impella cp implant. the patient's leg was cold to the touch at the time of the observation. the patient was monitored for two days on support, at which point the family decided to withdraw care. the patient passed away. there is not enough evidence to disassociate the patient's passing and the use of the impella. | [5 empty] | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. b.2 intervention was incorrectly reported on the initial manufacturer device report number 1220648-2025-27066. b.7 added information as it was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27066. e.1 revised us phone number as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-27066. added fax number as it was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27066. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-27066 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 added reporting contact facility name and fax number as they were inadvertently omitted on the initial manufacturer device report number 1220648-2025-27066. h.6 code 925 was reported incorrectly on the initial manufacturer device report number 1220648-2025-27066 and a new code was added. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)."
10762 | 1220648-2025-27067 | 21724348 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21724348&pc=OZD | 12/21/2024 | 04/03/2025 |  | 03/28/2025 |  |  | 03/03/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025546804 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 50 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complaint team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a 50 year old male patient that endured right axillary hematoma with a "definite" relationship to the impella device used: ¿patient developed large hematoma with associated pain and swelling at impella site (right axillary). required emergent axillary exploration and hematoma evacuation.¿ the patient recovered with no sequelae. additionally, the complaints team received a notification via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured ventricular tachycardia with a "possible" relationship to the impella device used: ¿multiple runs of vt. impella needed repositioning and vt resolved.¿. the patient recovered with no sequelae and bridged to transplant. | [5 empty] | The investigation of access site bleeding and ventricular tachycardia has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding cannot be determined since the product was not returned and insufficient clinical details were provided. as the necessary clinical information was not provided and the device was not returned, the root cause of the ventricular tachycardia was not determined. b.3 date of event was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-27067. e.1 added fax number and us phone number as they were inadvertently omitted from the initial manufacturer device report number 1220648-2025-27067. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-27067 as it is unknown if the initial reporter also sent the report to the food and drug administration. g.1 added reporting contact facility name and fax number as they were inadvertently omitted from the initial manufacturer device report number 1220648-2025-27067. g.2 report source: study was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27067 and was added. h.6 code 1888 was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-27067 and a new code was added to health effect - clinical code. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10763 | 1220648-2025-27068 | 21724351 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21724351&pc=OZD | 03/04/2025 | 06/02/2025 |  | 03/28/2025 |  |  | 03/04/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US | [5 empty] | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Male |  |  |  | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the impella cp stopped when performing an automated impella controller (aic) to aic transfer. the impella did not restart again even when it was reconnected to the original aic. the patient was reported to be stable on extracorporeal membrane oxygenation (ecmo). | [5 empty] | The investigation of automated impella controller (aic) - pump stop has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27068 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella console was not returned for investigation. should new information be received, a supplemental manufacturer device report will be filed. d4 product information is unknown.
10764 | 1220648-2025-27069 | 21724358 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21724358&pc=OZD | 03/04/2025 | 04/14/2025 |  | 03/28/2025 |  |  | 03/04/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1124473 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Read Input Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a patient was to undergo high risk percutaneous coronary intervention and was implanted with impella cp for mechanical circulatory support. the automated impella controller (aic) failed to recognize the pump. the pump was switched to a different aic and support was successfully initiated and used without issue for the duration of the procedure. no patient harm was reported. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of pump not detected has been completed. the data logs confirmed the reported failure mode. this could be most likely due to the crystal issue on the impellatronic (epm) board. device analysis: this console was tested and the reported failure mode "the console could not recognize the pump connection" was not reproduced. the console was tested with a known good pump and purge cassette, and it functioned as expected. the root cause for not recognizing the pump connection was likely an impellatronic malfunction. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27069 as it is unknown if the initial reporter also sent the report to the food and drug administration.
10765 | 1220648-2025-27070 | 21724361 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21724361&pc=OZD | 03/04/2025 | 03/28/2025 |  | 03/28/2025 |  |  | 03/04/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026622738 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 36 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the patient developed compartment syndrome. distal perfusion catheter to left femoral was placed and left leg fasciotomy was performed. the patient remained on support for an additional 2 days. the patient was reported as stable. | [5 empty] | The impella pump was not returned for investigation. should new information be received, a supplemental manufacturer device report will be filed.
10766 | 1220648-2025-27071 | 21724364 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21724364&pc=OZD | 03/04/2025 | 03/28/2025 |  | 03/28/2025 |  |  | 03/04/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025593723 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  |  | White | Death | Yes | Yes | Pumping Stopped | Infusion or Flow Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. it was reported that the following the implant, oozing was noted at the impella site. the heparin drip was temporarily halted to manage the condition. ten days into support, the motor current suddenly increased and saturated flow was encountered. the pump stopped after flows could not be maintained. the patient decompensated and the care was withdrawn at the request of the family. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. | [5 empty] | The impella pump was discarded. should new information be received, a supplemental manufacturer device report will be filed. instructions for use as follows impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention & impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
10767 | 1220648-2025-27073 | 21724471 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21724471&pc=OZD | 03/05/2025 | 03/28/2025 |  | 03/28/2025 |  |  | 03/05/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1392605 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | BIOMEDICAL ENGINEER | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant from biomedical team reported that automated impella controller (aic) displayed a failure alarm. there was no patient involvement. | [5 empty] | The impella console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
10768 | 1220648-2025-27049 | 21724923 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21724923&pc=OZD | 01/22/2024 | 03/28/2025 |  | 03/28/2025 |  |  | 01/22/2024 | 03/28/2025 |  |  | 11/30/2025 |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024427547 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 49 YR |  | Male |  | Hispanic |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient (unknown race) was implanted with an impella 5.5 device as an elective placement. during procedure the patient experienced ventricular tachycardia resolved by repositioning. post implant, same day, the patient's hemoglobin dropped from 8.2 to 7.9 requiring one unit of red blood cells and additionally, the patient experienced "slight" ectopy. the patient was ruled in for non-st-segment elevation myocardial infarction. cardiac catheterization showed triple vessel disease with 100% occlusion of all three main vessels filled by collaterals and a follow up transesophageal transthoracic echocardiogram showed an ejection fraction of 25% and normal right ventricle. the patient was scheduled for coronary artery bypass graft (cabg) surgery with impella 5.5 support. after placement of the pump, the patient had an episode of ventricular tachycardia possibly due to the pump interacting with the papillary muscle. the pump was repositioned and the ventricular tachycardia resolved. the cabg was completed, and the patient was transferred to the unit in stable condition. later, post implant same day, it was noted the patient's hemoglobin dropped from 8.2 to 7.9, and one unit of red blood cells was given. the cause of this event was unknown, but graft site bleeding could not be ruled out. additionally, approximately four after start of support, the performance level was reduced to p-6 due to slight ectopy noted due to the cannula interaction with the septal wall; however, there was no further issue since the flow reduction. support continued. after four days of support, the patient was successfully weaned, and the device was explanted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10769 | 1220648-2025-27072 | 21725095 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21725095&pc=OZD | 03/04/2025 | 05/19/2025 |  | 03/28/2025 |  |  | 03/05/2025 | 03/28/2025 |  |  |  |  | 03/28/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025598978 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | Company representation |  |  |  | 84 YR |  | Male |  |  |  | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Embolism/Embolus | No Clinical Signs, Symptoms or Conditions | [36 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a patient undergoing high risk percutaneous coronary intervention (hrpci) was implanted with an impella cp device for mechanical circulatory support. it was reported that shortly after the implant, on auto-mode with a performance of about 3.5 l/min the parameters showed a flow of 0.1 l/min and a motor current of about 300 ma. the pump was repositioned to no avail. after three tries of increasing the flow by elevating the p-level, the team decided to explant the pump. since the procedure had been almost finished, the team decided not to implant a second impella. the patient was stable, and no harm was reported. | [5 empty] | The impella pump was not returned for investigation. should new information be received, a supplemental manufacturer device report will be filed. | The investigation of low pump flow has been completed. the returned cp pump visually inspected. biomaterial observed wrapped around impeller and purge gap. no broken blade or cannula kink observed. the biomaterial was soaked and removed to run test the pump where low flows were not reproduced. data log reviewed: flows, motor current stable until sudden drop in flow with no motor current (mc) spike observed. suction alarm with impella in ventricle alarms. clinical states repositioning pump and giving blood volume did not resolve flow issue. low pump flow: the cause of the low pump flow issue was most likely biomaterial as observed on the returned product. failure mode thrombus added as a result of biomaterial observation. as the necessary information was not provided, the root cause of the thrombus was not determined. d9 date device returned to manufacturer was inadvertently omitted on the initial manufacturer device report number 1220648-2025-27072. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27072 as it is unknown if the initial reporter also sent the report to the food and drug administration. h3 was erroneously selected on the initial manufacturer device report number 1220648-2025-27072 as the device had been received. h6 type of investigation 4114 as erroneously entered on the initial manufacturer device report number 1220648-2025-27072. h10 narrative erroneously stated the device was not returned on the initial manufacturer device report number 1220648-2025-27072.
10770 | 1220648-2025-26989 | 21706121 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21706121&pc=OZD | 03/02/2025 | 07/21/2025 |  | 03/27/2025 |  |  | 03/02/2025 | 03/27/2025 |  |  |  |  | 03/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1098587 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 55 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Pumping Stopped | Defective Device | [5 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an automated impella controller (aic)/impella cp support ongoing and the patient was undergoing percutaneous coronary intervention for the thrombotic occlusion of the coronary. the console failed during the support and was smelling burnt. the console was replaced and support continued on the second aic. the patients condition continued to deteriorate over several hours on the unit post implant and the patient was made do not resuscitate and expired shortly after. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - damage during therapy has been completed. data analysis: console logs from the day of the incident align with the complaint, showing multiple motor error messages after approximately 1 hour and 16 minutes of pump operation. the pump stopped at (b)(6) 2025 at 18:55:06. the console was turned off at around 18:58:12, and the aic was exchanged. device analysis: the console was sent back for further analysis. during the investigation, a burnt capacitor was found on the impellatronic board, which led to the pump stop. upon examining the aic, it was noted that the lcd for battery 1 was not displaying any charge indication, and the power battery manager (pbm) was found to be defective, as it was unable to charge the other battery (batt 2) as well. this was not related to the reported complaint. after replacing the impellatronic pca, pbm pca and both batteries with a known-good, the aic operated correctly with both the pump and purge cassette connected, functioning without any issues during 2 hours of testing. root cause: the reported complaint of the waveform began to disappear and flow drops (i.e. pump stop) was due to a component malfunction (failed tantalum capacitor) on the impellatronic board. no issue related to the reported failure mode of "aic - damage during therapy" was seen within the aic.
10771 | 1220648-2025-26990 | 21706191 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21706191&pc=OZD | 03/01/2025 | 03/27/2025 |  | 03/27/2025 |  |  | 03/01/2025 | 03/27/2025 |  |  |  |  | 03/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025598906 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on impella cp support ongoing for cardiogenic shock. after just under 27 hours of impella cp support, the patient self explanted. the patient was brought back to the cardiac catheterization lab. the plan is to leave explanted. the patient survived the procedure. additional post event information was received. the patient is alive, stable, extubated and doing fine without any support. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
10772 | 1220648-2025-26991 | 21706434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21706434&pc=OZD | 03/11/2024 | 03/27/2025 |  | 03/27/2025 |  |  | 03/03/2025 | 03/27/2025 |  |  |  |  | 03/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1484655 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The customer reported a patient was on impella cp support due to a st-elevated myocardial infarction. it was reported that while doing groin dressing, a team member noticed that the placement signal had disappeared from automated impella controller (aic). after a few minutes, the placement signal returned. the patient was stable throughout. the medical doctor was made aware of the issue and no changes to the plan were made. upon investigation completion, the engineer found a loss of optical data. | [5 empty] | The investigation has been completed. console logs from the reported day of event were consistent with the complaint and showed placement signal disappearing and later returning (and associated with opm comm crc error). the issue described in the complaint was confirmed and determined to have been caused by optical bench firmware communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the automated impella controller software operated as designed. the cause of the placement signal loss was due to optical bench firmware communication protocol anomaly.
10773 | 1220648-2025-26992 | 21706814 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21706814&pc=OZD | 02/13/2025 | 03/27/2025 |  | 03/27/2025 |  |  | 03/03/2025 | 03/27/2025 |  |  |  |  | 03/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025594062 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was implanted with an impella cp due to acute myocardial infarction. the patient was unable to maintain circulation due to intraperitoneal bleeding. the bleeding was likely caused by injury during insertion of the extracorporeal membrane oxygenation or impella. there were no problems with the functionality of the pump. a blood transfusion was performed as a treatment for the bleeding but the amount of blood transfused is unknown. the patient expired on support. the cause of death was multiple organ failure due to bleeding.the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was discarded by the customer and therefore, an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed.
10774 | 1220648-2025-26993 | 21706956 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21706956&pc=OZD | 03/03/2025 | 04/11/2025 |  | 03/27/2025 |  |  | 03/03/2025 | 03/27/2025 |  |  |  |  | 03/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1889477 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 63 YR |  | Female |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported an automated impella controller (aic) was being use with an impella cp to a support a patient admitted with acute myocardial infarction. it was reported that the console displayed "impella position not reliable". values were in range but aic was switched to the backup. the wrror was resolved after switching to the backup aic. there was no patient harm reported. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - loss of opm data has been completed. the data log confirmed the complaint date, placement signal was confirmed to be unreliable. the console was returned for investigation and a 5s was performed with very low signal not reliable (snr). the console was cleaned but the snr was still quite low. root cause: the cause of the unreliable placement signal was due to the low snr optical bench.
10775 | 1220648-2025-26994 | 21707118 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21707118&pc=OZD | 03/03/2025 | 04/19/2025 |  | 03/27/2025 |  |  | 03/03/2025 | 03/27/2025 |  |  |  |  | 03/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026615986 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with an impella 5.5 for a cardiac procedure. it was reported that multiple complications were encountered during impella 5.5 support. five days into support, an impella in aorta alarm appeared. an echocardiogram was performed and verified that the device was in proper positioning despite the ongoing alarm. as support continued, the flow rate was also observably lower. with the patient deteriorating, the decision was made to replace the pump. the patient survived the procedure. | [5 empty] | The investigation of optical signal issue and low pump flow has been completed. optical signal issue: data log review confirmed false impella position alarms after motor current (mc) spike occurred. the cause of the issue was most likely biomaterial due to data log review showing an mc spike occurring before positioning alarms occur. low pump flow: data log review shows period of low flow occurs immediately preceding an mc spike. after the spike, low flow does not occur the remainder of the case. the cause of the issue was most likely biomaterial due to data log review showing mc spike following the period of low flow after which time the low flows are corrected. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10776 | 1220648-2025-26995 | 21707603 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21707603&pc=OZD | 03/02/2025 | 04/30/2025 |  | 03/27/2025 |  |  | 03/02/2025 | 03/27/2025 |  |  |  |  | 03/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025603927 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 61 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Pumping Stopped | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump placed to support a patient admitted in with cardiogenic shock and cardiomyopathy. after the impella cp was placed the impella cp limb had ischemia. the pump remained on for 12 days and then a pump stop was reported. the pump was explanted and the patient survived. | [5 empty] | The investigation of limb ischemia and pump stop has been completed. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. the pump had been running for 292 hours. however, without product returned, the root cause of the pump stop cannot be confirmed. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella rp with smartassist system catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella rp with smartassist system catheter. in this case, the motor is no longer being purged and may eventually stop. monitor motor current and consider replacing impella rp with smartassist system catheter whenever a rise in motor current is seen.¿ impella cp with smartassist for use during cardiogenic shock section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
10777 | 1220648-2025-26996 | 21708120 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21708120&pc=OZD | 12/10/2024 | 03/28/2025 |  | 03/27/2025 |  |  | 12/23/2024 | 03/27/2025 |  |  | 09/30/2025 |  | 03/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024409919 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Hospitalization | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | A notification was received via abiomed's long-term outcome and quality indicator impella registry regarding a patient that experienced hemolysis with a "definite" relationship to the impella device used. the patient with an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support, and one day after start of support, the patient developed hemolysis. the impella cp device placed supported the patient to recovery; the patient recovered from the event with no sequelae. the following day the patient was successfully weaned and the impella device was explanted with the patient having survived. per the study hospitalization was prolonged as a result of the event. no further information was provided. | [5 empty] | The impella device was not received from the customer and therefore, an evaluation of the device was not possible. upon investigation closure, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. the pump was not returned for investigation. the total case run time was 24 hours. data logs showed a trending placement signal during the case run. also motor current spike was noted at the end of the case run, with pump flow saturation. the cause of the hemolysis issue was most likely biomaterial. b.7 revised as information was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26996. e.1 added fax number as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26996. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26996 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.6 codes 3233 and 11 were entered incorrectly on the initial manufacturer device report number 1220648-2025-26996 as the investigation was completed. new codes have been added. h.10 revised as investigation results were inadvertently omitted on the initial manufacturer device report number 1220648-2025-26996.
10778 | 1220648-2025-26997 | 21708444 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21708444&pc=PYX | 02/02/2025 | 03/27/2025 |  | 03/27/2025 |  |  | 02/03/2025 | 03/27/2025 |  |  |  |  | 03/27/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025578207 |  | 1000323 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 73 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The user facility reported a patient with an acute myocardial infarction/cardiogenic shock implanted with an impella rp flex device experienced hemolysis and would later expire. the patient was admitted in with right ventricular dysfunction. the impella rp flex was implanted, the patient was placed on continuous renal replacement therapy and was subsequently transferred to the tertiary center for care. after approximately four days on support, the patient started having pink-tinged urine and labs came back hemolyzed; a chest x-ray the following day confirmed good impella rp flex placement. however, the hemolysis event was not resolved. the following day impella alarming purge system blocked; echocardiogram showed some improvement to right ventricle. however, the decision was made to remove the device and the patient expired. further details surrounding the event including treatment for the hemolysis along with the cause of death are unknown. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10779 | 1220648-2025-26998 | 21708687 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21708687&pc=OZD | 12/12/2024 | 04/14/2025 |  | 03/27/2025 |  |  | 12/12/2024 | 03/27/2025 |  |  |  |  | 03/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025557295 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Tachycardia | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a patient with cardiomyopathy was implanted with an impella 5.5 device for mechanical circulatory support (mcs) while pending transplant. the patient experienced intermittent low pump flow, thrombus and ventricular tachycardia leading to replacement of the device. approximately seven days after the start of impella mcs, low flows were reported intermittently on the pump. the patient was given hydralazine for high mean arterial pressure. again, five and seven days later, the pump had low flows, and the patient was given volume. the patient continued support. now 26 days later, biomaterial was noted near outflow and tissue plasma activator was added to the purge solution. thereafter, 15 days later, the pump had more low flows with ventricular tachycardia which prompted, along with implant duration, the plan to change out the device two days later. the impella 5.5 device was explanted and replaced without incident. the patient would continue impella mcs while awaiting a heart transplant. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the analysis, a supplemental report will be filed. instructions for use for the related event are as follows: impella 5.5 with smartassist system: section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: "in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ | The investigation of low pump flow, thrombus, vf/vt/af/at and has been completed. the returned complaint 5.5 device visually inspected. biomaterial observed around impeller and purge gap. no broken blade or cannula kink observed. pump flow confirmed after biomaterial removed. biomaterial removed after resolved it in tergazyme, and pump connected to aic and run for 10 mins at p-9. pump flow was within specified limits and no low pump flow reproduced. pump support was ongoing for 2 months while patient on transplant list. the cause of the low pump flow issue was biomaterial ingestion. as the necessary information was not provided, the root cause of the thrombus was not determined. as the necessary information was not provided, the root cause of the vt/vf was not determined. a5 race was erroneously selected on the initial manufacturer device report number 1220648-2025-26998. b2 product problem was inadvertently not selected on the initial manufacturer device report number 1220648-2025-26998. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26998 as it is unknown if the initial reporter also sent the report to the food and drug administration.
10780 | 1220648-2025-26999 | 21709536 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21709536&pc=OZD | 02/28/2025 | 04/09/2025 |  | 03/27/2025 |  |  | 03/01/2025 | 03/27/2025 |  |  |  |  | 03/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1470310 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 78 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the automated impella controller (aic) had placement signal issue while in use in the intensive care unit. the aic was exchanged, and alarm and placement signal was resolved. there was no patient harm. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of optical signal issue has been completed. the console was received for investigation. the console logs confirmed the reported issue. the device analysis was performed and the root cause for the controller error and loss of placement signal is due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic sw operated as designed.
10781 | 1220648-2025-27000 | 21710182 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21710182&pc=OZD | 03/02/2025 | 05/12/2025 |  | 03/27/2025 |  |  | 03/02/2025 | 03/27/2025 |  |  |  |  | 03/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025519154 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported that upon inserting the impella cp pump successfully in a a patient, it was recognized that the staff had never progressed past the start up screen to initiate the purge sequence. the pump was removed and replaced as a result of the issue. there was no noted patient harm. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. | The investigation of priming issue has been completed. no product was returned. the log review showed initial pump connection. there were no purge-related alarms in the logs. data log review showed purge pressure at 0 while placement signal showed the pump had been inserted in the patient, meaning that the pump had been placed before priming the pump. the clinical details say staff never progressed past start up screen to initiate purge sequence. the root cause of the priming issue was most likely use-related since clinical details say that the pump was inserted before progressing to the purge steps on the screen, which was confirmed with data logs.
10782 | 1220648-2025-27001 | 21710306 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21710306&pc=OZD | 03/04/2025 | 04/19/2025 |  | 03/27/2025 |  |  | 03/04/2025 | 03/27/2025 |  |  |  |  | 03/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC |  |  | US |  | 2025606973 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient was to undergo high-risk percutaneous coronary intervention and was implanted with an impella 5.5 pump for mechanical circulatory support. it was reported that the pump stopped unexpectedly during patient support. attempts to restart the device were unsuccessful and the pump was explanted. the patient was reported to be stable, and no patient harm was reported. | [5 empty] | The investigation of pump stop has been completed. no product was returned. log review shows the pump stop ~308 hours into the case with multiple failed restart attempts as reported in the clinical details. alarm 119 (impella stopped) triggered, indicating an electrical failure. the root cause of the pump stop was most likely an electrical open based on the alarm in the data logs but it is unknown how the failure occurred due to insufficient clinical details. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: direct aortic insertion ¿maintaining act at or above 250 seconds will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ impella 5.5 with smartassist section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ impella 5.5 with smartassist for use during cardiogenic shock section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
10783 | 1220648-2025-27002 | 21710349 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21710349&pc=OZD | 03/01/2025 | 03/27/2025 |  | 03/27/2025 |  |  | 03/01/2025 | 03/27/2025 |  |  |  |  | 03/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026612831 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that after successfully gaining access to the femoral artery, the physician was unable to advance the guidewire through the patient's vasculature for impella cp delivery. it was believed that the guidewire was in the false lumen of the aorta post dissection repair. as a result of the noted resistance, the decision was made to abort the implant. there was no patient harm. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella cp with smartassist system section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
10784 | 1220648-2025-27003 | 21710520 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21710520&pc=OZD | 12/21/2024 | 09/26/2025 |  | 03/27/2025 |  |  | 12/21/2024 | 03/27/2025 |  |  |  |  | 03/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025561736 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Decrease in Pressure | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that there was a low purge flow. the low purge flow was encountered and the purge fluid was changed to tissue plasminogen activator as a result, which increased the purge flow. the patient survived the event. | The complainant also reported placement signal issues. additionally, there was an active unreliable placement signal alarm which was disabled. |  |  |  |  | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist section: using the automated impella controller with the impella catheter ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue and high purge pressure. no product was returned for evaluation. data logs were reviewed, optical 5s parameters were consistent with an oscillating contrast issue. upon review of the data logs, it is apparent that the console is the potential cause of the issue. no problem was found with the pump during the case. no motor current spikes were observed at the start of the purge pressure rise. the root cause of the high purge pressure was most likely due to biomaterial buildup based on data log analysis. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-27003. b7 other relevant history, including preexisting medical conditions updated. e4 should have been left blank on the initial report. d10 concomitant medical products and therapy dates updated.
10785 | 1220648-2025-27004 | 21710574 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21710574&pc=OZD | 11/10/2024 | 03/27/2025 |  | 03/27/2025 |  |  | 12/25/2024 | 03/27/2025 |  |  | 11/30/2025 |  | 03/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2024434228 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 39 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Tachycardia | Aortic Valve Insufficiency/ Regurgitation | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | A notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured ventricular tachycardia (vt) with a "possible" relationship to the impella device used: ¿per operative note on 11/10/24: finally, the patient was in a stable vt rhythm for which we performed external electrical cardioversion.¿. the patient recovered with no sequelae. eleven days later, another notification was received via abiomed¿s long-term outcome and quality indicator (loqi) impella registry regarding a patient that endured aortic regurgitation with a "probable" relationship to the impella device used: ¿there was moderate eccentric aortic valve insufficiency that appeared after impella removal and as such i felt aortic valve replacement was necessary given this. impella was powered down and pulled back across the aortic valve into the aorta. at this point, we noticed that there was moderate aortic valve insufficiency and that aortic valve replacement would be mandatory.¿. the patient recovered with no sequelae. | [5 empty] | The investigation has been completed. the product was not returned. the cause of the heart valve damage was not determined due to lack of clinical details and no product returned. as the necessary information was not provided, the root cause of the ventricular tachycardia was not determined. as noted in the impella instructions for use: impella 5.5 with smart assist for use during cardiogenic shock section: potential adverse events (unites states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or cardiac or vascular injury (including ventricular perforation) (including ventricular perforation).¿ section: axillary insertion of impella 5.5 catheter ¿use fluoroscopy for placement impella catheter performance will be compromised if correct placement cannot be confirmed. while other imaging techniques, such as transesophageal echocardiography (tee), can help confirm the position of the impella catheter after placement, tee does not allow visualization of the entire catheter assembly and is inadequate for reliably placing the impella catheter across the aortic valve¿ section: warnings & cautions: warnings: section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis ¿unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death.¿ in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ section: use of echocardiography for positioning of impella catheter ¿if the inlet area of the impella catheter is too close to or entangled in the papillary muscle and/or subannular structures surrounding the mitral valve, it can affect mitral valve function and negatively impact catheter flow.¿ section: understanding and managing impella catheter position alarms. ¿if the impella catheter is completely out of the ventricle, do not attempt to reposition the catheter across the valve without a guidewire.¿
10786 | 1220648-2025-27005 | 21710710 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21710710&pc=OZD | 02/18/2025 | 03/27/2025 |  | 03/27/2025 |  |  | 03/01/2025 | 03/27/2025 |  |  |  |  | 03/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025584219 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that a patient was on impella cp support and expired from a left ventricle rupture. the left ventricle rupture was found after impella insertion. it is unclear whether impella insertion caused the problem or not. the relationship between the cause of death and the impella device is unknown. | [5 empty] | The impella device was discarded by the customer and therefore,¿an evaluation of the device was not possible. should any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation: section: axillary insertion of the impella 5.5 catheter: section: direct aortic insertion: section: use of impella with transcatheter aortic valves: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿
10787 | 1220648-2025-27006 | 21712762 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21712762&pc=OZD | 12/27/2024 | 03/27/2025 |  | 03/27/2025 |  |  | 12/27/2024 | 03/27/2025 |  |  | 12/31/2025 |  | 03/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025449546 |  | 1000459 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump could not be inserted due to calcification in the patient's vasculature. the implant was subsequently aborted and an impella cp device was placed for planned support. there was no resulting patient harm. | [5 empty] | The investigation has been completed. the product was not returned for review. based on clinical description, the root cause of delivery issue was most likely due to patient condition as the patient had severe calcified vessels. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿
10788 | 1220648-2025-27007 | 21712903 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21712903&pc=OZD | 01/14/2025 | 03/27/2025 |  | 03/27/2025 |  |  | 01/14/2025 | 03/27/2025 |  |  |  |  | 03/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025576538 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant had placed an impella 5.5 pump to support a patient admitted in with known cardiomyopathy and planning for transplant list. the team observed a possible purge leak. fluid was leaking from the purge housing on the impella. no observable active leak was seen but droplets of fluid were noted inside the plastic housing and droplets of fluid could be felt occasionally from the bottom of the housing. there was no replacement. the housing moisture was being monitored. the fluid was not sticky as would be expected with dextrose so it is unsure if this is an active leak or not. no harm or injury was reported. the pump remained on for just under 19 days til heart was transplanted. | [5 empty] | The investigation has been completed. the pump and purge cassette were returned for evaluation. the data logs were returned and there were no purge alarms seen. the pressure and purge flow were stable throughout support. the returned product was inspected and the catheter was cut and small droplets were seen in the housing. the purge line was reattached using a luer and needle tip connection to be able to purge the system. it was purged for approximately four days and purge alarms occurred and there were no leaks seen inside the sidearm. the sidearm was cut and it was pressurized with a syringe to be able to exacerbate the leak if there was one. no leaks were noted. images were taken under the keyence and no damages or buildup of fluid could be seen on the device. when changing the purge tubing to the sidearm, it is a wet-to-wet connection. this can lead to fluid into the sidearm. the root cause of the purge leak - pump is most likely due to use as the issue was not seen in the data logs and was not able to be reproduced in the lab.
10789 | 1220648-2025-27008 | 21712928 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21712928&pc=OZD | 01/26/2025 | 09/22/2025 |  | 03/27/2025 |  |  | 01/26/2025 | 03/27/2025 |  |  |  |  | 03/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025589664 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 67 YR |  | Male |  |  |  | Injury | Yes | Yes | Increase in Pressure | Device Sensing Problem | [5 empty] | Thrombocytopenia | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported the patient was on impella 5.5 support and during support, there were high purge pressures. lysis protocol was initiated. support continued. furthermore, the patient developed heparin-induced thrombocytopenia and was being treated systemically with argatroban. support continued until weaning was successful and the procedural outcome was the patient survived surgery. | The complainant also reported placement signal issues. the pump¿s position was confirmed, and support continued. |  |  |  |  | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. as noted in the impella instructions for use: impella 5.5 for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The root cause of the thrombocytopenia cannot be determined as insufficient clinical details were provided. in the data logs it can be seen that there are some downward ticks in the purge pressure and leading up to high purge pressure. purge pressure then begins to rise on 1/26 with no reported kinks in the line. it takes ~7 hours for the purge pressure to trigger a hpp alarm. it remains high for about 10 hours. the purge cassette was changed and there were no further ticks seen for the remaining of the case. the pc was changed from (b)(6). during this change, tpa was also added to the purge and said to resolve the issue. the purge pressure stabilized and there is a gap in the data logs as the console was changed. on 2/17 there is one more increase in pp over 4 minutes that lasts for 40, and then the purge pressure is stable for the rest of the data logs provided till 2/24. the rest of the data logs to 3/11 were not given. the cassette that was returned was not the cassette used at the provided rise. the returned product reproduced high purge pressure and would not run as there was biomaterial in the gap, and ingress into the purge line. since the data logs show the purge within specification until the 24th after it resolved on 1/26 and 2/17 and there were no further reported purge events, it cannot be determined what the root cause of the high purge pressure event was on 1/26. due to lack of logs returned and insufficient information provided, the root cause of the high purge pressure could not be determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27008 as it is unknown if the initial reporter also sent the report to the food and drug administration. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the 5s parameters are stable throughout the case. the returned product showed a significant kink in the fiber near the red pump plug. a light continuity test was performed and light leaked form that location. the root cause of the optical signal issue is most likely due to a damaged fiber line. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include optical signal issue description. h6 updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-27008. b7 other relevant history, including preexisting medical conditions updated. d10 concomitant medical products and therapy dates updated. h3 device evaluated by manufacturer updated.
10790 | 1220648-2025-27009 | 21712936 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21712936&pc=OZD | 01/26/2025 | 03/27/2025 |  | 03/27/2025 |  |  | 01/26/2025 | 03/27/2025 |  |  |  |  | 03/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025572125 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant had an impella cp pump placed to support a patient with cardiogenic shock/acute myocardial infarction and other underlying comorbidities. the impella cp was placed and was supporting with signs of hemolysis. creatinine was increasing and possible continous renal replacement therapy (crrt) to support the patient through acute state. urine output was less than 30 milliters per hours and tea colored. crrt was being seting up to be started. the pump was explanted after two days and then the patient was supported by an impella 5.5 pump. | [5 empty] | The investigation has been completed. no product was returned. per the clinical investigation, data log were reviewed: no mc spike outside of the p-level changes observed. no positioning issue seen. no suctions or placement signal trend observed. good device position confirmed while hemolysis issue reported. unclear if hemolysis resolved. the cause of the hemolysis issue cannot be determined since complaint device not returned for analysis and insufficient clinical data provided. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
10791 | 1220648-2025-27010 | 21712944 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21712944&pc=OZD | 12/31/2024 | 10/01/2025 |  | 03/27/2025 |  |  | 12/31/2024 | 03/27/2025 |  |  |  |  | 03/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025536661 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 80 YR |  | Male |  |  |  | Injury | Yes | No | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient was brought back to the catheterization lab for explant of the impella cp because the aneurysm site was oozing. the physician placed a long-covered stent for closure of both sites and post dilated for a good effect. the procedural outcome was the patient survived surgery. | The complainant reported that the device exhibited an optical sensor issue. |  |  |  |  | The investigation has been completed for access site bleeding. clinical information noted that there was an aneurysm and oozing at the access site. hemostasis was achieved with a long covered stent. the root cause of the access site bleeding was not determined since insufficient clinical details were provided. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. no product was returned. data logs were only available for the first 25 hours of the case, which includes the time of the complaint. log review showed no abnormalities with the 5s optical parameters. there were placement signal low alarms leading up to the positioning alarms along with a slight downward trend in ps. by the time of the alarms, the pump was running at p-4 or below. the majority of the alarms occurred while on p-2. impella in aorta and impella in ventricle were alternating while the pump was on p-2. no motor current spikes were seen. motor current leading up to the position alarms was not as pulsatile as it was initially. the root cause of the optical signal issue was most likely patient condition related. a review of the device history record (dhr) for the suspect product, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. h6: added optical signal issue "2917" as part of a retrospective review of optical signal issues in accordance with fda recommendation.
10792 | 1220648-2025-27011 | 21712950 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21712950&pc=OZD | 01/02/2025 | 09/22/2025 |  | 03/27/2025 |  |  | 01/02/2025 | 03/27/2025 |  |  |  |  | 03/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025573396 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Ischemia | [12 empty] | Required Intervention | [5 empty] | AUTOMATED IMPELLA CONTROLLER | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant also reported that the pump lost the optical signal but this did not prompt any pump replacement. | The complainant had an impella cp pump placed to support a acute myocardial infarction/cardiogenic shock patient. the team observed signs of limb ischemia a day after implant. the leg accessed by the impella cp had no dopplerable pulses. the team placed a distal perfusion. the patient remained stable and the procedural outcome was the patient survived surgery. |  |  |  |  | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the pump and data log were provided for investigation. no physical damage was observed on the returned pump, which passed the laser continuity test of the optical signal line. the pump's optical and placement signal functions performed as per specifications, with no issues reported during the test run. the data log indicates that the signal-to-noise ratio (snr) was low from the start of the case, randomly dropping to zero while the placement signal spiked to 3000, triggering a placement signal not reliable (psnr) alarm. the snr drop and placement signal trend were independent of the p-level change and any patient condition. additionally, the pumps were investigated before and after the run and were found to have a similar low snr trend during the case. upon reviewing the data logs and returned product, no problems were found with the pump and it is apparent that the console could be the potential cause of the issue. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b.1. product problem updated. b.5. updated to include optical signal issue description. h.6. updated to include optical signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-27011. d.10. concomitant medical products and therapy dates updated. e.4. erroneously entered as "no.". | The investigation has been completed. the pump and data log were provided for investigation. no physical damage was observed on the returned pump. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿.
10793 | 1220648-2025-27012 | 21712960 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21712960&pc=OZD | 01/06/2025 | 09/09/2025 |  | 03/27/2025 |  |  | 01/06/2025 | 03/27/2025 |  |  |  |  | 03/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025589257 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Tachycardia | [11 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a patient with impella cp support developed bleeding at the right axillary site. pressure dressing was applied, and heparin was held for 48 hours. bleeding was resolved. three days later, the patient had a 40-beat run of ventricular tachycardia and was converted. the patient was back in paced rhythm. the patient was made a do not resuscitate and expired on impella support. there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of access site bleeding and vf/vt/af/at has been completed. the impella device was not returned for evaluation. access site bleeding: the root cause of access site bleeding was most likely anticoagulation management since hemostasis was achieved by stopping the heparin. vt/vf: the root cause of the vt/vf was unable to be determined due to insufficient clinical details. b2 intervention was inadvertently not selected on the initial manufacturer device report number 1220648-2025-27012. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27012 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact name prefix was erroneously entered on the initial manufacturer device report number 1220648-2025-27012.
10794 | 1220648-2025-27013 | 21713590 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21713590&pc=OZD | 01/13/2025 | 04/02/2025 |  | 03/27/2025 |  |  | 01/13/2025 | 03/27/2025 |  |  | 09/30/2025 |  | 03/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024402525 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 85 YR |  | Male |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Low Blood Pressure/ Hypotension | [9 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that following impella cp implant, a patient experienced bleeding at the triple lumen and impella access site. a large hematoma was noted at the impella site. the patient had a drop in blood pressure and central venous pressure. manual pressure was held and five units of packed red blood cells with two units of fresh frozen plasma were given to counteract the blood loss. heparin was discontinued in response to the bleed and femstop was applied just above the access site. with the site redressed and surgiseal applied, the bleeding was contained. support was continued following the noted intervention. however, the patient succumbed to the their condition and passed away the following day. it was noted that the patient had begun to show signs of hemolysis prior to their passing for which support levels were turned down. it is unknown if altering the support level resolved the hemolysis. it is unknown if the noted issues contributed to the patient's passing. | [5 empty] | The investigation for access sit bleeding has been completed. the pump was returned for inspection. upon inspection, no abnormalities were noted with the pump nor the catheter. the root cause of the access site bleeding was determined to be anticoagulation management because the partial thromboplastin time values (200) were higher than therapeutic values and the heparin was discontinued and femstop was placed to control the bleeding. as noted in the impella instructions for use: impella cp with smart assist system, section: general patient care considerations, ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | E.4 should have been left blank on the initial manufacturer device report number 1220648-2025-27013 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.1 type of reportable event was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-27013.
10795 | 1220648-2025-27014 | 21713739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21713739&pc=OZD | 02/28/2025 | 04/23/2025 |  | 03/27/2025 |  |  | 02/28/2025 | 03/27/2025 |  |  |  |  | 03/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025582265 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 69 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Positioning Failure | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that impella cp was placed and due to left ventricle (lv) orientation, the cp had to be deeper in the lv causing hemolysis from the outlet/atrioventricular interaction. the physician attempted to reposition several times under echo unsuccessfully. no further information provided. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter. ¿evaluate the position of the impella catheter if the automated impella controller. displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle. ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis. ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis and positioning issues has been completed. the product and data logs was not returned for review. based on clinical description, the root cause of hemolysis was most likely due to pump was not in the proper position since the pump was placed deeper due to left ventricle (lv) orientation. the root cause of the positioning issue was most likely patient condition related since the pump was placed deeper purposely due to unfavorable lv orientation. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-27014 as it is unknown if the initial reporter also sent the report to the food and drug administration.
10796 | 1220648-2025-27015 | 21713748 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21713748&pc=OZD | 03/01/2025 | 03/27/2025 |  | 03/27/2025 |  |  | 03/01/2025 | 03/27/2025 |  |  |  |  | 03/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1205056 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Failure to Read Input Signal | [6 empty] | Insufficient Information | [12 empty] | Death | [24 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported the purge disc was not detected when changing purge cassette. after several attempts to slide it in, they switched to a new automated impella controller and there was no further issues. at day 55 of pump support care was withdrawn and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation has been completed. the automated impella controller was returned for investigation. the logs from the complaint date revealed that the console issued purge disc not detected alarm and occurred multiple times. the physical console was examined and found to have a broken purge disc flag. the purge disc flag was observed to be broken and was the root cause of the reported purge disc not detected alarm issue.
10797 | 1220648-2025-27016 | 21713777 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21713777&pc=OZD | 01/31/2025 | 03/27/2025 |  | 03/27/2025 |  |  | 03/01/2025 | 03/27/2025 |  |  |  |  | 03/27/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025584412 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 74 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that a patient on impella cp support was observed to have a cerebral hemorrhage. the patient later expired due to the cerebral hemorrhage. it was believed that a potential relationship existed between the cause of death and the device because activated clotting time (act) management was performed, but poor control of other coagulation systems was observed. the act values were controlled at 160-180. activated partial thromboplastin time was over 300. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10798 | 1220648-2025-26957 | 21695869 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21695869&pc=OZD | 02/28/2025 | 04/08/2025 |  | 03/26/2025 |  |  | 02/28/2025 | 03/26/2025 |  |  |  |  | 03/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025605549 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 60 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Thrombosis/Thrombus | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella cp support due to acute myocardial infarction. during support, there was bleeding at the impella site. heparin was withheld and support was continued. additionally, upon explant, a clot was evacuated from the iliac artery. the impella cp was placed in the femoral artery so the relationship to the clot and the impella is unknown. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system. section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output. use of the repositioning sheath and peel-away introducer. ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: impella cp with smart assist catheter insertion. section: axillary insertion of the impella cp with smart assist catheter. section: use of impella with transcatheter aortic valves. ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation of access site bleeding and thrombus has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding is determined to be anti coagulation management because the bleeding was resolved by stopping the heparin. as the necessary information was not provided, the root cause of the thrombus was not determined.
10799 | 1220648-2025-26958 | 21695944 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21695944&pc=OZD | 02/28/2025 | 04/11/2025 |  | 03/26/2025 |  |  | 02/28/2025 | 03/26/2025 |  |  |  |  | 03/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025598936 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella cp support and during support, there were positioning issues. the pump was explanted and replaced with another impella cp due to not being able to reposition the first pump. the patient survived the procedure. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. | The investigation of positioning issues has been completed. the root cause of the positioning issues cannot be determined since complaint device not returned for analysis and insufficient clinical data provided.
10800 | 1220648-2025-26959 | 21696002 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21696002&pc=OZD | 02/28/2025 | 04/17/2025 |  | 03/26/2025 |  |  | 02/28/2025 | 03/26/2025 |  |  |  |  | 03/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026631319 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 50 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Priming Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump being prepped for support when the pump failed to zero and could not be troubleshot. the pump was replaced and support proceeded without harm or injury. | [5 empty] | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of priming issue has been completed. per review of data log analysis and returned product evaluation, no problem was found during the case.
10801 | 1220648-2025-26960 | 21696048 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21696048&pc=OZD | 02/26/2025 | 04/11/2025 |  | 03/26/2025 |  |  | 02/26/2025 | 03/26/2025 |  |  |  |  | 03/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | GM |  | 1691229 |  | 0042-0040-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Select Signal | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported during the implant of an impella 5.5 device to a patient with cardiomyopathy for mechanical circulatory support (mcs), the impella 5.5 pump was displayed to be defective after being plugged into two different automated impella controllers (aic). controller alarm, "pump not detected" occurred. consequently, a new impella 5.5 device and a different aic were utilized for successful implant and start of mcs. there was no report or indication of patient harm as a result of the event. | [5 empty] | The automated impella controller device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. this is one of two devices associated with the event. refer to medical device report for the impella 5.5 serial number (b)(6) with event date 26-feb-2025 and patient identifier ending in (b)(6). | The investigation of pump not detected has been completed. on the complaint date, after the pump was initially connected via the ¿case start¿ wizard, the console displayed ¿impella defective (error reading eeprom) ¿ alarm,¿. the pump was reconnected several times, and the console was also power cycled manually several times. however, event was persistent, confirming the reported failure mode. after the third power cycle, the pump serial number was recorded. however, the console still displayed the event. the console was tested and was unable to reproduce the impella defective alarm using two different test pumps. both light emitting diodes on the impellatronic were on, and the 20 volt and 5 volt supply to the impellatronic board was good. the ribbon cable between the impellatronic board and the 4-in-1 was appropriately connected, and no hardware issues were found on the console. the root cause of the ¿impella defective¿ alarm was not determined due to the inability to reproduce.
10802 | 1220648-2025-26961 | 21696078 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21696078&pc=OZD | 02/19/2025 | 05/15/2025 |  | 03/26/2025 |  |  | 02/26/2025 | 03/26/2025 |  |  |  |  | 03/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1336513 |  | 0042-0000-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 55 YR |  | Female |  | Non Hispanic | Asian | Malfunction | No | Yes | Failure to Read Input Signal | Use of Device Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that upon inserting the cassette and purge reservoir into the automated impella controller (aic), the priming screen did not advance. the components were removed and reinserted multiple times, but the issue persisted. the aic was exchanged and priming was completed without issue. there was no resulting patient harm as a result of this event. | [5 empty] | The automated impella controller device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) ¿ purge issue - other has been completed. data analysis: on the complaint date (b)(6) 2025, in the ¿case start¿ wizard ¿ step 3, although there was an epm reset, the pump connection was not recognized. the console was then displayed ¿connect grey connectors screen¿ which likely led the customer to claim that the ¿priming screen did not proceed.¿ there was no issue related to the purge system was found in the logs. the ¿case start¿ wizard was reinitiated, however, the pump recognition issue persisted even after the epm reset. device analysis: this console was tested after arriving and it was able to recognize three different optical test pumps. unable to reproduce the reported failure mode. root cause: the pump detection issue was due to the rarely occurring epm crystal issue. the impellatronic pca does not have permanent damage, but this intermittent pump detection failure occurs infrequently and already has a sw mitigation in place.
10803 | 1220648-2025-26962 | 21696143 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21696143&pc=OZD | 02/26/2025 | 03/26/2025 |  | 03/26/2025 |  |  | 02/26/2025 | 03/26/2025 |  |  |  |  | 03/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1146769 |  | 0042-0000-US |  | No |  |  | No |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Pumping Stopped | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | [9 empty] | The complainant reported the automated impella controller was supporting a patient with an acute myocardial infarction/cardiogenic shock who was implanted with an impella cp device for mechanical circulatory support (mcs) when the pump stopped. following the implant, flow saturation and high purge pressure were observed. the pump subsequently stopped while components were being gathered for troubleshooting, replacement. the pump was removed and a new impella cp device was implanted for continuation of therapy. the patient was reported to be in stable condition following the event. | [5 empty] | The automated impella controller device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. the is one of two devices associated with the event. refer to the initial medical device report for the impella cp device sn (b)(6) with date of event 26-feb-2025 and patient identifier ending in (b)(6).
10804 | 1220648-2025-26963 | 21696243 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21696243&pc=OZD | 02/26/2025 | 04/09/2025 |  | 03/26/2025 |  |  | 02/26/2025 | 03/26/2025 |  |  |  |  | 03/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025605954 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hematuria | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and developed hemolysis the following day post implant and would subsequently expire. the patient presented with difficulty breathing and shortness of breath and was found to be in atrial fibrillation. catheterization was completed and revealed 90 percent occlusion of the left anterior descending, which was treated with balloon angioplasty and stenting. the decision was made to place an impella cp postoperatively and continue with transfer out to a higher level of care facility. after the pump was placed the following day there were signs of hemolysis. the patient had a traumatic foley at the first hospital before arriving to the tertiary center. the signs of hemolysis observed were hemolyzed labs and elevated plasma hemoglobin. the performance level for support was decreased. two days later, sedation was turned off. the patient was not responding, and status remained unchanged now four days later after start of mcs (still no patient response). hemolysis was again noted. the performance level had been increased, and urine and sputum became bloody. the performance level was reduced, to manage the complication. however, care was withdrawn, and the patient would expire this same day. although care was withdrawn there is not enough information to exclude the impella cp as an associated factor the event. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system, section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ | The investigation of hemolysis has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined.
10805 | 1220648-2025-26964 | 21696625 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21696625&pc=OZD | 02/26/2025 | 09/17/2025 |  | 03/26/2025 |  |  | 02/26/2025 | 03/26/2025 |  |  |  |  | 03/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025507090 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Detachment of Device or Device Component | [6 empty] | No Clinical Signs, Symptoms or Conditions | [18 empty] | 10MM VASCUTEK GELWEAVE GRAFT | [18 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella 5.5 device as elective placement in the setting of coronary artery bypass graft (cabg) surgery for mechanical circulatory support. the impella 5.5 was inserted direct to the aorta, and a 10mm vascutek gelweave graft was tunneled to right supraclavicular space. post cabg the patient was noted to be doing well. the patient continued on support for approximately two days per the care plan. at the time of explant minimal resistance was encountered, and upon removal of impella 5.5 from graft, only the motor housing was connected to the catheter with remainder of cannula inside the graft. the cannula could be felt inside the graft, and a pair of locking forceps was utilized to secure the cannula and remove from graft. back bleed of the graft performed. the graft was sutured and tucked under skin. the patient remained hemodynamically stable during procedure and was noted to be in stable condition following the event. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the analysis, a supplemental report will be filed. | The investigation for the reported product damage (detached inflow/outflow - great vessel) has been completed. there are not enough details surrounding the removal of the device to make a determination if excessive force was used to remove the product. the product had a clamp mark but it was stated that that was from retrieving the cannula. additionally, there was residual glue at the bond of the cannula and motor, indicating there was an intact bond. due to a lack of sufficient evidence during the removal of the 5.5, the root cause of the product damage (detached inflow/outflow) cannot be determined. corrections to the initial mfr report: h3 has been updated to device evaluation completed. h6 type of investigation code 10 added. | The investigation of cannula/pigtail detachment - great vessel/heart has been completed. the failure mode (fm) was changed product damage (detached inflow/outflow) to reflect the location of the cannula detachment. due to a lack of sufficient evidence during the removal of the 5.5, the root cause of the product damage (detached inflow/outflow) cannot be determined.
10806 | 1220648-2025-26965 | 21696782 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21696782&pc=OZD | 02/26/2025 | 04/29/2025 |  | 03/26/2025 |  |  | 02/26/2025 | 03/26/2025 |  |  |  |  | 03/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026618667 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the insertion of impella cp device to a patient undergoing a high-risk percutaneous coronary intervention was unsuccessful due to the anatomy. the physician was unable to advance the impella cp device due to the vascular anatomy. the case was aborted. the status of the patient following the event was noted to be unknown. however, there was no report or indication of any adverse effects to the patient. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details, the root cause of delivery issue is determined to be patient condition because the pump was unable to pass through the vasculature due to calcification in the vessels.
10807 | 1220648-2025-26966 | 21696795 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21696795&pc=OZD | 02/28/2025 | 04/10/2025 |  | 03/26/2025 |  |  | 02/28/2025 | 03/26/2025 |  |  |  |  | 03/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025583847 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 83 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support developed bleeding from the impella cp access site. the medical doctor placed a mattress suture and red blood cells were received. the patient remained stable and on support. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of access site bleeding was most likely use issue since the anticoagulant other than heparin or sodium bicarbonate was administered to the patient which lead to the bleed.
10808 | 1220648-2025-26967 | 21696916 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21696916&pc=OZD | 02/26/2025 | 03/26/2025 |  | 03/26/2025 |  |  | 02/26/2025 | 03/26/2025 |  |  | 01/31/2026 |  | 03/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025532864 |  | 0048-0047 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 61 YR |  | Male |  |  |  | Death | Yes | Yes | Restricted Flow rate | Malposition of Device | [5 empty] | Cardiac Arrest | Hemodynamic instability | [11 empty] | Required Intervention | Life Threatening | Death | [22 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient (unknown race and ethnicity) with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and experienced low pump flow and device malpositioning. the patient went into cardiac arrest and expired. the patient presented to hospital with four-day history of significant chest pain and found having an acute myocardial infarction. the patient was taken to the catheterization found an occluded left anterior descending artery, right coronary artery and 80% occluded left circumflex artery with extensive left ventricular wall ischemia and decreased right ventricular function. during the percutaneous coronary intervention, the patient started exhibiting shock signs and the decision made to implant an impella cp device. the implant was uncomplicated, and the patient was taken to the unit on mcs. over the next six hours the patients clinical state deteriorated, needing fluid resuscitation, inotropic support. the patient was determined to be unsuitable for ecmo, lvad or heart transplant given the presentation and clinical state. the patient had increasing suction alarms due to volume and right ventricular function and as the patient worsened. the patient had an episode of hematemesis resulting in excessive movement and malpositioning of the impella cp device into the ventricle. attempts were made to reposition impella cp pump under echocardiogram; however, the patient subsequently arrested with a rhythm of asystole. after 30 minutes of resuscitation attempts including rsi, the patient had never regained a heart rhythm and decision was made to withdraw care. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)."
10809 | 1220648-2025-26968 | 21697169 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21697169&pc=OZD | 02/28/2025 | 03/26/2025 |  | 03/26/2025 |  |  | 02/28/2025 | 03/26/2025 |  |  |  |  | 03/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026624504 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support and after the implant, the flows were unable to go above 1l. flows were much lower than expected so the device had to be switched out for a new cp which provided much better flows. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10810 | 1220648-2025-26969 | 21697210 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21697210&pc=OZD | 03/02/2025 | 05/05/2025 |  | 03/26/2025 |  |  | 03/02/2025 | 03/26/2025 |  |  |  |  | 03/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025530953 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 70 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support, the patient had limb ischemia. the pump was explanted and the leg appeared normal after explant. the procedural outcome was the patient survived the surgery. | [5 empty] | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | Reported insufficient leg perfusion. patient with known pvad. pump was weaned and explanted 5 hours later. leg looked better after explant. a cannula kink was observed on the returned product. no abnormalities were observed. the investigation of limb ischemia has been completed. limb ischemia can occur during impella support. when making a determination as to the cause of the limb ischemia, the following clinical information is necessary: patients pre-morbid condition and peri-procedural condition. any concomitant medication/ vascular size or variations thereof. detailed description of the symptoms experienced by the patient. physical examination findings, including pulse assessment and visual examination of the affected limb. diagnostic imaging results, such as doppler ultrasound, angiography, or magnetic resonance angiography. laboratory test results, including blood tests for markers of inflammation or clotting disorders. any relevant information about risk factors for peripheral arterial disease, such as smoking, diabetes, or hypertension. patient's response to previous treatments or interventions for vascular conditions. with a returned impella, the max od could be assessed to ensure it is within specification. the product could also be evaluated for any burrs or sharp edges. additionally, the clinical information above would need to be considered in the context of the returned product. to determine the true root cause of limb ischemia, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as all the necessary information was not provided, the root cause of the limb ischemia was not determined.
10811 | 1220648-2025-26970 | 21697430 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21697430&pc=OZD | 02/28/2025 | 03/26/2025 |  | 03/26/2025 |  |  | 02/28/2025 | 03/26/2025 |  |  |  |  | 03/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025523035 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 64 year old female patient with an impella cp being inserted for support during cardiogenic shock, liver shock, and septic shock, was found to have no distal flow to the leg with the sheath. external bypass with 5fr antegrade sheath was placed and the leg was warm with good color. it was further reported that care was withdrawn and patient expired while on cp support. no additional details were provided about the cause of death as it relates to the device. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10812 | 1220648-2025-26971 | 21697756 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21697756&pc=OZD | 03/03/2025 | 05/02/2025 |  | 03/26/2025 |  |  | 03/03/2025 | 03/26/2025 |  |  |  |  | 03/26/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC |  |  | US |  | 1292326 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | Yes | Application Program Problem | [6 empty] | Insufficient Information | [12 empty] | Death | [5 empty] | IMPELLA RP FLEX. | [18 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella rp flex and impella cp bipella device for mechanical circulatory support. the pump was placed and had an automated impella controller (aic) alarm for optical signal followed by a yellow controller alarm. after 96 hours of support, the decision was made by the family to withdraw care and the patient expired. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude the impella as an associated factor in the patient's outcome. | [5 empty] | The automated impella controller (aic) was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | G.4 revised pma/510(k) number as it was previously entered incorrectly on the initial manufacturer device report number 1220648-2025-26971. | The investigation of automated impella controller (aic) - controller error (yellow alarm) and aic - loss of opm data has been completed. data analysis: on 3/2 rp pump with console ic5202, immediately unplugged, then plugged back in. after the pump was plugged back in, on 3/3, the logs show two commands "processsamplingdatafromopticalbench: sentstopacquisitioncmd ack" then "clearallopticaldatasharedmemory: 0x4000". this results in "processdualsensorsampleforopmcommstoppederror: calculator placement signal 1058" which is indicative of the failure mode loss of optical data due to a software defect. the real time (rt) logs confirmed no optical data at this time. the pump stayed on this console until the end of the case on 3/6 despite the lack of optical signal. device analysis: the console was tested on an engineering lab bench. the failure mode was not reproduced while running an optical pump simulator. there was no issue with the console hardware. root cause: the root cause of the controller error was due to an aic software issue. this was entered into jama with defect/ bug id gid-dft-2927090.
10813 | 1220648-2025-26972 | 21698104 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21698104&pc=OZD | 02/17/2025 | 03/26/2025 |  | 03/26/2025 |  |  | 02/28/2025 | 03/26/2025 |  |  |  |  | 03/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025589281 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 57 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhagic Stroke | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | It was reported via abiomed's long-term outcome and quality indicator (loqi) database that a 57-year-old female patient implanted witha impella 5.5, had an adverse event of intracerebral hemorrhage with "probable" relatedness to the pump. it was reported that the patient had facial weakness after extubation, and tongue deviation to left. potentially in setting of traumatic extubation, left facial droop noted, slurred speech. patient was on aspirin and brilinta. mri and cta confirmed the patient had an evolving hemorrhagic stroke and patient had no improvement. family initiated withdrawing care due to being fully dependent on impella. eventually care was withdrawn and the patient expired. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)".
10814 | 21698166 | 21698166 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21698166&pc=OZD | 03/21/2025 | 03/24/2025 | 03/24/2025 | 03/26/2025 | 03/24/2025 | 03/26/2025 |  | 03/26/2025 |  |  |  |  | 03/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD | [5 empty] | ABIOMED, INC. |  |  | US | [5 empty] | Yes |  |  | Not Available |  |  | No |  |  |  | User Facility report | Yes | Other Health Care Professional | [6 empty] | 65 YR |  | Female | 68 KG |  |  | Malfunction | No | Yes | [7 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Unknown |  |  | Yes | Initial submission | [5 empty] | Y | HOSPITAL |  |  | Around 2200, purge pressure increased to >1000. multiple rns (registered nurses) to change purge tubing, de-air line, etc. when those measures did not fix the pressure, abiomed rep was called. they asked us to exchange the console as the next troubleshooting step. after switching the console, purge pressure returned to wdl [within defined limits] and impella was no longer alarming.
10815 | 1220648-2025-26973 | 21698364 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21698364&pc=OZD | 03/02/2025 | 03/28/2025 |  | 03/26/2025 |  |  | 03/02/2025 | 03/26/2025 |  |  |  |  | 03/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1627246 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump support ongoing for a surgical patient. the automated impella controller (aic) alarmed after 7 days with console failure. the aic was replaced and support proceeded without harm or injury. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - controller failure (red alarm) has been completed. the console logs from the reported day of event are consistent with the complaint and showed an controller error alarm triggered during the clinical procedure. the real time logs confirmed that all signals received from optical bench (placement, signal-to-noise ratio, contrast, lamp, gain) are represented as -16384 values due to the crc error. the console was returned for investigation and was tested. voltage from the power battery manager (pbm) at connector (j3) was 4.9 volts. the voltage on the optical bench lamp connector (p4) was 4.8 volts, and the lamp power was 3.51 uw. visible light was detected from the optical pump connector in the console, and "5s" parameter testing confirmed that the optical system was within specification. the controller error was due to an optical bench firmware (fw) communication protocol anomaly, resulting in misalignment of data communication packets received by epc. the aic software operated as designed. the root cause of the ¿controller error (opm comm crc invalid) [optical pump connected] - alarm issued (i.e. aic - loss of opm data)" was due to a firmware issue (optical bench fw anomaly). b.7 added information as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26973. h.3 revised was the device evaluated by the manufacturer? and evaluation summary attached fields as the investigation was completed when initial manufacturer device report number 1220648-2025-26973 was submitted. h.6 codes 3233, 3221, 4315 and 11 were entered incorrectly on initial manufacturer device report number 1220648-2025-26973 as the investigation was completed. new codes have been added. h.10 revised as investigation results were inadvertently omitted on initial manufacturer device report number 1220648-2025-26973.
10816 | 1220648-2025-26974 | 21698434 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21698434&pc=OZD | 03/03/2025 | 04/19/2025 |  | 03/26/2025 |  |  | 03/03/2025 | 03/26/2025 |  |  |  |  | 03/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025540202 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Restricted Flow rate | Positioning Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was on impella cp support and during support, the surgeon attempted to reposition the impella cp. however, any movement to pull back created suction alarms. the pump was kinked near the outflow, contributing to low flows and suction. the pump was removed and no patient harm was reported. the patient survived the procedure. | [5 empty] | The investigation of low pump flow and positioning issues has been completed. the impella device was not returned for evaluation. low pump flow: clinical information states upon initiating the pump, p-level quickly dropped from p-9 to p-1 and never had flow above 1.7l/min and angio revealed a kink near the outflow. data logs were reviewed and an auto-response p-level change from p-9 to p-2 was observed. a suction alarm was present at p-9 and ceased at p-2. pump flow increased from 1 l/min at p-9 to 1.5 l/min at p-2. purge pressure initially exceeded 1100 mmhg, then rapidly dropped to around 400 mmhg, stabilizing at approximately 600 mmhg during the p-level change. therefore, the root cause of the low pump flow issue was most likely a kinked cannula due to improper initial positioning resulting in the pump deep. positioning issue: clinical information states that the angiogram found pump seated deep at implant although imaging was used for pump implant. repositioning was failed as it led to suction alarms. no position alarms were observed in the logs. therefore, the root cause of the positioning issue was most likely improper initial positioning since pump was found to be deep upon activation with low pump flow. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10817 | 1220648-2025-26975 | 21699184 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21699184&pc=OZD | 03/03/2025 | 10/29/2025 |  | 03/26/2025 |  |  | 03/03/2025 | 03/26/2025 |  |  |  |  | 03/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC |  |  | US |  | 2025611282 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Male |  | Hispanic | White | Malfunction | No | Yes | Leak/Splash | Device Sensing Problem | Material Split, Cut or Torn |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [18 empty] | AUTOMATED IMPELLA CONTROLLER | IMPELLA CP SERIAL NUMBER (B)(6). | [17 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient was implanted with an impella cp device for mechanical circulatory support. during insertion it was reported that the first sheath was defective. there was a crack that caused a leak. the sheath was replaced and new sheath delivered. there was no patient harm reported. | The complainant also reported that, during support, there were placement signal issues but this did not prompt intervention. the pump was weaned and explanted at conclusion of the high risk percutaneous coronary intervention. |  |  |  |  | The investigation of introducer damage ¿ mechanical has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. | The impella device was not received from the customer. should the device or any new information be received, a supplemental mdr will be filed. | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed. the pump was not returned for investigation. the data log shows that signal-to-noise ratio (snr0 and contrast dropped close to 0 and 10 as soon as the pump was inserted into the patient's body and ramped up to p9 with a placement signal not reliable (psnr) alarm of 130. gain, light and lamp were not affected during case run. additionally, snr is periodic, but its period does not match the periodicity of the pulsatile placement signal during the case run. as per the clinical guidelines, a perclose suture was placed before inserting a short 14fr peel-away sheath, which had a cracked valve and began to leak after the dilator was removed. the ci was entered, and a new sheath was inserted without issue. additionally, psnr was reported during the case. the introducer damage failure mode was replaced with the access site bleeding failure mode since the bleeding was reported from the damaged introducer sheath. the cause of the access site bleeding issue was most likely damaged valve. the damage valve location and type could not be confirmed without product back. the cause of the optical signal issue was not determined since the product was not returned and sufficient clinical information was not provided. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time.
10818 | 1220648-2025-26976 | 21699663 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21699663&pc=OZD | 03/02/2025 | 03/26/2025 |  | 03/26/2025 |  |  | 03/02/2025 | 03/26/2025 |  |  |  |  | 03/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025507266 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that an impella 5.5 pump was unable to advance through the innominate artery during attempted delivery. upon meeting resistance, the catheter kinked and the decision was made to abort the implant. there was no resulting patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿
10819 | 1220648-2025-26977 | 21700085 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21700085&pc=OZD | 03/03/2025 | 05/08/2025 |  | 03/26/2025 |  |  | 03/03/2025 | 03/26/2025 |  |  |  |  | 03/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025558971 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient was to be implanted with impella 5.5 for mechanical support. during the attempted placement, the implant was aborted due to the patient¿s vascular anatomy. no patient harm was reported. | [5 empty] | The impella pump was not returned for investigation. should new information be received, a supplemental manufacturer device report will be filed. instructions for use as follows: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella 5.5 with smartassist for use during cardiogenic shock. section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings. section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella 5.5 with smartassist for use during cardiogenic shock with automated impella controller section: warnings & cautions: cautions: ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings: ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ impella 5.5 with smartassist for use during impella: section: warnings & cautions: warnings, section: pre-support evaluation, section: axillary insertion of the impella 5.5 catheter, section: direct aortic insertion, section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings: ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the delivery issue was most likely patient condition related, case aborted due to patient's vascular anatomy.
10820 | 1220648-2025-26978 | 21700353 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21700353&pc=OZD | 02/26/2025 | 05/08/2025 |  | 03/26/2025 |  |  | 02/26/2025 | 03/26/2025 |  |  |  |  | 03/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025484481 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes |  | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | Pumping Stopped | Infusion or Flow Problem |  |  |  |  | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs), and following the implant, flow saturation and high purge pressure were observed. the pump subsequently stopped while components were being gathered for troubleshooting, replacement. the pump was removed and a new impella cp device was implanted for continuation of therapy. the patient was reported to be in stable condition following the event. per the plan of care, the patient was successfully weaned from impella mcs with an uneventful explanted of the impella cp device. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of the analysis, a supplemental report will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: using the automated impella controller with the impella catheter ¿achieving an act = 250 seconds prior to removing the dilator will help prevent a thrombus from entering the catheter and causing a sudden stop on startup.¿ section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ | The investigation of pump stop, saturated flow, and high-pressure purge has been completed. product was returned and evaluated. biomaterial was found on the impeller and in the purge gap. no kinks were found in the pump catheter or purge cassette tubing. there was also blood ingress found in the purge line just above the motor housing. high purge pressure was reproduced while purging the pump in a bucket. when attempting to turn the pump on, there was an immediate pump stop due to high motor current. data log review shows high pump flow upon pump startup with an upward shift in motor current shortly after. purge pressure began to rise within the first 10 minutes of the case and then varied before eventually trending up past 1000 mmhg before the pump stop. just before the pump stop, motor current shifted upward again and pump flow was saturated. based on logs, the pump was able to be restarted after the first pump stop but stopped again shortly after and could not be restarted again. pump stop: the root cause of the pump stop was most likely biomaterial since biomaterial was found on the impeller and purge gap, blood was found in the purge line, data logs show high purge pressure and saturated flow just before the pump stop, and the pump stop reproduced in the lab. saturated flow: the root cause of the saturated flow was most likely biomaterial since biomaterial was found on the impeller and purge gap, blood was found in the purge line, and data logs show high purge pressure and upward shifts in motor current. high purge pressure: the root cause of the high purge pressure was most likely biomaterial since biomaterial was found on the impeller and purge gap, blood was found in the purge line, the high purge pressure reproduced in the lab, and data logs show upward shifts in motor current. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26978 as it is unknown if the initial reporter also sent the report to the food and drug administration. h6 medical device problem codes 1491 and 2964 were inadvertently omitted on the initial manufacturer device report number 1220648-2025-26978.
10821 | 1220648-2025-26979 | 21702254 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21702254&pc=OZD | 02/26/2025 | 03/27/2025 |  | 03/26/2025 |  |  | 02/26/2025 | 03/26/2025 |  |  |  |  | 03/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025569807 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 83 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Increase in Pressure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient undergoing a high-risk percutaneous coronary intervention was implanted with an impella cp device for mechanical circulatory support and following the implant, purge pressure high alarm occurred and was unable to be resolved. the impella cp device advanced across aortic valve and started. after starting, purge flow was 4.5 ml/hr. over the next ten minutes, the purge flow steadily decreased until alarming purge pressure high. purge line was traced, and no kinks were present. the decision made to swap for a new purge cassette. after cassette changed, purge flow did not normalize, and alarm persisted. another automated impella controller (aic) was brought in, and aic-to-aic transfer was conducted. again, alarm remained present. the physician continued with case with alarm present. intravascular ultrasound, ballooned, stented, and post dilated the proximal left circumflex and left anterior descending into the left main. physician displeased with alarm and distraction during the case. the physician stated the patient was appropriately anticoagulated prior to implant and does not believe thrombus could have been a factor in purge pressure failure. the impella cp device was explanted. there was no report or indication of harm to the patient as a result of the event. | [5 empty] | The investigation of high pressure purge has been completed. the returned pump was inspected and there were no physical abnormalities. there was biomaterial observed in the purge gap below the impeller. the pump was soaked and run in the lab. the high purge pressure did not reproduce. the pump passed electrical testing. there were no data logs provided for the case. as the high purge pressure was not reproduced and biomaterial was observed in the purge gap, the root cause of the high purge pressure was most likely biomaterial. b.7 information was added as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26979. e.4 should have been left blank on the initial manufacturer device report number 1220648-2025-26979 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.3 revised selections as investigation was completed at time of initial manufacturer device report number 1220648-2025-26979 submission. h.6 several investigation result codes were inadvertently omitted from initial manufacturer device report number 1220648-2025-26979 and were added. h.10 investigation results were inadvertently omitted from initial manufacturer device report number 1220648-2025-26979 and were added. additional manufacturer narrative instructions for use (ifus) were reported on the previously submitted manufacturer device report number 1220648-2025-26979 incorrectly. no ifus are needed to be reported for this report in accordance with updated procedures. | The impella device was returned, and an evaluation is underway. upon completion of the analysis, a supplemental report will be filed. instructions for use for the related event are as follows: impella cp with smartassist. section: troubleshooting the purge system: ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop. clinicians should monitor motor current and consider replacing the impella catheter whenever a rise in motor current is seen.¿
10822 | 1220648-2025-26980 | 21702262 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21702262&pc=OZD | 02/26/2025 | 06/06/2025 |  | 03/26/2025 |  |  | 02/26/2025 | 03/26/2025 |  |  |  |  | 03/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026615541 |  | 1000482 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 69 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Sense | Failure to Read Input Signal | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported during the implant of an impella 5.5 device to a patient with cardiomyopathy for mechanical circulatory support (mcs), the impella 5.5 pump was displayed to be defective after being plugged into two different automated impella controllers (aic). consequently, a new impella 5.5 device and different aic were utilized for successful implant and start of mcs. there was no report or indication of patient harm as a result of the event. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of pump not detected has been completed. the pump was returned and evaluated. there were no visual abnormalities with the patient cable plug pins. there were no signs of kinks or bond issues. the pump successfully connected to the lab console with no crc errors or issues in detection based on imc logs. the root cause of the pump not being detected was not determined since the returned product did not reproduce the failure. d9 device returned to manufacturer date added e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26980 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 manufacturing site name, address, country and fax number were erroneously entered on the initial manufacturer device report number 1220648-2025-26980. h6 medical device problem code and component code were erroneously entered on the initial manufacturer device report number 1220648-2025-26980. new codes entered.
10823 | 1220648-2025-26981 | 21702275 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21702275&pc=OZD | 02/26/2025 | 03/26/2025 |  | 03/26/2025 |  |  | 02/26/2025 | 03/26/2025 |  |  |  |  | 03/26/2025 |  | CONTROL, PUMP SPEED, CARDIOPULMONARY BYPASS |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025609471 |  | 0048-0045 |  | No |  |  | No |  |  | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 84 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs) and experienced bleeding at the site requiring intervention. the patient presented to the emergency department in cardiogenic shock. the patient experienced torsades and was defibrillated. the decision was made to intubate and activate the catheterization lab. pt was taken to the catheterization lab and experienced a couple more episodes of ventricular fibrillation, which required defibrillation. a left heart a catheterization showed significant multivessel disease in the left and right coronary arteries. the decision was made to implant impella cp device. the impella cp was eventually implanted with some difficulty, then started and maintained flows around 3.5 l/min on auto. single access technique was obtained for the percutaneous coronary intervention, which was completed to the left anterior descending, left main and left circumflex arteries. thereafter, the patient was transported to the unit on mcs. however, some time thereafter, "substantial" bleeding at the site was noted, and the impella cp device was explanted over-the-wire. an 18fr dry-seal sheath inserted to achieve hemostasis, and subsequently, a new impella cp device was inserted for continuation in therapy. the amount of blood loss and if blood products were administered are unknown. the patient was reported to be in stable condition following the event. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
10824 | 1220648-2025-26982 | 21702303 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21702303&pc=OZD | 02/26/2025 | 10/03/2025 |  | 03/26/2025 |  |  | 02/26/2025 | 03/26/2025 |  |  |  |  | 03/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025594939 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Device Sensing Problem | Infusion or Flow Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | Death |  |  |  |  | AUTOMATED IMPELLA CONTROLLER. | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was initially implanted with an impella cp device for mechanical circulatory support and transferred to an impella 5.5 device for escalation in care. following the impella 5.5 implant, access site bleeding and saturated flows were experienced requiring intervention. bleeding at anastomosis during insertion was noted. consequently, blood products were administered, and protamine was given. approximately five days later, there were signs of flow saturation. flows were increased by 1 liter. the impella 5.5 pump position was checked under transthoracic echocardiogram. replacement of the impella 5.5 device was contemplated but was waiting for confirmed neurological status. team acknowledged awareness of the potential of an acute pump stop, and the patient continued on support. approximately, 19 days later, it was noted heparin was off for potential bleeding concerns. no source was identified; no blood products were given. the latest update noted the patient to be stable continuing with impella 5.5 device support. | The complainant also reported that there were placement signal issues. impella position confirmed with point-of-care ultrasound. placement signal monitoring was then disabled. | The decision was made to withdraw care and the patient expired. it is unknown if the use of the impella was related to patient death, thus there is not enough evidence to exclude impella as an associated factor in the patient's outcome. |  |  |  | The device remains implanted, supporting the patient at the time of the mdr writing, and therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation). | This report is being submitted as part of a retrospective review of optical signal issues, per fda recommendation. the investigation has been completed for optical signal issue. data logs indicate several positions in aorta alarms and also several impella position unknown alarms. no other abnormalities seen from logs. the root cause of the optical signal issue was most likely patient condition related due to false position in aorta alarm. a review of the device history record (dhr) for the suspect product and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time the product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. b5 updated to include placement signal issue description. h6 updated to include placement signal coding. corrections that were made from the initial manufacturer device report number 1220648-2025-26982. d10 concomitant medical products and therapy dates updated h4 device manufacture date corrected. | The investigation of access site bleeding and saturated flow has been completed. no pump was returned for evaluation. as per clinical, bleeding from anastomosis during insertion and resolved with administering blood products/protamine. the root cause of the access site bleeding issue was not determined since product was not returned and due to insufficient clinical details. data logs were returned for review. the log indicating a rise in motor current (mc) from 450 to 550 ma at same p-level (p-8) between (b)(6) 2025. saturated flows observed with flow rising up to 5. l/min. purge parameters were normal. later p-level was dropped to p-6 and flows were partially low during the rest of case. also, flows were slightly saturated at p-6 during (b)(6) 2025 with flows up to 4.4l/min. rt log indicating saturated flows with diastolic flow waveform maxed at 5.6 l/min on p-8. mc rising from 418 to 523 per rt log min/max values. no other abnormalities. the root cause of the saturated issue was most likely biomaterial per log review. b2 outcome revised to reflect death based on the additional information received from the clinical site. b5 revised to reflect death based on the additional information received from the clinical site. d6b explant date added. h1 type of reportable event revised to reflect death based on the additional information received from the clinical site. h6 health effect - impact code revised to reflect death based on the additional information received from the clinical site. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26982 as it is unknown if the initial reporter also sent the report to the food and drug administration.
10825 | 1220648-2025-26983 | 21702313 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21702313&pc=OZD | 02/28/2025 | 04/08/2025 |  | 03/26/2025 |  |  | 02/28/2025 | 03/26/2025 |  |  |  |  | 03/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1127920 |  | 0042-0000-EU |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Defective Device | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the cable near the power plug on an automated impella controller was damaged. the insulation on the cable was noted to be slightly eroded. charging of the device was still possible. there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - damage during therapy has been completed. there was no evidence was found in the log to confirm the reported failure mode. the console was tested after arriving and a damaged cable was confirmed. the root cause for damaged power cord was most likely a use issue. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26983 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting manufacturing site fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26983.
10826 | 1220648-2025-26984 | 21702323 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21702323&pc=OZD | 03/02/2025 | 04/08/2025 |  | 03/26/2025 |  |  | 03/02/2025 | 03/26/2025 |  |  |  |  | 03/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1392605 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that an automated impella controller experienced two controller errors during the night shift. both errors self resolved and the aic was replaced. there was no patient harm. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of automated impella controller (aic) - controller error (yellow alarm) has been completed. the console logs from the reported day of event are consistent with complaint. the console was retuned for investigation. the voltage from the power battery manager (pbm) at connector (j3) was 4.9v and the voltage on the optical bench lamp connector (p4) was 4.8v. visible light was detected from the optical pump connector in the console. while the parameters testing signal not reliable (snr) was marginally passing the specification as well as some distortion in analog output i.e. fringe pattern. the controller error was due to an optical bench fw communication protocol anomaly, resulting in misalignment of data communication packets received by epc. d6a and d6b should have been left blank on the initial manufacturer device report number 1220648-2025-26984. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26984 as it is unknown if the initial reporter also sent the report to the food and drug administration. h4 erroneously selected on the initial manufacturer device report number 1220648-2025-26984. h8 erroneously selected on the initial manufacturer device report number 1220648-2025-26984.
10827 | 1220648-2025-26985 | 21702333 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21702333&pc=OZD | 03/02/2025 | 05/07/2025 |  | 03/26/2025 |  |  | 03/02/2025 | 03/26/2025 |  |  |  |  | 03/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025595789 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 90 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a hematoma developed at the impella cp site during patient support. manual pressure was applied and one unit of packed red blood cells was given to manage the condition. support continued uninterrupted. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output: use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of access site bleeding is determined to be patient movement because the hematoma was noted after the patient woke up. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26985 as it is unknown if the initial reporter also sent the report to the food and drug administration.
10828 | 1220648-2025-26986 | 21702360 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21702360&pc=OZD | 03/02/2025 | 05/06/2025 |  | 03/26/2025 |  |  | 03/02/2025 | 03/26/2025 |  |  |  |  | 03/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025610501 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 33 YR |  | Male |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a patient was on impella cp support. it was reported that during the peel-away sheath removal, the pump advanced into the ventricle and twisted upon itself. the medical doctor pulled the pump back across the valve and was able to untwist the impella. the pump was then able to advance across the atrioventricular into the ventricle again. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of the positioning issue most likely was use related since during the peel-away sheath removal, the pump advanced into the ventricle and twisted upon itself. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26986 as it is unknown if the initial reporter also sent the report to the food and drug administration.
10829 | 1220648-2025-26987 | 21702454 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21702454&pc=OZD | 03/03/2025 | 03/26/2025 |  | 03/26/2025 |  |  | 03/03/2025 | 03/26/2025 |  |  |  |  | 03/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1239161 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 38 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that the patient was on impella support when the automated impella controller (aic) had a purge disc not recognized error and multiple attempts were made to change the cassette. the aic was replaced. there was no patient harm reported. | [5 empty] | The investigation of automated impella controller (aic); purge disc/flag issues has been completed. the data logs from the complaint date revealed that the console issued purge disc not detected alarm and the alarm got resolved after the purge disk was inserted. e console was examined and a broken purge disc flag was identified. the purge disc flag was observed to be broken and was the root cause of the reported purge disc not detected alarm issue.
10830 | 1220648-2025-26988 | 21702457 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21702457&pc=OZD | 03/03/2025 | 05/08/2025 |  | 03/26/2025 |  |  | 03/03/2025 | 03/26/2025 |  |  |  |  | 03/26/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025571100 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 59 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the patient was to be implanted with impella 5.5 for mechanical support. during the attempted placement, the implant was aborted due to the patient¿s vascular anatomy. the physician attempted several times, but was still unsuccessful due to anatomy and spasm of artery, physician did not want to try another access point and decision was made to abort 5.5 and place on peripheral venoarterial extracorporeal membrane oxygenation. no patient harm was reported. | [5 empty] | The impella pump was not returned for investigation. should new information be received, a supplemental manufacturer device report will be filed. instructions for use as follows: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. impella 5.5 with smartassist for use during cardiogenic shock section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ impella 5.5 with smartassist for use during impella rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. impella 5.5 with smartassist for use during cardiogenic shock & impella rp smart assist system with automated impella controller section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ iimpella 5.5 with smartassist for use during impella & rp smart assist system with automated impella controller & rp flex with smart assist system with automated impella controller section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected: decreased ventricular cavity size, ventricular aneurysms, thin-walled ventricles due to chronic dilation, congenital heart disease, or compromised cardiac tissue quality.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance.¿ section: warnings & cautions: warnings ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing the risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly after cpr with chest x-ray guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue is determined to be patient condition because the pump was unable to pass through artery due spasm. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26988 as it is unknown if the initial reporter also sent the report to the food and drug administration.
10831 | 1220648-2025-26926 | 21686206 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21686206&pc=OZD | 02/27/2025 | 03/25/2025 |  | 03/25/2025 |  |  | 02/27/2025 | 03/25/2025 |  |  |  |  | 03/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025593472 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 74 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a patient was on impella cp support and during support, the patient developed a hematoma at the access site. femstop was place, packed red blood cells and platelets were given and heparin was withheld. support continued but the patient remained critical. however, care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10832 | 1220648-2025-26927 | 21686259 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21686259&pc=OZD | 02/27/2025 | 03/25/2025 |  | 03/25/2025 |  |  | 02/27/2025 | 03/25/2025 |  |  |  |  | 03/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025556297 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ventricular Fibrillation | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported that a patient had an impella cp inserted for non-st elevated myocardial infarction. during insertion, the patient went from asystole to ventricular fibrillation (vfib). the patient was defibrillated multiple times and finally came out of vfib into normal sinus rhythm. the pump was eventually escalated to an impella 5.5. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10833 | 1220648-2025-26928 | 21686365 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21686365&pc=OZD | 02/27/2025 | 04/15/2025 |  | 03/25/2025 |  |  | 02/27/2025 | 03/25/2025 |  |  |  |  | 03/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025553826 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Fluid/Blood Leak | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella 5.5 support and during support, liquid and sticky sugar crystals were found around the purge cassette and wetness inside the purge housing. the nurse was advised to change out the purge cassette and to watch over night. the purge continued to peak after a few hours so discussion with surgeon was needed to switch out the pump. the patient was brought to the operating room to switch out the impella 5.5 for a new pump due to the purge system leaking. the patient's outcome was stable and the procedure was successful with other device. | [5 empty] | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock; section: cleaning: ¿do not clean with or expose any part of the clear sidearm of the impella catheter (e.g. infusion filter, pressure reservoir) to alcohol. alcohol has been shown to cause cracks and leaks in these components. carefully read labels on common skin preps and lotions to avoid using any alcohol-containing products in the area of the infusion filter or pressure reservoir." ¿clean the connector cable with 70% isopropyl alcohol.¿ | The investigation of purge leak - pump has been completed. only the pump was returned for investigation. data showed no abnormalities related to purge pressure, motor current, or pump position. according to the clinical assessment, the pump was replaced with a new one due to a purge system leak. additionally, the doctor reported cassette change due to sugar crystals around the purge cassette. the returned pump had biomaterial observed on the outflow cage. the pump was primed successfully with a test purge cassette. the purge pressure values were within tolerance limits during the test. the pump was run in a bucket of water with observed saturated pump flow. the cause of the purge leak issue was not determined since the returned pump did not reproduce the issue, and the purge cassette was not returned for investigation.
10834 | 1220648-2025-26930 | 21686534 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21686534&pc=OZD | 02/26/2025 | 03/25/2025 |  | 03/25/2025 |  |  | 02/26/2025 | 03/25/2025 |  |  |  |  | 03/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026614916 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Female |  |  | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | Ischemia | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 68 year female patient had lost pulses in the lower extremity following impella cp implant. the patient's foot was cold to the touch and full strength flolan was added to assist with the condition. bilateral perfusion catheters were placed and support continued with the implanted pump. five days later, an emergent ct scan showed that the patient also suffered a brain bleed. the family decided to withdraw care and the patient passed away. there is no enough evidence to disassociate the use of the impella with the patient's passing. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
10835 | 1220648-2025-26931 | 21686656 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21686656&pc=PYX | 02/25/2025 | 10/16/2025 |  | 03/25/2025 |  |  | 02/25/2025 | 03/25/2025 |  |  |  |  | 03/25/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025577623 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Device Sensing Problem | Adverse Event Without Identified Device or Use Problem | [5 empty] | Hemorrhage/Blood Loss/Bleeding | Unspecified Vascular Problem | [11 empty] | Required Intervention | [5 empty] | IMPELLA CONTROLLER | IMPELLA CP | [17 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | Followup | [7 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella rp device for mechanical circulatory support (mcs). during support bleeding from all access sites, mouth and nose were noted. one unit of red blood cells was administered. additionally, the rp pump showed no placement signal. according to the report, the patient remained critically ill, and the following day it was noted the patient was still bleeding from all sites and was being transfused. | It was reported that a patient on bipella (cp/rp) experienced bleeding from all sites and was being transfused. additionally, it was noted that a x-ray was showing the pump slightly shallow. the rp continued to show no placement signal. ultimately care was withdrawn and the patient expired. | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was initially implanted with an impella cp followed by right-side support via an impella rp device for mechanical circulatory support (mcs). prior to the impellas being implanted, the patient was undergoing a transcatheter aortic valve replacement procedure and experienced a dissection of the right coronary artery. the patient rapidly decompensated, ventricular fibrillation arrested requiring at least 20 shocks. the valve was deployed in the descending aorta and covered with a stent. the impella cp device was then implanted via the right femoral artery. a stent was placed in right coronary artery but no flow. the patient was taken to the unit on impella cp mcs. later that day, bleeding from all access sites, mouth and nose were noted. one unit of red blood cells was administered. the patient had decreased flows noted and echocardiogram showed right ventricular failure. the patient went into ventricular fibrillation arrest requiring three shocks. the decision was made to place an impella rp device via the right femoral vein, which was successfully completed. the impella cp support was able to be increased without suction. according to the report, the patient remained critically ill, and the following day it was noted the patient was still bleeding from all sites and was being transfused. two days later, the patient had a temporary atrial pacemaker lead placed and five days thereafter the patient's outcome at explant was noted as "expired." no further details surrounding the demise outcome were made available. the cause of death was not provided. there is not enough information to exclude the impella cp device as an associated factor to the outcome (demise) given site bleeding, the period of time with decrease flows, and cardiac arrhythmia requiring cardioversion which may have added, to what appeared to have been, an already complex situation. |  |  |  | Based on the clinical details provided, the root cause of the access site bleed was most likely patient condition related since bleeding occurred at multiple sites simultaneously. b2: omit death, reported on the pump in error. b5: corrected statement to include all pertinent information. h6: corrected health effect-impact code. h10: added the investigation results. | This mdr is being submitted as part of a retrospective review of optical signal issues in accordance with fda recommendation. optical signal issue: product was not returned for investigation. data log review confirmed the placement signal not reliable (psnr) alarm (#1054) as snr was below 250 throughout the entire case. the first rt log of the case showed that snr was below 250 prior to the pump even being started. this caused all pressure measurement calculations based on the optical sensor to be dashed out/unreliable for the majority of the case (pap, papi, cvp). meanwhile, the dp ps, flow calculations, and motor current were stable. based on the findings from data log review, the failure mode was changed to optical signal issue. the two previous pumps run on the console were also checked and both also had consistently low snr, one of which being linked to an optical signal issue complaint ((b)(4)). upon review of data logs, it is apparent that the console is the potential cause of the issue. please refer to complaint # (b)(4) for an investigation into the console. *though the complaint intake mentions a different console (ic1292), there are only logs available for pump sn (b)(6) being run on ic1294, and the two other pumps with similar optical signal issues were also run on the same console, so it was likely a typo in the ci.* a review of the device history record (dhr) for the suspect device, applicable quality notifications, and historical complaint data was conducted. this review revealed that the product met all manufacturing release criteria, and no product deficiencies were identified at the time product was manufactured and released to the customer. all pertinent information available to abiomed has been submitted at this time. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. this is one of two medical device reports associated with this event. refer to initial medical device report for impella cp serial number (b)(6) with date of event 24-feb-2025 and identifier ending in (b)(6). as noted in the impella instructions for use: impella rp with smart assist system with automated impella controller. section: warnings and cautions: "make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound."; "be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open."; "make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound."
10836 | 1220648-2025-26932 | 21686739 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21686739&pc=OZD | 02/27/2025 | 03/25/2025 |  | 03/25/2025 |  |  | 02/27/2025 | 03/25/2025 |  |  |  |  | 03/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025569093 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 53 YR |  | Female |  | Non Hispanic | Black Or African American | Malfunction | No | Yes | Patient-Device Incompatibility | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 53 year old female patient had an attempted impella 5.5 implant and it was aborted due to inability to advance impella 5.5 past the axillary artery. the implant was aborted and an impella cp was implanted. the patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk pci states the following : section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance.
10837 | 1220648-2025-26933 | 21686778 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21686778&pc=OZD | 02/24/2025 | 04/21/2025 |  | 03/25/2025 |  |  | 02/24/2025 | 03/25/2025 |  |  |  |  | 03/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025535571 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | White | Death | Yes | Yes | Restricted Flow rate | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ventricular Fibrillation | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was initially implanted with an impella cp followed by right-side support via an impella rp device for mechanical circulatory support (mcs). prior to the impellas being implanted, the patient was undergoing a transcatheter aortic valve replacement procedure and experienced a dissection of the right coronary artery. the patient rapidly decompensated, ventricular fibrillation arrested requiring at least 20 shocks. the valve was deployed in the descending aorta and covered with a stent. the impella cp device was then implanted via the right femoral artery. a stent was placed in the right coronary artery but no flow. the patient was taken to the unit on impella cp mcs. later this day, bleeding from all access sites, mouth and nose was noted. one unit of red blood cells was administered. the patient had decreased flows noted and echocardiogram showed right ventricular failure. the patient went into ventricular fibrillation arrest requiring three shocks. the decision was made to place an impella rp device via the right femoral vein, which was successfully completed. the impella cp support was able to be increased without suction. according to the report, the patient remained critically ill, and the following day it was noted the patient was still bleeding for all sites and being transfused. two days later, the patient had a temporary atrial pacemaker lead placed and five days thereafter the patient's outcome at explant was noted as "expired." no further details surrounding the demise outcome were made available. the cause of death was not provided. there is not enough information to exclude the impella cp device as an associated factor to the outcome (demise) given site bleeding, the period of time with decrease flows, and cardiac arrhythmia requiring cardioversion which may have added, to what appeared to have been, an already complex situation. | [5 empty] | The investigation of access site bleeding, low pump flow, and vf/vt/af/at has been completed. the impella device was not returned for evaluation. clinical mentions bleeding was from multiple access sites including impella access site. no other additional information related to bleeding was provided. the root cause of the access site bleeding - major was most likely patient condition related as evidenced by the multiple-site bleeding information. as the necessary information was not provided, the root cause of the vt/vf was not determined. no product was returned for analysis. the data logs returned did not cover the entire cp support duration. the pump was explanted on 4/3 and logs were only provided from implant to 3/28. the logs do not show any motor current spikes and clinical did not report any positioning issues. there were suction alarms which resolved and the logs confirmed low flow for matching p-levels. the low flows resolved after pump turned to p7. the root cause of the low pump flow was not determined as no product was returned for analysis and limited clinical information related to failure was provided. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. this is one of two medical device reports associated with this event. refer to initial medical device report for impella rp serial number (b)(6) with date of event 25-feb-2025 and identifier ending in (b)(6). as noted in the impella instructions for use: impella cp with smart assist system: section: general patient care considerations: "assess access site for bleeding and hematoma." section: entering cardiac output: use of the repositioning sheath and peel-away introducer: "remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site." section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)."
10838 | 1220648-2025-26934 | 21686906 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21686906&pc=OZD | 02/26/2025 | 04/15/2025 |  | 03/25/2025 |  |  | 02/26/2025 | 03/25/2025 |  |  |  |  | 03/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025562336 |  | 1000100 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 81 YR |  | Male |  | Non Hispanic | White | Death | Yes | Yes | Infusion or Flow Problem | [6 empty] | Insufficient Information | [12 empty] | Death | Other | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella 5.5 support and during support, there was flow saturation. there were increased motor currents. support continued for two days at which time the care was withdrawn and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The investigation of saturated flow has been completed. the pump was returned cut at catheter. upon visual inspection, dried blood was present on impeller and within outflow cage. biomaterial was present within the purge gap. pump teardown was performed and inside of motor housing did not show any biomaterial or abnormalities within the motor. data logs were reviewed and lt log of case shows pump flows becoming saturated around day 5 of support after pump experienced an instance of high purge pressure. a slight trend up in motor current is observed with saturation of flows. no motor current spikes observed to indicate biomaterial ingestion. ¿placement signal low¿ alarms present with saturated flows. clinical details noted that the when saturated flows were observed, inverted ao/lv waveforms were noted by clinical team. it was noted that patient had stopped systemic heparin due to a retroperitoneal bleed. patient remained on support for an additional 2 days. the root cause of the saturated flow cannot be definitively determined as the pump was returned cut and was unable to be run in attempt to reproduce saturated flows. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
10839 | 1220648-2025-26935 | 21687193 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21687193&pc=OZD | 02/27/2025 | 04/15/2025 |  | 03/25/2025 |  |  | 02/27/2025 | 03/25/2025 |  |  |  |  | 03/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026612642 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 26 YR |  | Female |  |  | White | Death | Yes | Yes | Restricted Flow rate | Deformation Due to Compressive Stress | [5 empty] | Insufficient Information | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump placed to support a patient admitted with acute myocardial infarction/cardiogenic shock. the pump was explanted after seven hours due to a kinked cannula. the pump was explanted and replaced by a new impella cp. the patient did expire on the second pump. however, there is not enough evidence to exclude the first impella as an associated factor in the patient's outcome. | [5 empty] | The investigation has been completed. kinked cannula due to aortic arche and caused low flow. the pump was replaced. low flow was added since the kinked cannula which caused low flow per clinical description. the data logs were not returned for review. visual inspection found a kink on cannula about 15 mm from the of cage and cannula bonding site. no other issues were found on the rest of the pump. the root cause of low flow was most likely kinked cannula since the clinical team confirmed the kink caused low flow. the root cause of kinked cannula was most likely patient condition related since the clinical team confirmed the kink was due to the patient has short aortic arche. h.6 medical device problem code added due to investigation results. b.2 required intervention was inadvertently omitted from initial manufacturer device report number 1220648-2025-26935. | The impella pump and purge cassette were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
10840 | 1220648-2025-26936 | 21687226 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21687226&pc=OZD | 02/27/2025 | 03/25/2025 |  | 03/25/2025 |  |  | 02/27/2025 | 03/25/2025 |  |  |  |  | 03/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025600813 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Insufficient Information | [12 empty] | Required Intervention | Death | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 52-year-old male patient with an impella cp for a fall prompted by acute myocardial infarction where he had a visible hematoma and unknown neurological status. it was reported that the distal pulses in the patient's lower extremity was lost following impella cp implant. the physician was made aware and revascularization was performed. however, the family ultimately made the patient do not resuscitate (dnr) as a result of the overall prognosis. the patient passed away. there is not enough evidence to disassociate the patient's passing and the use of the device. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smartassist for use during cardiogenic shock and high risk pci states the following: section: potential adverse events (united states): "acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)".
10841 | 1220648-2025-26937 | 21687387 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21687387&pc=OZD | 11/29/2024 | 05/07/2025 |  | 03/25/2025 |  |  | 02/27/2025 | 03/25/2025 |  |  |  |  | 03/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025534726 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 50 YR |  | Male |  | Hispanic | White | Injury | Yes | Yes | Failure to Pump | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 50 year old male patient with a cp pump placed to support cardiogenic shock. the patient had previously been admitted and found not to be compliant with meds or lifevest use. the patient had persistent ventricular tachycardia (vt) and cardioversions and was placed on the cp for support. the cp supported for 6 days and was explanted and supported by ecmo. abiomed later received a notification from long-term outcome and quality indicator (loqi) reported upon a pump stop/no flows that could not be restarted and hemolysis/renal failure with a "possible" relationship to the cp use. the patient survived the procedure. | [5 empty] | The investigation of pump stop and hemolysis has been completed. the impella device was not returned for evaluation. data log reviewed: motor current (mc) spike observed before pump stopped, at same time purge pressure rose. pump wasn't able to restart. no placement signal trend observed. the cause of the pump stop issue most likely was biomaterial ingestion based on mc spike and high motor current pump stop observed in data logs. hemolysis issue reported on 11/29/2024. product not returned. data logs and implant/explant of pump 12/7/2024 to 12/13/2024. timing of hemolysis unclear and lack of clinical details. the cause of the hemolysis issue cannot be determined since complaint device not returned for analysis and insufficient clinical data provided. instructions for use: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿ b3 date of event erroneously entered on the initial manufacturer device report number 1220648-2025-26937. g1 reporting contact facility name was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26937. g2 report source; study was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26937. h10 incomplete instructions for use were provided on the initial manufacturer device report number 1220648-2025-26937. | The device was not received for evaluation. if the device or new information is received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: impella stopped: ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
10842 | 1220648-2025-26939 | 21687545 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21687545&pc=OZD | 02/28/2025 | 05/02/2025 |  | 03/25/2025 |  |  | 02/28/2025 | 03/25/2025 |  |  |  |  | 03/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1133370 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Other Health Care Professional | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Display or Visual Feedback Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | [8 empty] | The hospital had an automated impella controller (aic) in field service at abiomed and was seen to malfunction. there was no display and an inability to communicate with the aic. there was no patient involvement and the controller was returned for investigation. | [5 empty] | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. | The investigation of automated impella controller (aic) - display issue has been completed. there was no log evidence available to confirm the reported failure mode. the reported issue was identified during the service. the reported failure mode is reproducible ¿ upon pressing the power button, the green led above the rotary knob was on, but the display was off. the power from the pbm to the 4-in-1 is 23.8v under ac power, as intended. at connector x51 on the 4-in-1, there was no 5v between pins 4 & 3, and pins 4 & 2, which confirms the 4-in-1 malfunction. the root cause of the console display not turning on was a malfunction of the 4-in-1. h5 was erroneously selected on the initial manufacturer device report number 1220648-2025-26939. h8 was erroneously selected on the initial manufacturer device report number 1220648-2025-26939.
10843 | 1220648-2025-26938 | 21687554 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21687554&pc=OZD | 02/28/2025 | 05/09/2025 |  | 03/25/2025 |  |  | 02/28/2025 | 03/25/2025 |  |  |  |  | 03/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025609987 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 51 YR |  | Female |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, the pump ingested a thrombus and there was continuous suction. the pump was explanted and the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of low pump flow and thrombus has been completed. no product was returned for analysis. data logs reviewed: motor current (mc) spike observed on date low pump flow issue reported (b)(6) followed by drop in p-level and pump flow. suctions occurred, quick troubleshooting alarms in aorta and ventricle observed. no placement signal trend observed. clinical details confirmed observing thrombus ingestion during the impella support on (b)(6) which supporting the data logs reading. the cause of the low pump flow issue most likely was biomaterial ingestion as evidenced by the ingestion pattern observed in the logs. as the necessary information was not provided, the root cause of the thrombus was not determined. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation: section: impella cp with smart assist catheter insertion: section: axillary insertion of the impella cp with smart assist catheter: section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death."
10844 | 1220648-2025-26940 | 21687583 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21687583&pc=PYX | 02/24/2025 | 04/04/2025 |  | 03/25/2025 |  |  | 02/24/2025 | 03/25/2025 |  |  | 01/31/2026 |  | 03/25/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025517153 |  | 1000323 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | Other | Death |  |  |  | IMPELLA CP | [18 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock implanted with an impella rp flex device for mechanical circulatory support experienced limb ischemia and would subsequently expire. the patient underwent elective transcatheter aortic valve replacement. while still in the procedure room while attempting to extubate, the patient started having rhythm changes and decreasing sats eventually going into ventricular fibrillation and ventricular tachycardia requiring numerous defibrillations. coronary angiogram showed coronaries were not compromised. transesophageal echocardiogram showed a severely dilated right ventricle. therefore, the decision was made to place right-sided support as well. the patient was placed on an impella cp (via the left femoral artery) and impella rp flex (via the left femoral vein) for bipella support with a "very" poor prognosis. approximately four hours into mcs, it was noted that the patient had no distal flow in the lower left extremity. the persistent condition was monitored by the staff. after approximately two days of support, the patient was no longer responding to the maximum doses of vasopressors and inotropes. the patient passed away on support. | [5 empty] | The investigation of limb ischemia has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. after medical safety officer review, the complainant reported that the patient on bipella support had no distal flow in the left lower extremity and signs of mottling was noted. it was also reported that the patient prognosis was poor and expired while on support. the demise outcome is associated with the impella cp device in related manufacturer device report 1220648-2025-26903. b.2 revised from the initial manufacturer device report 1220648-2025-26940 due to medical safety officer review. e.4 should have been left blank on the initial manufacturer device report number 220648-2025-26940 as it is unknown if the initial reporter also sent the report to the food and drug administration. h.1 revised from the initial manufacturer device report 1220648-2025-26940 to reflect serious injury per medical safety officer review. h.6 health effect - impact code 1802 per medical safety officer review, death was related to the impella cp and will be reported in manufacturer device report 1220648-2025-26903. code 925 was previously reported incorrectly on the initial manufacturer device report 1220648-2025-26903. component code 3038 was added. | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this is one of two medical device reports associated with the event. refer to manufacturing report number 1220648-2025-26903 for the impella cp. instructions for use for the related event are as follows: impella rp flex with smartassist system. section: contraindications: ¿the impella rp flex with smartassist system is contraindicated for use with patients experiencing any of the following conditions: disorders of the pulmonary artery wall that would preclude placement or correct positioning of the impella rp flex with smartassist device; mechanical valves, severe valvular stenosis or valvular regurgitation of the tricuspid or pulmonary valve; mural thrombus of the right atrium or vena cava; anatomic conditions precluding insertion of the pump; presence of a vena cava filter or caval interruption device, unless there is clear access from the internal jugular vein to the right atrium that is large enough to accommodate a 22 fr catheter.¿
10845 | 1220648-2025-26941 | 21687634 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21687634&pc=PYX | 01/26/2025 | 03/25/2025 |  | 03/25/2025 |  |  | 02/24/2025 | 03/25/2025 |  |  |  |  | 03/25/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP FLEX |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025594213 |  |  |  | Device was returned to manufacturer |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility | Company representation |  |  | 56 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Required Intervention | Life Threatening | Hospitalization | Disability | [21 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | A notification was received via abiomed's long-term outcome and quality indication impella registry regarding a patient who experienced a cerebral vascular accident. the patient was a status 7 on the orthotopic heart transplant list and unable to move forward due to pulmonary hypertension. the decision was made to place and impella 5.5 for mechanical circulatory support during bridge to transplant. approximately 12 days later an impella rp flex was placed for bipella support. approximately 12 days thereafter bipella support initiation, a computed tomography head scan revealed a brain mass. anticoagulation was put on standby overnight and motor current of the impella rp flex was monitored. a repeat scan was completed. the etiology was being discussed as either hemorrhage or malignancy. patient survived the event, and the adverse event was determined to be "probably" related to the device(s) used. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. this is one of two devices associated with this event and patient's implant experience. refer to manufacturing report number 1220648-2025-25809 for the impella 5.5 device.
10846 | 1220648-2025-26942 | 21687721 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21687721&pc=OZD | 02/26/2025 | 03/25/2025 |  | 03/25/2025 |  |  | 02/28/2025 | 03/25/2025 |  |  |  |  | 03/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025595808 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | Asian | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported that the patient on impella cp support had a large hematoma, and it was noted upon arrival to the icu. pressure applied, angle matched, and site redressed. ooze continued to trickle around the sheath with a drop in hemoglobin and hematocrit. two units of rbc were administered. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10847 | 1220648-2025-26943 | 21689558 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21689558&pc=OZD | 04/15/2025 | 03/25/2025 |  | 03/25/2025 |  |  | 02/27/2025 | 03/25/2025 |  |  | 11/30/2025 |  | 03/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2024433483 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 76 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 76 year old female patient who had a cp pump for support for high risk percutaneous cardiac intervention (hrpci). the patient was noted to have expired as a do not resuscitate (dnr). abiomed's long-term outcome and quality indicator (loqi) database later reported that an adverse event of ventricular tachycardia (vt) with "possible" relationship to the pump use. no additional details were provided. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10848 | 1220648-2025-26944 | 21689986 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21689986&pc=OZD | 02/27/2025 | 03/25/2025 |  | 03/25/2025 |  |  | 02/28/2025 | 03/25/2025 |  |  |  |  | 03/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026617126 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  |  |  | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | Hemorrhage/Blood Loss/Bleeding | Perforation of Vessels | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant had a impella cp pump placed to support a patient admitted in with acute myocardial infarction and cardiogenic shock. the pump was supporting when there was a bleed at the access site that was due to a vascular damage during insertion. the surgeon did vascular repair. the team also noted signs of hemolysis. there was very little urine output and positive hemolysis labs. the pump remains on but the team was weaning and per the local team the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿ ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
10849 | 1220648-2025-26945 | 21690269 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21690269&pc=OZD | 02/28/2025 | 04/29/2025 |  | 03/25/2025 |  |  | 02/28/2025 | 03/25/2025 |  |  |  |  | 03/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026615073 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a 62 year old male patient that the patient on impella cp support developed a large hematoma after transfer to icu. angle matching, forward tension sutures were applied, and hematoma was controlled. three units of prbcs were administered. the patient was eventually explanted. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations. ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output, use of the repositioning sheath and peel-away introducer: ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of access site bleeding was most likely use issue since hemostasis was achieved with proper angle matching and forward tension sutures. g1 reporting contact facility name was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26945. h6 health effect - clinical code 1884 was inadvertently omitted on the initial manufacturer device report number.
10850 | 1220648-2025-26946 | 21690506 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21690506&pc=OZD | 02/28/2025 | 05/02/2025 |  | 03/25/2025 |  |  | 02/28/2025 | 03/25/2025 |  |  |  |  | 03/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025508164 |  | 0048-0024-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 29 YR |  | Female |  | Non Hispanic | Asian | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was on impella cp support and during support, there were low pump flows. thrombus was noted. the pump was explanted. | [5 empty] | The pump and purge cassette were was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings: section: pre-support evaluation section: impella cp with smart assist catheter insertion section: axillary insertion of the impella cp with smart assist catheter section: use of impella with transcatheter aortic valves ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The investigation of thrombus and low pump flow has been completed. as the necessary information was not provided, the root cause of the thrombus was not determined. clinical states that since start of the case flows observed were 0.4l/min at p2. thrombus was found on the inlet of the pump at explant which was extracted by the facility to analyze it. pump and purge cassette were returned and there was no trace of biomaterial on the pump. the impeller was not damaged and cannula had no kinks. data logs were not returned. as per the clinical information provided and product investigation, the root cause of the low pump flow issue was biomaterial which was found on the inlet of the pump at explant.
10851 | 1220648-2025-26947 | 21690559 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21690559&pc=OZD | 02/28/2025 | 04/15/2025 |  | 03/25/2025 |  |  | 02/28/2025 | 03/25/2025 |  |  |  |  | 03/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025606223 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that a patient on impella cp support had an echocardiogram that suspected clot on catheter shaft. the impella cp was removed and there was no patient harm. | [5 empty] | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. section: warnings & cautions: warnings: section: pre-support evaluation: section: impella cp with smart assist catheter insertion: section: axillary insertion of the impella cp with smart assist catheter: section: use of impella with transcatheter aortic valves: ¿in patients with transcatheter aortic valves position the impella system carefully to avoid interaction with the transcatheter aortic valve prosthesis. unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death. in this situation, avoid repositioning while the device is running; turn the device to p0 during repositioning or any movement that could bring the outlet windows into proximity to the valve stent structures. if there is low flow observed in a patient implanted with a transcatheter aortic valve prosthesis, consider damage of the impeller and replace the impella as soon as possible.¿ "unintentional interaction of the impella motor housing with the tavr device may result in destruction of the impeller blades. this can lead to systemic embolization, serious injury, or death." | The failure mode of ¿thrombus¿ is to be used when a thrombus is noted but cannot be connected to any other adverse events or decrease in pump performance. thrombus can be observed in the body or on the pump upon explant. when making a determination as to the cause of the thrombus, the following clinical information must be considered: ¿patient's medical history, including any previous thrombotic events or conditions ¿description of the location and characteristics of the thrombus (e.g., size, shape, consistency). ¿relevant laboratory test results, such as coagulation profiles and platelet counts ¿imaging studies, including ultrasound, ct scans, or angiography, to assess the extent and location of the thrombus. ¿any underlying medical conditions or risk factors that may contribute to thrombus formation (e.g., atrial fibrillation, hypercoagulable states). ¿medications or treatments that the patient was receiving at the time of thrombus formation. ¿surgical or procedural details if applicable (e.g, cardiac catheterization). ¿any symptoms or clinical manifestations associated with the thrombus (e.g., pain, swelling, ischemic events). ¿presence of any other indwelling cannulae, lines, or devices such as ecmo, dialysis catheters, swan gantz catheters, central lines, etc. ¿response to any previous anticoagulation or thrombolytic therapies, if applicable. with a returned impella, the device could be inspected for any burrs or rough edges that may promote thrombus growth. to determine the true root cause of the thrombus, the clinical information mentioned above would need to be assessed in conjunction with evidence from the returned device. as the necessary information was not provided, the root cause of the thrombus was not determined.
10852 | 1220648-2025-26948 | 21690576 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21690576&pc=PYX | 02/28/2025 | 04/15/2025 |  | 03/25/2025 |  |  | 02/28/2025 | 03/25/2025 |  |  | 08/31/2025 |  | 03/25/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2024395754 |  | 0046-0035 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 68 YR |  | Male |  | Non Hispanic | White | Injury | Yes | Yes | Restricted Flow rate | [6 empty] | Thrombosis/Thrombus | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed a impella rp to support a patient admitted in with postcardiotomy cardiogenic shock and was placed on the rp and 5.5 pump. the rp was having low flows and was removed after 12 minutes. the team observed there to be thrombus at the pump. the team decided not to replace the pump and the patient remained on the 5.5 pump alone. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed. | The investigation of thrombus and low pump flow has been completed. reported a spike in motor current and drop in flows shortly after initiation. pump was removed and biomaterial was noted on the inlet. pump and purge cassette returned. biomaterial found around impeller. no other damages or abnormalities found. data logs show motor current spike followed by low flows, step up in placement signal and suction shortly after the pump is started. the root cause of the low pump flow was most likely biomaterial since there was biomaterial found around the impeller and ingestion pattern in the data logs. as the necessary information was not provided, the root cause of the thrombus was not determined.
10853 | 1220648-2025-26949 | 21690939 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21690939&pc=OZD | 02/28/2025 | 03/25/2025 |  | 03/25/2025 |  |  | 02/28/2025 | 03/25/2025 |  |  |  |  | 03/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025564845 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 75 YR |  | Female |  | Non Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Retroperitoneal Hemorrhage | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was on impella cp support and the pump was explanted. two hours after the explant, the patient started to decompensate. five units of blood products were given. the physician suspected a retroperitoneal bleed following impella removal. the patient expired. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed.
10854 | 1220648-2025-26950 | 21691224 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21691224&pc=OZD | 02/28/2025 | 03/25/2025 |  | 03/25/2025 |  |  | 02/28/2025 | 03/25/2025 |  |  |  |  | 03/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025551718 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 70 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The user facility reported a 70 year old male patient was on impella cp support and during support, the patient had a hematoma at the access site. the hematoma was expressed and heparin was paused. support continued. after the impella was explanted, the patient developed limb ischemia in the patient's right femoral artery (impella cp access side). the patient was taken to the hybrid room to do an aoro of the right lower extremity. when ro completed, there was no distal flow to the rfa or rfv. a fasciotomy was completed. the doctor believed that this issue was caused by the original rp flex insertion. it is unknown if the impella cp contributed to the limb ischemia. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10855 | 1220648-2025-26951 | 21692843 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21692843&pc=OZD | 02/27/2025 | 04/25/2025 |  | 03/25/2025 |  |  | 02/27/2025 | 03/25/2025 |  |  |  |  | 03/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025539647 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Hispanic | White | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Renal Failure | Gastrointestinal Hemorrhage | [11 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | Followup | [7 empty] | The user facility reported a 64 year old male patient developed a gastrointestinal bleed during impella 5.5 support. heparin was being run through the purge line at the time of the bleed and anticoagulants were halted to manage the condition. support continued uninterrupted. nine days later, the family made the decision to withdraw care and the patient passed away. there is not enough evidence to disassociate the death with the use of the impella. | It was later reported via abiomed¿s long-term outcome and quality indicator (loqi) impella registry that the patient endured an acute kidney injury (aki) during impella support. the aki was believed to have a possible relationship to the impella device. continuous renal replacement therapy was initiated to treat the condition. |  |  |  |  | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella 5.5 with smart assist for use during cardiogenic shock states the following: section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿. | The investigation of vascular damage - peripheral vessel has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26951 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26951. h6 health effect - impact code 4641 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26951. | The product involved in the report has not been returned. a review of the device history record is in-progress. all information reasonably known as of 25 mar 2025 has been included in the health authority report. should additional information be obtained a follow-up health authority report will be provided. the information provided by airlife. represents all of the known information at this time. despite good faith efforts to obtain additional information, the complainant/reporter was unable or unwilling to provide any further patient, product, or procedural details to airlife. airlife has no independent knowledge of the event reported but is relaying the information that was provided by the user facility where the incident occurred. this product incident is documented in the airlife complaint database and is identified as complaint (b)(4). this information is submitted pursuant to 21cfr803, in compliance with the medical device reporting requirement and should not be considered to be an admission that a airlife product is defective or causes serious injury. the complaint of "spontaneous xxx etco2 module on screen" regarding parts 301-5107et and 301-5707f-se was not confirmed. the root cause was not determined but could possibly be a result of user error. a risk assessment was performed and the ultimate risk was determined to be low which does not require escalation to the capa review board. there have been 9 other complaints regarding the same part and a similar issue within the 24 months preceding this reported event. a resolution letter was sent to the customer. complaints will continue to be monitored for possible trends.
10856 | 1220648-2025-26952 | 21692850 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21692850&pc=OZD | 02/28/2025 | 05/16/2025 |  | 03/25/2025 |  |  | 02/28/2025 | 03/25/2025 |  |  |  |  | 03/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025611525 |  | 0048-0014 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 71 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Component Missing | Extra Components | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that a 14fr 25cm peel away sheath was not in the package of a received set. another impella cp set was opened to extract the necessary component for impella implant. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. | The investigation of packaging issues has been completed. no product or images were returned for evaluation. clinical details noted that when the pump packaging was opened to prep the pump, the 14frx25cm dilator was missing. instead, two 14frx13cm dilators were in the packaging. the root cause of the packaging issues was most likely due to a missing component of the 14frx25cm long dilator. a secondary failure mode of "packaging issues" was added to the investigation due to an extra 14frx13cm dilator present in the introducer kit. the root cause of the packaging issues was most likely due to an extra component as an extra 14frx13cm dilator was in the introducer kit. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26952 as it is unknown if the initial reporter also sent the report to the food and drug administration.
10857 | 1220648-2025-26953 | 21692869 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21692869&pc=OZD | 02/28/2025 | 03/25/2025 |  | 03/25/2025 |  |  | 02/28/2025 | 03/25/2025 |  |  |  |  | 03/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1824652 |  | 0042-0040-JP |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | Asian | Malfunction | No | Yes | Battery Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  | No | Initial submission | [9 empty] | The complainant reported that during impella support the automated impella controller (aic) power supply failed and the aic switched to battery power. the aic was exchanged and there was no patient harm. | [5 empty] | The investigation has been completed. the automated impella controller was returned for investigation. the console was tested after arriving in field service. the console indicated that it was running on battery power when it was booted up even when it was connected to ac power. the power cord and fuses were inspected but no issues were found. the voltage was measured between the power supply and the pbm, registering approximately 0.193v, which was way below the normal voltage of 24v. after swapping the power supply with a known good one, the reported power issue resolved. the root cause of the reported power failure issue was determined to be a power supply malfunction. .
10858 | 1220648-2025-26954 | 21692911 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21692911&pc=OZD | 02/28/2025 | 04/15/2025 |  | 03/25/2025 |  |  | 02/28/2025 | 03/25/2025 |  |  |  |  | 03/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025601952 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 77 YR |  | Female |  | Hispanic |  | Death | Yes | Yes | Failure to Pump | [6 empty] | Ischemia | [12 empty] | Death | Required Intervention | [23 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 77-year-old female patient on impella cp supportduring support, there was slow distal perfusion noted via angiogram from impella side port. the patient developed ischemia with a cold mottled foot. additionally, a pump stop occurred. there were low purge pressures with high motor current prior to the pump stop. care was withdrawn and the patient expired. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: troubleshooting the purge system ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿ | The investigation of limb ischemia and pump stop has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26954 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26954. h6 component code 3038 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26954.
10859 | 1220648-2025-26955 | 21692980 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21692980&pc=PYX | 02/28/2025 | 04/11/2025 |  | 03/25/2025 |  |  | 02/28/2025 | 03/25/2025 |  |  | 04/30/2025 |  | 03/25/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP |  |  | ABIOMED, INC. |  |  | US |  | 2025512773 |  | 0046-0035 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 54 YR |  | Male |  | Non Hispanic | Black Or African American | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 54 year old male patient with a rp pump to support due to acute myocardial infarction. the patient suffered a st elevated myocardial infarction (stemi) and ventricular tachycardia (vt) arrest and was transferred to the tertiary center. the pump groin site had to have extra suturing for the access site bleeding. the team then placed ecmo when the pump was explanted. the patient survived the procedure. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. impella rp flex with smartassist instructions for use & clinical reference manual section: using the automated impella controller with the impella rp flex smartassist system catheter use of the repositioning sheath and the 23 fr peel-away introducer ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound.¿ impella rp with smart assist system with automated impella controller section: warnings and cautions ¿make sure there is no bleeding at the transition from the repositioning sheath to the internal jugular vein. close and dress the wound. ¿be sure that the stopcock on the repositioning sheath is always kept in the closed position. significant bleed back can result if the stopcock is open.¿ ¿make sure there is no bleeding at the transition from the repositioning sheath to the femoral vein. close and dress the wound.¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding issue was most likely use related, a second mattress suture was placed, forward retention suture placed for the sheath due to the sheath backing out of the groin to achieve hemostasis. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26955 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26955.
10860 | 1220648-2025-26956 | 21693007 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21693007&pc=OZD | 02/28/2025 | 04/14/2025 |  | 03/25/2025 |  |  | 02/28/2025 | 03/25/2025 |  |  |  |  | 03/25/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2026614172 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported a 52 year old male patient on impella 5.5 support due to cardiomyopathy. during support, there was a lot of sanguineous drainage from the site. the doctor placed another stitch and a sandbag. later a washout was performed at the insertion site. three units of packed red blood cells were given. support continued. the patient was eventually weaned and explanted successfully. | [5 empty] | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smart assist system states the following: section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation). | The investigation of hemolysis and positioning issues has been completed. the impella device was not returned for evaluation. without additional clinical details, the root cause of the reported issues could not be determined. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26956 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26956. g1 manufacturing site fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26956. h10 instructions for use were erroneously on the initial manufacturer device report number 1220648-2025-26956.
10861 | 1220648-2025-26901 | 21675706 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21675706&pc=OZD | 02/24/2025 | 04/09/2025 |  | 03/24/2025 |  |  | 02/24/2025 | 03/24/2025 |  |  |  |  | 03/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025611349 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Perforation | [12 empty] | Life Threatening | Required Intervention | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient was implanted with an impella 5.5 device for mechanical circulatory support (mcs) and experienced a left ventricular perforation requiring intervention with a survived outcome. upon evaluation for transurethral resection of the prostate, cardiology evaluation was performed and was found to have severe left ventricle dysfunction. a transthoracic echocardiogram showed dilated cardiomyopathy with left ventricular ejection fraction of 22% with severe akinesis of septal and anterolateral walls. a left heart catheterization was performed showing diffuse multi-vessel disease to the right coronary artery, the proximal and mid left anterior descending artery, and the left circumflex artery with the left to right collaterals. left ventricle ejection fraction was 27%, right ventricular ejection fraction was 61%, and scar burden was 24%. viability was noted to be absent in the right coronary territory but present to the left. given newly diagnosed systolic heart failure and dilated ischemic cardiomyopathy, the decision was made to place preoperative axillary impella 5.5 to assist with optimization prior to coronary revascularization. the impella 5.5 device was successfully placed, and the patient was sent to the unit on impella mcs. on day eight of mcs, the patient underwent coronary artery bypass graft surgery x 4, mitral valve replacement with a 27 millimeter mosaic bioprosthetic valve, aortotomy for impella 5.5 intra-operative removal and direct placement post completion of mitral. median sternotomy was performed with left internal mammary artery takedown and harvesting of the saphenous vein. central cannulation of the superior vena cava and aorta due to difficulty advancing guidewire to left femoral vein. upon completion of antegrade cardioplegia via root vent, surgical mode was activated and aortotomy was performed. the impella 5.5 was removed from the left ventricle and secured in the chest with a ray-tech sponge. distal grafts and mitral valve were then completed. the impella 5.5 was then placed under direct visualization back into the left ventricle via the aortotomy. upon closure of aortotomy, the tip of the impella 5.5 perforated the left ventricle free wall requiring bovine pericardium repair patch. the surgeon stated this was an accident due to re-inserting the device deep into the left ventricle. the impella support initiated at support level p-1, de-airing techniques and volume resuscitation were performed. minimal pulse pressure was noted despite dual inotropes and vasopressor support. intra-aortic balloon pump (iabp) was placed. blood products were given: four units of packed red blood cells, three units of fresh frozen plasma and three units of platelets, cell saver, 250 colloid, 1.5 liter crystalloid, and methylene blue was given for gross vasoplegia. aortic valve wires and atrial pacing attempted to assist with aortic valve opening. defibrillation x 1 was administered coming off the impella 5.5 pump. coagulation was noted and chest closure was performed. impella position was optimal with inflow 5 centimeters below the aortic valve area. left ventricular ejection fraction was 15%, inferior wall was akinetic. right ventricular function was mildly reduced. no valvular leak was noted at bioprosthesis. surgical sites were dressed. the patient continued on impella support with the iabp was removed the following day. the patient was noted to be feeling "great" the day before planned explant of the impella 5.5 device. day 23 of support, the patient was successfully weaned from the impella 5.5 after what was noted to be successful cardiac surgery. | [5 empty] | The investigation of ventricle perforation has been completed. no product was returned for evaluation. clinical details noted that the pump perforated the ventricle at the apex during coronary artery bypass graft (cabg) and mitral valve replacement (mvr) procedure after re-inserting the pump via direct visualization and the surgeon noted that the pump was likely re-inserted too deep in the ventricle. ventricle perforation was repaired with a patch and the patient was able to remain on support for an additional 17 days. data logs were reviewed and long-term log of full case showed pump was turned down to p-0 at time of cabg procedure on (b)(6) 2025. after the pump was re-inserted and initiated, ¿impella in aorta¿ and ¿impella position unknown¿ alarms are triggered. additionally, placement signal and motor current after the pump was restarted showed very low pulsatility for approximately 30 minutes. the root cause of the ventricle perforation was most likely due to improper positioning after the pump was re-inserted into left ventricle during cabg and mvr procedure. h.5 added selection as it was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26901. | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock. section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings & cautions: warnings: section: pre-support evaluation. section: axillary insertion of the impella 5.5 catheter. section: direct aortic insertion. section: use of impella with transcatheter aortic valves. ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions: ¿handle with care. the impella catheter can be damaged during removal from packaging, preparation, insertion, and removal. do not bend, pull, or place excess pressure on the catheter or mechanical components at any time.¿ section: warnings & cautions: warnings: ¿avoid manual compression of the inlet and outlet areas of the cannula assembly.¿
10862 | 1220648-2025-26902 | 21675756 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21675756&pc=OZD | 02/24/2025 | 04/28/2025 |  | 03/24/2025 |  |  | 02/24/2025 | 03/24/2025 |  |  |  |  | 03/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025559758 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 64 YR |  | Male |  |  |  | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Ischemia | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient in acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support. during support, after approximately six hours after implant, the patient developed limb ischemia. care was escalated to an impella 5.5, and distal perfusion was reestablished after the removal of the impella cp. the patient was reported to be stable following the event. | [5 empty] | The investigation of limb ischemia has been completed. clinical reported malperfusion of the right leg after implant. the pump was explanted and distal perfusion was reestablished. pump was returned. no damages or abnormalities observed. limb ischemia can occur during impella support. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. | The impella device was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10863 | 1220648-2025-26903 | 21675862 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21675862&pc=OZD | 02/24/2025 | 04/14/2025 |  | 03/24/2025 |  |  | 02/24/2025 | 03/24/2025 |  |  |  |  | 03/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025599138 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Female |  | Hispanic | White | Death | Yes | Yes | Increase in Pressure | [6 empty] | Ischemia | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported a patient with acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support and experienced high purge pressure followed by limb ischemia with a demise outcome. the patient underwent elective transcatheter aortic valve replacement. while still in the procedure room while attempting to extubate, the patient started having rhythm changes and decreasing sats eventually going into ventricular fibrillation and ventricular tachycardia requiring numerous defibrillations. coronary angiogram showed coronaries were not compromised. transesophageal echocardiogram showed a severely dilated right ventricle. therefore, the decision was made to place right-sided support as well. the patient was placed on an impella cp and impella rp flex devices (bipella support) with a "very" poor prognosis. approximately four hours into support, the impella cp device experienced high purge pressure. tissue plasminogen activator was initiated in response to the purge pressure high alarm. it was additionally noted that the patient had no distal flow in their lower left extremity (impella cp leg). the persistent condition was monitored by the staff. after two days of support, the patient was no longer responding to the maximum doses of vasopressors and inotropes. the patient passed away on support. medical safety review of the event noted there is not enough evidence to exclude the impella cp device as an associated factor in the patient's outcome (demise). | [5 empty] | The investigation of high purge pressure and limb ischemia has been completed. the impella cp was returned and evaluated. no kinks were found on the pump catheter or purge cassette tubing. biomaterial was found in the purge gap. high purge pressure was reproduced while purging the pump in the lab. no blood ingress into the catheter purge line was observed. the high purge pressure resolved after cutting the catheter just above the motor housing. data logs showed that initial purge pressure was 700-800 mmhg and purge flow was below four milliliters per hour. about four hours into the case, the purge pressure rose from ~700 to greater than 1000 mmhg within about 30 minutes. the high purge pressure did not resolve by the end of the case, which was 43 hours total. no spikes in motor current were seen, but a slight increase throughout the case was seen. the root cause of the high purge pressure was most likely biomaterial buildup since the biomaterial was found on the purge gap of the returned pump, high purge pressure was reproduced, and data logs showed a gradual rise in purge pressure. as all the necessary information was not provided, the root cause of the limb ischemia was not determined. | The impella pump, purge cassette and data stick were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia.
10864 | 1220648-2025-26904 | 21675890 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21675890&pc=OZD | 02/06/2025 | 08/04/2025 |  | 03/24/2025 |  |  | 02/26/2025 | 03/24/2025 |  |  |  |  | 03/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025597136 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Study | Health Professional | User facility |  |  |  | 52 YR |  | Female |  |  |  | Injury | Yes | Yes | Increase in Pressure | Pumping Stopped | [5 empty] | Hematoma | Cardiogenic Shock | Pseudoaneurysm | Thrombocytopenia | [9 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | Months later in (b)(6) 2025 the long-term outcome and quality indicator (loqi) impella registry further reported upon the bleed and pseudoaneurysm and the blood products that were delivered to treat the patient. the patient was medically managed and imaging performed. | The user facility reported a 52-year-old female patient on a cp pump support going for 21 days for cardiogenic shock/cardiomyopathy. the pump was supporting when the pump stopped. the effect to the patient was not known. after the explant the loqi database reported upon various adverse events with "possible" relation to the use of the impella. the allegation was of rising purge pressure (pp), worsening cardiogenic shock which lead to pump explant and medical management, the team also alleged thrombocytopenia which may have been due to blood loss from a hematoma and pseudoanerysm. this was treated by blood products being infused. the patient bridged to left ventricle assist device (lvad). |  |  |  |  | B.5 additional information was received from the clinical site and was added. h.6 added codes due to new information received from the clinical site. the impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist for use during cardiogenic shock and high risk percutaneous coronary intervention section: warnings, cautions, and precautions ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿ impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ b.3 revised date of event as it was previously entered incorrectly on initial manufacturer device report number 1220648-2025-26904. e.1 revised facility name and phone number as they were previously entered incorrectly on initial manufacturer device report number 1220648-2025-26904. an incorrect email address was reported on initial manufacturer device report number 1220648-2025-26904. fax number was added as it was inadvertenly omitted from the initial manufacturer device report number 1220648-2025-26904. g.1 reporting contact facility name was added as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26904. g.2 report source; study was added as it was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26904. h.10 updated additional manufactrurer narrative and instructions for use since the initial manufacturer device report number 1220648-2025-26904. | The impella device will not be received from the customer. therefore, investigation of the device is not possible. should the device or any new information be received, a supplemental mdr will be filed. the instructions for use for the impella cp with smartassist for use during cardiogenic shock states the following: section: potential adverse events (united states). ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: troubleshooting the purge system. ¿unresolved purge pressure high alarm if not resolved by the recommendations provided, high purge pressure¿which triggers the ¿purge pressure high¿ alarm message¿could be an indication of a kink in the impella catheter. in this case, the motor is no longer being purged and may eventually stop.¿ section: impella stopped. ¿if the impella catheter has stopped suddenly: 1. try to restart the catheter at previous p-level. 2. if the impella does not restart, try to restart at p-2. 3. if the impella does not restart or stops again, wait 1 minute and try to restart again. 4. if the impella restarts, wean down to p-2 as the patient can tolerate. under these circumstances, catheter function is not reliable and the impella may stop again. 5. if the impella does not restart, remove the impella from the ventricle as soon as possible to avoid aortic insufficiency.¿
10865 | 1220648-2025-26905 | 21676008 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21676008&pc=OZD | 02/25/2025 | 05/07/2025 |  | 03/24/2025 |  |  | 02/25/2025 | 03/24/2025 |  |  |  |  | 03/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2026613476 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 65 YR |  | Female |  | Non Hispanic | White | Malfunction | No | Yes | Positioning Failure | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported that the patient on impella cp support was being transferred when impella in aorta alarm occurred. attempts were made to repositioned but unsuccessful. the impella cp was exchanged for a new one. the patient remained in stable condition until successful completion of the case. | [5 empty] | The investigation of positioning issues has been completed. the impella device was not returned for evaluation. reported impella in aorta alarms and patient movement just prior to the positioning issue. the root cause of the positioning issue was most likely patient movement since patient movement was reported just prior to the issue. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed.
10866 | 1220648-2025-26906 | 21676321 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21676321&pc=OZD | 02/25/2025 | 04/03/2025 |  | 03/24/2025 |  |  | 02/25/2025 | 03/24/2025 |  |  |  |  | 03/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025610843 |  | 0048-0045 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 79 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Hemorrhage/Blood Loss/Bleeding | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had an impella cp pump placed along with extracorporeal membrane oxygenation for the support of a patient admitted in with acute myocardial infarction/cardiogenic shock. the patient had a hematoma at the femoral site along with some oozing. the team troubleshot with pressure holding, knee immobilizer, and chemical paralysis to reduce the patient movement. additionally, the patient did get red blood cells and albumin. the patient remained stable until successful case completion. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ | The investigation of access site bleeding has been completed. the impella device was not returned for evaluation. the root cause of the access site bleeding was unable to be determined as insufficient clinical details were provided and no product was returned for analysis. h.6 code 4641 was inadvertently omitted from the initial manufacturer device report number 1220648-2025-26906.
10867 | 1220648-2025-26907 | 21676892 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21676892&pc=OZD | 02/24/2025 | 04/16/2025 |  | 03/24/2025 |  |  | 02/24/2025 | 03/24/2025 |  |  |  |  | 03/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1114573 |  | 0042-0000-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Female |  |  | White | Death | Yes | Yes | Application Program Problem | [6 empty] | Cardiac Arrest | Hemodynamic instability | [11 empty] | Death | Life Threatening | Required Intervention | [22 empty] | R |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported a patient (unknown age and ethnicity) with an ejection fraction of 20% (decrease from 45% one day prior) in cardiogenic shock from an acute myocardial infarction was implanted with an impella cp device for mechanical circulatory support. thereafter, it was noted the automated impella controller (aic) screen was black while patient was being actively supported. assistance was called and the team was instructed to turn the aic back on using the power button. upon power up, ¿controller failure¿ and ¿unexpected console shutdown¿ alarms presented. the team then advised to do an aic-to-aic transfer, and the case completed without issue on secondary controller. the patient was sent to the unit on bilevel positive airway pressure (bipap) for recovery and subsequent transfer. however, within three hours of start of support, while on the unit, the patient expired. the patient was noted to be doing fine on bipap on arrival to unit; however, the patient's oxygen saturation started dropping. during intubation the patient was given a large amount of sedatives. the patient's heart stopped, and cardiopulmonary resuscitation was started; the patient coded for approximately 40 minutes until time of demise was called. | [5 empty] | The investigation of automated impella controller (aic) - controller failure (red alarm) has been completed. the console logs from the reported day of event show that about half an hour after pump start, imc log entries end abruptly, and then aic is power-cycled, with the first boot-up unsuccessful and second boot-up concludes with controller failure alarm and unexpected controller shutdown. although the pump was automatically restarted within 30 seconds, the purge system was no longer operational. four minutes later the pump was transferred to a backup console, where the support continued without any issues for about 2.5 hours. console was tested and controller failure alarm was reproduced. replacing purge pressure sensor did not resolve the issue, however when pud board was replaced, the alarm resolved even with the original purge pressure sensor. console shutdown/restart was not reproduced during testing. reported controller failure alarm was caused by malfunction of purge pressure measurement system, due to pud board defect. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26907 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The automated impella controller device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
10868 | 1220648-2025-26908 | 21676988 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21676988&pc=OZD | 02/25/2025 | 03/24/2025 |  | 03/24/2025 |  |  | 02/25/2025 | 03/24/2025 |  |  |  |  | 03/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025600674 |  | 0048-0045 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 72 YR |  | Female |  | Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | [11 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant had a impella cp pump placed to support a patient admitted in with acute myocardial infarction and cardiogenic shock. the pump was placed but there was an access site bleed/hematoma. the team applied manual pressure but bleeding continued. additionally distal limb noted to have decreased perfusion as evidenced by decrease capillary refill, cool, and discolored. decision made by the doctor to explant impella device at bedside. once explanted doppler pulses confirmed for right pedal pulse. the procedural outcome was the patient survived surgery. | [5 empty] | The impella device was discarded by the customer, therefore, investigation of the device was not possible. should any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿ section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿
10869 | 1220648-2025-26909 | 21677005 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21677005&pc=OZD | 02/26/2025 | 03/24/2025 |  | 03/24/2025 |  |  | 02/26/2025 | 03/24/2025 |  |  |  |  | 03/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025496172 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 48 YR |  | Male |  | Hispanic | White | Death | Yes | No | Restricted Flow rate | [6 empty] | Ischemia | Renal Failure | Tachycardia | [10 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant reported the patient was on impella cp support and during support, the patient had acute kidney injury (aki). the patient's urine was wine color. the patient's cr was 1.18 on admission and was 2.18 while on support. dialysis was started. the relationship between the aki and the impella is unknown. additionally, there was low pump flow. one unit of packed red blood cells were given to the patient and the suction events resolved. furthermore, the patient had limb ischemia. a fem bypass was performed. distal limb still cold, ischemic and mottled. the pump was attempted to be weaned but the patient went into ventricular tachycardia. the patient was shocked many times. care was withdrawn and the patient expired. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation.¿
10870 | MW5168015 | 21677013 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21677013&pc=QOC | 08/09/2024 | 03/17/2025 |  | 03/24/2025 |  |  |  | 03/24/2025 |  |  |  |  | 03/24/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA SMART ASSIST.5 |  |  | QOC | [5 empty] | ABIOMED, INC. | [8 empty] | No | [6 empty] | Health Professional |  |  | Voluntary report | Not Available | Physician | [6 empty] | 72 YR |  | Male |  | Non Hispanic | White | Death | No | Yes | [7 empty] | Insufficient Information | [12 empty] | Hospitalization | Required Intervention | Life Threatening | Disability | Other | Death | [23 empty] | Initial submission | [9 empty] | Impella placed and dad never recovered and got extremely worse.
10871 | 1220648-2025-26910 | 21677374 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21677374&pc=OZD | 02/26/2025 | 11/25/2025 |  | 03/24/2025 |  |  | 02/26/2025 | 03/24/2025 |  |  |  |  | 03/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025608857 |  | 0048-0045 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 66 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemolysis | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant had placed an impella cp for support of a patient admitted in for ablation. the patient had signs of hemolysis and was explanted after 25 hours and replaced with an intra-aortic balloon pump. the procedural outcome was the patient survived the surgery. | [5 empty] | The investigation has been completed. the pump was returned for investigation. the data log shows no signals related to biomaterial interaction, positioning issues, or the patient condition during the case run. no other physical damage was observed on the returned product other than a kink in the cannula. a tiny amount of blood was seen in the delo joint between the cannula and the cage. the pump was then tested in a bucket of water, and it ran according to specifications. the pump was subsequently sent for hemolysis testing and failed the test with an mih > 20. the pump was closely examined after a tear down and no impeller damage was observed. the root cause of the hemolysis issue was unable to be definitively determined based on returned product and log analysis. d10 concomitant medical products and therapy dates updated. | The impella pump was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
10872 | 1220648-2025-26911 | 21677524 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21677524&pc=OZD | 02/26/2025 | 05/07/2025 |  | 03/24/2025 |  |  | 02/26/2025 | 03/24/2025 |  |  |  |  | 03/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025610901 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Male |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Arrhythmia | Hemolysis | Ischemia | [10 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The complainant reported the patient was on impella cp support and during support, the patient had limb ischemia. distal pulses were not able to dopplered. support continued. additionally, the patient had signs of hemolysis and the patient was in route for vt ablation at this time. the pump was repositioned. however, eventually the pump was explanted. the procedural outcome was the patient survived surgery. | [5 empty] | The investigation of limb ischemia, vf/vt/af/at, and hemolysis has been completed. the impella device was not returned for evaluation. based on clinical description, the root cause of hemolysis was most likely due to pump was not in the proper position. as all the necessary information was not provided, the root cause of the limb ischemia and vf/vt/af/at was not determined. h6 component code 3038 was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26911. | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smartassist system: section: potential adverse events (united states): ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: use of echocardiography for positioning of the impella catheter: ¿evaluate the position of the impella catheter if the automated impella controller displays position alarms or if you observe lower than expected flows or signs of hemolysis. if the catheter does not appear to be correctly positioned, initiate steps to reposition it.¿ impella catheter too far into the left ventricle: ¿obstruction of the impella catheter inlet area can lead to increased mechanical forces on blood cell walls and subsequent hemolysis, which often presents as dark or blood colored urine.¿ section: hemolysis: ¿hemolysis should be monitored during support. patients who develop high levels of hemolysis may show signs of decreased hemoglobin levels, dark or blood-colored urine, and in some cases, acute renal failure. plasma-free hemoglobin (pfhgb) is the best indicator to confirm whether a patient is exposed to an unacceptable level of hemolysis¿. ¿patient conditions¿ including catheter position, pre-existing medical conditions, and small left ventricular volumes¿ may also play a role in patient susceptibility to hemolysis.¿ section: suction: ¿suction may occur if the blood volume available for the impella catheter is inadequate or restricted. suction limits the amount of support that the impella catheter can provide to the patient and results in a decrease in arterial pressure and cardiac output. it can damage blood cells, leading to hemolysis. it may also be an indicator of right heart failure.¿
10873 | 1220648-2025-26912 | 21677963 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21677963&pc=OZD | 02/16/2025 | 03/24/2025 |  | 03/24/2025 |  |  | 02/26/2025 | 03/24/2025 |  |  |  |  | 03/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025584192 |  | 0048-0024-JP |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 73 YR |  | Male |  | Non Hispanic | Asian | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Stroke/CVA | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported that while a patient was on impella cp support they suffered a cerebral hemorrhage and died with no signs of improvement in bilateral heart failure. the doctors commented that the cause of death was not related to the impella and the impella functioned normally without any issues. however, the device could not conclusively be disassociated from the outcome. | [5 empty] | The impella device was not received from the customer. therefore, the investigation of the device was not possible. should the device or any new information be received, a supplemental manufacturer device report will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿ section: warnings, cautions & precautions: ¿to reduce the possibility of fibers being drawn into the impella, customers should avoid exposing the inlet and cannula section of the impella heart pumps to any surfaces or fluid baths where the device can come into contact with loose or floating fibers.¿
10874 | 1220648-2025-26914 | 21678702 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21678702&pc=OZD | 02/24/2025 | 03/24/2025 |  | 03/24/2025 |  |  | 02/24/2025 | 03/24/2025 |  |  |  |  | 03/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025603452 |  | 1000080 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 62 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hematoma | Ischemia | Unspecified Vascular Problem | [10 empty] | Required Intervention | Disability | [23 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient with an acute myocardial infarction/cardiogenic shock was implanted with an impella cp device for mechanical circulatory support (mcs). the patient presented to the emergency department and was found to have significant 4-5 mm st elevation anteriorly with reciprocal changes. the patient was taken emergently to the catheterization lab, coronary angiogram showed multi-vessel disease (proximal left anterior descending (lad) 100% occlusion, 95% circumflex (lcx), and 40 % right coronary artery [rca]- right side dominant). left ventricular gram showed severely depressed ejection fraction estimated at 20%. percutaneous coronary intervention (pci) was performed to the lad and lcx arteries. the patient was intubated during pci for hypoxemia. post pci, the patient began decompensating. decision made to exchange sheath for impella peel away and place urgently place the impella cp via the right femoral artery, which was completed without complications independently. the patient was transferred to the unit on mcs. same day later that afternoon, a large hematoma with fem-stop was noted at the access site with positive distal pulses. subsequently, the patient was taken to operating for pulmonary artery catheter placement, right femoral artery repair, and venoarterial extracorporeal membrane oxygenation (va ecmo) cannulation. the right femoral artery was repaired, but the impella cp was found to be occluding the right artery. right-sided superficial femoral artery was configured during cannulation to provide distal perfusion. after initiation of ecmo flow, p-level was decreased to p-5 and two units packed red blood cells and 1500ml lr was given. heparin was reversed per surgeon preference due to surgical bleeding. the impella cp was repositioned under transesophageal echocardiogram to 3.8 cm. patient determined to be hemodynamically stable and transferred to back to the unit. the patient continued support for two additional days before the device was explanted. the outcome at explant is unknown to note. | [5 empty] | The impella device was not received from the customer; therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. instructions for use for the related event are as follows: impella cp with smartassist system. section: potential adverse events: ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury¿ (including ventricular perforation).
10875 | 1220648-2025-26913 | 21678703 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21678703&pc=OZD | 01/24/2025 | 03/24/2025 |  | 03/24/2025 |  |  | 01/24/2025 | 03/24/2025 |  |  |  |  | 03/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025555856 |  | 1000100 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 52 YR |  | Female |  | Non Hispanic | White | Injury | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced serious drainage from their impella access site during support. a suture was placed at the site and the condition resolved. support continued uninterrupted until the patient was bridged to a transplant. | [5 empty] | The investigation into the reported issue has been completed. no device was received for evaluation. device history record review confirmed that the pump passed all post sterile inspection checks. clinical information was sufficient to determine the root cause of the reported issue. the root cause was determined to be use issue because the bleeding was resolved by adding an additional suture at the access site.
10876 | MW5168064 | 21678753 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21678753&pc=OZD | 03/11/2025 | 03/18/2025 |  | 03/24/2025 |  |  |  | 03/24/2025 |  |  |  |  | 03/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA |  |  | OZD | [5 empty] | ABIOMED, INC. | [8 empty] | No | [6 empty] | Health Professional |  |  | Voluntary report | Yes | Risk Manager | [8 empty] | Female |  |  |  | Malfunction | No | Yes | [7 empty] | No Clinical Signs, Symptoms or Conditions | [41 empty] | Initial submission | [9 empty] | Abiomed impella cp set catheter advanced into patient's body. catheter attached to impella machine and failed to activate. catheter was removed, a new impella cp set was opened and advanced, the new catheter worked appropriately.
10877 | 1220648-2025-26915 | 21679318 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21679318&pc=OZD | 02/26/2025 | 04/11/2025 |  | 03/24/2025 |  |  | 02/26/2025 | 03/24/2025 |  |  |  |  | 03/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025536647 |  | 1000080 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 58 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Restricted Flow rate | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  |  | Initial submission | Followup | [8 empty] | The user facility reported a patient was implanted with an impella cp without any issues. when securing the groin, the impella cp began to have waveform issues, low flows and suction events. echocardiogram was brought in and the impella cp position was stable. it was decided to exchange the impella cp for a new one. the patient survived the procedure. | [5 empty] | The investigation of low pump flow has been completed. the returned product was inspected and there were no abnormalities. the pump had dried dextrose around the impeller which cleared after soaking the pump. the pump was run in the lab and low pump flow did not reproduce. the data logs from the case were reviewed and low pump flow for matching p-level was confirmed. there were no ingestion patterns or motor current spikes observed. there were suction alarms which resolved when the pump was turned down to p0. there were low flows after the pump was turned back to p8. clinical details mentioned blood volume was given with no resolution. the root cause of the low pump flow issue was not determined as the issue did not reproduce on the returned product and limited clinical information related to failure was provided. | The impella pump, purge cassette and data stick were returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
10878 | 1220648-2025-26916 | 21679390 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21679390&pc=OZD | 02/27/2025 | 04/23/2025 |  | 03/24/2025 |  |  | 02/27/2025 | 03/24/2025 |  |  |  |  | 03/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | OZD |  |  | AUTOMATED IMPELLA CONTROLLER |  |  | ABIOMED, INC. |  |  | US |  | 1651078 |  | 0042-0040-US |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | No | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 0 DA |  | Unknown |  |  |  | Malfunction | No | Yes | Application Program Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | R |  |  |  | Initial submission | Followup | Followup | [7 empty] | The customer reported that they sent an automated impella controller back due to repeated fail attempts to start. the console was investigated by the post market engineer team and advised the technician to change the sd flash cards due to corruption, and a new set of battery. both components were replaced and the console booted up fine, as were the new battery charging to spec. upon receiving the console, the customer stated that the console still remains the same having problems starting up. there was no patient involvement. | [5 empty] | The investigation of automated impella controller (aic) - power on issues/blank screen has been completed. the console logs were not available for analysis, since the console would not boot up. after arriving for inspection, the issue was reproduced: console would not boot up, being ¿stuck¿ at the initial stage of the process, presumably before os was accessed. replacing cf cards with known-good ones did not resolve the issue. after replacing entire 4-in-1 assembly (including cf cards), the console was able to boot up. analysis of original cf cards using diagnostic software indicated that cards were corrupt. console¿s inability to boot up was caused by defective 4-in-1 and corrupted cf cards (damaged by malfunctioning 4-in-1.) | The investigation of automated impella controller (aic) - power on issues/blank screen has been completed. the console logs were not available for analysis, since the console would not boot up. during the investigation, the issue was reproduced: console would not boot up, being ¿stuck¿ at the initial stage of the process, presumably before os was accessed. replacing cf cards with known-good ones did not resolve the issue. after replacing entire 4-in-1 assembly (including cf cards), the console was able to boot up. analysis of original cf cards using diagnostic software indicated that cards were corrupt. the root cause of the aic issue was a defective 4in1 board. | The console was returned, and an evaluation is underway. upon investigation closure, a supplemental manufacturer device report will be filed.
10879 | 1220648-2025-26918 | 21681485 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21681485&pc=OZD | 12/24/2024 | 03/24/2025 |  | 03/24/2025 |  |  | 12/25/2024 | 03/24/2025 |  |  |  |  | 03/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025571375 |  | 0048-0014 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 77 YR |  | Male |  |  |  | Injury | Yes | Yes | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Required Intervention | [24 empty] | Invalid/Incomplete |  |  |  | Initial submission | [9 empty] | The complainant reported a patient was implanted with an impella device for mechanical circulatory support. the patient experienced access site bleeding. best practices were discussed and some were applied. additionally, there was discussion on administering blood products. however, the bleeding was resolved and no further intervention was specified. it was also reported that there were positioning issues. the device was repositioned. the patient continued on support until the device was successfully weaned. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella ifu: impella cp with smartassist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation).¿
10880 | 1220648-2025-26920 | 21682609 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21682609&pc=OZD | 02/26/2025 | 03/24/2025 |  | 03/24/2025 |  |  | 02/26/2025 | 03/24/2025 |  |  |  |  | 03/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025606263 |  | 1000080 |  | No |  |  | No |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 64 YR |  | Female |  | Non Hispanic | Black Or African American | Death | Yes | No | Adverse Event Without Identified Device or Use Problem | [6 empty] | Hemorrhage/Blood Loss/Bleeding | [12 empty] | Death | [24 empty] | Invalid/Incomplete |  |  | No | Initial submission | [9 empty] | The complainant had placed a impella cp pump for the support of a patient admitted in with st-segment elevation myocardial infarction and coding during cardiac percutaneous coronary intervention. the cp was placed and supported for 51 hours til pump removed. the access site was oozing and blood products were given, manual hold and heparin titration. it was reported that care was withdrawn and the patient expired. the relationship between the impella and cause of death is unknown. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella cp with smart assist system section: general patient care considerations ¿assess access site for bleeding and hematoma.¿ section: entering cardiac output use of the repositioning sheath and peel-away introducer ¿remove the peel-away introducer completely from the artery over the catheter shaft to prevent trauma and significant bleeding and apply manual pressure above the puncture site.¿ section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)."
10881 | 1220648-2025-26921 | 21682630 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21682630&pc=OZD | 02/26/2025 | 04/25/2025 |  | 03/24/2025 |  |  | 02/27/2025 | 03/24/2025 |  |  |  |  | 03/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA 5.5 |  |  | OZD |  |  | IMPELLA 5.5 |  |  | ABIOMED, INC. |  |  | US |  | 2025567999 |  | 1000100 |  | No |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 67 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Advance | Patient-Device Incompatibility | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant reported that the impella 5.5 was advanced through the graft anastomosis, but the tip was unable to be advanced beyond a certain point in the patient¿s anatomy. after multiple attempts, the decision was to abort the 5.5 and place a cp through the axillary graft. there was no patient harm. | [5 empty] | The impella device was not received from the customer, therefore, investigation of the device was not possible. should the device or any new information be received, a supplemental mdr will be filed. as noted in the impella instructions for use: impella 5.5 with smartassist for use during cardiogenic shock section: potential adverse events (united states) ¿acute renal dysfunction, aortic valve injury, bleeding, cardiogenic shock, cerebral vascular accident/stroke, death, hemolysis, limb ischemia, myocardial infarction, renal failure, thrombocytopenia and cardiac or vascular injury (including ventricular perforation)¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position. [addition for 5.0, 5.5 only] in instances where the impella pump has been placed prior to performing cardiac surgery with aortic cross clamping and cardioplegic arrest, care should be taken when manipulating the heart when the pump position is fixed with application of the aortic cross clamp across the pump catheter.¿ ¿to reduce the risk of cardiac or vascular injury (including ventricular perforation) when advancing or torquing the impella, adjustments should be performed under imaging guidance. section: warnings & cautions: warnings section: pre-support evaluation section: axillary insertion of the impella 5.5 catheter section: direct aortic insertion section: use of impella with transcatheter aortic valves ¿to reduce the risk of cardiac injury (including ventricular perforation), physicians should exercise special care when inserting the impella catheter in patients with complex anatomy. this includes patients with known or suspected decreased ventricular cavity size, ventricular aneurysms, congenital heart disease, or compromised cardiac tissue quality in the settings of acute infarction with tissue necrosis.¿ ¿to reduce the risk of vascular injury, physicians should exercise caution when inserting the impella catheter in patients with complex peripheral vascular anatomy. this includes patients with known or suspected: unrepaired abdominal aortic aneurysm, significant descending thoracic aortic aneurysm, dissection of the ascending/ transverse/descending aorta, chronic anatomical changes in the relationship of the aorta/aortic valve/ventricular alignment, significant mobile atheromatous disease in the thoracic or abdominal aorta or peripheral vessels.¿ ¿physicians should exercise special care when inserting the impella catheter during active cardiopulmonary resuscitation (cpr). in addition, active cpr maneuvers may change the position of the impella device, introducing a risk of cardiac or vascular injury (including ventricular perforation). check that the pump is positioned correctly in the left ventricle after cpr with echocardiography guidance. section: warnings & cautions: cautions ¿dilators and catheters should be removed slowly from the sheath. rapid removal may damage the valve, resulting in blood flow through the valve.¿ section: warnings & cautions: warnings ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance prior to manipulating the heart, and monitor position.¿ ¿to reduce the risk of cardiac or vascular injury (including perforation) when manipulating the heart during cardiac surgery, evaluate the position of the pump using imaging guidance.¿ | The investigation of delivery issues has been completed. the impella device was not returned for evaluation. based on the clinical details the root cause of delivery issue was most likely patient condition since the tip was unable to be advanced beyond a certain point in the patient¿s anatomy. b3 date of event was erroneously entered on the initial manufacturer device report number 1220648-2025-26921. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26921 as it is unknown if the initial reporter also sent the report to the food and drug administration.
10882 | 1220648-2025-26922 | 21682644 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21682644&pc=OZD | 02/27/2025 | 05/28/2025 |  | 03/24/2025 |  |  | 02/27/2025 | 03/24/2025 |  |  |  |  | 03/24/2025 |  | TEMPORARY NON-ROLLER TYPE LEFT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA CP |  |  | OZD |  |  | IMPELLA CP |  |  | ABIOMED, INC. |  |  | US |  | 2025609316 |  | 0048-0014 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Foreign | Health Professional | User facility | Company representation |  |  | 65 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Failure to Pump | Device Sensing Problem | [5 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The complainant had a impella cp pump being placed to support a patient admitted in with cardiogenic shock and cardiomyopathy. it was reported that impella showed no placement of left ventricle curves. the connection of the plug has been checked. the medical doctor reports that the activated clotting time is in range and they found no thrombus in the ventricle. it was further reported the pump start not possible, medical stuff try to prep impella cp again however not successful. the controller was replaced but same problem occurred. they decide to use a new impella cp set. the patient was stable. | [5 empty] | The investigation of optical signal issue and pump did not start has been completed. the data logs were not returned. pump did not start. there were no damages to the returned product. the pump was purged with both purge cassettes and was able to run for 42 minutes in the lab with no issues. due to lack of data logs returned and clinical information provided, the root cause of the pump did not start cannot be determined. optical signal issue. the 5s parameters were taken from the pump and they were within specification. when connecting the pump to the console, there were placement signal not reliable (psnr) alarms that reproduced. the imc logs were checked for the pump connected to the lab console and the calibrated g0=0. the root cause of the optical signal issue is most likely due to an eeprom corruption error, however the cause is unknown. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26922 as it is unknown if the initial reporter also sent the report to the food and drug administration. | The impella device was returned, and an evaluation is underway. upon completion of our analysis, a supplemental report will be filed.
10883 | 1220648-2025-26923 | 21682670 | https://www.accessdata.fda.gov/scripts/cdrh/cfdocs/cfmaude/detail.cfm?mdrfoi__id=21682670&pc=PYX | 02/27/2025 | 04/14/2025 |  | 03/24/2025 |  |  | 02/27/2025 | 03/24/2025 |  |  |  |  | 03/24/2025 |  | TEMPORARY NON-ROLLER TYPE RIGHT HEART SUPPORT BLOOD PUMP |  |  | IMPELLA RP |  |  | PYX |  |  | IMPELLA RP FLEX |  |  | ABIOMED, INC. |  |  | US |  | 2025587516 |  | 1000323 |  | Device was returned to manufacturer |  |  | Returned to Manufacturer |  | Yes | No | Health Professional |  |  | Manufacturer report | Yes | Physician | Health Professional | User facility | Company representation |  |  |  | 56 YR |  | Male |  | Non Hispanic | White | Malfunction | No | Yes | Device Sensing Problem | [6 empty] | No Clinical Signs, Symptoms or Conditions | [37 empty] | Invalid/Incomplete |  |  | No | Initial submission | Followup | [8 empty] | The user facility reported that a 56 year old male patient had a rp flex placed after the patient was having ventricular tachycardia (vt) when being weaned from anesthesia. there was right ventricular (rv) failure and patient was being supported by the rp and 5.5 for the coronary artery bypass graft (cabg). the rp had signal issues and was removed and replaced. the patient survived the procedure. | [5 empty] | The investigation of optical signal issue has been completed. the product was returned and evaluated. there were no abnormalities found with the optical sensor from visual inspection and 5s parameter check. there were also no abnormalities found with the dp sensor from visual inspection and electrical testing. the pump was run in a bucket of water and there were no abnormalities related to the sensor seen during this time. pump flow was high during the test run but that is expected when running the pump in water per input from r&d. field data log review of ~5 minutes while the pump was connected to the console but outside of the patient showed the pap read ~100 mmhg and ps read ~50 mmhg. a reading from the sensors while outside of the patient does not indicate any sensor issues per r&d input. no abnormalities were found with the data logs. the dp sensor measures the pressure difference between the inside and outside of the cannula, and the pressure value is used to monitor flow and is an input for the placement signal per the instructions for use there were also no alarms/events associated with the sensors during this time. there was no problem found during the case nor with the returned pump. the pump functioned to specification d4 model number was erroneously entered on the initial manufacturer device report number 1220648-2025-26923. e4 should have been left blank on the initial manufacturer device report number 1220648-2025-26923 as it is unknown if the initial reporter also sent the report to the food and drug administration. g1 reporting contact facility name and fax number was inadvertently omitted on the initial manufacturer device report number 1220648-2025-26923. h6 investigation conclusions code 4315 as erroneously entered on the initial manufacturer device report number 1220648-2025-26923. | The impella device has been received but has not yet been evaluated. when the device investigation has been completed or any new information has been received, a supplemental mdr will be filed.